## Sheet1
index | CMPLNT_HDR_NUM | CMPLNT_HDR_DESC | SR_CREATE_TS_UTC_parsed | SYS_TYP_DESC | Code | CMPLNT_CD_DESC | CMPLNT_ITEM_NUM | MATL_NUM_ID | CMPLNT_ITEM_DESC | Append_Text | EVENT_DESCRIPTION | Append_Text_cleaned | Top_1_Prediction | Top_2_Prediction | Top_3_Prediction | Top_4_Prediction | Top_5_Prediction | Top5_Combined | Top5_Accuracy
3581 | 703380728 | XSE - Software and Camera Issues | 2019-09-25T15:53:50.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Software and Camera Issues 2019-09-24 SCL1902893, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during training. It was reported that when opening the door on the workstation, the power switch was still fully in the "ON" position even though the system was unplugged. The representative started up the system, but stealth was not an option. When clicking on setup process, the Nav code and instruments were not there. The navigation camera was then plugged in and there was a set of beeps for power, but not a second set. All different user types including user, demo, and super user were tried, but stealth was still not an option. Two hard restarts of the system were done, but the issue was not resolved. While demoing the system for the a surgeon, both X-Eye and 3Define cameras were not working. After the demo, the representative went into the softare utilities for the cameras and all cameras were set to off. All cameras were enabled in the utilities. After doing a soft restart and unplugging and plugging the navigation camera in, stealth was an option. The representative indicated the system had not been shut down properly. There was no pat ient involvement in the case. 2019-09-26 email 2019-09-26 18:00 (Rep): Additional information received from a manufacturer representative reported that demo mode was being used during the event. No error messages were displayed. 2019-10-10 email 2019-10-10 07:30, email 2019-10-10 09:15 (Rep): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-09-24 SCL1902893, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during training. It was reported that when opening the door on the workstation, the power switch was still fully in the "ON" position even though the system was unplugged. The representative started up the system, but stealth was not an option. When clicking on setup process, the Nav code and instruments were not there. The navigation camera was then plugged in and there was a set of beeps for power, but not a second set. All different user types including user, demo, and super user were tried, but stealth was still not an option. Two hard restarts of the system were done, but the issue was not resolved. While demoing the system for the a surgeon, both X-Eye and 3Define cameras were not working. After the demo, the representative went into the softare utilities for the cameras and all cameras were set to off. All cameras were enabled in the utilities. After doing a soft restart and unplugging and plugging the navigation camera in, stealth was an option. The representative indicated the system had not been shut down properly. There was no pat ient involvement in the case. 2019-09-26 email 2019-09-26 18:00 (Rep): Additional information received from a manufacturer representative reported that demo mode was being used during the event. No error messages were displayed. 2019-10-10 email 2019-10-10 07:30, email 2019-10-10 09:15 (Rep): No new information. | xse software camera issues training opening door workstation power switch fully position unplugged started stealth option clicking setup process nav code instruments navigation camera plugged set beeps power second set different user types including user demo super user tried stealth option hard restarts issue resolved demoing surgeon xeye cameras working demo went softare utilities cameras cameras set cameras enabled utilities doing soft restart unplugging plugging navigation camera stealth option indicated shut properly pat ient case email additional demo mode error messages displayed email email new mazor x software installation kit | NAV3246 | NAV2058 | NAV4110 | NAV2056 | NAV2054 | NAV3246NAV2058NAV4110NAV2056NAV2054 | True
6857 | 704051241 | XSE - Error 293 during SRA Setup | 2020-11-04T14:43:04.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 30 | KIT0830 | USA REGIONAL KIT KIT0830 | XSE - Error 293 during SRA Setup 2020-11-04 00700186, Word Doc (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was failing setup and displayed a error message 293. A hard reboot of the system was done. The surgical arm was setup and the arm moved for about five minutes, but the same error message was displayed. The system software was recently updated. When the manufacturer representative turned the system on, there were no user profiles or previous scans on the system. A new profile was created and the patient scans were loaded. The use of the guidance system was aborted and the surgeon proceeded with the case using navigation. There was no patient involvement. 2020-11-04 IUD x7 (other): No new information. 2021-02-10 IUD (other): No new information. USA REGIONAL KIT KIT0830 | 2020-11-04 00700186, Word Doc (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was failing setup and displayed a error message 293. A hard reboot of the system was done. The surgical arm was setup and the arm moved for about five minutes, but the same error message was displayed. The system software was recently updated. When the manufacturer representative turned the system on, there were no user profiles or previous scans on the system. A new profile was created and the patient scans were loaded. The use of the guidance system was aborted and the surgeon proceeded with the case using navigation. There was no patient involvement. 2020-11-04 IUD x7 (other): No new information. 2021-02-10 IUD (other): No new information. | xse error sra setup word doc outside procedure surgical arm failing setup displayed error message hard reboot surgical arm setup arm moved minutes error message displayed software recently updated turned user profiles previous scans new profile created scans loaded use aborted surgeon proceeded case using navigation iud new iud new usa regional kit | NAV3246 | NAV3249 | NAV3248 | NAV4095 | NAV1006 | NAV3246NAV3249NAV3248NAV4095NAV1006 | True
3419 | 703342710 | XSE - unable to register | 2019-09-03T20:22:03.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - unable to register 2019-Aug-29 SCL1902507 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that segmentation was successful on the first try from T7-L2, however on fluoro to CT matching there were yellow values at T7, T8, T9 and L2, and there were red values at T10, T11, T12, and L1. They attempted to adjust segmentation lines several times and manages to turn the yellow value at T9 to red, they also took a few more x- rays and tried to re register several times. The two L2 trajectories were sent and the Right sided trajectory hit the BMB, while the left side was way outside the open incision area. The T7 and T8 trajectories on the left side were sent and drilled accurately. The surgeon decided to unmount the robot and to try for the more superior levels. He was instructed on the best places to place the dual clamp to best reach the superior trajectories, but he determined that due to the patients' coronal imbalance that it would be too difficult to send the trajectories around the BMB. The surgeon decided to free hand all of the other screws. It was noted that it was impossible to say how long exactly the case was delayed, but due to the fact that they spent 30 mins active with the robot and only successfully sent 2 of the 30 planned trajectories, the rep would guess it was at least over an hour. There was no patient harm or additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM TPL0059 | 2019-Aug-29 SCL1902507 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that segmentation was successful on the first try from T7-L2, however on fluoro to CT matching there were yellow values at T7, T8, T9 and L2, and there were red values at T10, T11, T12, and L1. They attempted to adjust segmentation lines several times and manages to turn the yellow value at T9 to red, they also took a few more x- rays and tried to re register several times. The two L2 trajectories were sent and the Right sided trajectory hit the BMB, while the left side was way outside the open incision area. The T7 and T8 trajectories on the left side were sent and drilled accurately. The surgeon decided to unmount the robot and to try for the more superior levels. He was instructed on the best places to place the dual clamp to best reach the superior trajectories, but he determined that due to the patients' coronal imbalance that it would be too difficult to send the trajectories around the BMB. The surgeon decided to free hand all of the other screws. It was noted that it was impossible to say how long exactly the case was delayed, but due to the fact that they spent 30 mins active with the robot and only successfully sent 2 of the 30 planned trajectories, the rep would guess it was at least over an hour. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse unable register spinal procedure segmentation successful try fluoro ct matching yellow values red values attempted adjust segmentation lines times manages turn yellow value red took x rays tried register times trajectories sent right sided trajectory hit bmb left way outside open incision area trajectories left sent drilled accurately surgeon decided unmount robot try superior levels instructed best places place dual clamp best reach superior trajectories determined patients coronal imbalance difficult send trajectories bmb surgeon decided free hand screws noted impossible say long exactly case delayed fact spent mins active robot successfully sent planned trajectories guess hour additional complications anticipated mazor x | NAV3243 | NAV2051 | NAV5000 | NAV3250 | NAV1012 | NAV3243NAV2051NAV5000NAV3250NAV1012 | False
3300 | 703306859 | XSE - No Registration - Aborted | 2019-08-13T12:25:10.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059-04 | MAZOR X SYSTEM (REV 4) | XSE - No Registration - Aborted 2019-08-08 SCL1902347, CCR x2 (HCP, REP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that previous hardware was present at T3-T8. The previous hardware was removed prior to mounting and registration (Hooks and Rods). A midline incision was open to expose the spine. 3Define and Draw Spine were both successful. A dual clamp was attached to T3 and T5. Validation was successful off T6, however acceptable values were not achieved. Multiple manual adjustment and registration attempts were attempted. New AP and OBL images were taken. T2 was the only body to achieve an acceptable value. Visual verification confirmed no shift. Arm sent to T2 on the left. Surgeon said the trajectory was off and pointing at T1. Surgeon made the decision to remount and attempt registration at the bottom on the construct. The dual clamp was attached to T12 and T10. Validation was successful off T11, however acceptable values were not achieved. Multi ple manual adjustment and registration attempts were attempted. New AP and OBL images were taken and registration failed again. A new 3D Marker was used and registration again failed. The surgeon made the decision to abort the use of the guidance system and freehand the screws. There was known impact on the patient outcome. Information unrelated to this event was included in a separate PE; See PE 703310421 for arm guide issues. 2019-08-14 email 2019-08-14 10:15 (Rep): Additional information received from a manufacturer representative reported the procedure was delayed roughly one hour due to the registration issues. There was no patient harm and the all of the screws were successfully placed freehand. MAZOR X SYSTEM (REV 4) | 2019-08-08 SCL1902347, CCR x2 (HCP, REP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that previous hardware was present at T3-T8. The previous hardware was removed prior to mounting and registration (Hooks and Rods). A midline incision was open to expose the spine. 3Define and Draw Spine were both successful. A dual clamp was attached to T3 and T5. Validation was successful off T6, however acceptable values were not achieved. Multiple manual adjustment and registration attempts were attempted. New AP and OBL images were taken. T2 was the only body to achieve an acceptable value. Visual verification confirmed no shift. Arm sent to T2 on the left. Surgeon said the trajectory was off and pointing at T1. Surgeon made the decision to remount and attempt registration at the bottom on the construct. The dual clamp was attached to T12 and T10. Validation was successful off T11, however acceptable values were not achieved. Multi ple manual adjustment and registration attempts were attempted. New AP and OBL images were taken and registration failed again. A new 3D Marker was used and registration again failed. The surgeon made the decision to abort the use of the guidance system and freehand the screws. There was known impact on the patient outcome. Information unrelated to this event was included in a separate PE; See PE 703310421 for arm guide issues. 2019-08-14 email 2019-08-14 10:15 (Rep): Additional information received from a manufacturer representative reported the procedure was delayed roughly one hour due to the registration issues. There was no patient harm and the all of the screws were successfully placed freehand. | xse registration aborted ccr spinal procedure treat degenerative disc fixation previous hardware present previous hardware removed prior mounting registration midline incision open expose spine draw spine successful dual clamp attached validation successful acceptable values achieved multiple manual adjustment registration attempts attempted new ap obl images taken body achieve acceptable value visual verification confirmed shift arm sent left surgeon said trajectory pointing surgeon decision remount attempt registration construct dual clamp attached validation successful acceptable values achieved multi ple manual adjustment registration attempts attempted new ap obl images taken registration failed new marker registration failed surgeon decision abort use freehand screws known impact outcome unrelated included separate pe pe arm guide issues email additional procedure delayed roughly hour registration issues screws successfully placed freehand mazor x | NAV4112 | NAV3242 | NAV4091 | NAV5000 | NAV3243 | NAV4112NAV3242NAV4091NAV5000NAV3243 | True
8694 | 704463061 | XSE - Navigation Inaccurate | 2021-06-22T19:54:26.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 20 | M728894B252 | NAVIGATION SOFTWARE MODULE 1.2 | XSE - Navigation Inaccurate 2021-Jun-22 Salesforce C00228904, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a T12-L2 open procedure was done. The surgical system was mounted to the patient using a screw in the PSIS. The surgeon started at left T12 and worked down and then moved to right T12 and worked down. When the surgeon was navigating with the drill on the right hand side, the trajectory of the drill looked more than 2 mm superior. Due to the inaccuracy, the surgeon applied extra pressure to match their planned screws. Lateral images were taken after the case and the screw appeared accurate. The manufacturer representative completed a 3 point and 10 point accuracy test after the case and the system passed with no issue. There was no patient harm and the procedure was delayed less than an hour. 2021-07-01 IUD x3 (Other): No new information. NAVIGATION SOFTWARE MODULE 1.2 | 2021-Jun-22 Salesforce C00228904, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a T12-L2 open procedure was done. The surgical system was mounted to the patient using a screw in the PSIS. The surgeon started at left T12 and worked down and then moved to right T12 and worked down. When the surgeon was navigating with the drill on the right hand side, the trajectory of the drill looked more than 2 mm superior. Due to the inaccuracy, the surgeon applied extra pressure to match their planned screws. Lateral images were taken after the case and the screw appeared accurate. The manufacturer representative completed a 3 point and 10 point accuracy test after the case and the system passed with no issue. There was no patient harm and the procedure was delayed less than an hour. 2021-07-01 IUD x3 (Other): No new information. | xse navigation inaccurate salesforce iud spinal procedure open procedure surgical mounted using screw psis surgeon started left worked moved right worked surgeon navigating drill right hand trajectory drill looked mm superior inaccuracy surgeon applied extra pressure match planned screws lateral images taken case screw appeared accurate completed point point accuracy test case passed issue procedure delayed hour iud new navigation software module | NAV5001 | NAV5000 | NAV3264 | NAV1012 | NAV3243 | NAV5001NAV5000NAV3264NAV1012NAV3243 | True
5020 | 703591547 | X - Mounting/Arm issues | 2020-01-31T16:22:54.000000+0000 | MazorX | NAV2051 | INTRA-OP ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Mounting/Arm issues 2018-04-19 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the advanced accuracy test completed. Shoulder calibration and camera function checked. Surgical plan was reviewed with surgeon prior to surgery and no issues were noted. The system was attached to the bed by one person with no issues. Representative had ancillary staff monitor system indication for left plate sensor to turn green. Plate sensor indicating light still not fixed. 120 mm schanz pin placed in patient's right PSIS. BMB was attached and the guidance system mounted. 3Define scan and draw spine were completed. Middle body labeled L5. Two x-rays were taken in AP. Guidance system bumped with C-arm when obtaining the oblique image. The system got an arm shift error. The representative checked the stability of the schanz pin. The guidance system was remou nted to the patient. 3Define scan and draw spine were repeated. One x-ray was taken to obtain AP. Two x-rays were taken to obtain oblique image. Registration was successful for bodies L3 - S1. Arm was sent to mark skin in all trajectories. Surgeon exposed on left side and arm was re-sent and k wires placed. Surgeon repeated procedure on the right side. Lateral x-rays were taken throughout to verify. AP image was taken once all k wires were inserted. Surgeon was pleased with placement but was unsure whether k wire was in pedicle at L5 on the right. All trajectories tapped and stimmed at or greater than 20. Surgeon proceeded with decompression and interbody placement. The guidance system was removed from the bed with no issues prior to decompression per the surgeon’s request. There was no patient harm. MAZOR X SYSTEM TPL0059 | 2018-04-19 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the advanced accuracy test completed. Shoulder calibration and camera function checked. Surgical plan was reviewed with surgeon prior to surgery and no issues were noted. The system was attached to the bed by one person with no issues. Representative had ancillary staff monitor system indication for left plate sensor to turn green. Plate sensor indicating light still not fixed. 120 mm schanz pin placed in patient's right PSIS. BMB was attached and the guidance system mounted. 3Define scan and draw spine were completed. Middle body labeled L5. Two x-rays were taken in AP. Guidance system bumped with C-arm when obtaining the oblique image. The system got an arm shift error. The representative checked the stability of the schanz pin. The guidance system was remou nted to the patient. 3Define scan and draw spine were repeated. One x-ray was taken to obtain AP. Two x-rays were taken to obtain oblique image. Registration was successful for bodies L3 - S1. Arm was sent to mark skin in all trajectories. Surgeon exposed on left side and arm was re-sent and k wires placed. Surgeon repeated procedure on the right side. Lateral x-rays were taken throughout to verify. AP image was taken once all k wires were inserted. Surgeon was pleased with placement but was unsure whether k wire was in pedicle at L5 on the right. All trajectories tapped and stimmed at or greater than 20. Surgeon proceeded with decompression and interbody placement. The guidance system was removed from the bed with no issues prior to decompression per the surgeon’s request. There was no patient harm. | x mountingarm issues ccr spinal procedure advanced accuracy test completed shoulder calibration camera function checked surgical plan reviewed surgeon prior surgery issues noted attached bed person issues ancillary staff monitor indication left plate sensor turn green plate sensor indicating light fixed mm schanz pin placed patients right psis bmb attached mounted scan draw spine completed middle body labeled xrays taken ap bumped carm obtaining oblique image got arm shift error checked stability schanz pin remou nted scan draw spine repeated xray taken obtain ap xrays taken obtain oblique image registration successful bodies arm sent mark skin trajectories surgeon exposed left arm resent k wires placed surgeon repeated procedure right lateral xrays taken verify ap image taken k wires inserted surgeon pleased placement unsure k wire pedicle right trajectories tapped stimmed greater surgeon proceeded decompression interbody placement removed bed issues prior decompression surgeon’s request mazor x | NAV3242 | NAV4091 | NAV3250 | NAV2051 | NAV3254 | NAV3242NAV4091NAV3250NAV2051NAV3254 | True
13 | 702971082 | X - Comm Issue with Surgical Sys.-Case 1 | 2019-01-23T10:24:04.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Comm Issue with Surgical Sys.-Case 1 2019-01-21 SC SCL1900205, PDF 2019-01-23 05:16 (Rep, HCP): Medtronic received information regarding a guidance system being used for level two extreme lateral interbody fusion (XLIF) with lateral posterior spinal function (PSF) . It was reported that all trajectories were instrumented and when the surgeon was introducing screws, the manufacturer representative noticed the light on the shoulder was blinking; however, the shoulder was not unlocked. The manufacturer representative thought there was a miscommunication between the encoders and workstation. There was no delay in the procedure over an hour. The case was not altered and was successfully completed. After the case, the manufacturer representative did some troubleshooting and found the arm, shoulder and workstation were not communicating normally. Sometimes the arm would lock and sometimes it would not. The light continually blinked regardless of the lock status. No patient harm was reported. Information unrelated to this event was included in separate PEs. See PE 702971083 for Case 2, 702971084 for X-Eye Issues, 702971085 for Arm Vibration, and 702971376 for Schanz Arm Malfunction. 2019-02-06 e1 (Rep): No new information. 2019-02-26 e2 (Rep): No new information. UNK MAZOR X SHOULDER | 2019-01-21 SC SCL1900205, PDF 2019-01-23 05:16 (Rep, HCP): Medtronic received information regarding a guidance system being used for level two extreme lateral interbody fusion (XLIF) with lateral posterior spinal function (PSF) . It was reported that all trajectories were instrumented and when the surgeon was introducing screws, the manufacturer representative noticed the light on the shoulder was blinking; however, the shoulder was not unlocked. The manufacturer representative thought there was a miscommunication between the encoders and workstation. There was no delay in the procedure over an hour. The case was not altered and was successfully completed. After the case, the manufacturer representative did some troubleshooting and found the arm, shoulder and workstation were not communicating normally. Sometimes the arm would lock and sometimes it would not. The light continually blinked regardless of the lock status. No patient harm was reported. Information unrelated to this event was included in separate PEs. See PE 702971083 for Case 2, 702971084 for X-Eye Issues, 702971085 for Arm Vibration, and 702971376 for Schanz Arm Malfunction. 2019-02-06 e1 (Rep): No new information. 2019-02-26 e2 (Rep): No new information. | x comm issue surgical syscase sc pdf level extreme lateral interbody fusion lateral posterior spinal function trajectories instrumented surgeon introducing screws noticed light shoulder blinking shoulder unlocked thought miscommunication encoders workstation delay procedure hour case altered successfully completed case did troubleshooting arm shoulder workstation communicating normally arm lock light continually blinked regardless lock status unrelated included separate pes pe case xeye issues arm vibration schanz arm malfunction new new unk mazor x shoulder | NAV4107 | NAV2058 | NAV4101 | NAV2047 | NAV4108 | NAV4107NAV2058NAV4101NAV2047NAV4108 | True
10329 | 704855855 | XSE - Frequent Low Pressure Message | 2022-02-09T21:09:36.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE - Frequent Low Pressure Message 2022-Feb-09 Salesforce C00242976 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system displayed a low pressure error message. The compressor would not turn on to build pressure, which caused the manufacturer representative not to be able to move the positioner. The issue occurred when attempting to move the surgical system away from the patient so the representative could remove it from the bed. There was no patient harm and the procedure was not delayed. 2022-02-18 IUD (Other): No new information. 2022-02-22 email 2022-02-22 05:46 (Rep): No new information. 2022-04-19 email 2022-04-19 11:02 (Rep): No new information. 2022-04-22 email 2022-04-22 13:16 (Rep): No new information. Information unrelated to this event was included in a separate PE; See PE 704904769. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2022-Feb-09 Salesforce C00242976 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system displayed a low pressure error message. The compressor would not turn on to build pressure, which caused the manufacturer representative not to be able to move the positioner. The issue occurred when attempting to move the surgical system away from the patient so the representative could remove it from the bed. There was no patient harm and the procedure was not delayed. 2022-02-18 IUD (Other): No new information. 2022-02-22 email 2022-02-22 05:46 (Rep): No new information. 2022-04-19 email 2022-04-19 11:02 (Rep): No new information. 2022-04-22 email 2022-04-22 13:16 (Rep): No new information. Information unrelated to this event was included in a separate PE; See PE 704904769. | xse frequent low pressure message salesforce spinal p rocedure displayed low pressure error message compressor turn build pressure caused able positioner issue occurred attempting surgical away remove bed procedure delayed iud new email new email new email new unrelated included separate pe pe kit svc surgical new | NAV4104 | NAV3264 | NAV2054 | NAV3250 | NAV3251 | NAV4104NAV3264NAV2054NAV3250NAV3251 | True
3166 | 703268324 | XSE - surgeon screen won't turn on | 2019-07-19T15:21:29.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - surgeon screen won't turn on 2019-Jul-18 SCL1902133 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon screen would not turn on when the system was booted up. It had not worked for the last two cases. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. SURGEON SCREEN | 2019-Jul-18 SCL1902133 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon screen would not turn on when the system was booted up. It had not worked for the last two cases. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse surgeon screen wont turn spinal procedure surgeon screen turn booted worked cases additional complications anticipated surgeon screen | NAV3253 | NAV2054 | NAV3295 | NAV5000 | NAV1006 | NAV3253NAV2054NAV3295NAV5000NAV1006 | True
9854 | 704734936 | XSE - SRA Failed Advanced Accuracy Test | 2021-11-26T00:18:35.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - SRA Failed Advanced Accuracy Test 2021-Nov-02 Salesforce C00238821, 00726473, IUD x2 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed an advanced accuracy test after the system software was updated. The surgical arm was replaced. The replacement surgical arm failed the stress test due to max current being exceeded for J3 and position errors for J4 and J5. A field motor calibration was done and the surgical arm passed the stress test. There was no patient involvement. 2022-01-30 IUD (Other, For): No new information. 2022-02-14 IUD (Other, For): No new information. 2022-03-08 IUD (Other, For): No new information. 2022-03-29 302769001 (Other, For): No new information. 2022-03-30 IUD x4 (Other, For): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2021-Nov-02 Salesforce C00238821, 00726473, IUD x2 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed an advanced accuracy test after the system software was updated. The surgical arm was replaced. The replacement surgical arm failed the stress test due to max current being exceeded for J3 and position errors for J4 and J5. A field motor calibration was done and the surgical arm passed the stress test. There was no patient involvement. 2022-01-30 IUD (Other, For): No new information. 2022-02-14 IUD (Other, For): No new information. 2022-03-08 IUD (Other, For): No new information. 2022-03-29 302769001 (Other, For): No new information. 2022-03-30 IUD x4 (Other, For): No new information. | xse sra failed advanced accuracy test salesforce iud outside procedure surgical arm failed advanced accuracy test software updated surgical arm replaced replacement surgical arm failed stress test max current exceeded position errors field motor calibration surgical arm passed stress test iud new iud new iud new new iud new kit svc surgical arm refurb | NAV3249 | NAV4093 | NAV2054 | NAV4092 | NAV4095 | NAV3249NAV4093NAV2054NAV4092NAV4095 | True
7528 | 704223449 | XSE - CANBUS Error | 2021-02-10T17:45:26.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 40 | KIT1237-01 | SVC KIT1237-01 CABLE MAIN SYSTEM | XSE - CANBUS Error 2021-Feb-10 Salesforce C00220783 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Mazor system was showing a CANBUS error 10251 stating fatal error connection lost on nodes 1-6. No patient present. 2021-02-11 IUD x9 (Other): No new information. 2021-03-27 IUD x3 (Other): No new information. 2021-05-08 IUD (Other): No new information. SVC KIT1237-01 CABLE MAIN SYSTEM | 2021-Feb-10 Salesforce C00220783 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Mazor system was showing a CANBUS error 10251 stating fatal error connection lost on nodes 1-6. No patient present. 2021-02-11 IUD x9 (Other): No new information. 2021-03-27 IUD x3 (Other): No new information. 2021-05-08 IUD (Other): No new information. | xse canbus error salesforce outside procedure mazor showing canbus error stating fatal error connection lost nodes present iud new iud new iud new svc cable main | NAV2047 | NAV2054 | NAV3256 | NAV3262 | NAV3261 | NAV2047NAV2054NAV3256NAV3262NAV3261 | True
4530 | 703549098 | XSE - Monitor Issue | 2020-01-07T20:00:14.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 30 | ASM0275-01S | KIT SVC ASM0275-01S MONITOR | XSE - Monitor Issue 2020-01-07 00681501 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the work station monitor has the front plastic perimeter piece of the monitor damaged and it is currently fastened to the system with tape. No patient was present. 2020-01-28 IUD x6 (Other): No new information. 2020-02-10 IUD (Rep): No new information. 2022-05-16 IUD (Other): No new information. KIT SVC ASM0275-01S MONITOR | 2020-01-07 00681501 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the work station monitor has the front plastic perimeter piece of the monitor damaged and it is currently fastened to the system with tape. No patient was present. 2020-01-28 IUD x6 (Other): No new information. 2020-02-10 IUD (Rep): No new information. 2022-05-16 IUD (Other): No new information. | xse monitor issue outside procedure work station monitor plastic perimeter piece monitor damaged currently fastened tape present iud new iud new iud new kit svc monitor | NAV3261 | NAV2054 | NAV2058 | NAV3253 | NAV3251 | NAV3261NAV2054NAV2058NAV3253NAV3251 | True
6743 | 704020777 | XSE - Compressor Running Often | 2020-10-19T20:17:40.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE - Compressor Running Often 2020-10-19 00699071 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the compressor in the workstation was running more often. The system built pressure, but the compressor had to come on to maintain pressure in the system. There was no patient involvement. 2020-10-20 IUD x2 (Other): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2020-10-19 00699071 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the compressor in the workstation was running more often. The system built pressure, but the compressor had to come on to maintain pressure in the system. There was no patient involvement. 2020-10-20 IUD x2 (Other): No new information. | xse compressor running outside procedure compressor workstation running built pressure compressor come maintain pressure iud new kit svc pneumatic panel | NAV3264 | NAV1006 | NAV2054 | NAV2047 | NAV2058 | NAV3264NAV1006NAV2054NAV2047NAV2058 | True
1222 | 703193365 | X - Shoulder Issue | 2019-06-06T17:44:46.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Shoulder Issue Screw deviations; Shoulder doesn’t lock on startup from track ball – had to lock from software – was discussed on Whatsapp group. Plan: Surgeon planned to do case MIS but had to switch to open due to lack of correct hardware -surgeon planned case with RTC. Mount: Dual clamp. Sync: Surface scan was accurate -draw spine didn't work because we couldn't move the arm any further up on the bed due to hip pads so the angle was too large for eyes to see reference marker. Registration was very difficult due to multiple reasons -chest pad created too much of a shadow -had to remove to finally get an AP that worked -arm pads were maxed out but created a lot of issues trying to get OB image -segmentation was finally achieved, L4 didnt match, went back to adjust and lost L5. Decided to proceed forward and instrument available trajectories and the went back and re-registered to get L5. Operate: Multiple medial deviations -soft tissue caused a lot of trouble especially at the superior and inferior portion -the clamp and bridge got in the way for multiple trajectories due to rotated SPs. L5 left was completely unreachable so surgeon opted to freehand MAZOR X SOFTWARE KIT VER-3.0.2 | Screw deviations; Shoulder doesn’t lock on startup from track ball – had to lock from software – was discussed on Whatsapp group. Plan: Surgeon planned to do case MIS but had to switch to open due to lack of correct hardware -surgeon planned case with RTC. Mount: Dual clamp. Sync: Surface scan was accurate -draw spine didn't work because we couldn't move the arm any further up on the bed due to hip pads so the angle was too large for eyes to see reference marker. Registration was very difficult due to multiple reasons -chest pad created too much of a shadow -had to remove to finally get an AP that worked -arm pads were maxed out but created a lot of issues trying to get OB image -segmentation was finally achieved, L4 didnt match, went back to adjust and lost L5. Decided to proceed forward and instrument available trajectories and the went back and re-registered to get L5. Operate: Multiple medial deviations -soft tissue caused a lot of trouble especially at the superior and inferior portion -the clamp and bridge got in the way for multiple trajectories due to rotated SPs. L5 left was completely unreachable so surgeon opted to freehand | x shoulder issue screw deviations; shoulder doesn’t lock startup track ball – lock software – discussed whatsapp group plan surgeon planned case mis switch open lack correct hardware surgeon planned case rtc mount dual clamp sync surface scan accurate draw spine didnt work arm bed hip pads angle large eyes reference marker registration difficult multiple reasons chest pad created shadow remove finally ap worked arm pads maxed created lot issues trying ob image segmentation finally achieved didnt match went adjust lost decided proceed forward instrument available trajectories went reregistered operate multiple medial deviations soft tissue caused lot trouble especially superior inferior portion clamp bridge got way multiple trajectories rotated sps left completely unreachable surgeon opted freehand mazor x software kit | NAV4091 | NAV3242 | NAV3246 | NAV4107 | NAV4095 | NAV4091NAV3242NAV3246NAV4107NAV4095 | True
3417 | 703342708 | X - Dual Clamp Screw Stripped | 2019-09-03T20:05:33.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1355-03 | DUAL CLAMP | X - Dual Clamp Screw Stripped 2019-Aug-28 SCL1902501 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the screw in the Dual Clamp and Clamp link was stripped and could not be removed. There was no patient harm or additional complications reported or anticipated as a result of the event. DUAL CLAMP | 2019-Aug-28 SCL1902501 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the screw in the Dual Clamp and Clamp link was stripped and could not be removed. There was no patient harm or additional complications reported or anticipated as a result of the event. | x dual clamp screw stripped spinal procedure screw dual clamp clamp link stripped removed additional complications anticipated dual clamp | NAV3299 | NAV3295 | NAV4090 | NAV4100 | NAV4095 | NAV3299NAV3295NAV4090NAV4100NAV4095 | True
8041 | 704340699 | XSE - Boom Monitor Not Working | 2021-04-14T16:14:40.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Boom Monitor Not Working 2021-Apr-14 Salesforce C00224634 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system would not connect to the external boom monitors. There was no patient involvement. 2021-04-21 IUD x4 (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Apr-14 Salesforce C00224634 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system would not connect to the external boom monitors. There was no patient involvement. 2021-04-21 IUD x4 (Other): No new information. | xse boom monitor working salesforce outside procedure connect external boom monitors iud new unk mazor x workstation | NAV2058 | NAV3261 | NAV3264 | NAV2047 | NAV3262 | NAV2058NAV3261NAV3264NAV2047NAV3262 | True
1896 | 703220835 | X- Trajectories Deviation | 2019-06-21T17:20:06.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X- Trajectories Deviation Dr. Vaught-Southeast Hospital L4-L5 PSF 8th case at Southeast System #68 . Planning was completed the day before the procedure. Surgeon was shown axial angles and soft tissue pressure was discussed. Surgeon was happy with axial angles and was shown how far off midline the trajectories were at skin level. Surgeon used single schantz pin for single level L4-L5 PSF. 8:25-Arm attach, one person 30sec. 8:42-Robot drape-1 person, 1.5min. 9:17-Platform mounted (single schantz pin) 9:18-3Define 9:19-Draw spine 9:20-AP, 1 image 9:24-Oblique, 3 images.-Generic AP and OB were used to save time because there was not an initial path to OB image. 9:28-registration was achieved on basic with all green values. All levels were confirmed by surgeon.-All trajectories were sent initially but not drilled at the request of the surgeon to get an idea how far he would have to bilaterally retract for his midline incision. -Once retractors were in place trajectories were resent. -L4 left sent and trajectory would not fit into incision. CSR talked about tran smucular approach but the easiest option would be to go back in planning and change the axial angles a few degrees more medial. Surgeon was very firm on decision to not change his "ideal" plan. -Surgeon then pulled out retractors and lengthened his incision. -Re tractors were placed again once incision was lengthened. -L4 left was resent and drilled and K-wire placed.-L5 left was sent and drilled and K-wire placed. -L4 right was sent and did not fit into incision. Instrumentation would collide with bilateral retractor. -CSR suggested changing the axial angles slightly and this would likely result in success, but surgeon would not accept anything other than his original plan. -Surgeon said he would expose more but again did not have success fitting the L4 right trajectory within the incision. Before any further suggestions could be made, the surgeon aborted the case and decided just to use navigation. -A c-arm image was then taken of the remaining L4 and L5 left K-wires and both K-wires had superior deviation from the planned trajectory. Also the L5 left trajectory does not appear to have the same lateral angle as the plan.-Surgeon was not happy with results. His biggest complaint was the fact that when looking at skin incisions the distance off midline was not even close to reality when trying to place trajectories inside an open midline incision. The plan accounted for the distance off midline being about 4 -4.5cm in reality he said that it was greater than 6.5cm for both sides.-A possible reason the trajectories at L4 and L5 left had deviated may very well be related to the retraction used. Although bilateral retraction was used at all times, the surgeon placed and replaced the bilateral retraction system many times in his effort to lengthen the incision, as he would not change the axial angles of the original plan. MAZOR X SOFTWARE KIT VER-2.3 | Dr. Vaught-Southeast Hospital L4-L5 PSF 8th case at Southeast System #68 . Planning was completed the day before the procedure. Surgeon was shown axial angles and soft tissue pressure was discussed. Surgeon was happy with axial angles and was shown how far off midline the trajectories were at skin level. Surgeon used single schantz pin for single level L4-L5 PSF. 8:25-Arm attach, one person 30sec. 8:42-Robot drape-1 person, 1.5min. 9:17-Platform mounted (single schantz pin) 9:18-3Define 9:19-Draw spine 9:20-AP, 1 image 9:24-Oblique, 3 images.-Generic AP and OB were used to save time because there was not an initial path to OB image. 9:28-registration was achieved on basic with all green values. All levels were confirmed by surgeon.-All trajectories were sent initially but not drilled at the request of the surgeon to get an idea how far he would have to bilaterally retract for his midline incision. -Once retractors were in place trajectories were resent. -L4 left sent and trajectory would not fit into incision. CSR talked about tran smucular approach but the easiest option would be to go back in planning and change the axial angles a few degrees more medial. Surgeon was very firm on decision to not change his "ideal" plan. -Surgeon then pulled out retractors and lengthened his incision. -Re tractors were placed again once incision was lengthened. -L4 left was resent and drilled and K-wire placed.-L5 left was sent and drilled and K-wire placed. -L4 right was sent and did not fit into incision. Instrumentation would collide with bilateral retractor. -CSR suggested changing the axial angles slightly and this would likely result in success, but surgeon would not accept anything other than his original plan. -Surgeon said he would expose more but again did not have success fitting the L4 right trajectory within the incision. Before any further suggestions could be made, the surgeon aborted the case and decided just to use navigation. -A c-arm image was then taken of the remaining L4 and L5 left K-wires and both K-wires had superior deviation from the planned trajectory. Also the L5 left trajectory does not appear to have the same lateral angle as the plan.-Surgeon was not happy with results. His biggest complaint was the fact that when looking at skin incisions the distance off midline was not even close to reality when trying to place trajectories inside an open midline incision. The plan accounted for the distance off midline being about 4 -4.5cm in reality he said that it was greater than 6.5cm for both sides.-A possible reason the trajectories at L4 and L5 left had deviated may very well be related to the retraction used. Although bilateral retraction was used at all times, the surgeon placed and replaced the bilateral retraction system many times in his effort to lengthen the incision, as he would not change the axial angles of the original plan. | x trajectories deviation dr vaughtsoutheast hospital psf case southeast planning completed day procedure surgeon shown axial angles soft tissue pressure discussed surgeon happy axial angles shown far midline trajectories skin level surgeon single schantz pin single level psf attach person person mounted spine image imagesgeneric ap ob save time initial path ob image achieved basic green values levels confirmed surgeonall trajectories sent initially drilled request surgeon idea far bilaterally retract midline incision retractors place trajectories resent left sent trajectory fit incision csr talked tran smucular approach easiest option planning change axial angles degrees medial surgeon firm decision change ideal plan surgeon pulled retractors lengthened incision tractors placed incision lengthened left resent drilled kwire left sent drilled kwire placed right sent did fit incision instrumentation collide bilateral retractor csr suggested changing axial angles slightly likely success surgeon accept original plan surgeon said expose did success fitting right trajectory incision suggestions surgeon aborted case decided just use navigation carm image taken remaining left kwires kwires superior deviation planned trajectory left trajectory does appear lateral angle plansurgeon happy results biggest complaint fact looking skin incisions distance midline close reality trying place trajectories inside open midline incision plan accounted distance midline reality said greater sidesa possible reason trajectories left deviated related retraction bilateral retraction times surgeon placed replaced bilateral retraction times effort lengthen incision change axial angles original plan mazor x software kit | NAV4091 | NAV4110 | NAV5000 | NAV2044 | NAV3241 | NAV4091NAV4110NAV5000NAV2044NAV3241 | True
8651 | 704457516 | XSE - Not Booting Up/Error Messages | 2021-06-18T15:52:04.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 130 | ASM0270 | KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | XSE - Not Booting Up/Error Messages 2021-Jun-18 Salesforce C00228729, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the guidance system would not boot up past the splash screen. The manufacturer representative powered down the system for five minutes and then power the system back up, but the issue was not resolved. The PC power was cycled and Error Message 4496 Robot Calibration Not Found was displayed. The representative was able to log into the software, but Error Message 3344 Emergency Stop Button Pressed was displayed. The system would not allow the representative to reset the emergency stop button. Another hard restart was done and the issue was not resolved. There was no patient involvement. 2021-06-25 IUD x6 (Other): No new information. 2021-07-18 IUD x5 (Other): No new information. 2021-08-02 IUD (Other): No new information. 2021-08-02 IUD x6 (Other): No new information. 2021-08-13 IUD x11 (Other): No new information. 2021-09-12 302559211, IUD x2 (Other): No new information. 2021-09-20 IUD (Other): No new information. 2021-10-04 IUD x2 (Other): No new information. 2021-10-24 IUD (Other): No new information. 2021-10-31 IUD x3 (Other): No new information. 2021-11-08 IUD (Other): No new information. 2021-11-08 IUD x5, 302621970 (Other): No new information. 2021-12-08 IUD (Other): No new information. KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | 2021-Jun-18 Salesforce C00228729, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the guidance system would not boot up past the splash screen. The manufacturer representative powered down the system for five minutes and then power the system back up, but the issue was not resolved. The PC power was cycled and Error Message 4496 Robot Calibration Not Found was displayed. The representative was able to log into the software, but Error Message 3344 Emergency Stop Button Pressed was displayed. The system would not allow the representative to reset the emergency stop button. Another hard restart was done and the issue was not resolved. There was no patient involvement. 2021-06-25 IUD x6 (Other): No new information. 2021-07-18 IUD x5 (Other): No new information. 2021-08-02 IUD (Other): No new information. 2021-08-02 IUD x6 (Other): No new information. 2021-08-13 IUD x11 (Other): No new information. 2021-09-12 302559211, IUD x2 (Other): No new information. 2021-09-20 IUD (Other): No new information. 2021-10-04 IUD x2 (Other): No new information. 2021-10-24 IUD (Other): No new information. 2021-10-31 IUD x3 (Other): No new information. 2021-11-08 IUD (Other): No new information. 2021-11-08 IUD x5, 302621970 (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse booting uperror messages salesforce iud outside p rocedure boot past splash screen powered minutes power issue resolved pc power cycled error message robot calibration displayed able log software error message emergency stop button pressed displayed allow reset emergency stop button hard restart issue resolved iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new kit svc compressor kit new | NAV2047 | NAV3246 | NAV3248 | NAV3264 | NAV1006 | NAV2047NAV3246NAV3248NAV3264NAV1006 | True
283 | 703066180 | XSE - Burned Up PC/Monitor Connector | 2019-03-21T13:59:54.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | UNK_MAZORX_PC | UNK MAZOR X PC | XSE - Burned Up PC/Monitor Connector 2019-03-15 SCL1900666, Intake Form, Image (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative plugged the monitor HDMI cable into a silver box in the operating room that provided 5 volts of power. The power caused the graphics bard in the PC and the HDMI connector of the monitor to burn up. There was no patient harm, but the use of the robot was aborted and the case was not completed. There was a delay in the procedure of more than an hour. 2019-04-22 e1 (Rep): Additional information received from a manufacturer representative reported the use of the guidance system was completely aborted and they did not get a second chance to use the system as the monitor went out. The case was delayed about 45 minutes and the patient was operated on using stealth navigation. UNK MAZOR X PC | 2019-03-15 SCL1900666, Intake Form, Image (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative plugged the monitor HDMI cable into a silver box in the operating room that provided 5 volts of power. The power caused the graphics bard in the PC and the HDMI connector of the monitor to burn up. There was no patient harm, but the use of the robot was aborted and the case was not completed. There was a delay in the procedure of more than an hour. 2019-04-22 e1 (Rep): Additional information received from a manufacturer representative reported the use of the guidance system was completely aborted and they did not get a second chance to use the system as the monitor went out. The case was delayed about 45 minutes and the patient was operated on using stealth navigation. | xse burned pcmonitor connector image spinal procedure plugged monitor hdmi cable silver box operating room provided volts power power caused graphics bard pc hdmi connector monitor burn use robot aborted case completed delay procedure hour additional use completely aborted did second chance use monitor went case delayed minutes operated using stealth navigation unk mazor x pc | NAV2049 | NAV3261 | NAV3262 | NAV2058 | NAV3263 | NAV2049NAV3261NAV3262NAV2058NAV3263 | True
6685 | 704004895 | XSE - Screw Deviation at L4 Left | 2020-10-09T20:12:18.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 30 | ASM0206-01 | SURGICAL ARM | XSE - Screw Deviation at L4 Left 2020-10-07 00698521, CCR, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that the left L4 screw was deviated during an open L4-L5 TLIF. The surgeon operated on the right L4 and L5 without issue so they proceeded to the left side. L4 was first, followed by L5. When executing L5, the surgeon found that the incision was not large enough. This caused excessive lateral pressure on the surgical arm from the soft tissue and led to the shoulder unlocking. The surgeon decided to abort the use of the guidance system for left L5. During the procedure, there was an issue with air pressure in the guidance system. The pneumatic motor would not boot on by itself. Once during the operation step and once after the procedure, a low pressure error message was displayed and the system had to be rebooted to get the pressure to be restored. After the procedure, the confirmation images showed that L4 left was medial by 3 mm. The surgeon believed this was due to the excessive pressure placed on the surgical arm. L5 left was accurate. The patient was experiencing nerve damage. The procedure was delayed less than an hour. 2020-10-22 IUD x3, email 2020-10-19 15:00 (Rep): Additional information received from a manufacturer representative reported that the L4 screw was repositioned freehand. Confirmation images were taken and the screw was accurately positioned. It was unknown if any action was taken due to the nerve damage. 2020-10-26 IUD (Other): No new information. SURGICAL ARM | 2020-10-07 00698521, CCR, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that the left L4 screw was deviated during an open L4-L5 TLIF. The surgeon operated on the right L4 and L5 without issue so they proceeded to the left side. L4 was first, followed by L5. When executing L5, the surgeon found that the incision was not large enough. This caused excessive lateral pressure on the surgical arm from the soft tissue and led to the shoulder unlocking. The surgeon decided to abort the use of the guidance system for left L5. During the procedure, there was an issue with air pressure in the guidance system. The pneumatic motor would not boot on by itself. Once during the operation step and once after the procedure, a low pressure error message was displayed and the system had to be rebooted to get the pressure to be restored. After the procedure, the confirmation images showed that L4 left was medial by 3 mm. The surgeon believed this was due to the excessive pressure placed on the surgical arm. L5 left was accurate. The patient was experiencing nerve damage. The procedure was delayed less than an hour. 2020-10-22 IUD x3, email 2020-10-19 15:00 (Rep): Additional information received from a manufacturer representative reported that the L4 screw was repositioned freehand. Confirmation images were taken and the screw was accurately positioned. It was unknown if any action was taken due to the nerve damage. 2020-10-26 IUD (Other): No new information. | xse screw deviation left ccr iud spinal procedure treat degenerative disc fixation left screw deviated open tlif surgeon operated right issue proceeded left followed executing surgeon incision large caused excessive lateral pressure surgical arm soft tissue led shoulder unlocking surgeon decided abort use left procedure issue air pressure pneumatic motor boot operation step procedure low pressure error message displayed rebooted pressure restored procedure confirmation images showed left medial mm surgeon believed excessive pressure placed surgical arm left accurate experiencing nerve damage procedure delayed hour iud email additional screw repositioned freehand confirmation images taken screw accurately positioned unknown action taken nerve damage iud new surgical arm | NAV3264 | NAV4098 | NAV1012 | NAV5000 | NAV4095 | NAV3264NAV4098NAV1012NAV5000NAV4095 | True
4231 | 703516238 | X - OOBF/SRA Failed Accuracy Test | 2019-12-12T23:24:40.000000+0000 | MazorX | NAV4093 | INACCURACY DURING ADVANCED SETUP | 20 | ASM0206-01 | SURGICAL ARM | X - OOBF/SRA Failed Accuracy Test 2019-12-12 00680310 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the advanced accuracy test at eight of the 10 points during install. There was no patient involvement. 2019-12-23 IUD x6 (Other): Additional information received from the field service engineer that the arm was replaced on 2019 December 23 and returned for analysis. SURGICAL ARM | 2019-12-12 00680310 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the advanced accuracy test at eight of the 10 points during install. There was no patient involvement. 2019-12-23 IUD x6 (Other): Additional information received from the field service engineer that the arm was replaced on 2019 December 23 and returned for analysis. | x oobfsra failed accuracy test outside procedure surgical arm failed advanced accuracy test points install iud additional field service engineer arm replaced december returned analysis surgical arm | NAV2054 | NAV4093 | NAV3249 | NAV4092 | NAV1006 | NAV2054NAV4093NAV3249NAV4092NAV1006 | True
2492 | 703236591 | X - not locked message | 2019-07-01T18:44:18.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | ASM0214-02 | SSY | X - not locked message 2019-Jun-27 SCL1901868 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the shoulder was physically locked however system gave message shoulder not locked. The hood was closed but system gave a message hood not closed. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. SSY | 2019-Jun-27 SCL1901868 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the shoulder was physically locked however system gave message shoulder not locked. The hood was closed but system gave a message hood not closed. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x locked message spinal procedure shoulder physically locked gave message shoulder locked hood closed gave message hood closed additional complications anticipated ssy | NAV2058 | NAV2054 | NAV4108 | NAV3295 | NAV1012 | NAV2058NAV2054NAV4108NAV3295NAV1012 | True
6572 | 703979446 | XSE - Stripped Screw On the Arm Guide | 2020-09-26T00:03:03.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS1650-03S | KIT SVC MAS1650-03S UNIV SCREWDRIVER | XSE - Stripped Screw On the Arm Guide 2020-Sep-25 Salesforce C00213671 (rep): Medtronic received information regarding a guidance system being outside of a procedure. It was reported that the screw on the arm guide was stripped out. There was no patient involvement. 2020-09-30 IUD (Other): No new information. 2020-10-21 email 2020-10-21 06:15 (Rep): The small screw on the universal screw driver has been an re-occurring problem of coming so mewhat loose or falling out entirely. 2020-11-09 email 2020-11-09 13:45 (Rep): No new information. 2020-12-09 email 2020-12-09 17:30, email 2020-12-09 09:45, PDF, 00702419, IUD (Rep):Additional information received from a manufacturer representative reported that the arm guide was lost. 2021-01-12 email 2021-01-12 15:00 (Rep): Additional information received from a manufacturer representative reported that the screw driver was discarded. 2021-01-19 IUD (Other): No new information. 2021-Feb-09 IUD (other): No new information. KIT SVC MAS1650-03S UNIV SCREWDRIVER | 2020-Sep-25 Salesforce C00213671 (rep): Medtronic received information regarding a guidance system being outside of a procedure. It was reported that the screw on the arm guide was stripped out. There was no patient involvement. 2020-09-30 IUD (Other): No new information. 2020-10-21 email 2020-10-21 06:15 (Rep): The small screw on the universal screw driver has been an re-occurring problem of coming so mewhat loose or falling out entirely. 2020-11-09 email 2020-11-09 13:45 (Rep): No new information. 2020-12-09 email 2020-12-09 17:30, email 2020-12-09 09:45, PDF, 00702419, IUD (Rep):Additional information received from a manufacturer representative reported that the arm guide was lost. 2021-01-12 email 2021-01-12 15:00 (Rep): Additional information received from a manufacturer representative reported that the screw driver was discarded. 2021-01-19 IUD (Other): No new information. 2021-Feb-09 IUD (other): No new information. | xse stripped screw arm guide salesforce outside procedure screw arm guide stripped iud new email small screw universal screw driver reoccurring problem coming mewhat loose falling entirely email new email email pdf iud additional arm guide lost email additional screw driver discarded iud new iud new kit svc univ screwdriver | NAV3299 | NAV3294 | NAV4090 | NAV2054 | NAV3246 | NAV3299NAV3294NAV4090NAV2054NAV3246 | True
7526 | 704221830 | XSE - Unable to Register | 2021-02-09T21:45:29.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Unable to Register 2021-Feb-09 Salesforce C00220760, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat trauma with screws and rods. It was reported that during a procedure on 9-FEB-2021, they were unable to register the final 3 levels for their lumbar case. The rep stated they registered the first 3 levels without issue and were able to place wires. They attempted to register the bottom 3 and it would not register. The Mazor stated there was registration problems and to check labeling, which the rep had done. They decided to switch to scan and plan to continue the case. Unable to confirm why registration would not pass. There was no patient harm and the procedure was delayed less than an hour. 2021-02-18 IUD (Other): No new information. 2021-02-26 email 2021-02-26 14:15 (Rep): Additional information was received stating that troubleshooting involved recalibration of the c-arm, re-imaging, and re-registration. There was no positioning change and no work was done between segments. 2021-03-08 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-Feb-09 Salesforce C00220760, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat trauma with screws and rods. It was reported that during a procedure on 9-FEB-2021, they were unable to register the final 3 levels for their lumbar case. The rep stated they registered the first 3 levels without issue and were able to place wires. They attempted to register the bottom 3 and it would not register. The Mazor stated there was registration problems and to check labeling, which the rep had done. They decided to switch to scan and plan to continue the case. Unable to confirm why registration would not pass. There was no patient harm and the procedure was delayed less than an hour. 2021-02-18 IUD (Other): No new information. 2021-02-26 email 2021-02-26 14:15 (Rep): Additional information was received stating that troubleshooting involved recalibration of the c-arm, re-imaging, and re-registration. There was no positioning change and no work was done between segments. 2021-03-08 IUD (Other): No new information. | xse unable register salesforce ccr spinal procedure treat trauma screws rods procedure unable register final levels lumbar case stated registered levels issue able place wires attempted register register mazor stated registration problems check labeling decided switch scan plan continue case unable confirm registration pass procedure delayed hour iud new email additional stating troubleshooting involved recalibration carm reimaging reregistration positioning change work segments iud new mazor x | NAV4112 | NAV3242 | NAV1012 | NAV3243 | NAV3244 | NAV4112NAV3242NAV1012NAV3243NAV3244 | True
2724 | 703248844 | XSE - Nav inaccurate/Malleted Off Trajec | 2019-07-09T13:17:36.000000+0000 | MazorX Stealth Edition | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Nav inaccurate/Malleted Off Trajec 2019-Jul-03 SCL1901926, CCR, Intake Form (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat degenerative discs using fixation. It was reported the plan was reviewed with the surgeon in the operating room prior to the start of the operation. They paid special attention to the docking surfaces and skin incisions due to the percutaneous approach. The bone mount bridge was chosen based on clinical indication and surgeon preference. The 120 mm Schanz pin was mounted in the right PSIS for stabilization. Accurate registration was achieved on the first attempt using the basic algorithm. The surgeon verified there was no shift from CT to fluoro. The surgeon completed the discectomy, decompression and insertion of the interbody spacer prior to attaching the surgical system to the patient. Navigated instruments were verified prior to attaching the system to the patient. An accurate 3Define and Draw spine were done on the first attempt. The rep sent the surgical arm to the snapshot position and the surgeon attached the snapshot tracker. The rep recalculated and resent snapshot tracker prior to acquiring. Surgeon requested the rep to send to left L4 and left L5 to ensure incision was acceptable prior to drilling. Surgeon made initial incision for left L4 and left L5 trajectories. The rep then sent to Left L4 and had surgeon check navigation accuracy utilizing the passive planar probe and divot on the arm guide. The navigation was slightly off and the rep recommended acquiring a new spapshot prior to drilling the first trajectory, and the surgeon agreed. The rep sent to snapshot, an HCP attached snapshot tracker and the rep recalculated and resent to snapshot before acquiring. Surgeon verified navigation was accurate with "PPB" and Divot at that time. The surgeon inserted screws at Left L4 then Left L5. The rep then sent to Right L4. Surgeon malleted drill guide slightly too hard and knocked the arm off of trajectory. the rep asked surgeon to remove drill guide and cannula in order to resend to Right L4. Once arm was back in position, surgeon inserted screws at Right L4 and Right L5. All trajectories were accurate and the surgeon was pleased with the final construct. Trajectories were verified with in tra-op fluoro and an O-arm scan. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jul-03 SCL1901926, CCR, Intake Form (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat degenerative discs using fixation. It was reported the plan was reviewed with the surgeon in the operating room prior to the start of the operation. They paid special attention to the docking surfaces and skin incisions due to the percutaneous approach. The bone mount bridge was chosen based on clinical indication and surgeon preference. The 120 mm Schanz pin was mounted in the right PSIS for stabilization. Accurate registration was achieved on the first attempt using the basic algorithm. The surgeon verified there was no shift from CT to fluoro. The surgeon completed the discectomy, decompression and insertion of the interbody spacer prior to attaching the surgical system to the patient. Navigated instruments were verified prior to attaching the system to the patient. An accurate 3Define and Draw spine were done on the first attempt. The rep sent the surgical arm to the snapshot position and the surgeon attached the snapshot tracker. The rep recalculated and resent snapshot tracker prior to acquiring. Surgeon requested the rep to send to left L4 and left L5 to ensure incision was acceptable prior to drilling. Surgeon made initial incision for left L4 and left L5 trajectories. The rep then sent to Left L4 and had surgeon check navigation accuracy utilizing the passive planar probe and divot on the arm guide. The navigation was slightly off and the rep recommended acquiring a new spapshot prior to drilling the first trajectory, and the surgeon agreed. The rep sent to snapshot, an HCP attached snapshot tracker and the rep recalculated and resent to snapshot before acquiring. Surgeon verified navigation was accurate with "PPB" and Divot at that time. The surgeon inserted screws at Left L4 then Left L5. The rep then sent to Right L4. Surgeon malleted drill guide slightly too hard and knocked the arm off of trajectory. the rep asked surgeon to remove drill guide and cannula in order to resend to Right L4. Once arm was back in position, surgeon inserted screws at Right L4 and Right L5. All trajectories were accurate and the surgeon was pleased with the final construct. Trajectories were verified with in tra-op fluoro and an O-arm scan. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse nav inaccuratemalleted trajec ccr spinal procedure treat degenerative discs using fixation plan reviewed surgeon operating room prior start operation paid special attention docking surfaces skin incisions percutaneous approach bone mount bridge chosen based clinical indication surgeon preference mm schanz pin mounted right psis stabilization accurate registration achieved attempt using basic algorithm surgeon verified shift ct fluoro surgeon completed discectomy decompression insertion interbody spacer prior attaching surgical navigated instruments verified prior attaching accurate draw spine attempt sent surgical arm snapshot position surgeon attached snapshot tracker recalculated resent snapshot tracker prior acquiring surgeon requested send left left ensure incision acceptable prior drilling surgeon initial incision left left trajectories sent left surgeon check navigation accuracy utilizing passive planar probe divot arm guide navigation slightly recommended acquiring new spapshot prior drilling trajectory surgeon agreed sent snapshot hcp attached snapshot tracker recalculated resent snapshot acquiring surgeon verified navigation accurate ppb divot time surgeon inserted screws left left sent right surgeon malleted drill guide slightly hard knocked arm trajectory asked surgeon remove drill guide cannula order resend right arm position surgeon inserted screws right right trajectories accurate surgeon pleased final construct trajectories verified traop fluoro oarm scan additional complications anticipated mazor x software installation kit | NAV3246 | NAV2056 | NAV4091 | NAV3253 | NAV2051 | NAV3246NAV2056NAV4091NAV3253NAV2051 | False
886 | 703179120 | X- Inaccurate Trajectories/Deviations | 2019-05-29T21:34:44.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059-04 | MAZOR X SYSTEM (REV 4) | X- Inaccurate Trajectories/Deviations 2018-10-15 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that before the system was rolled into the operating room, the manufacturer representative conducted 10 point test along with stress test. System passed both. The Pre-operative plan was reviewed the morning of the procedure. T2-T6 were planned for an open procedure. Screws would be placed at T2, T3, skip T4, potentially only on one side of T5 and at T6. There were no major concerns discussed. The disposable clamp was utilized for the procedure and placed at T4. There were no issues mounting the platform. Registration was achieved on the first attempt with no complications. After registration, there were some issues throughout the procedure. Trajectories were sent and wires were placed at T2 left and right Trajectories were sent and wires were placed at T3 left and right After the first 4 wires, the surgeon stopped to observe the anatomy a little more closely and felt like T2 and T3 on the right were very medial. An AP and lateral x-ray were taken and confirmed both T3 and T2 on the right were medial. The platform was observed and there was no opportunity for soft tissue pressure pushing the tools medial as the tools fell straight down with no contact with the skin. The plan was then reviewed and there was no opportunity to skive medial. The surgeon did not want to adjust the plan and re-send the trajectory and pointed out that the arm was nowhere near where it was planned. The suggestion was then made to re-register and re-send the trajectories. The second registration was achieved on the second attempt after getting new oblique images to clear up the end plates. The surgeon requested that T3 be sent first since any deviation could seen best at that body. T3 left was sent first and the wire was placed back in the original hole. When T3 on the right was re-sent, the surgeon noted that the arm tube was lined up very medial again. T4 on the right was also sent so that the surgeon could compare the trajectory to the one above. The surgeon noted that this was also very medial. The plan was adjusted at T3 right to translate the screw out laterally and re-sent but the arm went back to the same medial trajectory. The surgeon removed the robot and proceeded freehand. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM (REV 4) | 2018-10-15 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that before the system was rolled into the operating room, the manufacturer representative conducted 10 point test along with stress test. System passed both. The Pre-operative plan was reviewed the morning of the procedure. T2-T6 were planned for an open procedure. Screws would be placed at T2, T3, skip T4, potentially only on one side of T5 and at T6. There were no major concerns discussed. The disposable clamp was utilized for the procedure and placed at T4. There were no issues mounting the platform. Registration was achieved on the first attempt with no complications. After registration, there were some issues throughout the procedure. Trajectories were sent and wires were placed at T2 left and right Trajectories were sent and wires were placed at T3 left and right After the first 4 wires, the surgeon stopped to observe the anatomy a little more closely and felt like T2 and T3 on the right were very medial. An AP and lateral x-ray were taken and confirmed both T3 and T2 on the right were medial. The platform was observed and there was no opportunity for soft tissue pressure pushing the tools medial as the tools fell straight down with no contact with the skin. The plan was then reviewed and there was no opportunity to skive medial. The surgeon did not want to adjust the plan and re-send the trajectory and pointed out that the arm was nowhere near where it was planned. The suggestion was then made to re-register and re-send the trajectories. The second registration was achieved on the second attempt after getting new oblique images to clear up the end plates. The surgeon requested that T3 be sent first since any deviation could seen best at that body. T3 left was sent first and the wire was placed back in the original hole. When T3 on the right was re-sent, the surgeon noted that the arm tube was lined up very medial again. T4 on the right was also sent so that the surgeon could compare the trajectory to the one above. The surgeon noted that this was also very medial. The plan was adjusted at T3 right to translate the screw out laterally and re-sent but the arm went back to the same medial trajectory. The surgeon removed the robot and proceeded freehand. No patient harm was reported and surgery was delayed less than an hour. | x inaccurate trajectoriesdeviations ccr spinal procedure rolled operating room conducted point test stress test passed preoperative plan reviewed morning procedure planned open procedure screws placed skip potentially major concerns discussed disposable clamp utilized procedure placed issues mounting platform registration achieved attempt complications registration issues procedure trajectories sent wires placed left right trajectories sent wires placed left right wires surgeon stopped observe anatomy little closely felt like right medial ap lateral xray taken confirmed right medial platform observed opportunity soft tissue pressure pushing tools medial tools fell straight contact skin plan reviewed opportunity skive medial surgeon did want adjust plan resend trajectory pointed arm near planned suggestion reregister resend trajectories second registration achieved second attempt getting new oblique images clear end plates surgeon requested sent deviation seen best body left sent wire placed original hole right resent surgeon noted arm tube lined medial right sent surgeon compare trajectory surgeon noted medial plan adjusted right translate screw laterally resent arm went medial trajectory surgeon removed robot proceeded freehand surgery delayed hour mazor x | NAV3243 | NAV4091 | NAV5000 | NAV3242 | NAV3241 | NAV3243NAV4091NAV5000NAV3242NAV3241 | True
170 | 703023824 | X - Hair in Sterile Drape | 2019-02-24T21:26:48.000000+0000 | MazorX | NAV3232 | DRAPE ISSUES | 20 | MAS2696-01 | MAZOR X SS STERILE SLEEVE | X - Hair in Sterile Drape 2019-Feb-15 SCL1900411 intake form, (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Steril Drape had a Hair inside the peel pack. Drape was opened to be passed off and a piece of hair was noticed. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SS STERILE SLEEVE | 2019-Feb-15 SCL1900411 intake form, (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Steril Drape had a Hair inside the peel pack. Drape was opened to be passed off and a piece of hair was noticed. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x hair sterile drape spinal procedure steril drape hair inside peel pack drape opened passed piece hair noticed noted additional complications anticipated mazor x ss sterile sleeve | NAV3232 | NAV2054 | NAV2046 | NAV5000 | NAV1006 | NAV3232NAV2054NAV2046NAV5000NAV1006 | True
7395 | 704195249 | XSE - E-Stop Issues | 2021-01-26T17:35:37.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - E-Stop Issues 2021-01-26 00705165 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while booting the system up, the system was showing that the arm e-stop was on. The manufacturer representative reported that the arm e-stop was red but the e-stop button was not pressed down. Doing a soft restart of the system showed "Gains lost on driver 1, maintenance required." No patient was present. 2021-Feb-05 IUDx2 (other): No new information. 2021-03-04 IUD (Other): No new information. 2021-04-06 email 2021-04-06 09:31, email 2021-04-06 10:31, email 2021-04-06 11:16 (Rep): Additional information received from a manu facturer representative reported that the E-stop button was not stuck. The button was not pressed down and the issue was determined to be a software defect. KIT SVC ASM0206-01S SURGICAL ARM | 2021-01-26 00705165 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while booting the system up, the system was showing that the arm e-stop was on. The manufacturer representative reported that the arm e-stop was red but the e-stop button was not pressed down. Doing a soft restart of the system showed "Gains lost on driver 1, maintenance required." No patient was present. 2021-Feb-05 IUDx2 (other): No new information. 2021-03-04 IUD (Other): No new information. 2021-04-06 email 2021-04-06 09:31, email 2021-04-06 10:31, email 2021-04-06 11:16 (Rep): Additional information received from a manu facturer representative reported that the E-stop button was not stuck. The button was not pressed down and the issue was determined to be a software defect. | xse estop issues outside procedure booting showing arm estop arm estop red estop button pressed doing soft restart showed gains lost driver maintenance required present new iud new email email email additional manu facturer estop button stuck button pressed issue determined software defect kit svc surgical arm | NAV2047 | NAV4095 | NAV3252 | NAV2053 | NAV3246 | NAV2047NAV4095NAV3252NAV2053NAV3246 | True
10035 | 704778253 | XSE ¦ Side Cover Rubber Grommets Worn | 2021-12-22T16:35:17.000000+0000 | MazorX Stealth Edition | NAV2057 | SURGICAL SYSTEM ISSUES | 20 | MAS2105 | KIT SVC MAS2105 RUBBER NIP ASSY 03S NEW | XSE ¦ Side Cover Rubber Grommets Worn 2021-Dec-21 Salesforce C00240079 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the rubber grommets for the side covers are worn. Discovered during PM with no patient present. 2022-01-05 IUD x2 (Other): The issue was likely due to service related activities. 2022-01-14 IUD x2 (Other): No new information. 2022-02-02 IUD (Other): No new information. KIT SVC MAS2105 RUBBER NIP ASSY 03S NEW | 2021-Dec-21 Salesforce C00240079 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the rubber grommets for the side covers are worn. Discovered during PM with no patient present. 2022-01-05 IUD x2 (Other): The issue was likely due to service related activities. 2022-01-14 IUD x2 (Other): No new information. 2022-02-02 IUD (Other): No new information. | xse cover rubber grommets worn salesforce outside procedure rubber grommets covers worn discovered pm present iud issue likely service related activities iud new iud new kit svc rubber nip assy new | NAV2058 | NAV3299 | NAV3264 | NAV3249 | NAV4099 | NAV2058NAV3299NAV3264NAV3249NAV4099 | False
7409 | 704197845 | XSE - Inaccurate Tracking | 2021-01-27T17:37:06.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Inaccurate Tracking 2021-01-27 00705267, 00705265, 00705268, IUD x11 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the Midas drill was showing lower when navigating than the actual physical location. The surgeon was able to place all screws accurately with the navigation being off. Troubleshooting involved trying long and short attachments, changing the camera angle, and taking a new snapshot. During the procedure, the surgeon unscrewed the screw from the driver and the surgical arm shoulder shifted. The surgeon decided to abort the use of the guidance system and complete the procedure freehand. The manufacturer representative further noted that the manipulator was hard to move and the hood slammed shut when closing on the workstation. There was no patient harm and the procedure was delayed less than an hour. 2021-02-22 email 2021-02-22 09:16, email 2021-02-22 08:16 (Rep): No new information. 2021-03-14 email 2021-03-14 17:15 (Rep): Additional information received from a manufacturer representative reported that the issue was resolved. UNK MAZOR X WORKSTATION | 2021-01-27 00705267, 00705265, 00705268, IUD x11 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the Midas drill was showing lower when navigating than the actual physical location. The surgeon was able to place all screws accurately with the navigation being off. Troubleshooting involved trying long and short attachments, changing the camera angle, and taking a new snapshot. During the procedure, the surgeon unscrewed the screw from the driver and the surgical arm shoulder shifted. The surgeon decided to abort the use of the guidance system and complete the procedure freehand. The manufacturer representative further noted that the manipulator was hard to move and the hood slammed shut when closing on the workstation. There was no patient harm and the procedure was delayed less than an hour. 2021-02-22 email 2021-02-22 09:16, email 2021-02-22 08:16 (Rep): No new information. 2021-03-14 email 2021-03-14 17:15 (Rep): Additional information received from a manufacturer representative reported that the issue was resolved. | xse inaccurate tracking iud spinal procedure midas drill showing lower navigating actual physical location surgeon able place screws accurately navigation troubleshooting involved trying long short attachments changing camera angle taking new snapshot procedure surgeon unscrewed screw driver surgical arm shoulder shifted surgeon decided abort use complete procedure freehand noted manipulator hard hood slammed shut closing workstation procedure delayed hour email email new email additional issue resolved unk mazor x workstation | NAV2058 | NAV1012 | NAV4108 | NAV3260 | NAV3246 | NAV2058NAV1012NAV4108NAV3260NAV3246 | False
311 | 703076602 | X - Failed Registration | 2019-03-27T17:39:07.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - Failed Registration 2019-02-25 SCL1900732, Intake Form, CCR, e1, e2, e3, e4 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure for deformity being treated with fixation. It was reported that there were registration difficulties during the procedure. The surgeon planned the L2-L4 XLIF, L4-S1 ALIF, and L2-S2 PSF the day prior to the case. The surgeon wanted to use navigation as a verification to the guidance system planning. The procedure was setup to use Scan and Plan with navigation for the upper segment and CT to fluoro for the lower segment. Both of the spins were planned to be taken at the same time. The first spin would be for the lower section, S2-L5, and the second spin for the upper section, L5-L2. During the procedure, the LPSIS was used for the navigation reference frame and the RPSIS was used for the Schanz arm. The 3Define scan, draw spine, and spins were completed. The surgical system arm was sent out to tap the first trajectory and navigation showed to be in disc space. Navigation verification was attempted off of the spinous process, but was not accurate. The manufacturer representative thought it was possibly the reference frame had been moved. The surgeon was unsure which system was accurate so fluoro images were taken. Several fluoro attempts were made due to rotation in the patient's anatomy, but the images verified the guidance system was accurate. Surgeon proceed with the case using the guidance system. For the next trajectory, an anti-skive tool was used and the reduction tube was placed. Fluoro verification in the lateral view showed the reduction tube was 20 mm anterior of vertebral body. A prior AP image showed the wires to be placed accurately. The surgeon proceeded with tapping wires and placing screws for the remaining trajectories L2 to L5. For the next segment using CT-Fluoro, the pre-operative CT scans were loaded and stacked with the O-arm spins. The 3Define scan and draw spine were completed followed by taking AP/oblique images. A titanium cage at L5-S1 and screws placed in L2-L5 were not in the pre-op CT scans. Manual registration of S1 and S2 was attempted, but registration failed. New lateral images were taken and a shift was noticed in review of the registration failure. The manufacturer representative suggested calibration of the C-arm, but the surgeon decided to freehand the S1 trajectories. The S2 trajectories were not placed. There was no patient harm, but the procedure was delayed more than an hour due to the O-arm registration and CT-Fluoro registration issues. Information unrelated to this event was included in a separate PE; See PE 703031554 for Navigation Inaccuracies. 2019-04-01 e5, e6 (Rep): No new information. 2019-04-08 e7 (Rep): Additional information received from a manufacturer representative reported the procedure was delayed more than an hour, but there was no harm to the patient. All of the screws placed during the procedure were accurate. It was unknown if another procedure would be done to place the S2 screws. The trajectory had not deviated based on the the fluoro, but the reduction tube on L3 had passed approximately 10 mm anterior. The vertebra did not have a shift. A shift was seen on the system software when attempting registration. The second segment was instrumented free hand. MAZOR X SOFTWARE KIT VER-4.0.0 | 2019-02-25 SCL1900732, Intake Form, CCR, e1, e2, e3, e4 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure for deformity being treated with fixation. It was reported that there were registration difficulties during the procedure. The surgeon planned the L2-L4 XLIF, L4-S1 ALIF, and L2-S2 PSF the day prior to the case. The surgeon wanted to use navigation as a verification to the guidance system planning. The procedure was setup to use Scan and Plan with navigation for the upper segment and CT to fluoro for the lower segment. Both of the spins were planned to be taken at the same time. The first spin would be for the lower section, S2-L5, and the second spin for the upper section, L5-L2. During the procedure, the LPSIS was used for the navigation reference frame and the RPSIS was used for the Schanz arm. The 3Define scan, draw spine, and spins were completed. The surgical system arm was sent out to tap the first trajectory and navigation showed to be in disc space. Navigation verification was attempted off of the spinous process, but was not accurate. The manufacturer representative thought it was possibly the reference frame had been moved. The surgeon was unsure which system was accurate so fluoro images were taken. Several fluoro attempts were made due to rotation in the patient's anatomy, but the images verified the guidance system was accurate. Surgeon proceed with the case using the guidance system. For the next trajectory, an anti-skive tool was used and the reduction tube was placed. Fluoro verification in the lateral view showed the reduction tube was 20 mm anterior of vertebral body. A prior AP image showed the wires to be placed accurately. The surgeon proceeded with tapping wires and placing screws for the remaining trajectories L2 to L5. For the next segment using CT-Fluoro, the pre-operative CT scans were loaded and stacked with the O-arm spins. The 3Define scan and draw spine were completed followed by taking AP/oblique images. A titanium cage at L5-S1 and screws placed in L2-L5 were not in the pre-op CT scans. Manual registration of S1 and S2 was attempted, but registration failed. New lateral images were taken and a shift was noticed in review of the registration failure. The manufacturer representative suggested calibration of the C-arm, but the surgeon decided to freehand the S1 trajectories. The S2 trajectories were not placed. There was no patient harm, but the procedure was delayed more than an hour due to the O-arm registration and CT-Fluoro registration issues. Information unrelated to this event was included in a separate PE; See PE 703031554 for Navigation Inaccuracies. 2019-04-01 e5, e6 (Rep): No new information. 2019-04-08 e7 (Rep): Additional information received from a manufacturer representative reported the procedure was delayed more than an hour, but there was no harm to the patient. All of the screws placed during the procedure were accurate. It was unknown if another procedure would be done to place the S2 screws. The trajectory had not deviated based on the the fluoro, but the reduction tube on L3 had passed approximately 10 mm anterior. The vertebra did not have a shift. A shift was seen on the system software when attempting registration. The second segment was instrumented free hand. | x failed registration ccr spinal procedure deformity treated fixation registration difficulties procedure surgeon planned xlif alif psf day prior case surgeon wanted use navigation verification planning procedure setup use scan plan navigation upper segment ct fluoro lower segment spins planned taken time spin lower section second spin upper section procedure lpsis navigation reference frame rpsis schanz arm scan draw spine spins completed surgical arm sent tap trajectory navigation showed disc space navigation verification attempted spinous process accurate thought possibly reference frame moved surgeon unsure accurate fluoro images taken fluoro attempts rotation patients anatomy images verified accurate surgeon proceed case using trajectory antiskive tool reduction tube placed fluoro verification lateral view showed reduction tube mm anterior vertebral body prior ap image showed wires placed accurately surgeon proceeded tapping wires placing screws remaining trajectories segment using ctfluoro preoperative ct scans loaded stacked oarm spins scan draw spine completed followed taking apoblique images titanium cage screws placed preop ct scans manual registration attempted registration failed new lateral images taken shift noticed review registration failure suggested calibration carm surgeon decided freehand trajectories trajectories placed procedure delayed hour oarm registration ctfluoro registration issues unrelated included separate pe pe navigation inaccuracies new additional procedure delayed hour screws placed procedure accurate unknown procedure place screws trajectory deviated based fluoro reduction tube passed approximately mm anterior vertebra did shift shift seen software attempting registration second segment instrumented free hand mazor x software kit | NAV4112 | NAV5001 | NAV5000 | NAV4091 | NAV2051 | NAV4112NAV5001NAV5000NAV4091NAV2051 | True
10345 | 704858886 | XSE - Alleged Inaccuracy | 2022-02-11T01:11:04.000000+0000 | MazorX Stealth Edition | NAV1012 | SYSTEM REPORTED | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Alleged Inaccuracy 2022-Feb-10 Salesforce C00243532, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were deviations during a one level L4/L5 case. The surgeon used a Scan & Plan approach for the procedure. The surgical arm was sent to the trajectories on the left side and the system was showing off medially and superiorly. The accuracy pointer was checked on the divot and it was good. The surgeon decided to place the screw and move to the right side. On the right side, the trajectories were off laterally and superiorly. A new snapshot was taken and the divot was checked. The surgeon moved ahead with placing screws on the right side. C-arm and post-op images were taken and all screws were within the pedicle, but off superiorly. There was no patient harm and the procedure was delayed less than an hour. 2022-03-04 email 2022-03-04 12:15, email 2022-03-07 12:01, image 2022-03-07 12:01, IUD x2 (Rep): Additional information received from a manufacturer representative reported the snapshot tracker and reference frame were not defective and passed testing after the surgical arm was replaced. 2022-03-17 e1 (Other): No new information. 2022-03-24 e2, e3, e4 , IUD (Rep): Additional information received from a manufacturer representative reported that the stealth integration test was used to verify the snapshot tracker and reference frame were accurate. 2022-03-28 e5, e6 (Rep): Additional information received from a manufacturer representative reported that a single Schanz pin placed in the patient's left side and a Schanz arm was used to mount the surgical system to the patient. MAZOR X SYSTEM TPL0059 | 2022-Feb-10 Salesforce C00243532, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were deviations during a one level L4/L5 case. The surgeon used a Scan & Plan approach for the procedure. The surgical arm was sent to the trajectories on the left side and the system was showing off medially and superiorly. The accuracy pointer was checked on the divot and it was good. The surgeon decided to place the screw and move to the right side. On the right side, the trajectories were off laterally and superiorly. A new snapshot was taken and the divot was checked. The surgeon moved ahead with placing screws on the right side. C-arm and post-op images were taken and all screws were within the pedicle, but off superiorly. There was no patient harm and the procedure was delayed less than an hour. 2022-03-04 email 2022-03-04 12:15, email 2022-03-07 12:01, image 2022-03-07 12:01, IUD x2 (Rep): Additional information received from a manufacturer representative reported the snapshot tracker and reference frame were not defective and passed testing after the surgical arm was replaced. 2022-03-17 e1 (Other): No new information. 2022-03-24 e2, e3, e4 , IUD (Rep): Additional information received from a manufacturer representative reported that the stealth integration test was used to verify the snapshot tracker and reference frame were accurate. 2022-03-28 e5, e6 (Rep): Additional information received from a manufacturer representative reported that a single Schanz pin placed in the patient's left side and a Schanz arm was used to mount the surgical system to the patient. | xse alleged inaccuracy salesforce iud spinal procedure deviations level case surgeon scan plan approach procedure surgical arm sent trajectories left showing medially superiorly accuracy pointer checked divot good surgeon decided place screw right right trajectories laterally superiorly new snapshot taken divot checked surgeon moved ahead placing screws right carm postop images taken screws pedicle superiorly procedure delayed hour email email image iud additional snapshot tracker reference frame defective passed testing surgical arm replaced new iud additional stealth integration test verify snapshot tracker reference frame accurate additional single schanz pin placed patients left schanz arm mount surgical mazor x | NAV5000 | NAV1012 | NAV3246 | NAV3243 | NAV3299 | NAV5000NAV1012NAV3246NAV3243NAV3299 | True
7661 | 704245369 | XSE - Multiple errors | 2021-02-24T01:07:29.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 40 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Multiple errors 2021-02-23 00707147 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that during a surgeon demo they received multiple errors - CANBUS errors on nodes 1-8, "Incorrect driver firmware identified" and "Root calibration not found on driver." No patient was present. 2021-02-23 IUD x3, 00707694 (Rep): Additional information received from a manufacturer representative reported that the controller kept losing connection and there was a loud whine and hissing sound coming from it. The representative noted that the surgical system became damaged when being removed from the bed. 2021-03-04 IUD x3 (Other): No new information. 2021-03-09 x10 (Other): No new information. 2021-04-06 IUD (Other): No new information. 2021-05-10 IUD x2 (Other): No new information. 2021-05-18 IUD x2 (Other): No new information. 2021-05-26 IUD (Other): No new information. 2021-05-31 IUD (Other): No new information. 2021-06-06 IUD (Other): No new information. 2021-07-13 IUD (Other): No new information. 2021-08-04 IUD (Other): No new information. 2022-03-23 IUD (Other): No new information. 2022-03-24 IUD x2 (Other): No new information. 2022-04-03 IUD (Other): No new information. 2022-04-10 IUD x2 (Other): No new information. 2022-05-12 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-02-23 00707147 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that during a surgeon demo they received multiple errors - CANBUS errors on nodes 1-8, "Incorrect driver firmware identified" and "Root calibration not found on driver." No patient was present. 2021-02-23 IUD x3, 00707694 (Rep): Additional information received from a manufacturer representative reported that the controller kept losing connection and there was a loud whine and hissing sound coming from it. The representative noted that the surgical system became damaged when being removed from the bed. 2021-03-04 IUD x3 (Other): No new information. 2021-03-09 x10 (Other): No new information. 2021-04-06 IUD (Other): No new information. 2021-05-10 IUD x2 (Other): No new information. 2021-05-18 IUD x2 (Other): No new information. 2021-05-26 IUD (Other): No new information. 2021-05-31 IUD (Other): No new information. 2021-06-06 IUD (Other): No new information. 2021-07-13 IUD (Other): No new information. 2021-08-04 IUD (Other): No new information. 2022-03-23 IUD (Other): No new information. 2022-03-24 IUD x2 (Other): No new information. 2022-04-03 IUD (Other): No new information. 2022-04-10 IUD x2 (Other): No new information. 2022-05-12 IUD (Other): No new information. | xse multiple errors outside procedure surgeon demo multiple errors canbus errors nodes incorrect driver firmware identified root calibration driver present iud additional controller kept losing connection loud whine hissing sound coming noted surgical damaged removed bed iud new new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new kit svc cart elec panel assy | NAV2047 | NAV2058 | NAV2054 | NAV2057 | NAV1006 | NAV2047NAV2058NAV2054NAV2057NAV1006 | True
4746 | 703573579 | XSE - CD Drive Not Working | 2020-01-21T20:05:44.000000+0000 | MazorX Stealth Edition | NAV3257 | CD DRIVE ISSUES | 30 | DSD0038-01 | DVD, PORTABLE DVD REWRITER, GP70NS50 | XSE - CD Drive Not Working 2020-01-21 00682433 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that the CD drive was not functioning. It will load the disk and spin, but it will not read or show anything in the file system. This occurred outside of surgery with no patient present. There was no patient involvement. 2020-01-21 IUD x4 (Other): Additional information received from a field service engineer reporting that the system was checked on-site and the complaint could not be confirmed. The issue could not be reproduced. The drive was working correctly and the HUB USB cable was also plugged into the appropriate port. The system was operational. 2020-02-26 IUD (Other): No new information. 2020-05-11 IUD (Other): No new information. DVD, PORTABLE DVD REWRITER, GP70NS50 | 2020-01-21 00682433 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that the CD drive was not functioning. It will load the disk and spin, but it will not read or show anything in the file system. This occurred outside of surgery with no patient present. There was no patient involvement. 2020-01-21 IUD x4 (Other): Additional information received from a field service engineer reporting that the system was checked on-site and the complaint could not be confirmed. The issue could not be reproduced. The drive was working correctly and the HUB USB cable was also plugged into the appropriate port. The system was operational. 2020-02-26 IUD (Other): No new information. 2020-05-11 IUD (Other): No new information. | xse cd drive working use cd drive functioning load disk spin read file occurred outside surgery present iud additional field service engineer reporting checked onsite complaint confirmed issue reproduced drive working correctly hub usb cable plugged appropriate port operational iud new iud new dvd portable dvd rewriter | NAV3257 | NAV3249 | NAV3253 | NAV1006 | NAV4092 | NAV3257NAV3249NAV3253NAV1006NAV4092 | True
6721 | 704012874 | XSE - CANBUS Error | 2020-10-14T20:14:47.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CANBUS Error 2020-10-14 00698773 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a 295 Fatal Error CANBUS connection lost error message was displayed. A gains lost on driver error was also displayed. A hard reboot of the system was done and the system functioned as intended. UNK MAZOR X WORKSTATION | 2020-10-14 00698773 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a 295 Fatal Error CANBUS connection lost error message was displayed. A gains lost on driver error was also displayed. A hard reboot of the system was done and the system functioned as intended. | xse canbus error outside procedure fatal error canbus connection lost error message displayed gains lost driver error displayed hard reboot functioned intended unk mazor x workstation | NAV2047 | NAV4095 | NAV3261 | NAV3262 | NAV2054 | NAV2047NAV4095NAV3261NAV3262NAV2054 | True
6637 | 703995213 | XSE - 3Define Unsuccessful | 2020-10-06T00:09:05.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 120 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - 3Define Unsuccessful 2020-09-30 00698147 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were issues getting the 3Define camera to work during the procedure. There was a green light on the camera, but a 3Define scan was unable to be taken. There was no patient harm and the procedure was delayed less than hour. 2020-10-06 IUD x7 (Other): No new information. 2020-10-20 IUD x11 (Other): No new information. 2020-01-05 IUD (Other): No new information. 2021-06-08 IUD (Other): No new information. 2021-06-09 IUDx3 (other): No new information. 2021-07-15 IUD (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2020-09-30 00698147 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were issues getting the 3Define camera to work during the procedure. There was a green light on the camera, but a 3Define scan was unable to be taken. There was no patient harm and the procedure was delayed less than hour. 2020-10-06 IUD x7 (Other): No new information. 2020-10-20 IUD x11 (Other): No new information. 2020-01-05 IUD (Other): No new information. 2021-06-08 IUD (Other): No new information. 2021-06-09 IUDx3 (other): No new information. 2021-07-15 IUD (Other): No new information. | xse unsuccessful spinal procedure issues getting camera work procedure green light camera scan unable taken procedure delayed hour iud new iud new iud new iud new new iud new mazor x pc station | NAV3244 | NAV3262 | NAV4112 | NAV1006 | NAV3246 | NAV3244NAV3262NAV4112NAV1006NAV3246 | True
1265 | 703194984 | X - Manipulator/Positioner issues | 2019-06-07T12:35:10.000000+0000 | MazorX | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 30 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X - Manipulator/Positioner issues It was reported that there were difficulties mounting the robot to the bed. once mounted and locked to the bed adapter, the air pressure was then locked but there was an issue removing the positioner. There were multiple attempts to remove the positioner by raising and lowering the bed, moving the work station further and closer to the bed and rotating the bed away and towards me to see if that would help. There were also 2-3 attempts of dismounting and remount the robot to the bed. This was later the solution to remove the positioner. registration was difficult as all of the bodies could not be registered. L4 and L5 were registered instantly but S1 would not. in our images, there seemed to be something completely blocking out S1. new images were acquired after manipulating the c-arm with wag and decreasing our angle to see if we can get a clearer image at S1 but nothing helped. The surgeon did not want to take any more images or take the time to locate what was blocking S1 under the table and decided to free hand S1. UNK MAZOR X POSITIONER | It was reported that there were difficulties mounting the robot to the bed. once mounted and locked to the bed adapter, the air pressure was then locked but there was an issue removing the positioner. There were multiple attempts to remove the positioner by raising and lowering the bed, moving the work station further and closer to the bed and rotating the bed away and towards me to see if that would help. There were also 2-3 attempts of dismounting and remount the robot to the bed. This was later the solution to remove the positioner. registration was difficult as all of the bodies could not be registered. L4 and L5 were registered instantly but S1 would not. in our images, there seemed to be something completely blocking out S1. new images were acquired after manipulating the c-arm with wag and decreasing our angle to see if we can get a clearer image at S1 but nothing helped. The surgeon did not want to take any more images or take the time to locate what was blocking S1 under the table and decided to free hand S1. | x manipulatorpositioner issues difficulties mounting robot bed mounted locked bed adapter air pressure locked issue removing positioner multiple attempts remove positioner raising lowering bed moving work station closer bed rotating bed away help attempts dismounting remount robot bed later solution remove positioner registration difficult bodies registered registered instantly images completely blocking new images acquired manipulating carm wag decreasing angle clearer image helped surgeon did want images time locate blocking table decided free hand unk mazor x positioner | NAV3250 | NAV4103 | NAV3264 | NAV4105 | NAV3242 | NAV3250NAV4103NAV3264NAV4105NAV3242 | True
8320 | 704396387 | XSE - Schanz Arm Loose | 2021-05-14T23:52:15.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 50 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Schanz Arm Loose 2021-May-14 Salesforce C00226497 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there were four schanz arms loose and a schanz bond screw was stripped. There was no impact on the patient outcome. The procedure was not delayed. 2021-05-26 IUD x6 (Other): No new information. 2021-06-28 email 2021-06-28 14:01 (Rep): Troubleshooting involved taking the Shanz arms apart and reassembling them. The surgeon observed rust inside the connection point and believed that the age of the arms had caused them to loosen. The shanz bond screw was stripped due to use. The cause of loose Schanz arms or stripped screw was due to being used many times and tightening it repeatedly. KIT SVC MAS2351-04S SCHANZ ARM | 2021-May-14 Salesforce C00226497 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there were four schanz arms loose and a schanz bond screw was stripped. There was no impact on the patient outcome. The procedure was not delayed. 2021-05-26 IUD x6 (Other): No new information. 2021-06-28 email 2021-06-28 14:01 (Rep): Troubleshooting involved taking the Shanz arms apart and reassembling them. The surgeon observed rust inside the connection point and believed that the age of the arms had caused them to loosen. The shanz bond screw was stripped due to use. The cause of loose Schanz arms or stripped screw was due to being used many times and tightening it repeatedly. | xse schanz arm loose salesforce spinal p rocedure schanz arms loose schanz bond screw stripped impact outcome procedure delayed iud new email troubleshooting involved taking shanz arms apart reassembling surgeon observed rust inside connection point believed age arms caused loosen shanz bond screw stripped use cause loose schanz arms stripped screw times tightening repeatedly kit svc schanz arm | NAV3299 | NAV3296 | NAV4109 | NAV3298 | NAV4097 | NAV3299NAV3296NAV4109NAV3298NAV4097 | True
2860 | 703256441 | X - Software Issue / X eye Issue | 2019-07-12T17:03:35.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Software Issue / X eye Issue System continues to crash with bovie use, continued X eye communication errors with bovie use. XCSR met with hospital Biomed to get grounding cord approved, due to short cord XCSR had to plug secondary ground into an outlet on same circuit RBT was plugged into. This continues to be an issue at MidState. XCSR also noticed that system takes longer to change between screens while "processing" than other systems.I would like someone to come and check the system out. Also any info on secondary grounding would be helpful. UNK MAZOR X X-EYE CAMERA | System continues to crash with bovie use, continued X eye communication errors with bovie use. XCSR met with hospital Biomed to get grounding cord approved, due to short cord XCSR had to plug secondary ground into an outlet on same circuit RBT was plugged into. This continues to be an issue at MidState. XCSR also noticed that system takes longer to change between screens while "processing" than other systems.I would like someone to come and check the system out. Also any info on secondary grounding would be helpful. | x software issue x eye issue continues crash bovie use continued x eye communication errors bovie use xcsr met hospital biomed grounding cord approved short cord xcsr plug secondary ground outlet circuit rbt plugged continues issue midstate xcsr noticed takes longer change screens processing systemsi like come check info secondary grounding helpful unk mazor x xeye camera | NAV4110 | NAV3256 | NAV3259 | NAV3244 | NAV3248 | NAV4110NAV3256NAV3259NAV3244NAV3248 | False
4721 | 703569598 | X-Registration issues/X-eye not tracking | 2020-01-20T23:56:53.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X-Registration issues/X-eye not tracking 2018-03-06 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that surgeon, fellow and representative planned case prior to procedure and approved implant placement and trajectory plan. Single Arm Schanz pin was agreed to be placed in left PSIS for the MIS procedure. Cage was placed in the lateral position at L3 - L4. Surgeon opted to add screws to L1 - L2 and L5 - S1 for stability. System was mounted and placed onto bed without any issue. Arm was draped and pin placed in Left PSIS. The guidance system was mounted to patient without a problem. While draping, scrubbed tech wrapped red rubber bands for X Eye cameras around the neck of the camera, causing inability for the camera to move properly. 3Define was executed properly as planned. X Eyes were unable to consistently pick up the reference marker and stylus. Representative opted to move forward to generic AP and Oblique location for 3D Marker acquisition. Significant manual adjustments were needed in order to place 3D marker in proper location. L1 - S1 was unable to be achieved in the imaging due to larger patient. Surgeon requested to register twice to achieve L5 - S1. First before registering L1 - L2, Surgeon also requested to take the images for the upper portion of the registration ahead of time and save in case we needed them later to which the representative explained it was not possible regardless. L5 - S1 screws were placed with accuracy bilaterally and checked with neuromonitoring successfully. Surgeon requested to bring up the L1 - L2 images to try for registration. Representative explained why it was not possible and proceeded to accept the images to walk the surgeon through the process of why it will not work. (After accepting saved images and moving to segmentation) When placing reference marker in saved oblique Image, the cross-view line was vertical to the vertebral bodies instead of horizontal as normal. Representative explained the cross- view problems and why registration could not be achieved. Surgeon agreed to take additional AP and OB image to capture L1 - L2. After achieving RBCs, representative placed reference marker in oblique image to achieve registration and again saw the cross-view lines running vertical to the spine as opposed to horizontal across a vertebral body. Surgeon opted to freehand the top screws due to lengthy duration of the procedure. No patient harm was reported. MAZOR X SYSTEM TPL0059 | 2018-03-06 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that surgeon, fellow and representative planned case prior to procedure and approved implant placement and trajectory plan. Single Arm Schanz pin was agreed to be placed in left PSIS for the MIS procedure. Cage was placed in the lateral position at L3 - L4. Surgeon opted to add screws to L1 - L2 and L5 - S1 for stability. System was mounted and placed onto bed without any issue. Arm was draped and pin placed in Left PSIS. The guidance system was mounted to patient without a problem. While draping, scrubbed tech wrapped red rubber bands for X Eye cameras around the neck of the camera, causing inability for the camera to move properly. 3Define was executed properly as planned. X Eyes were unable to consistently pick up the reference marker and stylus. Representative opted to move forward to generic AP and Oblique location for 3D Marker acquisition. Significant manual adjustments were needed in order to place 3D marker in proper location. L1 - S1 was unable to be achieved in the imaging due to larger patient. Surgeon requested to register twice to achieve L5 - S1. First before registering L1 - L2, Surgeon also requested to take the images for the upper portion of the registration ahead of time and save in case we needed them later to which the representative explained it was not possible regardless. L5 - S1 screws were placed with accuracy bilaterally and checked with neuromonitoring successfully. Surgeon requested to bring up the L1 - L2 images to try for registration. Representative explained why it was not possible and proceeded to accept the images to walk the surgeon through the process of why it will not work. (After accepting saved images and moving to segmentation) When placing reference marker in saved oblique Image, the cross-view line was vertical to the vertebral bodies instead of horizontal as normal. Representative explained the cross- view problems and why registration could not be achieved. Surgeon agreed to take additional AP and OB image to capture L1 - L2. After achieving RBCs, representative placed reference marker in oblique image to achieve registration and again saw the cross-view lines running vertical to the spine as opposed to horizontal across a vertebral body. Surgeon opted to freehand the top screws due to lengthy duration of the procedure. No patient harm was reported. | xregistration issuesxeye tracking ccr spinal procedure surgeon fellow planned case prior procedure approved implant placement trajectory plan single arm schanz pin agreed placed left psis mis procedure cage placed lateral position surgeon opted add screws stability mounted placed bed issue arm draped pin placed left psis mounted problem draping scrubbed tech wrapped red rubber bands x eye cameras neck camera causing inability camera properly executed properly planned x eyes unable consistently pick reference marker stylus opted forward generic ap oblique location marker acquisition significant manual adjustments needed order place marker proper location unable achieved imaging larger surgeon requested register twice achieve registering surgeon requested images upper portion registration ahead time save case needed later explained possible regardless screws placed accuracy bilaterally checked neuromonitoring successfully surgeon requested bring images try registration explained possible proceeded accept images walk surgeon process work placing reference marker saved oblique image crossview line vertical vertebral bodies instead horizontal normal explained cross view problems registration achieved surgeon agreed additional ap ob image capture achieving rbcs placed reference marker oblique image achieve registration saw crossview lines running vertical spine opposed horizontal vertebral body surgeon opted freehand screws lengthy duration procedure mazor x | NAV4110 | NAV4112 | NAV3233 | NAV3244 | NAV2051 | NAV4110NAV4112NAV3233NAV3244NAV2051 | False
345 | 703078970 | X - Registration Difficulties/Inaccuracy | 2019-03-28T18:13:18.000000+0000 | MazorX | NAV4110 | NO TRACKING | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Registration Difficulties/Inaccuracy 2019-02-18 CCR, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that there we difficulties with registration and the L4 and L5 trajectories were medial. The surgeon planned an open revision for L3-S2 using two segments on 2019-02-17. Setup for the procedure was completed with no issues. During the procedure, the X-Eye would not pick up the reference marker. The manufacturer representative checked the USB cable, but the issue was not resolved. A default AP and oblique image were used. Registration was successful and L3-L5 was instrumented. New AP and oblique images were taken for the second segment. Registration was attempted, but could not be completed using the automatic, semi-automatics and manual registration processes. New images were taken, but automatic and semi-automatic registration failed. The manufacturer representative was able to complete registration using the manual process, but the values were not good. The surgeon had decided to proceed with the case and had free handed S1 on the left side. After additional attempts, the manufacturer representative was able to complete manual registration with good values so the surgeon was able to instrument screw number 9 and 10 with the guidance system. The surgeon thought the trajectory for L4 right was to medial so they did not drill. The manufacturer representative had gone into the plan and moved the trajectory lateral. The surgeon did not drill L5 left for the same reason. The case was able to be completed. There was no patient harm reported. 2019-04-29 e2 (Rep): Additional information received from a manufacturer representative reported the X-eye camera was functioning properly prior to the procedure. The representative clarified that the surgeon did not like the trajectory in the open case since they thought it looked either medial or lateral, but there was noting wrong with the trajectory. The surgeon thought the trajectory looked medial and decided not to instrument. The representative confirmed the USB cord was plugged in. The issue was determined to be the angle of the arm and it not being in an ideal position to pick up the reference marker. All instrumented trajectories were accurate when checked with fluoro and the system had been used successfully with no issue since the procedure. UNK MAZOR X X-EYE CAMERA | 2019-02-18 CCR, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that there we difficulties with registration and the L4 and L5 trajectories were medial. The surgeon planned an open revision for L3-S2 using two segments on 2019-02-17. Setup for the procedure was completed with no issues. During the procedure, the X-Eye would not pick up the reference marker. The manufacturer representative checked the USB cable, but the issue was not resolved. A default AP and oblique image were used. Registration was successful and L3-L5 was instrumented. New AP and oblique images were taken for the second segment. Registration was attempted, but could not be completed using the automatic, semi-automatics and manual registration processes. New images were taken, but automatic and semi-automatic registration failed. The manufacturer representative was able to complete registration using the manual process, but the values were not good. The surgeon had decided to proceed with the case and had free handed S1 on the left side. After additional attempts, the manufacturer representative was able to complete manual registration with good values so the surgeon was able to instrument screw number 9 and 10 with the guidance system. The surgeon thought the trajectory for L4 right was to medial so they did not drill. The manufacturer representative had gone into the plan and moved the trajectory lateral. The surgeon did not drill L5 left for the same reason. The case was able to be completed. There was no patient harm reported. 2019-04-29 e2 (Rep): Additional information received from a manufacturer representative reported the X-eye camera was functioning properly prior to the procedure. The representative clarified that the surgeon did not like the trajectory in the open case since they thought it looked either medial or lateral, but there was noting wrong with the trajectory. The surgeon thought the trajectory looked medial and decided not to instrument. The representative confirmed the USB cord was plugged in. The issue was determined to be the angle of the arm and it not being in an ideal position to pick up the reference marker. All instrumented trajectories were accurate when checked with fluoro and the system had been used successfully with no issue since the procedure. | x registration difficultiesinaccuracy ccr spinal procedure treat degenerative discs fixation difficulties registration trajectories medial surgeon planned open revision using segments setup procedure completed issues procedure xeye pick reference marker checked usb cable issue resolved default ap oblique image registration successful instrumented new ap oblique images taken second segment registration attempted completed using automatic semiautomatics manual registration processes new images taken automatic semiautomatic registration failed able complete registration using manual process values good surgeon decided proceed case free handed left additional attempts able complete manual registration good values surgeon able instrument screw number surgeon thought trajectory right medial did drill gone plan moved trajectory lateral surgeon did drill left reason case able completed additional xeye camera functioning properly prior procedure clarified surgeon did like trajectory open case thought looked medial lateral noting wrong trajectory surgeon thought trajectory looked medial decided instrument confirmed usb cord plugged issue determined angle arm ideal position pick reference marker instrumented trajectories accurate checked fluoro successfully issue procedure unk mazor x xeye camera | NAV4110 | NAV5000 | NAV3241 | NAV4095 | NAV4094 | NAV4110NAV5000NAV3241NAV4095NAV4094 | True
340 | 703078751 | XSE - Issues Building Pressure | 2019-03-28T16:54:56.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-01 | RNX-WS-1040 | XSE - Issues Building Pressure 2019-03-25 SCL1900733, Intake Form (Rep): Medtronic received information regarding a guidance system. It was reported that the system was having issues with building pressure during installation. The manufacturer representative thought the issue may be due to the high altitude. There was no patient involved in the event. 2019-04-24 e1 (Rep): Additional information received from a manufacturer representative reported the workstation was able to build pressure, but took four minutes to do some after being completely empty. The representative indicated that this was within specification, but longer compared to other system. RNX-WS-1040 | 2019-03-25 SCL1900733, Intake Form (Rep): Medtronic received information regarding a guidance system. It was reported that the system was having issues with building pressure during installation. The manufacturer representative thought the issue may be due to the high altitude. There was no patient involved in the event. 2019-04-24 e1 (Rep): Additional information received from a manufacturer representative reported the workstation was able to build pressure, but took four minutes to do some after being completely empty. The representative indicated that this was within specification, but longer compared to other system. | xse issues building pressure having issues building pressure installation thought issue high altitude involved additional workstation able build pressure took minutes completely indicated specification longer compared | NAV3264 | NAV2054 | NAV4095 | NAV3249 | NAV2058 | NAV3264NAV2054NAV4095NAV3249NAV2058 | True
7462 | 704210882 | XSE - Registration Difficulties | 2021-02-03T17:41:44.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Difficulties 2021-Feb-02 Salesforce C00220394 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when doing a L4-S1 case, she was unable to get registration on L5. This was the vertebrae that was most visible in the scan and the one she was using for the one-dot registration. L4 and S1 were registering with no issues. There were peak cages between L4-L5 and L5-S1, but the rep could barely see them in the c-arm images and they were not obstructing the bony anatomy. She tried to do semi-manual and manual registrations, but was unsuccessful. After 10 minutes, they decided to bring an o-arm to do scan and plan. There was a delay of more than 1 hour but no patient harm reported. 2021-Feb-18 IUDx3 (other): No new information. 2021-05-20 e1 (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-Feb-02 Salesforce C00220394 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when doing a L4-S1 case, she was unable to get registration on L5. This was the vertebrae that was most visible in the scan and the one she was using for the one-dot registration. L4 and S1 were registering with no issues. There were peak cages between L4-L5 and L5-S1, but the rep could barely see them in the c-arm images and they were not obstructing the bony anatomy. She tried to do semi-manual and manual registrations, but was unsuccessful. After 10 minutes, they decided to bring an o-arm to do scan and plan. There was a delay of more than 1 hour but no patient harm reported. 2021-Feb-18 IUDx3 (other): No new information. 2021-05-20 e1 (Other): No new information. | xse registration difficulties salesforce spinal p rocedure doing case unable registration vertebrae visible scan using onedot registration registering issues peak cages barely carm images obstructing bony anatomy tried semimanual manual registrations unsuccessful minutes decided bring oarm scan plan delay hour new new mazor x | NAV4112 | NAV3242 | NAV5000 | NAV3244 | NAV3261 | NAV4112NAV3242NAV5000NAV3244NAV3261 | True
2663 | 703245740 | X-X-Eyes communication error | 2019-07-06T17:17:55.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X-X-Eyes communication error X-Eyes communication error" message displayed. X-Eye camera non-functional. UNK MAZOR X X-EYE CAMERA | X-Eyes communication error" message displayed. X-Eye camera non-functional. | xxeyes communication error xeyes communication error message displayed xeye camera nonfunctional unk mazor x xeye camera | NAV4110 | NAV2047 | NAV3254 | NAV2054 | NAV4109 | NAV4110NAV2047NAV3254NAV2054NAV4109 | True
6163 | 703877870 | X - OOBF/PC Handle Damaged | 2020-07-29T23:21:59.000000+0000 | MazorX | NAV3262 | PC ISSUES | 20 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | X - OOBF/PC Handle Damaged 2020-07-29 00693679 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the PC was opened, the handle at the top of the PC was damaged. There was no patient involvement. 2020-08-06 IUD x3 (Other): No new information. 2020-08-11 IUD x2 (Other): The PC was still functioning. 2020-08-19 IUD (Other): No new information. 2020-08-30 IUD x4 (Other): No new information. 2020-10-27 IUD (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2020-07-29 00693679 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the PC was opened, the handle at the top of the PC was damaged. There was no patient involvement. 2020-08-06 IUD x3 (Other): No new information. 2020-08-11 IUD x2 (Other): The PC was still functioning. 2020-08-19 IUD (Other): No new information. 2020-08-30 IUD x4 (Other): No new information. 2020-10-27 IUD (Other): No new information. | x oobfpc handle damaged outside procedure pc opened handle pc damaged iud new iud pc functioning iud new iud new iud new mazor x pc station | NAV3262 | NAV2055 | NAV3251 | NAV1006 | NAV2054 | NAV3262NAV2055NAV3251NAV1006NAV2054 | True
1482 | 703202278 | X- Cart Handle Issues | 2019-06-12T12:43:47.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | MEC5416-01 | CART MAIN HANDLE | X- Cart Handle Issues Cart Handle is broken, right side snapped off. Missing arm guide screw. Accuracy pointer is bent b/c it was used to pry the positioner's teeth off the bed frame adapter. All these issues were found by Leafar when he was performing the PM CART MAIN HANDLE | Cart Handle is broken, right side snapped off. Missing arm guide screw. Accuracy pointer is bent b/c it was used to pry the positioner's teeth off the bed frame adapter. All these issues were found by Leafar when he was performing the PM | x cart handle issues cart handle broken right snapped missing arm guide screw accuracy pointer bent bc pry positioners teeth bed frame adapter issues leafar performing pm cart main handle | NAV2058 | NAV3295 | NAV4091 | NAV3246 | NAV2053 | NAV2058NAV3295NAV4091NAV3246NAV2053 | True
4410 | 703541536 | XSE - CANBUS Error | 2020-01-02T19:58:20.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | MAS1574-01 | BED FRAME ADAPTOR | XSE - CANBUS Error 2020-1-2 00681269(Rep): Medtronic received information regarding a guidance system being used during service. It was reported that the bed frame adaptor needed to be replaced because it was displaying a canbus error. No patient involvement. 2020-01-10 IUD x5 (Other): No new information. 2020-02-10 IUD x2 (Other): Additional information received reporting that analysis was completed on the bed frame adaptor. BED FRAME ADAPTOR | 2020-1-2 00681269(Rep): Medtronic received information regarding a guidance system being used during service. It was reported that the bed frame adaptor needed to be replaced because it was displaying a canbus error. No patient involvement. 2020-01-10 IUD x5 (Other): No new information. 2020-02-10 IUD x2 (Other): Additional information received reporting that analysis was completed on the bed frame adaptor. | xse canbus error service bed frame adaptor needed replaced displaying canbus error iud new iud additional reporting analysis completed bed frame adaptor bed frame adaptor | NAV3255 | NAV2047 | NAV2054 | NAV3295 | NAV3299 | NAV3255NAV2047NAV2054NAV3295NAV3299 | True
4777 | 703575865 | X - Registration Issues | 2020-01-22T18:37:26.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issues 2018-05-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done by the surgeon and manufacturer representative for an MIS fusion at L2-L4. Significant rotation and small pedicles were noted. All screws were planned according to best fixation trajectory options. A Schanz pin was placed at the right PSIS. Draw spine was utilized. The 3D marker was sent to AP and OBL for imaging. There were significant issues visualizing endplates at L3 due to titanium cages at L2/3 and L3/4, Image quality was extremely low as few endplates were visible. The initial registration attempt failed on the complex algorithm. Recalibration was attempted. New AP and OBL images were taken. The C-arm setting were switched from spine to orthopedic. While gathering new images, the tech again manipulated image and needed another recalibration. On the final attempt, AP image was acquired and OBL image still did not have clear endplates. The final registration attempt was done with the basic algorithm but it was unable to verify L3 when verifying accuracy. The representative moved back to adjust the segmentation lines but again was unable to verify L3, with red values. The surgeon agreed to move forward with L2 and L4. Trajectories were sent to left L2 and L4, then right L2 and L4. K wires were placed accurately and visual verification was approved. While instrumenting L3 by hand, the surgeon pulled the L2 K-wires, noting that they were out if the bone, despite strong tactile feedback and fluoro verification. The surgeon opted to place L2 freehand as well. Placement of L4 was verified to be accurate. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-05-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done by the surgeon and manufacturer representative for an MIS fusion at L2-L4. Significant rotation and small pedicles were noted. All screws were planned according to best fixation trajectory options. A Schanz pin was placed at the right PSIS. Draw spine was utilized. The 3D marker was sent to AP and OBL for imaging. There were significant issues visualizing endplates at L3 due to titanium cages at L2/3 and L3/4, Image quality was extremely low as few endplates were visible. The initial registration attempt failed on the complex algorithm. Recalibration was attempted. New AP and OBL images were taken. The C-arm setting were switched from spine to orthopedic. While gathering new images, the tech again manipulated image and needed another recalibration. On the final attempt, AP image was acquired and OBL image still did not have clear endplates. The final registration attempt was done with the basic algorithm but it was unable to verify L3 when verifying accuracy. The representative moved back to adjust the segmentation lines but again was unable to verify L3, with red values. The surgeon agreed to move forward with L2 and L4. Trajectories were sent to left L2 and L4, then right L2 and L4. K wires were placed accurately and visual verification was approved. While instrumenting L3 by hand, the surgeon pulled the L2 K-wires, noting that they were out if the bone, despite strong tactile feedback and fluoro verification. The surgeon opted to place L2 freehand as well. Placement of L4 was verified to be accurate. There was no patient harm and the procedure was not delayed over an hour. | x registration issues ccr spinal procedure planning surgeon mis fusion significant rotation small pedicles noted screws planned according best fixation trajectory options schanz pin placed right psis draw spine utilized marker sent ap obl imaging significant issues visualizing endplates titanium cages image quality extremely low endplates visible initial registration attempt failed complex algorithm recalibration attempted new ap obl images taken carm setting switched spine orthopedic gathering new images tech manipulated image needed recalibration final attempt ap image acquired obl image did clear endplates final registration attempt basic algorithm unable verify verifying accuracy moved adjust segmentation lines unable verify red values surgeon agreed forward trajectories sent left right k wires placed accurately visual verification approved instrumenting hand surgeon pulled kwires noting bone despite strong tactile feedback fluoro verification surgeon opted place freehand placement verified accurate procedure delayed hour mazor x | NAV4112 | NAV3237 | NAV4091 | NAV3242 | NAV3299 | NAV4112NAV3237NAV4091NAV3242NAV3299 | True
8094 | 704349379 | XSE - Deviation | 2021-04-20T00:17:39.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2021-Apr-15 Salesforce C00224942, IUD x3, email 2021-04-21 12:31 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that two segments were done during a T10-S2 open fusion case. The first segment was T10-L3 and a dual clamp and bone mount bridge was used to mount the surgical system to the patient. The second segment was L4-S2 and a Schanz arm and pin were used to mount the surgical system to the patient. The manufacturer representative noted that a drill was used when placing screws during the procedure. Imaging was done during the case and the surgeon thought the screws were placed well. Four days after the procedure, the surgeon told the representative that the S2 screw on the right side was deviated more than 10 mm. The patient experienced slight discomfort after the procedure. The representative indicated that the first 30 mm of the entry point looked identical to the planned trajectory, but the tap looked to have taken the path of least resistance down the joint. The surgery was delayed less than an hour. 2021-05-13, email 2021-05-13 17:31, e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the surgeon felt the screw was loose. The surgeon used a balltip probe and felt a slight inferior breach so they decided to re-instrument the trajectory during the procedure. No revision was completed. 2021-05-20 e1 (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-Apr-15 Salesforce C00224942, IUD x3, email 2021-04-21 12:31 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that two segments were done during a T10-S2 open fusion case. The first segment was T10-L3 and a dual clamp and bone mount bridge was used to mount the surgical system to the patient. The second segment was L4-S2 and a Schanz arm and pin were used to mount the surgical system to the patient. The manufacturer representative noted that a drill was used when placing screws during the procedure. Imaging was done during the case and the surgeon thought the screws were placed well. Four days after the procedure, the surgeon told the representative that the S2 screw on the right side was deviated more than 10 mm. The patient experienced slight discomfort after the procedure. The representative indicated that the first 30 mm of the entry point looked identical to the planned trajectory, but the tap looked to have taken the path of least resistance down the joint. The surgery was delayed less than an hour. 2021-05-13, email 2021-05-13 17:31, e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the surgeon felt the screw was loose. The surgeon used a balltip probe and felt a slight inferior breach so they decided to re-instrument the trajectory during the procedure. No revision was completed. 2021-05-20 e1 (Other): No new information. | xse deviation salesforce iud email spinal procedure segments open fusion case segment dual clamp bone mount bridge mount surgical second segment schanz arm pin mount surgical noted drill placing screws procedure imaging case surgeon thought screws placed days procedure surgeon told screw right deviated mm experienced slight discomfort procedure indicated mm entry point looked identical planned trajectory tap looked taken path resistance joint surgery delayed hour email additional surgeon felt screw loose surgeon balltip probe felt slight inferior breach decided reinstrument trajectory procedure revision completed new mazor x | NAV5000 | NAV1012 | NAV3243 | NAV3253 | NAV4097 | NAV5000NAV1012NAV3243NAV3253NAV4097 | True
403 | 703097662 | XSE - failed registration | 2019-04-09T18:09:39.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - failed registration 2019-Apr-05 SCL1900836 intake form, .msg x3 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that scan and plan registration failed. There were also two broken dual clamps. It was noted that the surgical time was delayed for over an hour and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Apr-05 SCL1900836 intake form, .msg x3 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that scan and plan registration failed. There were also two broken dual clamps. It was noted that the surgical time was delayed for over an hour and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse failed registration msg spinal procedure scan plan registration failed broken dual clamps noted surgical time delayed hour additional complications anticipated mazor x software installation kit | NAV5000 | NAV3246 | NAV3242 | NAV4112 | NAV2054 | NAV5000NAV3246NAV3242NAV4112NAV2054 | True
8317 | 704396387 | XSE - Schanz Arm Loose | 2021-05-14T23:52:15.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Schanz Arm Loose 2021-May-14 Salesforce C00226497 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there were four schanz arms loose and a schanz bond screw was stripped. There was no impact on the patient outcome. The procedure was not delayed. 2021-05-26 IUD x6 (Other): No new information. 2021-06-28 email 2021-06-28 14:01 (Rep): Troubleshooting involved taking the Shanz arms apart and reassembling them. The surgeon observed rust inside the connection point and believed that the age of the arms had caused them to loosen. The shanz bond screw was stripped due to use. The cause of loose Schanz arms or stripped screw was due to being used many times and tightening it repeatedly. KIT SVC MAS2351-04S SCHANZ ARM | 2021-May-14 Salesforce C00226497 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there were four schanz arms loose and a schanz bond screw was stripped. There was no impact on the patient outcome. The procedure was not delayed. 2021-05-26 IUD x6 (Other): No new information. 2021-06-28 email 2021-06-28 14:01 (Rep): Troubleshooting involved taking the Shanz arms apart and reassembling them. The surgeon observed rust inside the connection point and believed that the age of the arms had caused them to loosen. The shanz bond screw was stripped due to use. The cause of loose Schanz arms or stripped screw was due to being used many times and tightening it repeatedly. | xse schanz arm loose salesforce spinal p rocedure schanz arms loose schanz bond screw stripped impact outcome procedure delayed iud new email troubleshooting involved taking shanz arms apart reassembling surgeon observed rust inside connection point believed age arms caused loosen shanz bond screw stripped use cause loose schanz arms stripped screw times tightening repeatedly kit svc schanz arm | NAV3299 | NAV3296 | NAV4109 | NAV3298 | NAV4097 | NAV3299NAV3296NAV4109NAV3298NAV4097 | True
6848 | 704049287 | XSE - Inaccurate Navigation | 2020-11-03T18:42:33.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Inaccurate Navigation 2020-11-03 00700107, email 2020-11-05 09:01, IUD x5, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using screws and rods. It was reported that navigation was inaccurate by more than 3 mm and inferior on all levels during a procedure. This included the sagittal views as well. After taking confirmation images, all screws were found to be placed accurately. The surgeon took a second snapshot, checked the plan and checked the reference frame placement. The cause of navigation being off was found to be the fastener on the surgical arm where the robotic reference frame attaches was loose. The loose fastener caused the reference frame to be loose and the navigation to be off. This issue was able to be reproduced after the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-01 IUD x4 (Other): No new information. 2020-Dec-28 IUD (other): No new information. 2021-02-01 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-11-03 00700107, email 2020-11-05 09:01, IUD x5, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using screws and rods. It was reported that navigation was inaccurate by more than 3 mm and inferior on all levels during a procedure. This included the sagittal views as well. After taking confirmation images, all screws were found to be placed accurately. The surgeon took a second snapshot, checked the plan and checked the reference frame placement. The cause of navigation being off was found to be the fastener on the surgical arm where the robotic reference frame attaches was loose. The loose fastener caused the reference frame to be loose and the navigation to be off. This issue was able to be reproduced after the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-01 IUD x4 (Other): No new information. 2020-Dec-28 IUD (other): No new information. 2021-02-01 IUD (Other): No new information. | xse inaccurate navigation email iud ccr spinal procedure treat degenerative discs using screws rods navigation inaccurate mm inferior levels procedure included sagittal views taking confirmation images screws placed accurately surgeon took second snapshot checked plan checked reference frame placement cause navigation fastener surgical arm robotic reference frame attaches loose loose fastener caused reference frame loose navigation issue able reproduced procedure procedure delayed hour iud new iud new iud new kit svc surgical arm refurb | NAV5000 | NAV4097 | NAV3249 | NAV5001 | NAV4100 | NAV5000NAV4097NAV3249NAV5001NAV4100 | False
1000 | 703183454 | X - Deviation Issue | 2019-05-31T15:52:09.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviation Issue S1, S2 trajectories believed to be inaccurate. “Green” registration values, no apparent skive or mechanical deviations. MAZOR X SOFTWARE KIT VER-3.0.2 | S1, S2 trajectories believed to be inaccurate. “Green” registration values, no apparent skive or mechanical deviations. | x deviation issue trajectories believed inaccurate “green” registration values apparent skive mechanical deviations mazor x software kit | NAV4091 | NAV3242 | NAV3248 | NAV3247 | NAV3246 | NAV4091NAV3242NAV3248NAV3247NAV3246 | True
8128 | 704358429 | XSE - Multiple System Failures | 2021-04-23T16:20:14.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 30 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - Multiple System Failures 2021-Apr-22 Salesforce C00225180, C00225183, C00225182, C00225184 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there were multiple failures during training. The 3Define camera failed. The camera was red in the setup screen and an Error 10273 was displayed. During a scan, a Error 4469 was displayed. Troubleshooting involved checking the cables and they were attached well. During setup of the surgical arm, an Error 8047 Arm Not Reached message was displayed after pressing the emergency button. The surgical arm setup was repeated and the issue was resolved. The shoulder was moving significantly without any warning about shoulder shift being displayed. The surgeon screen was mirroring the workstation, but the screen was not responding to touch. There was no patient involvement. 2021-04-29 email 2021-04-29 13:31 (Rep, For): No new information. 2021-05-05 IUD (Other, For): No new information. 2021-05-20 e2, email 2021-05-20 1346 (Rep, For): No new information. 2021-07-05 e3 (Other, For): No new information. 2021-07-11 e4 (Rep, For): No new information. 2021-07-15 e5, e6, e7 (Rep, For): No new information. 2021-08-05 e8 (Other, For): No new information. 2021-08-29 e9 (Other, For): No new information. 2021-11-22 IUD (Other, For): No new information. 2022-01-06 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Apr-22 Salesforce C00225180, C00225183, C00225182, C00225184 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there were multiple failures during training. The 3Define camera failed. The camera was red in the setup screen and an Error 10273 was displayed. During a scan, a Error 4469 was displayed. Troubleshooting involved checking the cables and they were attached well. During setup of the surgical arm, an Error 8047 Arm Not Reached message was displayed after pressing the emergency button. The surgical arm setup was repeated and the issue was resolved. The shoulder was moving significantly without any warning about shoulder shift being displayed. The surgeon screen was mirroring the workstation, but the screen was not responding to touch. There was no patient involvement. 2021-04-29 email 2021-04-29 13:31 (Rep, For): No new information. 2021-05-05 IUD (Other, For): No new information. 2021-05-20 e2, email 2021-05-20 1346 (Rep, For): No new information. 2021-07-05 e3 (Other, For): No new information. 2021-07-11 e4 (Rep, For): No new information. 2021-07-15 e5, e6, e7 (Rep, For): No new information. 2021-08-05 e8 (Other, For): No new information. 2021-08-29 e9 (Other, For): No new information. 2021-11-22 IUD (Other, For): No new information. 2022-01-06 IUD (Other): No new information. | xse multiple failures salesforce outside procedure multiple failures training camera failed camera red setup screen error displayed scan error displayed troubleshooting involved checking cables attached setup surgical arm error arm reached message displayed pressing emergency button surgical arm setup repeated issue resolved shoulder moving significantly warning shoulder shift displayed surgeon screen mirroring workstation screen responding touch email new iud new email new new new new new new iud new iud new kit svc surgical arm refurb | NAV3246 | NAV2054 | NAV3253 | NAV3244 | NAV4108 | NAV3246NAV2054NAV3253NAV3244NAV4108 | False
774 | 703175835 | X- Workstation Issues | 2019-05-28T18:18:54.000000+0000 | MazorX | NAV3262 | PC ISSUES | 20 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | X- Workstation Issues Monitor is displaying, "No Signal Detected" UNK MAZOR X MONITOR | Monitor is displaying, "No Signal Detected" | x workstation issues monitor displaying signal detected unk mazor x monitor | NAV3261 | NAV2058 | NAV2054 | NAV3264 | NAV4101 | NAV3261NAV2058NAV2054NAV3264NAV4101 | False
1659 | 703208446 | X- Accessory Issue | 2019-06-14T20:38:39.000000+0000 | MazorX | NAV2053 | MISSING PARTS | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | X- Accessory Issue Sterile Processing Department lost 2 Items out the the Mazor X tray: the target extender and long peek screw for the target extender. Both pieces need to be replaced. UNK MAZOR X ACCESSORIES | Sterile Processing Department lost 2 Items out the the Mazor X tray: the target extender and long peek screw for the target extender. Both pieces need to be replaced. | x accessory issue sterile processing department lost items mazor x tray target extender long peek screw target extender pieces need replaced unk mazor x accessories | NAV3299 | NAV2053 | NAV3233 | NAV3295 | NAV2044 | NAV3299NAV2053NAV3233NAV3295NAV2044 | True
9038 | 704534784 | XSE ¦ Intermittent Black Monitors | 2021-08-02T20:18:28.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 20 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE ¦ Intermittent Black Monitors 2021-Aug-02 Salesforce C00231219 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that both the workstation and surgeon monitor would intermittently go black and then return after anywhere between 2 and 20 seconds. There was no impact on the patient outcome. 2021-08-03 IUD x7 (Other): No new information. 2021-08-24 IUD x2 (Other): No new information. 2021-09-22 IUD x2 (Other): No new information. 2021-10-03 IUD (Other): No new information. 2022-01-05 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2021-Aug-02 Salesforce C00231219 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that both the workstation and surgeon monitor would intermittently go black and then return after anywhere between 2 and 20 seconds. There was no impact on the patient outcome. 2021-08-03 IUD x7 (Other): No new information. 2021-08-24 IUD x2 (Other): No new information. 2021-09-22 IUD x2 (Other): No new information. 2021-10-03 IUD (Other): No new information. 2022-01-05 IUD (Other): No new information. | xse intermittent black monitors salesforce spinal p rocedure workstation surgeon monitor intermittently black return seconds impact outcome iud new iud new iud new iud new iud new mazor x pc station refurbished | NAV3261 | NAV3262 | NAV3253 | NAV5000 | NAV3247 | NAV3261NAV3262NAV3253NAV5000NAV3247 | False
7927 | 704308657 | XSE - Deviations | 2021-03-28T20:05:39.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviations 2021-Mar-26 Salesforce C00223400 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a L4-L5 case, the surgeon performed a partial hemilaminectomy on both L4 and L5 and then placed a cage. The surgeon then placed the clamp on the spinous process of L4. The surgeon confirmed the spinous process and platform were rigid and stable. The screws on the right side were done first. During instrumentation on the left L4 screws, the surgical arm moved and a "Arm not on trajectory" error message was displayed. The surgical arm was resent to the trajectory and the surgeon placed the screw. Left L5 was completed next. After placed screws, imaging was done. Both screws on the right were accurate, but the screws on the left were superolateral. A new registration was and the surgical arm was sent to left L4. The surgeon used fluoro to control the process. The drill was accurate was confirmed to be accurate. When trying to instrument the screw, fluoro showed that the screw would start in the pedicle accurately and then move superiorly. The surgeon did not want to adjust the planning to drill a new pilot hole so they aborted the use of the guidance system and completed the case freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-04-07 email 2021-04-07 08:01 (Rep, For): Additional information received from a manufacturer representative reported that it was unknown how much the trajectories were deviated and the cause of the deviations was not determined. MAZOR X SYSTEM TPL0059 | 2021-Mar-26 Salesforce C00223400 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a L4-L5 case, the surgeon performed a partial hemilaminectomy on both L4 and L5 and then placed a cage. The surgeon then placed the clamp on the spinous process of L4. The surgeon confirmed the spinous process and platform were rigid and stable. The screws on the right side were done first. During instrumentation on the left L4 screws, the surgical arm moved and a "Arm not on trajectory" error message was displayed. The surgical arm was resent to the trajectory and the surgeon placed the screw. Left L5 was completed next. After placed screws, imaging was done. Both screws on the right were accurate, but the screws on the left were superolateral. A new registration was and the surgical arm was sent to left L4. The surgeon used fluoro to control the process. The drill was accurate was confirmed to be accurate. When trying to instrument the screw, fluoro showed that the screw would start in the pedicle accurately and then move superiorly. The surgeon did not want to adjust the planning to drill a new pilot hole so they aborted the use of the guidance system and completed the case freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-04-07 email 2021-04-07 08:01 (Rep, For): Additional information received from a manufacturer representative reported that it was unknown how much the trajectories were deviated and the cause of the deviations was not determined. | xse deviations salesforce spinal procedure case surgeon performed partial hemilaminectomy placed cage surgeon placed clamp spinous process surgeon confirmed spinous process platform rigid stable screws right instrumentation left screws surgical arm moved arm trajectory error message displayed surgical arm resent trajectory surgeon placed screw left completed placed screws imaging screws right accurate screws left superolateral new registration surgical arm sent left surgeon fluoro control process drill accurate confirmed accurate trying instrument screw fluoro showed screw start pedicle accurately superiorly surgeon did want adjust planning drill new pilot hole aborted use completed case freehand procedure delayed hour email additional unknown trajectories deviated cause deviations determined mazor x | NAV5000 | NAV4101 | NAV4098 | NAV3246 | NAV3299 | NAV5000NAV4101NAV4098NAV3246NAV3299 | True
6952 | 704075310 | XSE - Deviation | 2020-11-17T22:50:33.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2020-11-12 00701073, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that during an open case, an XLIF was completed first. The patient was then flipped and the open incision was made. Registration was successfully completed with 2 AP and one oblique images. The surgeon used the guidance system to place screws at L3. Hardware at L4 and L5 was replaced and S1 was removed. Right L3 was deviated medially due to soft tissue pressure. Left L3 was accurate. Prior to the case, the system passed a setup test. Post op images were taken and the surgeon felt the placement was good; however, the patient had pain when they woke up. A revision was done on the following day. The screw was repositioned freehand. The initial procedure was delayed less than an hour. 2020-12-10 email 2020-12-10 13:15, IUD (Rep): Additional information received from a manufacturer representative reported that the screw was misplaced by 1 cm. MAZOR X SYSTEM TPL0059 | 2020-11-12 00701073, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that during an open case, an XLIF was completed first. The patient was then flipped and the open incision was made. Registration was successfully completed with 2 AP and one oblique images. The surgeon used the guidance system to place screws at L3. Hardware at L4 and L5 was replaced and S1 was removed. Right L3 was deviated medially due to soft tissue pressure. Left L3 was accurate. Prior to the case, the system passed a setup test. Post op images were taken and the surgeon felt the placement was good; however, the patient had pain when they woke up. A revision was done on the following day. The screw was repositioned freehand. The initial procedure was delayed less than an hour. 2020-12-10 email 2020-12-10 13:15, IUD (Rep): Additional information received from a manufacturer representative reported that the screw was misplaced by 1 cm. | xse deviation ccr spinal procedure treat degenerative discs fixation open case xlif completed flipped open incision registration successfully completed ap oblique images surgeon place screws hardware replaced removed right deviated medially soft tissue pressure left accurate prior case passed setup test post op images taken surgeon felt placement good pain woke revision following day screw repositioned freehand initial procedure delayed hour email iud additional screw misplaced cm mazor x | NAV5000 | NAV3242 | NAV3264 | NAV3249 | NAV3243 | NAV5000NAV3242NAV3264NAV3249NAV3243 | True
3873 | 703460566 | XSE - Surgical Arm failed Advanced Acc. | 2019-11-11T16:07:00.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 30 | ASM0206-01 | SURGICAL ARM | XSE - Surgical Arm failed Advanced Acc. 2019-Nov-11 00678044 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the surgical arm failed all left points of the advanced accuracy test. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-2-4 IUD x7(Other): no new information. SURGICAL ARM | 2019-Nov-11 00678044 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the surgical arm failed all left points of the advanced accuracy test. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-2-4 IUD x7(Other): no new information. | xse surgical arm failed advanced acc spinal procedure surgical arm failed left points advanced accuracy test additional complications anticipated iud new surgical arm | NAV4093 | NAV2054 | NAV3249 | NAV4101 | NAV3244 | NAV4093NAV2054NAV3249NAV4101NAV3244 | True
2887 | 703258664 | X - X-Eye and Clamp Link Stripped | 2019-07-15T14:02:31.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - X-Eye and Clamp Link Stripped Spoke about problems picking up Stylus Marker and with the clamp link stripping. Duration 10 mins UNK MAZOR X X-EYE CAMERA | Spoke about problems picking up Stylus Marker and with the clamp link stripping. Duration 10 mins | x xeye clamp link stripped spoke problems picking stylus marker clamp link stripping duration mins unk mazor x xeye camera | NAV4110 | NAV3299 | NAV2054 | NAV3295 | NAV4109 | NAV4110NAV3299NAV2054NAV3295NAV4109 | True
3781 | 703441860 | XSE - Damaged Long Drill Guide | 2019-10-30T20:02:02.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | KIT0736 | LONG DRILL GUIDE AND CANNULA KIT0736 | XSE - Damaged Long Drill Guide 2019-Oct-30 00677226, IUD (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the sites long drill guide and cannula were not sliding together smoothly as they should due to being old. The teeth on the drill guide were starting to bend. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-12-03 IUD x2 (Other): No new information. 2020-02-26 IUD (Other): No new information. LONG DRILL GUIDE AND CANNULA KIT0736 | 2019-Oct-30 00677226, IUD (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the sites long drill guide and cannula were not sliding together smoothly as they should due to being old. The teeth on the drill guide were starting to bend. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-12-03 IUD x2 (Other): No new information. 2020-02-26 IUD (Other): No new information. | xse damaged long drill guide iud outside procedure sites long drill guide cannula sliding smoothly old teeth drill guide starting bend additional complications anticipated iud new iud new long drill guide cannula | NAV3299 | NAV3295 | NAV2054 | NAV3262 | NAV3264 | NAV3299NAV3295NAV2054NAV3262NAV3264 | True
7119 | 704123074 | XSE - Damaged Instrumentation Trays | 2020-12-14T17:13:08.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 100 | 9734924 | FRAME 9734924 ROBOTIC REFERENCE | XSE - Damaged Instrumentation Trays 2020-12-11 00702590, e1, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instrumentation trays were damaged and the instruments failed tracking during preventative maintenance. The trajectory error ranged between 0.554 and 0.976 which was above the maximum amount allowed. Troubleshooting involved trying different spheres, trying different Navlock trackers and snapshot tracker, using long and short tools, adjusting the mounted position of the surgical system, hard restarts of the guidance system and trying different surgical arms. The stealth camera was viewing all of the spheres well during testing. The most probable cause was noted to be the washing and cleaning by the sterile processing department. The instruments were in rough condition and the edges. The edges of the trackers were chipped and the paint was faded. The trackers were difficult to detect during testing. There was no patient involvement. 2021-01-10 IUD (Other): No new information. 2021-01-10 IUD x2 (Other): No new information. 2021-02-02 email 2021-02-02 13:45, IUD (Other): No new information. 2021-02-24 IUD x2, email 2021-02-23 11:31 (Rep): No new information. 2021-03-09 IUD (Rep): Additional information received from a manufacturer representative reported that the issue was determined to be the camera. The camera was replaced and the issue was resolved. 2021-03-16 email 2021-03-16 20:45 (Rep): Additional information received from a manufacturer representative reported that all of the instruments were tested and found to be functioning as designed. FRAME 9734924 ROBOTIC REFERENCE | 2020-12-11 00702590, e1, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instrumentation trays were damaged and the instruments failed tracking during preventative maintenance. The trajectory error ranged between 0.554 and 0.976 which was above the maximum amount allowed. Troubleshooting involved trying different spheres, trying different Navlock trackers and snapshot tracker, using long and short tools, adjusting the mounted position of the surgical system, hard restarts of the guidance system and trying different surgical arms. The stealth camera was viewing all of the spheres well during testing. The most probable cause was noted to be the washing and cleaning by the sterile processing department. The instruments were in rough condition and the edges. The edges of the trackers were chipped and the paint was faded. The trackers were difficult to detect during testing. There was no patient involvement. 2021-01-10 IUD (Other): No new information. 2021-01-10 IUD x2 (Other): No new information. 2021-02-02 email 2021-02-02 13:45, IUD (Other): No new information. 2021-02-24 IUD x2, email 2021-02-23 11:31 (Rep): No new information. 2021-03-09 IUD (Rep): Additional information received from a manufacturer representative reported that the issue was determined to be the camera. The camera was replaced and the issue was resolved. 2021-03-16 email 2021-03-16 20:45 (Rep): Additional information received from a manufacturer representative reported that all of the instruments were tested and found to be functioning as designed. | xse damaged instrumentation trays iud outside procedure instrumentation trays damaged instruments failed tracking preventative maintenance trajectory error ranged maximum allowed troubleshooting involved trying different spheres trying different navlock trackers snapshot tracker using long short tools adjusting mounted position surgical hard restarts trying different surgical arms stealth camera viewing spheres testing probable cause noted washing cleaning sterile processing department instruments rough condition edges edges trackers chipped paint faded trackers difficult detect testing iud new iud new email iud new iud email new iud additional issue determined camera camera replaced issue resolved email additional instruments tested functioning designed frame robotic reference | NAV3299 | NAV3244 | NAV3295 | NAV3249 | NAV3262 | NAV3299NAV3244NAV3295NAV3249NAV3262 | True
2868 | 703256697 | X- Tablet/SW/Screws Issues | 2019-07-12T17:40:21.000000+0000 | MazorX | NAV3246 | DEFECTS | 40 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X- Tablet/SW/Screws Issues Tablet not responding or charging under the hood. Surgeon wants his tablet working ASAP. Believe it is an older model tablet. Stylus is an older model stylus, it was getting picked up by the x-eyes but no consistent beeps, and it would not give an OBL trajectory for registration. Disposable clamp the bridge that attaches to it, the screw keeps stripping from the square screw driver. UNKNOWN MAZOR X SW | Tablet not responding or charging under the hood. Surgeon wants his tablet working ASAP. Believe it is an older model tablet. Stylus is an older model stylus, it was getting picked up by the x-eyes but no consistent beeps, and it would not give an OBL trajectory for registration. Disposable clamp the bridge that attaches to it, the screw keeps stripping from the square screw driver. | x tabletswscrews issues tablet responding charging hood surgeon wants tablet working asap believe older model tablet stylus older model stylus getting picked xeyes consistent beeps obl trajectory registration disposable clamp bridge attaches screw keeps stripping square screw driver unknown mazor x sw | NAV3253 | NAV3246 | NAV3299 | NAV3242 | NAV4112 | NAV3253NAV3246NAV3299NAV3242NAV4112 | True
4353 | 703536975 | XSE - Missing Iron Man | 2019-12-27T23:56:28.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | MAS2497-01S | KIT SVC MAS2497-01S IRON MAN | XSE - Missing Iron Man 2019-12-27 00681140 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the iron man was caught in the drawer and had fallen into the system. No patient was present. 2020-01-07 IUD x7 (Other): No new information. 2020-03-09 IUD x3 (Other): No new information. 2020-03-31 IUD (Other): No new information. 2020-04-17 IUD x3 (Other): No new information. 2020-05-08 IUD (Other): No new information. KIT SVC MAS2497-01S IRON MAN | 2019-12-27 00681140 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the iron man was caught in the drawer and had fallen into the system. No patient was present. 2020-01-07 IUD x7 (Other): No new information. 2020-03-09 IUD x3 (Other): No new information. 2020-03-31 IUD (Other): No new information. 2020-04-17 IUD x3 (Other): No new information. 2020-05-08 IUD (Other): No new information. | xse missing iron man outside procedure iron man caught drawer fallen present iud new iud new iud new iud new iud new kit svc iron man | NAV2053 | NAV2054 | NAV3299 | NAV3249 | NAV2047 | NAV2053NAV2054NAV3299NAV3249NAV2047 | True
6518 | 703963517 | XSE - SRS Mounting Issues | 2020-09-17T19:58:18.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0214-02 | SSY | XSE - SRS Mounting Issues 2020-09-17 00696929 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the surgical system was being put on the bed by the surgeon, one of the teeth locked on and the other did not. The surgeon was unable to unlock the tooth in the override panel and the pressure would not increase. The bed frame adapter was manually detached from the bed and then the surgical system was able to be removed. The pressure in the system then started to build and the representative was able to mount the system to the bed. The surgeon did not feel comfortable using the system so they decided to abort the case and reschedule it for a later date. The procedure was delayed over an hour. The patient was no patient harm, but the patient was under anesthesia. 2020-09-23 IUD x4 (Other): The manufacturer representative was able to run through system checks and was able to place the robot on and off the bed 10 times with little effort, the system had no issues moving after attached and had no issues going back into storage mode. Ran through all accuracy checks with no issues, all cameras were working, everything seemed to be in working order. Leak was due to teeth not being mounted properly on BFA. SSY | 2020-09-17 00696929 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the surgical system was being put on the bed by the surgeon, one of the teeth locked on and the other did not. The surgeon was unable to unlock the tooth in the override panel and the pressure would not increase. The bed frame adapter was manually detached from the bed and then the surgical system was able to be removed. The pressure in the system then started to build and the representative was able to mount the system to the bed. The surgeon did not feel comfortable using the system so they decided to abort the case and reschedule it for a later date. The procedure was delayed over an hour. The patient was no patient harm, but the patient was under anesthesia. 2020-09-23 IUD x4 (Other): The manufacturer representative was able to run through system checks and was able to place the robot on and off the bed 10 times with little effort, the system had no issues moving after attached and had no issues going back into storage mode. Ran through all accuracy checks with no issues, all cameras were working, everything seemed to be in working order. Leak was due to teeth not being mounted properly on BFA. | xse srs mounting issues spinal procedure surgical bed surgeon teeth locked did surgeon unable unlock tooth override panel pressure increase bed frame adapter manually detached bed surgical able removed pressure started build able mount bed surgeon did feel comfortable using decided abort case reschedule later date procedure delayed hour anesthesia iud able run checks able place robot bed times little effort issues moving attached issues going storage mode ran accuracy checks issues cameras working working order leak teeth mounted properly bfa ssy | NAV3250 | NAV3264 | NAV4103 | NAV3246 | NAV3253 | NAV3250NAV3264NAV4103NAV3246NAV3253 | True
9102 | 704551890 | XSE - SRA Failed Accuracy Test | 2021-08-11T17:39:30.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Failed Accuracy Test 2021-Aug-11 Salesforce C00232694 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the accuracy test. Multiple accuracy pointers were tried with the surgical arm. A field motor calibration was done on the surgical arm and the issue was resolved. There was no patient involvement. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Aug-11 Salesforce C00232694 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the accuracy test. Multiple accuracy pointers were tried with the surgical arm. A field motor calibration was done on the surgical arm and the issue was resolved. There was no patient involvement. | xse sra failed accuracy test salesforce outside procedure surgical arm failed accuracy test multiple accuracy pointers tried surgical arm field motor calibration surgical arm issue resolved kit svc surgical arm refurb | NAV4092 | NAV4093 | NAV3249 | NAV2054 | NAV4096 | NAV4092NAV4093NAV3249NAV2054NAV4096 | True
7256 | 704157699 | XSE - Damaged Instruments | 2021-01-06T17:25:00.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 80 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Damaged Instruments 2021-Jan-06 Salesforce C00218834, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that when going through the instrument trays at the hospital, the manufacturer representative discovered that three Schanz bonds had stripped screws, two target extenders were lost and two Schanz arms were easily able to be placed over the Schanz screws. There was no patient involvement. 2021-Jan-21 IUDx2 (other): No new information. 2021-02-01 IUD x3 (Other): No new information. 2021-03-10 email 2021-03-10 20:30 (Rep): No new information. 2021-05-05 Attachment 2021-04-08 (Rep): No new information. 2021-05-09 e1 (Rep): No new information. 2021-05-18 e2 (Rep): No new information. 2021-05-25 e3 (Rep): No new information. 2021-06-15 e4 (Rep): No new information. KIT SVC MAS2351-04S SCHANZ ARM | 2021-Jan-06 Salesforce C00218834, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that when going through the instrument trays at the hospital, the manufacturer representative discovered that three Schanz bonds had stripped screws, two target extenders were lost and two Schanz arms were easily able to be placed over the Schanz screws. There was no patient involvement. 2021-Jan-21 IUDx2 (other): No new information. 2021-02-01 IUD x3 (Other): No new information. 2021-03-10 email 2021-03-10 20:30 (Rep): No new information. 2021-05-05 Attachment 2021-04-08 (Rep): No new information. 2021-05-09 e1 (Rep): No new information. 2021-05-18 e2 (Rep): No new information. 2021-05-25 e3 (Rep): No new information. 2021-06-15 e4 (Rep): No new information. | xse damaged instruments salesforce iud outside proc edure going instrument trays hospital discovered schanz bonds stripped screws target extenders lost schanz arms easily able placed schanz screws new iud new email new attachment new new new new new kit svc schanz arm | NAV3299 | NAV2053 | NAV3295 | NAV2054 | NAV3296 | NAV3299NAV2053NAV3295NAV2054NAV3296 | True
5883 | 703828253 | XSE - SRA Hit Disposable Clamp | 2020-06-30T15:06:06.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | ASM0206-01 | SURGICAL ARM | XSE - SRA Hit Disposable Clamp 2020-06-29 00691818, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the 3Define scan was completed during the procedure, but the scan was a little grainy. When sending the surgical arm to the oblique trajectory, the arm ran into the disposable clamp. The manufacturer representative noted that the s urgical arm seemed to keep getting closer to the bone mount bridge or Schanz pin. The emergency stop button was hit right away. There was no patient harm and the procedure was delayed less than an hour. 2020-07-01 email 2020-07-01 06:45, IUD, e1 (Rep): Additional information received from a manufacturer representative reported that no troubleshooting was done. The manufacturer representative just closely watched the surgical arm during the procedure. Two towels were placed over the platform when taking the 3Define scan. The case was completed successfully with no other issues. SURGICAL ARM | 2020-06-29 00691818, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the 3Define scan was completed during the procedure, but the scan was a little grainy. When sending the surgical arm to the oblique trajectory, the arm ran into the disposable clamp. The manufacturer representative noted that the s urgical arm seemed to keep getting closer to the bone mount bridge or Schanz pin. The emergency stop button was hit right away. There was no patient harm and the procedure was delayed less than an hour. 2020-07-01 email 2020-07-01 06:45, IUD, e1 (Rep): Additional information received from a manufacturer representative reported that no troubleshooting was done. The manufacturer representative just closely watched the surgical arm during the procedure. Two towels were placed over the platform when taking the 3Define scan. The case was completed successfully with no other issues. | xse sra hit disposable clamp spinal p rocedure scan completed procedure scan little grainy sending surgical arm oblique trajectory arm ran disposable clamp noted s urgical arm getting closer bone mount bridge schanz pin emergency stop button hit right away procedure delayed hour email iud additional troubleshooting just closely watched surgical arm procedure towels placed platform taking scan case completed successfully issues surgical arm | NAV4096 | NAV3249 | NAV3243 | NAV2054 | NAV3239 | NAV4096NAV3249NAV3243NAV2054NAV3239 | False
1449 | 703200412 | X - 3Define error/manupulator/connector | 2019-06-11T19:20:08.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 40 | ASM0356-01 | SURGEON SCREEN | X - 3Define error/manupulator/connector 2019-Jun-04 SCL1901664, SCL1901665, SCL1901666 intake form x4, msg x2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the manipulator rubber base was worn out and the Station screen connector was broken. It was also noted that during the case Setup, all cameras worked properly, and no malfunctions were detected. Follows bed mounting, the surgery workflow began and when the 3Dfine step was performed, an error message indicated a 3Dfine camera malfunction. Short troubleshooting did not resolve the problem. The functionality of the camera came back intermittently when moving the system shoulder and touching the 3Dfine camera cable but did not hold for the 3Dfine scan to be done. It was noted that this was an ongoing issue. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. SURGEON SCREEN | 2019-Jun-04 SCL1901664, SCL1901665, SCL1901666 intake form x4, msg x2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the manipulator rubber base was worn out and the Station screen connector was broken. It was also noted that during the case Setup, all cameras worked properly, and no malfunctions were detected. Follows bed mounting, the surgery workflow began and when the 3Dfine step was performed, an error message indicated a 3Dfine camera malfunction. Short troubleshooting did not resolve the problem. The functionality of the camera came back intermittently when moving the system shoulder and touching the 3Dfine camera cable but did not hold for the 3Dfine scan to be done. It was noted that this was an ongoing issue. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x errormanupulatorconnector msg spinal procedure manipulator rubber base worn station screen connector broken noted case setup cameras worked properly malfunctions detected follows bed mounting surgery workflow began step performed error message indicated camera malfunction short troubleshooting did resolve problem functionality camera came intermittently moving shoulder touching camera cable did hold scan noted ongoing issue additional complications anticipated surgeon screen | NAV3244 | NAV2054 | NAV3260 | NAV3295 | NAV3253 | NAV3244NAV2054NAV3260NAV3295NAV3253 | False
9462 | 704644379 | XSE ¦ Target Extended Screw Stripped | 2021-10-04T22:09:41.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2310-04S | KIT SVC MAS2310-04S TARGET EXTENDER | XSE ¦ Target Extended Screw Stripped 2021-Oct-04 Salesforce C00235764 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that before a case the manufacturer representative found the screw stripped on the target extender. No patient was present. 2021-10-11 IUD x2 (Other): No new information. 2021-10-19 email 2021-10-19 13:01 (Rep): The cause of the damage to the target extender was due to over tightening during the cases. KIT SVC MAS2310-04S TARGET EXTENDER | 2021-Oct-04 Salesforce C00235764 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that before a case the manufacturer representative found the screw stripped on the target extender. No patient was present. 2021-10-11 IUD x2 (Other): No new information. 2021-10-19 email 2021-10-19 13:01 (Rep): The cause of the damage to the target extender was due to over tightening during the cases. | xse target extended screw stripped salesforce outside procedure case screw stripped target extender present iud new email cause damage target extender tightening cases kit svc target extender | NAV3299 | NAV4090 | NAV3252 | NAV2054 | NAV3295 | NAV3299NAV4090NAV3252NAV2054NAV3295 | True
10433 | 704879658 | XSE - Damaged Camera Handle Clip | 2022-02-23T17:10:07.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | KIT1247-01 | SVC KIT1247-01 NAV HNDL CLIP W SCWS | XSE - Damaged Camera Handle Clip 2022-Feb-22 Salesforce C00244074 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the camera clip was found to be damaged during preventative maintenance. There was no patient involvement. 2022-03-11 IUD x2 (Other): No new information. 2022-03-25 IUD (Other): No new information. SVC KIT1247-01 NAV HNDL CLIP W SCWS | 2022-Feb-22 Salesforce C00244074 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the camera clip was found to be damaged during preventative maintenance. There was no patient involvement. 2022-03-11 IUD x2 (Other): No new information. 2022-03-25 IUD (Other): No new information. | xse damaged camera handle clip salesforce outside procedure camera clip damaged preventative maintenance iud new iud new svc nav hndl clip w scws | NAV3299 | NAV1013 | NAV2058 | NAV3244 | NAV2053 | NAV3299NAV1013NAV2058NAV3244NAV2053 | True
3884 | 703461336 | XSE - ATP failure, Pressure Leak | 2019-11-12T00:07:43.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - ATP failure, Pressure Leak 2019-11-11 00678108 (Rep): Medtronic received information regarding a guidance system being used during service. It was reported that there was a an ATP failure with the system. The system also had no arm, slow pneumatic leak, no power cord, and no BNC cable. No patient involvement. 2019-11-14 IUD x12 (Other): No new information. 2020-02-25 IUD x2 (Other): No new information. 2020-06-12 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-07-12 IUD x2 (Other): No new information. 2022-02-19 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2019-11-11 00678108 (Rep): Medtronic received information regarding a guidance system being used during service. It was reported that there was a an ATP failure with the system. The system also had no arm, slow pneumatic leak, no power cord, and no BNC cable. No patient involvement. 2019-11-14 IUD x12 (Other): No new information. 2020-02-25 IUD x2 (Other): No new information. 2020-06-12 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-07-12 IUD x2 (Other): No new information. 2022-02-19 IUD (Other): No new information. | xse atp failure pressure leak service atp failure arm slow pneumatic leak power cord bnc cable iud new iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV3256 | NAV3264 | NAV3249 | NAV4095 | NAV2054 | NAV3256NAV3264NAV3249NAV4095NAV2054 | False
951 | 703181421 | XSE encoder installed incorrect/iron man | 2019-05-30T19:27:51.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | MEC0625-01 | IRON MAN. | XSE encoder installed incorrect/iron man 2019-May-23 SCL1901337 msg (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used in outside of a procedure. It was reported that the front shoulder encoder was installed incorrectly. The iron man was missing upon delivery. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. IRON MAN. | 2019-May-23 SCL1901337 msg (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used in outside of a procedure. It was reported that the front shoulder encoder was installed incorrectly. The iron man was missing upon delivery. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. | xse encoder installed incorrectiron man msg outside procedure shoulder encoder installed incorrectly iron man missing delivery additional complications anticipated iron man | NAV2054 | NAV2053 | NAV3249 | NAV3299 | NAV3295 | NAV2054NAV2053NAV3249NAV3299NAV3295 | True
7899 | 704302709 | XSE - Unexpected Restart | 2021-03-25T00:04:03.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Unexpected Restart 2021-Mar-24 Salesforce C00223210 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system unexpectedly restarted when using navigation during the procedure. When navigating the interbody on the posterior prep instruments, the system did not show the interbody CAD model at the end of the inserter. The system then became unresponsive, the screen went black, a soft reboot started automatically and the screen went back to the home page where a patient is selected. The screen then went back to the operate mode where they were previously without the representative selecting anything. There were no further issues and the next case was completed without issue. There was no patient harm and the procedure was delayed less than an hour. 2021-04-07 IUD, e1, e2, e3, e4, e5, e6 (Rep): No new information. 2021-06-23 e7, e8 (Rep): No new information. 2021-06-29 e9 (Other): No new information. 2021-08-29 e10 (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Mar-24 Salesforce C00223210 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system unexpectedly restarted when using navigation during the procedure. When navigating the interbody on the posterior prep instruments, the system did not show the interbody CAD model at the end of the inserter. The system then became unresponsive, the screen went black, a soft reboot started automatically and the screen went back to the home page where a patient is selected. The screen then went back to the operate mode where they were previously without the representative selecting anything. There were no further issues and the next case was completed without issue. There was no patient harm and the procedure was delayed less than an hour. 2021-04-07 IUD, e1, e2, e3, e4, e5, e6 (Rep): No new information. 2021-06-23 e7, e8 (Rep): No new information. 2021-06-29 e9 (Other): No new information. 2021-08-29 e10 (Other): No new information. | xse unexpected restart salesforce spinal p rocedure unexpectedly restarted using navigation procedure navigating interbody posterior prep instruments did interbody cad model end inserter unresponsive screen went black soft reboot started automatically screen went home page selected screen went operate mode previously selecting issues case completed issue procedure delayed hour iud new new new new mazorx sw installation kit | NAV3248 | NAV3246 | NAV3247 | NAV2047 | NAV3235 | NAV3248NAV3246NAV3247NAV2047NAV3235 | True
1547 | 703205282 | X - Software issue | 2019-06-13T13:49:41.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Software issue 2018-03-12 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that shoulder calibration was checked. Advance accuracy test and surgical arm setup were completed. The schantz pin was placed into the right PSIS. 3Define scan was complete with good quality. Draw spine was successful and labeled off of L5. AP and oblique images were taken. Trajectories were sent to left L4, left L5, left S1, right L4, right L5 and right S1. Left sided screws were executed accurately and according to plan. Right L4 was checked under lateral fluoroscopy and a K-Wire placed. Right L5 was checked under lateral fluoroscopy and showed that the dilator was sitting low. After several attempts the surgeon decided to move on to right S1 where a wire was placed under lateral fluoroscopy. An AP fluoroscopic X-ray was taken to confirm wire placement and revealed that the right L4 wire had deviated laterally. Both right L4 and right L5 were then placed manually with a needle under fluoroscopy. There was no patient harm. MAZOR X SYSTEM TPL0059 | 2018-03-12 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that shoulder calibration was checked. Advance accuracy test and surgical arm setup were completed. The schantz pin was placed into the right PSIS. 3Define scan was complete with good quality. Draw spine was successful and labeled off of L5. AP and oblique images were taken. Trajectories were sent to left L4, left L5, left S1, right L4, right L5 and right S1. Left sided screws were executed accurately and according to plan. Right L4 was checked under lateral fluoroscopy and a K-Wire placed. Right L5 was checked under lateral fluoroscopy and showed that the dilator was sitting low. After several attempts the surgeon decided to move on to right S1 where a wire was placed under lateral fluoroscopy. An AP fluoroscopic X-ray was taken to confirm wire placement and revealed that the right L4 wire had deviated laterally. Both right L4 and right L5 were then placed manually with a needle under fluoroscopy. There was no patient harm. | x software issue ccr spinal procedure shoulder calibration checked advance accuracy test surgical arm setup completed schantz pin placed right psis scan complete good quality draw spine successful labeled ap oblique images taken trajectories sent left left left right right right left sided screws executed accurately according plan right checked lateral fluoroscopy kwire placed right checked lateral fluoroscopy showed dilator sitting low attempts surgeon decided right wire placed lateral fluoroscopy ap fluoroscopic xray taken confirm wire placement revealed right wire deviated laterally right right placed manually needle fluoroscopy mazor x | NAV4091 | NAV3251 | NAV3250 | NAV3254 | NAV4095 | NAV4091NAV3251NAV3250NAV3254NAV4095 | True
9564 | 704673052 | XSE ¦ Over Current Error Joint 6 | 2021-10-20T18:19:57.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 30 | UK |  | XSE ¦ Over Current Error Joint 6 2021-Oct-20 Salesforce C00236542 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when trying to move the surgical arm to the next position, the healthcare professionals got an error for Over current at joint 6 and were unable to send to the final destination. They resent the surgical arm and it went all the say. This occurred a few times throughout the case each time they resent and it worked. The issues extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-10-29 IUD (Other): No new information. | 2021-Oct-20 Salesforce C00236542 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when trying to move the surgical arm to the next position, the healthcare professionals got an error for Over current at joint 6 and were unable to send to the final destination. They resent the surgical arm and it went all the say. This occurred a few times throughout the case each time they resent and it worked. The issues extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-10-29 IUD (Other): No new information. | xse current error joint salesforce spinal procedure trying surgical arm position professionals got error current joint unable send final destination resent surgical arm went say occurred times case time resent worked issues extended surgical time hour impact outcome iud new | NAV4095 | NAV4101 | NAV4098 | NAV3249 | NAV3244 | NAV4095NAV4101NAV4098NAV3249NAV3244 | True
9112 | 704556022 | XSE ¦ Mazor sensor connection lost | 2021-08-13T09:40:08.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE ¦ Mazor sensor connection lost 2021-Aug-12 Salesforce C00232891 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the first time after the V5 installation and starting the system, then doing the surgical arm setup there was an error message "sensor connection failed" popped up and the x-eye camera was red. When going into the camera settings the same error message was shown and everything was blank. The x-eye camera was not working. A full dry run was done and there was no issue during the dry run. Several hard reboots were done, cable connections were all checked but still same error message occurred. No patient was present. 2021-09-16 03:01 email 2021-09-16 03:01 (Rep): The cause of the error messages was due to the cable not being plugged in. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Aug-12 Salesforce C00232891 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the first time after the V5 installation and starting the system, then doing the surgical arm setup there was an error message "sensor connection failed" popped up and the x-eye camera was red. When going into the camera settings the same error message was shown and everything was blank. The x-eye camera was not working. A full dry run was done and there was no issue during the dry run. Several hard reboots were done, cable connections were all checked but still same error message occurred. No patient was present. 2021-09-16 03:01 email 2021-09-16 03:01 (Rep): The cause of the error messages was due to the cable not being plugged in. | xse mazor sensor connection lost salesforce outside proc edure time installation starting doing surgical arm setup error message sensor connection failed popped xeye camera red going camera settings error message shown blank xeye camera working dry run issue dry run hard reboots cable connections checked error message occurred present email cause error messages cable plugged kit svc surgical arm refurb | NAV4110 | NAV3249 | NAV2047 | NAV3256 | NAV2054 | NAV4110NAV3249NAV2047NAV3256NAV2054 | True
3435 | 703347183 | X - schanz bond screw stripped | 2019-09-05T17:46:07.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC4530-01 | SLOTTED BALL ARM TARGET SCREW | X - schanz bond screw stripped 2019-Aug-30 SCL1902519 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the schanz bond screw was stripped and needed to be replaced. There was no patient harm or additional complications reported or anticipated as a result of the event. SLOTTED BALL ARM TARGET SCREW | 2019-Aug-30 SCL1902519 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the schanz bond screw was stripped and needed to be replaced. There was no patient harm or additional complications reported or anticipated as a result of the event. | x schanz bond screw stripped spinal procedure schanz bond screw stripped needed replaced additional complications anticipated slotted ball arm target screw | NAV3295 | NAV3299 | NAV4090 | NAV4100 | NAV3251 | NAV3295NAV3299NAV4090NAV4100NAV3251 | True
7990 | 704325810 | XSE - Noisy X Eye Camera Motor | 2021-04-07T00:10:28.000000+0000 | MazorX Stealth Edition | NAV4109 | LOOSE | 20 | ASM0386-02S | KIT SVC ASM0386-02S X-EYE MOTOR ASSY | XSE - Noisy X Eye Camera Motor 2021-Apr-06 Salesforce C00224127 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the motor of the X-Eye camera was noisy. The camera was working normally. There was no patient involvement. 2021-04-13 IUD x3 (Other): No new information. 2021-04-28 IUD (Other): No new information. 2021-05-25 email 2021-05-25 12:01, email 2021-05-25 12:45 (Rep): Additional information received from a manufacturer representative reported the issue was a defect in the X-Eye camera motor. The motor was replaced to resolve the issue. 2021-06-16 email 2021-06-16 17:01 (Rep): No new information. KIT SVC ASM0386-02S X-EYE MOTOR ASSY | 2021-Apr-06 Salesforce C00224127 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the motor of the X-Eye camera was noisy. The camera was working normally. There was no patient involvement. 2021-04-13 IUD x3 (Other): No new information. 2021-04-28 IUD (Other): No new information. 2021-05-25 email 2021-05-25 12:01, email 2021-05-25 12:45 (Rep): Additional information received from a manufacturer representative reported the issue was a defect in the X-Eye camera motor. The motor was replaced to resolve the issue. 2021-06-16 email 2021-06-16 17:01 (Rep): No new information. | xse noisy x eye camera motor salesforce outside procedure motor xeye camera noisy camera working normally iud new iud new email email additional issue defect xeye camera motor motor replaced resolve issue email new kit svc xeye motor assy | NAV4110 | NAV3299 | NAV4109 | NAV2054 | NAV3254 | NAV4110NAV3299NAV4109NAV2054NAV3254 | True
5350 | 703657321 | XSE - L3/L4 All Wires Deviated Left | 2020-03-09T19:04:13.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - L3/L4 All Wires Deviated Left 2020-03-09 00685920, CCR (rep, hcp): Medtronic received information regarding a guidance system being used for a L3-L4 Oblique Lumbar Interbody Fusion (OLIF) procedure. It was reported that the site had placed wires in the patient. When verifying the registration and accuracy, the left side wires had shifted laterally and the right side wires had shifted medial. The surgeon removed all the wires and another registration was performed. The wires were noted to be fine after verifying. All 4 screws were then re-instrumented accurately and verified with intra-operative fluoroscopy. It was noted that the only difference during the registrations was that the elbow was low in the first registration for the oblique image and the elbow was high for the second registration oblique image. The procedure was delayed by less than 1 hour and there was no impact to patient outcome. 2020-03-10 email 2020-03-10 20:15 (Rep): Additional information received from a manufacturer representative reported that the cause of the deviation was not determined. The deviation was between 3.5 and 10 mm. 2020-03-17 e1, e2 (Rep): Additional information received from a manufacturer representative reported that the surgeon was on the left side of the patient when placing the screws for the left side and on the right side of the patient for the right side screws. The platform was checked and was rigid. A flat jackson table with a wilson frame was used. The patient had a BMI of over 40. No changes were made to the bed after registration and no platform or patient shifts were noted. 2020-03-18 IUD x2 (Other): Additional information received from a manufacturer representative. Patient information was updated MAZOR X SYSTEM TPL0059 | 2020-03-09 00685920, CCR (rep, hcp): Medtronic received information regarding a guidance system being used for a L3-L4 Oblique Lumbar Interbody Fusion (OLIF) procedure. It was reported that the site had placed wires in the patient. When verifying the registration and accuracy, the left side wires had shifted laterally and the right side wires had shifted medial. The surgeon removed all the wires and another registration was performed. The wires were noted to be fine after verifying. All 4 screws were then re-instrumented accurately and verified with intra-operative fluoroscopy. It was noted that the only difference during the registrations was that the elbow was low in the first registration for the oblique image and the elbow was high for the second registration oblique image. The procedure was delayed by less than 1 hour and there was no impact to patient outcome. 2020-03-10 email 2020-03-10 20:15 (Rep): Additional information received from a manufacturer representative reported that the cause of the deviation was not determined. The deviation was between 3.5 and 10 mm. 2020-03-17 e1, e2 (Rep): Additional information received from a manufacturer representative reported that the surgeon was on the left side of the patient when placing the screws for the left side and on the right side of the patient for the right side screws. The platform was checked and was rigid. A flat jackson table with a wilson frame was used. The patient had a BMI of over 40. No changes were made to the bed after registration and no platform or patient shifts were noted. 2020-03-18 IUD x2 (Other): Additional information received from a manufacturer representative. Patient information was updated | xse wires deviated left ccr oblique lumbar interbody fusion procedure site placed wires verifying registration accuracy left wires shifted laterally right wires shifted medial surgeon removed wires registration performed wires noted fine verifying screws reinstrumented accurately verified intraoperative fluoroscopy noted difference registrations elbow low registration oblique image elbow high second registration oblique image procedure delayed hour impact outcome email additional cause deviation determined deviation mm additional surgeon left placing screws left right right screws platform checked rigid flat jackson table wilson frame bmi changes bed registration platform shifts noted iud additional updated mazor x | NAV5000 | NAV3237 | NAV3258 | NAV3242 | NAV2051 | NAV5000NAV3237NAV3258NAV3242NAV2051 | True
6327 | 703912955 | XSE - Abnormal Monitor Behavior | 2020-08-20T01:39:05.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | UNK_MXSE_SW | UNKNOWN MAZOR X SE SW | XSE - Abnormal Monitor Behavior 2020-08-19 00694921 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative observed abnormal behavior of the system during the fusion procedure. After registration and identification of the vertebral levels, the main monitor and surgeon screen went black with a spinning circle on the screen. The images came back up spontaneously after 30 seconds. Later in the procedure, the monitor moved to the highest position without anyone touching the software or system. The procedure was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-24 IUD x2 (Other): No new information. UNKNOWN MAZOR X SE SW | 2020-08-19 00694921 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative observed abnormal behavior of the system during the fusion procedure. After registration and identification of the vertebral levels, the main monitor and surgeon screen went black with a spinning circle on the screen. The images came back up spontaneously after 30 seconds. Later in the procedure, the monitor moved to the highest position without anyone touching the software or system. The procedure was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-24 IUD x2 (Other): No new information. | xse abnormal monitor behavior spinal procedure observed abnormal behavior fusion procedure registration identification vertebral levels main monitor surgeon screen went black spinning circle screen images came spontaneously seconds later procedure monitor moved highest position touching software procedure able completed procedure delayed hour iud new unknown mazor x se sw | NAV3248 | NAV3261 | NAV3246 | NAV4112 | NAV5000 | NAV3248NAV3261NAV3246NAV4112NAV5000 | True
8380 | 704404118 | XSE - Unable to Register | 2021-05-19T23:54:39.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 20 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - Unable to Register 2021-May-19 Salesforce C00226716, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation using screws and rods. It was reported that there were registration difficulties during a L3-S1 case. A pre-op CT was done with no interbodies present. The surgeon decided to perform multiple ALIFS with titanium cages with screws to hold them in place. Three cages were placed from L3-L4, L4-L5 and L5-S1. One second of C-arm images were acquired and registration was done for L3-S1; however, when approving registration, there was a shift in L3 and L4. The surgeon operated on L5 and S1 and placed screws. Using the same images, registration was tried again using one-dot, semi-manual and manual methods. L3 was finally able to be registered without a shift, but L4 could not be registered without a shift. The surgeon decided to abort the use of the guidance system and place the remaining screws freehand. In all of the registrations, the representative got to the approve step, but they were seeing shifts. There was no patient harm and the procedure was delayed less than an hour. 2021-05-26 IUD x4 (Other): No new information. 2021-06-24 IUD x4 (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2021-May-19 Salesforce C00226716, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation using screws and rods. It was reported that there were registration difficulties during a L3-S1 case. A pre-op CT was done with no interbodies present. The surgeon decided to perform multiple ALIFS with titanium cages with screws to hold them in place. Three cages were placed from L3-L4, L4-L5 and L5-S1. One second of C-arm images were acquired and registration was done for L3-S1; however, when approving registration, there was a shift in L3 and L4. The surgeon operated on L5 and S1 and placed screws. Using the same images, registration was tried again using one-dot, semi-manual and manual methods. L3 was finally able to be registered without a shift, but L4 could not be registered without a shift. The surgeon decided to abort the use of the guidance system and place the remaining screws freehand. In all of the registrations, the representative got to the approve step, but they were seeing shifts. There was no patient harm and the procedure was delayed less than an hour. 2021-05-26 IUD x4 (Other): No new information. 2021-06-24 IUD x4 (Other): No new information. | xse unable register salesforce ccr spinal procedure treat degenerative discs fixation using screws rods registration difficulties case preop ct interbodies present surgeon decided perform multiple alifs titanium cages screws hold place cages placed second carm images acquired registration approving registration shift surgeon operated placed screws using images registration tried using onedot semimanual manual methods finally able registered shift registered shift surgeon decided abort use place remaining screws freehand registrations got approve step seeing shifts procedure delayed hour iud new iud new mazor x pc station refurbished | NAV4112 | NAV4108 | NAV5000 | NAV4098 | NAV3242 | NAV4112NAV4108NAV5000NAV4098NAV3242 | True
6477 | 703946896 | X - SRA Didn't Reach Destination Case 1 | 2020-09-09T13:50:36.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - SRA Didn't Reach Destination Case 1 2020-09-08 00696076, 00696081, IUD x2 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after starting the 3Define scan, the surgical arm started to move, but stopped before or when reaching the first point. An error was displayed, which indicated "3Define Scan Failed. Rescan or Delete the Set Default Work Volume." Another 3Define scan was attempted, but the same thing occurred. The surgical arm was sent to the clear up position and a third scan was started. The 3Define scan was able to be successfully completed. At the end of the procedure, there was a pressure issue that caused the manipulator to tilt to the right. This occurred after successfully connecting the manipulator into the positioner and unmounting the surgical system from the bed. The representative had to physically support the manipulator to make sure it did not drop over the patient. After taking the system out, the issue was resolved and the surgical system was able to be stored. There was no patient harm and the procedure was delayed less than an hour. 2021-Feb-10 IUDx2 (other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-09-08 00696076, 00696081, IUD x2 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after starting the 3Define scan, the surgical arm started to move, but stopped before or when reaching the first point. An error was displayed, which indicated "3Define Scan Failed. Rescan or Delete the Set Default Work Volume." Another 3Define scan was attempted, but the same thing occurred. The surgical arm was sent to the clear up position and a third scan was started. The 3Define scan was able to be successfully completed. At the end of the procedure, there was a pressure issue that caused the manipulator to tilt to the right. This occurred after successfully connecting the manipulator into the positioner and unmounting the surgical system from the bed. The representative had to physically support the manipulator to make sure it did not drop over the patient. After taking the system out, the issue was resolved and the surgical system was able to be stored. There was no patient harm and the procedure was delayed less than an hour. 2021-Feb-10 IUDx2 (other): No new information. | x sra didnt reach destination case iud spinal procedure starting scan surgical arm started stopped reaching point error displayed indicated scan failed rescan delete set default work volume scan attempted thing occurred surgical arm sent clear position scan started scan able successfully completed end procedure pressure issue caused manipulator tilt right occurred successfully connecting manipulator positioner unmounting surgical bed physically support manipulator make sure did drop taking issue resolved surgical able stored procedure delayed hour new unk mazor x camera | NAV3244 | NAV3260 | NAV3264 | NAV3239 | NAV3246 | NAV3244NAV3260NAV3264NAV3239NAV3246 | False
357 | 703086078 | XSE - Joint shift errors | 2019-04-02T16:00:26.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Joint shift errors 2019-Mar-28 SCL1900772 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were repeated Joint shift errors in J5. No movement in the encoder readings when the joint moves. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Mar-28 SCL1900772 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were repeated Joint shift errors in J5. No movement in the encoder readings when the joint moves. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | xse joint shift errors spinal procedure repeated joint shift errors movement encoder readings joint moves additional complications anticipated surgical arm | NAV4098 | NAV2054 | NAV5000 | NAV4097 | NAV4095 | NAV4098NAV2054NAV5000NAV4097NAV4095 | True
7402 | 704195992 | XSE ¦ Lost Instrument | 2021-01-26T21:37:14.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MAS2310-04S | KIT SVC MAS2310-04S TARGET EXTENDER | XSE ¦ Lost Instrument 2021-01-26 00705211 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the target extender was lost. No patient was present. 2021-01-26 IUD x4 (Other): No new information. KIT SVC MAS2310-04S TARGET EXTENDER | 2021-01-26 00705211 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the target extender was lost. No patient was present. 2021-01-26 IUD x4 (Other): No new information. | xse lost instrument accessory outside procedure target extender lost present iud new kit svc target extender | NAV2053 | NAV3299 | NAV3295 | NAV2054 | NAV2047 | NAV2053NAV3299NAV3295NAV2054NAV2047 | True
610 | 703155429 | X - Registration difficulty | 2019-05-15T18:31:39.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Registration difficulty 2019-May-14 SCL1901235 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the rep took multiple AP & OBL images and labeled the vertebral bodies. When trying to place the reference marker on the back of the body it kept showing one level up. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-May-14 SCL1901235 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the rep took multiple AP & OBL images and labeled the vertebral bodies. When trying to place the reference marker on the back of the body it kept showing one level up. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x registration difficulty spinal procedure took multiple ap obl images labeled vertebral bodies trying place reference marker body kept showing level additional complications anticipated mazor x software kit | NAV3242 | NAV3246 | NAV4112 | NAV4091 | NAV5000 | NAV3242NAV3246NAV4112NAV4091NAV5000 | True
5942 | 703842706 | XSE - Unable to Register Patient | 2020-07-08T23:10:52.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Unable to Register Patient 2020-07-08 00692302 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were multiple issues during an L4-S2 open PSF. The reference frame and snapshot tracker were seen by the navigation, but they were unable to acquire images. The manufacturer representative was unable to highlight the acquire button. A soft restart of the system was done and the they were able to proceed, but they were unable to leave oblique. The work volume was cleared and a generic work volume was used to proceed with the case, but there was then issues with registration. The surgeon decided not to use the guidance system and to use navigation instead. There was no patient harm and the procedure was delayed over an hour. 2020-07-09 email 2020-07-09 16:45 (Rep): Additional information received from a manufacturer representative reported that a pediatrics spondylolisthesis case was done. The screws for L4-S2 were planned by the surgeon and skive potentials were noted. A Schanz pin was placed in the right PSIS and a clamp was placed on the spinous process of L4 to mount the surgical system to the patient. During the procedure, exposure of the patient's vertebrae was completed and the patient reference frame was placed on the L3 spinous process. The 3Define scan was completed and a snapshot was attempted, but the button could not be highlighted. A soft restart was done, but the 3Define scan was lost so it was redone. AP and oblique images were taken for registration. The first registration attempt failed and the software indicated that a new AP image was needed. The surgical arm was in the oblique position and when attempting to resend the arm to the AP position, the 8798 Unable to Reach Destination error code was displayed. Multiple attempts were made to send the surgical arm to the clear up, clear right, and clear left positions, but the issue was not resolved. A soft restart of the system was done to resolve the issue. The 3Define scan was was deleted so a generic work volume was used to resolve the issue. A new AP image was acquired, but the second registration attempt failed for all levels. Multiple new images were taken and multiple manual registration attempts were made. The representative requested the retractors be removed, but registration failed with yellow indication for S1/S2. No shift was detected. The surgeon decided to proceed with verified levels and work back to L4/L5. New 3Define scan was required for a third time. The surgical arm was sent to the left L2 trajectory and navigation showed to be medial to plan. The representative suggested that the reference frame moved due to the retractors. S2 and S1 screws were placed. When the surgical arm was sent to S2 right, navigation showed the tools to be more anterior in navigation than demonstrated. The tools were shown to be 20 mm in depth when they were only on the surface of the tissue. Navigation was verified with the placing the passive planar on the bony anatomy. A new snapshot was taken, but the length of tools continued to be anterior on navigation. The representative went back to registration with the previous images, but registration failed for L4/L5. The superior retractor and spinous process clamp were removed and new images were taken, but registration failed. The cause of the registration issues were not determined. The representative attempted registration after the case and they were able to achieve registration for L3-S2, but the registration for L5 showed both trajectories as unreachable with only the right screw having a solution. S1 and S2 were able to be registered successfully and placed accurately according to plan. 2020-07-15 IUD x2 (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-07-08 00692302 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were multiple issues during an L4-S2 open PSF. The reference frame and snapshot tracker were seen by the navigation, but they were unable to acquire images. The manufacturer representative was unable to highlight the acquire button. A soft restart of the system was done and the they were able to proceed, but they were unable to leave oblique. The work volume was cleared and a generic work volume was used to proceed with the case, but there was then issues with registration. The surgeon decided not to use the guidance system and to use navigation instead. There was no patient harm and the procedure was delayed over an hour. 2020-07-09 email 2020-07-09 16:45 (Rep): Additional information received from a manufacturer representative reported that a pediatrics spondylolisthesis case was done. The screws for L4-S2 were planned by the surgeon and skive potentials were noted. A Schanz pin was placed in the right PSIS and a clamp was placed on the spinous process of L4 to mount the surgical system to the patient. During the procedure, exposure of the patient's vertebrae was completed and the patient reference frame was placed on the L3 spinous process. The 3Define scan was completed and a snapshot was attempted, but the button could not be highlighted. A soft restart was done, but the 3Define scan was lost so it was redone. AP and oblique images were taken for registration. The first registration attempt failed and the software indicated that a new AP image was needed. The surgical arm was in the oblique position and when attempting to resend the arm to the AP position, the 8798 Unable to Reach Destination error code was displayed. Multiple attempts were made to send the surgical arm to the clear up, clear right, and clear left positions, but the issue was not resolved. A soft restart of the system was done to resolve the issue. The 3Define scan was was deleted so a generic work volume was used to resolve the issue. A new AP image was acquired, but the second registration attempt failed for all levels. Multiple new images were taken and multiple manual registration attempts were made. The representative requested the retractors be removed, but registration failed with yellow indication for S1/S2. No shift was detected. The surgeon decided to proceed with verified levels and work back to L4/L5. New 3Define scan was required for a third time. The surgical arm was sent to the left L2 trajectory and navigation showed to be medial to plan. The representative suggested that the reference frame moved due to the retractors. S2 and S1 screws were placed. When the surgical arm was sent to S2 right, navigation showed the tools to be more anterior in navigation than demonstrated. The tools were shown to be 20 mm in depth when they were only on the surface of the tissue. Navigation was verified with the placing the passive planar on the bony anatomy. A new snapshot was taken, but the length of tools continued to be anterior on navigation. The representative went back to registration with the previous images, but registration failed for L4/L5. The superior retractor and spinous process clamp were removed and new images were taken, but registration failed. The cause of the registration issues were not determined. The representative attempted registration after the case and they were able to achieve registration for L3-S2, but the registration for L5 showed both trajectories as unreachable with only the right screw having a solution. S1 and S2 were able to be registered successfully and placed accurately according to plan. 2020-07-15 IUD x2 (Other): No new information. | xse unable register spinal procedure multiple issues open psf reference frame snapshot tracker seen navigation unable acquire images unable highlight acquire button soft restart able proceed unable leave oblique work volume cleared generic work volume proceed case issues registration surgeon decided use use navigation instead procedure delayed hour email additional pediatrics spondylolisthesis case screws planned surgeon skive potentials noted schanz pin placed right psis clamp placed spinous process mount surgical procedure exposure patients vertebrae completed reference frame placed spinous process scan completed snapshot attempted button highlighted soft restart scan lost redone ap oblique images taken registration registration attempt failed software indicated new ap image needed surgical arm oblique position attempting resend arm ap position unable reach destination error code displayed multiple attempts send surgical arm clear clear right clear left positions issue resolved soft restart resolve issue scan deleted generic work volume resolve issue new ap image acquired second registration attempt failed levels multiple new images taken multiple manual registration attempts requested retractors removed registration failed yellow indication shift detected surgeon decided proceed verified levels work new scan required time surgical arm sent left trajectory navigation showed medial plan suggested reference frame moved retractors screws placed surgical arm sent right navigation showed tools anterior navigation demonstrated tools shown mm depth surface tissue navigation verified placing passive planar bony anatomy new snapshot taken length tools continued anterior navigation went registration previous images registration failed superior retractor spinous process clamp removed new images taken registration failed cause registration issues determined attempted registration case able achieve registration registration showed trajectories unreachable right screw having solution able registered successfully placed accurately according plan iud new mazor x | NAV3246 | NAV5001 | NAV4095 | NAV4112 | NAV3242 | NAV3246NAV5001NAV4095NAV4112NAV3242 | True
6501 | 703957060 | XSE - Incorrect Firmware | 2020-09-15T11:56:25.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - Incorrect Firmware 2020-09-15 00696609, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the incorrect firmware was on the system during setup. It was noted that the 3Define camera firmware was not 2.60.0.0 and it was a bunch of negative numbers. The manufacturer representative walked through reinstalling the firmware. The representative was going to use a generic workflow if they were not able to re-install the firmware. Additional information received reported that the correct firmware version for the 3Define camera was able to be installed. After the installation, the camera started going in the wrong direction. The representative set the camera to go up, but it moved in the wrong direction. The camera started left and only went left until an error was displayed. The representative tried to resend the camera multiple times, but a Error 3716 (Failed to Settle Under Max Tolerance) was displayed. A soft restart of the system was done, but a error 4476 "Not enough suface volume" was displayed. The camera was still going in the wrong direction. A hard restart of the system was done. The firmware was still correct. A shoulder calibration was done, but the system was still going in the wrong direction even though all testing was passed. The shoulder was reset in the super user settings and another hard restart was done, but the screen was blue. The system was restarted again. A three point accuracy test was then passed, but the 3Define scan showed a surface volume error. Another hard reboot of the system was done and the 3Define camera was working. There was no patient involvement.. 2020-09-29 IUD (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-09-15 00696609, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the incorrect firmware was on the system during setup. It was noted that the 3Define camera firmware was not 2.60.0.0 and it was a bunch of negative numbers. The manufacturer representative walked through reinstalling the firmware. The representative was going to use a generic workflow if they were not able to re-install the firmware. Additional information received reported that the correct firmware version for the 3Define camera was able to be installed. After the installation, the camera started going in the wrong direction. The representative set the camera to go up, but it moved in the wrong direction. The camera started left and only went left until an error was displayed. The representative tried to resend the camera multiple times, but a Error 3716 (Failed to Settle Under Max Tolerance) was displayed. A soft restart of the system was done, but a error 4476 "Not enough suface volume" was displayed. The camera was still going in the wrong direction. A hard restart of the system was done. The firmware was still correct. A shoulder calibration was done, but the system was still going in the wrong direction even though all testing was passed. The shoulder was reset in the super user settings and another hard restart was done, but the screen was blue. The system was restarted again. A three point accuracy test was then passed, but the 3Define scan showed a surface volume error. Another hard reboot of the system was done and the 3Define camera was working. There was no patient involvement.. 2020-09-29 IUD (Other): No new information. | xse incorrect firmware iud outside procedure incorrect firmware setup noted camera firmware bunch negative numbers walked reinstalling firmware going use generic workflow able reinstall firmware additional correct firmware version camera able installed installation camera started going wrong direction set camera moved wrong direction camera started left went left error displayed tried resend camera multiple times error displayed soft restart error suface volume displayed camera going wrong direction hard restart firmware correct shoulder calibration going wrong direction testing passed shoulder reset super user settings hard restart screen blue restarted point accuracy test passed scan showed surface volume error hard reboot camera working iud new unk mazor x camera | NAV3244 | NAV3248 | NAV2047 | NAV3246 | NAV1013 | NAV3244NAV3248NAV2047NAV3246NAV1013 | True
7182 | 704134753 | XSE - Navigation Showing Off | 2020-12-19T01:17:21.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 20 | M728894B252 | NAVIGATION SOFTWARE MODULE 1.2 | XSE - Navigation Showing Off 2020-12-18 00703082, e1, e2, e3, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the navigation for the dilator and powered drill appeared to be off both lateral and medial by a few millimeters during a L2-S2 case. The tap and screws were accurate. The manufacturer representative believed the only difference was the larger Navlock trackers were used with the tap and screws and the smaller trackers were used with the powered drill. Navigation depth looked accurate on all instruments. The surgeons technique was to place the dilator, powered drill, tap and then place the screw. The representative verified that the camera and tools had a clear line of sight. All of the screws were accurately placed and confirmed to be accurate with an O-arm spin after the case. There was no patient harm and the procedure was delayed less than an hour. 2021-Feb-01 e4, e5 (rep): The trajectories were off between 3.5-10 mm. After the 5.0 upgrade, this patient is no longer in the software and the export was not available. 2021-03-09 IUD (Other): No new information. NAVIGATION SOFTWARE MODULE 1.2 | 2020-12-18 00703082, e1, e2, e3, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the navigation for the dilator and powered drill appeared to be off both lateral and medial by a few millimeters during a L2-S2 case. The tap and screws were accurate. The manufacturer representative believed the only difference was the larger Navlock trackers were used with the tap and screws and the smaller trackers were used with the powered drill. Navigation depth looked accurate on all instruments. The surgeons technique was to place the dilator, powered drill, tap and then place the screw. The representative verified that the camera and tools had a clear line of sight. All of the screws were accurately placed and confirmed to be accurate with an O-arm spin after the case. There was no patient harm and the procedure was delayed less than an hour. 2021-Feb-01 e4, e5 (rep): The trajectories were off between 3.5-10 mm. After the 5.0 upgrade, this patient is no longer in the software and the export was not available. 2021-03-09 IUD (Other): No new information. | xse navigation showing iud spinal procedure navigation dilator powered drill appeared lateral medial millimeters case tap screws accurate believed difference larger navlock trackers tap screws smaller trackers powered drill navigation depth looked accurate instruments surgeons technique place dilator powered drill tap place screw verified camera tools clear line sight screws accurately placed confirmed accurate oarm spin case procedure delayed hour trajectories mm upgrade longer software export available iud new navigation software module | NAV5001 | NAV1012 | NAV3246 | NAV5002 | NAV3235 | NAV5001NAV1012NAV3246NAV5002NAV3235 | True
6071 | 703864093 | XSE - Failed Advanced Accuracy Test | 2020-07-21T23:17:51.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Failed Advanced Accuracy Test 2020-07-21 00693133 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed advanced accuracy test at point 8. There was no patient involvement. 2020-07-24 IUD x3 (Other): No new information. 2020-08-11 IUD x2 (Other): No new information. 2020-09-09 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-07-21 00693133 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed advanced accuracy test at point 8. There was no patient involvement. 2020-07-24 IUD x3 (Other): No new information. 2020-08-11 IUD x2 (Other): No new information. 2020-09-09 IUD (Other): No new information. | xse failed advanced accuracy test outside procedure surgical arm failed advanced accuracy test point iud new iud new iud new kit svc surgical arm refurb | NAV4093 | NAV3249 | NAV2054 | NAV4095 | NAV4092 | NAV4093NAV3249NAV2054NAV4095NAV4092 | True
959 | 703181477 | X - Surgical arm/Software issues | 2019-05-30T19:40:38.000000+0000 | MazorX | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical arm/Software issues It was reported that Matt Walters called the CSC to report a handful of different error messages. While the system was still on the bed he was using the workstaion and noticed a J3/J5 shift error pop up on the screen. When he went to look at the surgical arm he noticed the Emergency stop was engaged and the arm was illuminated Red. He alledges that the arm was then pushed manually out of the way of the patient. During this he also received CANBUS errors in Nodes 1-8 and RBT calibration file error and 3D scan not found in RBT driver. Once a hard restart was performed the E-stop was still showing as engaged even though it wasn't and now the arm would not illuminate at all and the system did not recognize it. SURGICAL ARM | It was reported that Matt Walters called the CSC to report a handful of different error messages. While the system was still on the bed he was using the workstaion and noticed a J3/J5 shift error pop up on the screen. When he went to look at the surgical arm he noticed the Emergency stop was engaged and the arm was illuminated Red. He alledges that the arm was then pushed manually out of the way of the patient. During this he also received CANBUS errors in Nodes 1-8 and RBT calibration file error and 3D scan not found in RBT driver. Once a hard restart was performed the E-stop was still showing as engaged even though it wasn't and now the arm would not illuminate at all and the system did not recognize it. | x surgical armsoftware issues matt walters called csc report handful different error messages bed using workstaion noticed shift error pop screen went look surgical arm noticed emergency stop engaged arm illuminated red alledges arm pushed manually way canbus errors nodes rbt calibration file error scan rbt driver hard restart performed estop showing engaged wasnt arm illuminate did recognize surgical arm | NAV2047 | NAV3249 | NAV4098 | NAV4101 | NAV3246 | NAV2047NAV3249NAV4098NAV4101NAV3246 | True
1612 | 703207913 | X-Stylus Marker/Scalpel Handle Issues | 2019-06-14T17:49:41.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2176-03 | STYLUS MARKER | X-Stylus Marker/Scalpel Handle Issues The following issues with the instrument tray are causing issues with the doctors particularly Dr. Villalobos who is still not using the X 100% of the time do to these issues. 1.The scalpel handle is offset and hard to place in the arm, this has caused the surgeon to struggle getting it in and turning it the way he wants. 2. The stylus is an older model and is not picked up by the camera without major manipulations of the lights and the stylus itself, even then the system has a hard time picking it up. This has caused Drs. Villalobos and Ramirez to abort using draw spine until it is corrected. STYLUS MARKER | The following issues with the instrument tray are causing issues with the doctors particularly Dr. Villalobos who is still not using the X 100% of the time do to these issues. 1.The scalpel handle is offset and hard to place in the arm, this has caused the surgeon to struggle getting it in and turning it the way he wants. 2. The stylus is an older model and is not picked up by the camera without major manipulations of the lights and the stylus itself, even then the system has a hard time picking it up. This has caused Drs. Villalobos and Ramirez to abort using draw spine until it is corrected. | xstylus markerscalpel handle issues following issues instrument tray causing issues doctors particularly dr villalobos using x time issues scalpel handle offset hard place arm caused surgeon struggle getting turning way wants stylus older model picked camera major manipulations lights stylus hard time picking caused drs villalobos ramirez abort using draw spine corrected stylus marker | NAV3299 | NAV3295 | NAV4095 | NAV3246 | NAV3249 | NAV3299NAV3295NAV4095NAV3246NAV3249 | True
5840 | 703818945 | XSE - Stripped Set Screw | 2020-06-24T15:58:53.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MEC4530-02S | KIT SVC MEC4530-02S BALL ARM TRGT SCREW | XSE - Stripped Set Screw 2020-06-24 00691416 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz bond set screw heads were stripped out of two of the screws. There was no patient involvement. 2020-07-01 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. 2020-11-05 email 2020-11-05 14:15 (Rep): Additional information received from a manufacturer representative reported that the parts were discarded. KIT SVC MEC4530-02S BALL ARM TRGT SCREW | 2020-06-24 00691416 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz bond set screw heads were stripped out of two of the screws. There was no patient involvement. 2020-07-01 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. 2020-11-05 email 2020-11-05 14:15 (Rep): Additional information received from a manufacturer representative reported that the parts were discarded. | xse stripped set screw outside procedure schanz bond set screw heads stripped screws iud new iud new email additional parts discarded kit svc ball arm trgt screw | NAV3299 | NAV2054 | NAV4090 | NAV4100 | NAV3295 | NAV3299NAV2054NAV4090NAV4100NAV3295 | True
4546 | 703550645 | X - 3Define Issues | 2020-01-08T13:36:48.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define Issues 2018-05-11 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that procedure was planned for fixation at L5-S1. The system was attached to the patient by a schanz pin in the patient's left iliac crest. The manufacturer representative present for the case was not able to get the 3Define scan to work. The arm was functioning and accurately sent to different positions but the 3Define scan never started or showed an image. A soft restart was performed and this resolved the issue. Registration was accomplished on the first attempt. All trajectories were placed accurately. No patient harm was reported and surgery was delayed less than an hour. UNK MAZOR X 3DEFINE CAMERA | 2018-05-11 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that procedure was planned for fixation at L5-S1. The system was attached to the patient by a schanz pin in the patient's left iliac crest. The manufacturer representative present for the case was not able to get the 3Define scan to work. The arm was functioning and accurately sent to different positions but the 3Define scan never started or showed an image. A soft restart was performed and this resolved the issue. Registration was accomplished on the first attempt. All trajectories were placed accurately. No patient harm was reported and surgery was delayed less than an hour. | x issues ccr spinal procedure procedure planned fixation attached schanz pin patients left iliac crest present case able scan work arm functioning accurately sent different positions scan started showed image soft restart performed resolved issue registration accomplished attempt trajectories placed accurately surgery delayed hour unk mazor x camera | NAV3244 | NAV3242 | NAV2047 | NAV3261 | NAV4096 | NAV3244NAV3242NAV2047NAV3261NAV4096 | True
365 | 703086496 | X - bone mount screw stripped | 2019-04-02T19:57:32.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2188-01 | BM BRIDGE | X - bone mount screw stripped 2019-Mar-29 SCL1900779 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Bone Mount Bridge plate where the bone mount bridge screws into was stripped. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. BM BRIDGE | 2019-Mar-29 SCL1900779 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Bone Mount Bridge plate where the bone mount bridge screws into was stripped. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x bone mount screw stripped spinal procedure bone mount bridge plate bone mount bridge screws stripped noted additional complications anticipated bm bridge | NAV3299 | NAV4100 | NAV3295 | NAV3298 | NAV2054 | NAV3299NAV4100NAV3295NAV3298NAV2054 | True
6814 | 704040200 | XSE - System Shut Down | 2020-10-28T22:39:55.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT0830 | USA REGIONAL KIT KIT0830 | XSE - System Shut Down 2020-10-28 00699776, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while in the middle of a procedure, the guidance system shut down and the screen went black. After a couple of minutes, the screen booted up to a white screen with a symbol. After a couple of minutes, the software started to the same point as before. The case was completed without issue. There was no patient harm and the procedure was delayed less than an hour.. Information unrelated to the event was included in a separate PE; See PE 704040203 for the surgeon screen. 2020-Nov-17 IUD (rep): No new information. 2020-11-19 email 2020-11-19 12:15 (Rep): No new information. USA REGIONAL KIT KIT0830 | 2020-10-28 00699776, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while in the middle of a procedure, the guidance system shut down and the screen went black. After a couple of minutes, the screen booted up to a white screen with a symbol. After a couple of minutes, the software started to the same point as before. The case was completed without issue. There was no patient harm and the procedure was delayed less than an hour.. Information unrelated to the event was included in a separate PE; See PE 704040203 for the surgeon screen. 2020-Nov-17 IUD (rep): No new information. 2020-11-19 email 2020-11-19 12:15 (Rep): No new information. | xse shut iud spinal procedure middle procedure shut screen went black couple minutes screen booted white screen symbol couple minutes software started point case completed issue procedure delayed hour unrelated included separate pe pe surgeon screen iud new email new usa regional kit | NAV3247 | NAV3246 | NAV3253 | NAV3248 | NAV3260 | NAV3247NAV3246NAV3253NAV3248NAV3260 | True
10439 | 704880642 | XSE - Incorrect Level Operated On | 2022-02-24T01:10:58.000000+0000 | MazorX Stealth Edition | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Incorrect Level Operated On 2022-Feb-23 Salesforce C00244127 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgeon operated on the wrong level during an open procedure with a fractured T10 vertebrae. The plan was to fuse T8-L1 while skipping T10 due to the fracture. Two segments were planned since the anatomy was too long. The first segment used a region of interest from T7-T11 with mount placed at T9. The plan was to place screws at T8 and T9. Registration was completed and matched to what they thought was T8 and T9. T7-T9 had green values, T10 had red values due to the fracture and T11 had yellow registration values. The surgeon matched what they thought was T8 in the software and began inserting screws, but they were actually at T7. The surgeon did not realize the error until they moved to register the second segment. The level tested fine with neuromonitoring after the screws were implanted. The second segment had a region of interest from T11-L1 with a mount placed at T12. The plan was to place screws at T11-L1. A k-wire was placed at T11 and a 2D lateral shot was taken. The image showed that they were two vertebrae above T11 without screws. The surgeon confirmed T11 anatomically and noted that they had placed screws at T7-T8 instead of T8-T9 as intended. The second segment was completed without issue. The surgeon then went back to T9, placed the mount, re-registered and placed the screw at that level. The fusion was extended by on vertebrae and included T7. There was no patient harm and the procedure was was delayed less than hour. 2022-03-01 IUD (Other): No new information. 2022-03-04 IUD x3, e1 (Rep): The right screws on T7 and T8 were shifted medially from the plan. 2022-03-15 IUD (Other): No new information. 2022-04-11 IUD (Other): No new information. 2022-04-13 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2022-Feb-23 Salesforce C00244127 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgeon operated on the wrong level during an open procedure with a fractured T10 vertebrae. The plan was to fuse T8-L1 while skipping T10 due to the fracture. Two segments were planned since the anatomy was too long. The first segment used a region of interest from T7-T11 with mount placed at T9. The plan was to place screws at T8 and T9. Registration was completed and matched to what they thought was T8 and T9. T7-T9 had green values, T10 had red values due to the fracture and T11 had yellow registration values. The surgeon matched what they thought was T8 in the software and began inserting screws, but they were actually at T7. The surgeon did not realize the error until they moved to register the second segment. The level tested fine with neuromonitoring after the screws were implanted. The second segment had a region of interest from T11-L1 with a mount placed at T12. The plan was to place screws at T11-L1. A k-wire was placed at T11 and a 2D lateral shot was taken. The image showed that they were two vertebrae above T11 without screws. The surgeon confirmed T11 anatomically and noted that they had placed screws at T7-T8 instead of T8-T9 as intended. The second segment was completed without issue. The surgeon then went back to T9, placed the mount, re-registered and placed the screw at that level. The fusion was extended by on vertebrae and included T7. There was no patient harm and the procedure was was delayed less than hour. 2022-03-01 IUD (Other): No new information. 2022-03-04 IUD x3, e1 (Rep): The right screws on T7 and T8 were shifted medially from the plan. 2022-03-15 IUD (Other): No new information. 2022-04-11 IUD (Other): No new information. 2022-04-13 IUD (Other): No new information. | xse incorrect level operated salesforce spinal p rocedure surgeon operated wrong level open procedure fractured vertebrae plan fuse skipping fracture segments planned anatomy long segment region mount placed plan place screws registration completed matched thought green values red values fracture yellow registration values surgeon matched thought software began inserting screws actually surgeon did realize error moved register second segment level tested fine neuromonitoring screws implanted second segment region mount placed plan place screws kwire placed lateral shot taken image showed vertebrae screws surgeon confirmed anatomically noted placed screws instead intended second segment completed issue surgeon went placed mount reregistered placed screw level fusion extended vertebrae included procedure delayed hour iud new iud right screws shifted medially plan iud new iud new iud new mazor x | NAV5000 | NAV4112 | NAV3246 | NAV3242 | NAV3243 | NAV5000NAV4112NAV3246NAV3242NAV3243 | True
10461 | 704884956 | XSE ¦ New Install Finding - Old Style | 2022-02-25T21:13:31.000000+0000 | MazorX Stealth Edition | NAV2057 | SURGICAL SYSTEM ISSUES | 20 | KIT0838-01 | SVC KIT0838-01 FRONT PLATE UPGRADE | XSE ¦ New Install Finding - Old Style 2022-Feb-25 Salesforce C00244262 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative was the in-service for the new install of the system when he noticed that the front plate was the old style plate. No patient was present. 2022-03-02 IUD (Other): No new information. 2022-03-09 IUD (Other): No new information. SVC KIT0838-01 FRONT PLATE UPGRADE | 2022-Feb-25 Salesforce C00244262 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative was the in-service for the new install of the system when he noticed that the front plate was the old style plate. No patient was present. 2022-03-02 IUD (Other): No new information. 2022-03-09 IUD (Other): No new information. | xse new install finding old style salesforce outside procedure inservice new install noticed plate old style plate present iud new iud new svc plate upgrade | NAV2054 | NAV2058 | NAV3249 | NAV2047 | NAV4100 | NAV2054NAV2058NAV3249NAV2047NAV4100 | False
5895 | 703831766 | XSE - Deviated Screws | 2020-07-01T23:07:17.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Screws 2020-07-01 00691939 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a percutaneous case, the surgical arm was sent to left L4 and navigation looked to be correct. The screw was drilled and placed. Left L5 looked good on navigation so the screw was drilled and tapped. The surgical arm was then sent to L4 and L5 right and navigation looked good at both locations so screws were placed. L4 left and right and L5 left were all superior to plan. The manufacturer representative believed the the deviation was due to skiving. L5 left was accurate. A new snapshot was done and the anatomy was off when checked. The representative believed there was a platform or patient shift. New AP and oblique images were taken and the patient was re-registered. All screws were able to be placed accurately. The representative believed that when placing the screws with the drill, the surgeon was using a lot of force to get the screw to go down the hole and this caused the shift or skive. There was no patient harm and the procedure was delayed more than an hour. 2020-07-02 email 2020-07-02 16:15 (Rep): Additional information received from a manufacturer representative reported that the trajectories were deviated less than 3.5 mm. There was no patient harm or symptoms. 2020-07-07 IUD (Other): No new information. 2020-07-14 CCR, Attachment (Rep): The 3Define scan, snapshot and draw spine were all accurate. During instrumentation, there was slight variance in the navigation, but not more than experienced in typical cases. When placing the Left L4 screw (first trajectory), the surgeon applied downward force on the powerease to keep the ATS screw moving forward. After all screws were placed, it was determined LL4, LL5 and RL4 were clearly superior compared to pre-op plan. RL5 seemed to be in a safe trajectory. The surgeon requested that the manufacturer representative update the plan and change angulation of LL4 and LL5 screws along with placing them lower in the pedicles. Once the new trajectories were sent, a lateral x-ray was taken and the trajectories still looked high. The manufacturer representative recommended surgeon complete an accuracy check and found that the navigation was slightly off from the accuracy check divot, indicating the need for a new snapshot to be acquired. Following the new snapshot, surgeon completed an accuracy check with a known anatomical landmark (transverse process of L4) and found that there was a platform shift or patient movement. The manufacturer representative recommended a second registration to be completed due to platform shift. All trajectories were placed accurately following the new registration. Trajectories were verified with intra op fluoro and neuromonitoring. MAZOR X SYSTEM TPL0059 | 2020-07-01 00691939 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a percutaneous case, the surgical arm was sent to left L4 and navigation looked to be correct. The screw was drilled and placed. Left L5 looked good on navigation so the screw was drilled and tapped. The surgical arm was then sent to L4 and L5 right and navigation looked good at both locations so screws were placed. L4 left and right and L5 left were all superior to plan. The manufacturer representative believed the the deviation was due to skiving. L5 left was accurate. A new snapshot was done and the anatomy was off when checked. The representative believed there was a platform or patient shift. New AP and oblique images were taken and the patient was re-registered. All screws were able to be placed accurately. The representative believed that when placing the screws with the drill, the surgeon was using a lot of force to get the screw to go down the hole and this caused the shift or skive. There was no patient harm and the procedure was delayed more than an hour. 2020-07-02 email 2020-07-02 16:15 (Rep): Additional information received from a manufacturer representative reported that the trajectories were deviated less than 3.5 mm. There was no patient harm or symptoms. 2020-07-07 IUD (Other): No new information. 2020-07-14 CCR, Attachment (Rep): The 3Define scan, snapshot and draw spine were all accurate. During instrumentation, there was slight variance in the navigation, but not more than experienced in typical cases. When placing the Left L4 screw (first trajectory), the surgeon applied downward force on the powerease to keep the ATS screw moving forward. After all screws were placed, it was determined LL4, LL5 and RL4 were clearly superior compared to pre-op plan. RL5 seemed to be in a safe trajectory. The surgeon requested that the manufacturer representative update the plan and change angulation of LL4 and LL5 screws along with placing them lower in the pedicles. Once the new trajectories were sent, a lateral x-ray was taken and the trajectories still looked high. The manufacturer representative recommended surgeon complete an accuracy check and found that the navigation was slightly off from the accuracy check divot, indicating the need for a new snapshot to be acquired. Following the new snapshot, surgeon completed an accuracy check with a known anatomical landmark (transverse process of L4) and found that there was a platform shift or patient movement. The manufacturer representative recommended a second registration to be completed due to platform shift. All trajectories were placed accurately following the new registration. Trajectories were verified with intra op fluoro and neuromonitoring. | xse deviated screws spinal procedure percutaneous case surgical arm sent left navigation looked correct screw drilled placed left looked good navigation screw drilled tapped surgical arm sent right navigation looked good locations screws placed left right left superior plan believed deviation skiving left accurate new snapshot anatomy checked believed platform shift new ap oblique images taken reregistered screws able placed accurately believed placing screws drill surgeon using lot force screw hole caused shift skive procedure delayed hour email additional trajectories deviated mm symptoms iud new ccr attachment scan snapshot draw spine accurate instrumentation slight variance navigation experienced typical cases placing left screw surgeon applied downward force powerease ats screw moving forward screws placed determined clearly superior compared preop plan safe trajectory surgeon requested update plan change angulation screws placing lower pedicles new trajectories sent lateral xray taken trajectories looked high recommended surgeon complete accuracy check navigation slightly accuracy check divot indicating need new snapshot acquired following new snapshot surgeon completed accuracy check known anatomical landmark platform shift movement recommended second registration completed platform shift trajectories placed accurately following new registration trajectories verified intra op fluoro neuromonitoring mazor x | NAV5000 | NAV5001 | NAV4091 | NAV4098 | NAV4108 | NAV5000NAV5001NAV4091NAV4098NAV4108 | True
10257 | 704842995 | XSR - Shoulder Lock Not Disengage | 2022-02-02T17:04:50.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSR - Shoulder Lock Not Disengage 2022-Feb-02 Salesforce C00242579, 00731249 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the shoulder lock would not disengage when when pressed. The light was just blinking. The manufacturer representative made sure the emergency stop button was not pressed and they cycled the button a couple of times. A hard reboot was also done twice. The surgical arm also failed during setup. The representative went into superuser and reset the arm movement. The surgical arm then passed a 10 point accuracy check. When the representative logged back into the surgeon profile and completed the accuracy check, the surgical arm failed. There was no patient involvement. 2022-02-08 IUD x2 (Other): No new information. 2022-03-01 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2022-Feb-02 Salesforce C00242579, 00731249 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the shoulder lock would not disengage when when pressed. The light was just blinking. The manufacturer representative made sure the emergency stop button was not pressed and they cycled the button a couple of times. A hard reboot was also done twice. The surgical arm also failed during setup. The representative went into superuser and reset the arm movement. The surgical arm then passed a 10 point accuracy check. When the representative logged back into the surgeon profile and completed the accuracy check, the surgical arm failed. There was no patient involvement. 2022-02-08 IUD x2 (Other): No new information. 2022-03-01 IUD (Other): No new information. | xsr shoulder lock disengage salesforce outside procedure shoulder lock disengage pressed light just blinking sure emergency stop button pressed cycled button couple times hard reboot twice surgical arm failed setup went superuser reset arm movement surgical arm passed point accuracy check logged surgeon profile completed accuracy check surgical arm failed iud new iud new kit svc surgical arm refurb | NAV4107 | NAV2047 | NAV4101 | NAV4095 | NAV4092 | NAV4107NAV2047NAV4101NAV4095NAV4092 | True
6513 | 703961050 | XSE - Positioner Hard to Move | 2020-09-16T19:57:30.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 40 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - Positioner Hard to Move 2020-09-16 00696824, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the positioner got stuck in the z-axis when the surgical system was mounted to the bed. The manufacturer representative had to put all of there weight on the positioner to get it to move. There was no patient harm and the procedure was delayed less than an hour. 2020-09-18 IUD x5, Image x4 (Rep): Additional information received from a manufacturer representative reported that the surgical system failed at install. The pneumatic tubing leading to the bottom coupler located in the teeth assembly was not properly fastened so air was leaking out. The pneumatic tube was reseated into the coupler to resolve the issue. 2020-10-05 IUD x4 (Other): No new information. 2021-01-13 e1 (Other): No new information. 2021-Mar-16 IUD (other): No new information. 2021-05-05 IUD (Other): No new information. 2021-05-12 IUD (Other): No new information. 2021-05-26 IUD x2 (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2020-09-16 00696824, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the positioner got stuck in the z-axis when the surgical system was mounted to the bed. The manufacturer representative had to put all of there weight on the positioner to get it to move. There was no patient harm and the procedure was delayed less than an hour. 2020-09-18 IUD x5, Image x4 (Rep): Additional information received from a manufacturer representative reported that the surgical system failed at install. The pneumatic tubing leading to the bottom coupler located in the teeth assembly was not properly fastened so air was leaking out. The pneumatic tube was reseated into the coupler to resolve the issue. 2020-10-05 IUD x4 (Other): No new information. 2021-01-13 e1 (Other): No new information. 2021-Mar-16 IUD (other): No new information. 2021-05-05 IUD (Other): No new information. 2021-05-12 IUD (Other): No new information. 2021-05-26 IUD x2 (Other): No new information. | xse positioner hard iud spinal procedure positioner got stuck zaxis surgical mounted bed weight positioner procedure delayed hour iud image additional surgical failed install pneumatic tubing leading coupler located teeth assembly properly fastened air leaking pneumatic tube reseated coupler resolve issue iud new new iud new iud new iud new iud new kit svc surgical syst type | NAV4104 | NAV3251 | NAV2054 | NAV3264 | NAV4103 | NAV4104NAV3251NAV2054NAV3264NAV4103 | True
7195 | 704137619 | XSE - Deviation/Stripped Target Extender | 2020-12-22T01:18:21.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2310-04S | KIT SVC MAS2310-04S TARGET EXTENDER | XSE - Deviation/Stripped Target Extender 2020-12-21 00703174, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that screw placement was superior and higher in the pedicle during the procedure when compared to the plan. The planned trajectories were adjusted to lower the screws and the screws were able to be placed successfully. The system was working as intended. The manufacturer representative noted that the target extender may be stripped. There was no patient harm and the procedure was delayed less than an hour. 2021-Jan-12 IUD (other): No new information. 2021-Feb-03 email 2/3 (rep): Additional information received indicated that the deviation was less than 3.5mm, k-wires were still in the pedicle. All of the screws were deviated. The cause of the deviation was believed to be the star market slightly stripped the extender piece, so it didn’t sit perfectly flush when we registered. There was no patient harm. 2021-03-03 email 2021-03-03 14:01 (Rep): No new information. KIT SVC MAS2310-04S TARGET EXTENDER | 2020-12-21 00703174, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that screw placement was superior and higher in the pedicle during the procedure when compared to the plan. The planned trajectories were adjusted to lower the screws and the screws were able to be placed successfully. The system was working as intended. The manufacturer representative noted that the target extender may be stripped. There was no patient harm and the procedure was delayed less than an hour. 2021-Jan-12 IUD (other): No new information. 2021-Feb-03 email 2/3 (rep): Additional information received indicated that the deviation was less than 3.5mm, k-wires were still in the pedicle. All of the screws were deviated. The cause of the deviation was believed to be the star market slightly stripped the extender piece, so it didn’t sit perfectly flush when we registered. There was no patient harm. 2021-03-03 email 2021-03-03 14:01 (Rep): No new information. | xse deviationstripped target extender iud spinal procedure screw placement superior higher pedicle procedure compared plan planned trajectories adjusted lower screws screws able placed successfully working intended noted target extender stripped procedure delayed hour iud new email additional indicated deviation kwires pedicle screws deviated cause deviation believed star market slightly stripped extender piece didn’t sit perfectly flush registered email new kit svc target extender | NAV3299 | NAV5000 | NAV3295 | NAV5001 | NAV4096 | NAV3299NAV5000NAV3295NAV5001NAV4096 | True
10473 | 704887516 | XSE - Unable to Load Exams | 2022-02-28T21:15:11.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Unable to Load Exams 2022-Feb-28 Salesforce C00244341 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when loading an an axial exam with 280 slices onto the laptop, the progress bar stopped at 279 slices and the software crashed. The issue occurred on multiple laptops. When importing the exam to an older software version, the import was successful. The exam was unable to be exported from the older software and imported to newer software. The manufacturer representative noted that not all exams were like this. There was no patient involvement. 2022-03-01 IUD (Other): No new information. 2022-03-08 IUD, e1 (Rep): No new information. 2022-03-13 e2, e3 (Rep): Additional information received from a manufacturer representative reported that the issue was replicated on multiple laptops and desktop PCs. 2022-03-14 e4, e5, e6 (Rep): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2022-Feb-28 Salesforce C00244341 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when loading an an axial exam with 280 slices onto the laptop, the progress bar stopped at 279 slices and the software crashed. The issue occurred on multiple laptops. When importing the exam to an older software version, the import was successful. The exam was unable to be exported from the older software and imported to newer software. The manufacturer representative noted that not all exams were like this. There was no patient involvement. 2022-03-01 IUD (Other): No new information. 2022-03-08 IUD, e1 (Rep): No new information. 2022-03-13 e2, e3 (Rep): Additional information received from a manufacturer representative reported that the issue was replicated on multiple laptops and desktop PCs. 2022-03-14 e4, e5, e6 (Rep): No new information. | xse unable load exams salesforce outside procedure loading axial exam slices laptop progress bar stopped slices software crashed issue occurred multiple laptops importing exam older software version import successful exam unable exported older software imported newer software noted exams like iud new iud new additional issue replicated multiple laptops desktop pcs new mazorx sw installation kit | NAV3235 | NAV3247 | NAV2056 | NAV3246 | NAV3248 | NAV3235NAV3247NAV2056NAV3246NAV3248 | True
56 | 702979108 | X- SW Bug | 2019-01-28T17:07:51.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X- SW Bug After performing 3 point test and I send the arm to clear up the error shown up "PLLlocked error of drive 4" . Resending surgical arm to any other predefined location resulted in the same error and arm stoped moving. The system was restarted then the error has cleared and back to normal operation. MAZOR X SOFTWARE KIT VER-4.0.0 | After performing 3 point test and I send the arm to clear up the error shown up "PLLlocked error of drive 4" . Resending surgical arm to any other predefined location resulted in the same error and arm stoped moving. The system was restarted then the error has cleared and back to normal operation. | x sw bug performing point test send arm clear error shown plllocked error drive resending surgical arm predefined location resulted error arm stoped moving restarted error cleared normal operation mazor x software kit | NAV3246 | NAV4098 | NAV3249 | NAV3248 | NAV2047 | NAV3246NAV4098NAV3249NAV3248NAV2047 | True
1948 | 703223002 | X - Registration issues | 2019-06-24T12:43:56.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Registration issues Event Description: Registration issues due to large titanium cages, and rotation of spine made it difficult to get clear endplates from L2-L4. Two registrations were required to capture all bodies. MAZOR X SOFTWARE KIT VER-3.0.2 | Event Description: Registration issues due to large titanium cages, and rotation of spine made it difficult to get clear endplates from L2-L4. Two registrations were required to capture all bodies. | x registration issues description registration issues large titanium cages rotation spine difficult clear endplates registrations required capture bodies mazor x software kit | NAV3242 | NAV3246 | NAV4091 | NAV3247 | NAV4112 | NAV3242NAV3246NAV4091NAV3247NAV4112 | True
9389 | 704622969 | XSE - Alleged Inaccuracy | 2021-09-22T02:01:37.000000+0000 | MazorX Stealth Edition | NAV5002 | MXSE - VARIANCE | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Alleged Inaccuracy 2021-Sep-21 Salesforce C00234779, email 2021-09-22 17:31, email 2021-09-24 19:45, email 2021-09-27 12:45, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were two inaccurate screws during the procedure. Prior to placing screws, a laminectomy was done. The patient was able to be registered without retractors placed. The manufacturer representative noted that the patient had a high BMI and there was an observable variance on navigation, but it was within tolerance. The surgeon did not complete any accuracy checks on landmarks and declined to complete a second snapshot prior to inserting screws. After the procedure, an AP image was taken and two screws were found to be deviated. An O-arm spin was done and the screws were estimated to be deviated 3-5 mm medial from the planned position. The screws were removed and one side of the construct was expanded due to the deviated screws. There was no patient harm and the procedure was delayed over an hour. 2021-10-04 email 2021-10-04 16:01 (Rep): Additional information was received stating that the screws were explanted in the same procedure before the patient left the operating room. 2021-10-04 email x2 2021-10-04 16:45 (Rep): No new information 2021-10-04 email 2021-10-04 17:01 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2021-Sep-21 Salesforce C00234779, email 2021-09-22 17:31, email 2021-09-24 19:45, email 2021-09-27 12:45, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were two inaccurate screws during the procedure. Prior to placing screws, a laminectomy was done. The patient was able to be registered without retractors placed. The manufacturer representative noted that the patient had a high BMI and there was an observable variance on navigation, but it was within tolerance. The surgeon did not complete any accuracy checks on landmarks and declined to complete a second snapshot prior to inserting screws. After the procedure, an AP image was taken and two screws were found to be deviated. An O-arm spin was done and the screws were estimated to be deviated 3-5 mm medial from the planned position. The screws were removed and one side of the construct was expanded due to the deviated screws. There was no patient harm and the procedure was delayed over an hour. 2021-10-04 email 2021-10-04 16:01 (Rep): Additional information was received stating that the screws were explanted in the same procedure before the patient left the operating room. 2021-10-04 email x2 2021-10-04 16:45 (Rep): No new information 2021-10-04 email 2021-10-04 17:01 (Rep): No new information. | xse alleged inaccuracy salesforce email email email iud spinal procedure inaccurate screws procedure prior placing screws laminectomy able registered retractors placed noted high bmi observable variance navigation tolerance surgeon did complete accuracy checks landmarks declined complete second snapshot prior inserting screws procedure ap image taken screws deviated oarm spin screws estimated deviated mm medial planned position screws removed construct expanded deviated screws procedure delayed hour email additional stating screws explanted procedure left operating room email new email new mazor x | NAV5000 | NAV1012 | NAV4112 | NAV5001 | NAV3299 | NAV5000NAV1012NAV4112NAV5001NAV3299 | False
2596 | 703244322 | X - Positioner/Manipulator issues | 2019-07-05T20:33:18.000000+0000 | MazorX | NAV3260 | MANIPULATOR ISSUES | 30 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | X - Positioner/Manipulator issues It was reported that system #24 was placed and locked on the bed, positioner and manipulator LED then turned off and were unable to be engaged. The manipulator was unable to be unlocked from the positioner, the positioner lock/unlock LED was pressed multiple times and continued to stay on and the system could not be removed from the bed. The hood LED was on and could be opened/closed along with full function of the arm, no change occurred in PSI. Storage stated was changed from ON to OFF, then ON and the system indicated it was in storage state. A hard restart was attempted with no change, Leafar and Ryan were called. The side electrical panel was partially removed and grossly inspected and no wires/connections appeared loose or disconnected. It was now noticed that the positioner lock button was turned off, a warning stating “shoulder not successfully locked to bed” occurred. After further inspection the left “claw” was not fully engaged on the bed frame adapter and positioned on the lip of the frame. A small gap was also noticed between the positioner and bed frame adapter on the corresponding side, large enough to except the head of a Philipps screw driver. The manipulator was then manually moved forward toward the left side to close the above mentioned gap, this allowed the left “claw” to fully engage/lock on the bed frame adapter which allowed the system to gain control of the positioner and manipulator. The system was then unlocked and taken of the bed with no issue and returned to workstation. During planning (on a laptop) all Red Reference Markers were placed at the back of the body/base of the pedicle, once the plan was uploaded to the workstation the Red Reference Markers were not in the same positions on some of the planned bodies, after registration and moving back to planning it was notice the RED Reference Markers were now all not in the same place as the plan from the laptop. UNK MAZOR X MANIPULATOR | It was reported that system #24 was placed and locked on the bed, positioner and manipulator LED then turned off and were unable to be engaged. The manipulator was unable to be unlocked from the positioner, the positioner lock/unlock LED was pressed multiple times and continued to stay on and the system could not be removed from the bed. The hood LED was on and could be opened/closed along with full function of the arm, no change occurred in PSI. Storage stated was changed from ON to OFF, then ON and the system indicated it was in storage state. A hard restart was attempted with no change, Leafar and Ryan were called. The side electrical panel was partially removed and grossly inspected and no wires/connections appeared loose or disconnected. It was now noticed that the positioner lock button was turned off, a warning stating “shoulder not successfully locked to bed” occurred. After further inspection the left “claw” was not fully engaged on the bed frame adapter and positioned on the lip of the frame. A small gap was also noticed between the positioner and bed frame adapter on the corresponding side, large enough to except the head of a Philipps screw driver. The manipulator was then manually moved forward toward the left side to close the above mentioned gap, this allowed the left “claw” to fully engage/lock on the bed frame adapter which allowed the system to gain control of the positioner and manipulator. The system was then unlocked and taken of the bed with no issue and returned to workstation. During planning (on a laptop) all Red Reference Markers were placed at the back of the body/base of the pedicle, once the plan was uploaded to the workstation the Red Reference Markers were not in the same positions on some of the planned bodies, after registration and moving back to planning it was notice the RED Reference Markers were now all not in the same place as the plan from the laptop. | x positionermanipulator issues placed locked bed positioner manipulator led turned unable engaged manipulator unable unlocked positioner positioner lockunlock led pressed multiple times continued stay removed bed hood led openedclosed function arm change occurred psi storage stated changed indicated storage state hard restart attempted change leafar ryan called electrical panel partially removed grossly inspected wiresconnections appeared loose disconnected noticed positioner lock button turned warning stating “shoulder successfully locked bed” occurred inspection left “claw” fully engaged bed frame adapter positioned lip frame small gap noticed positioner bed frame adapter corresponding large head philipps screw driver manipulator manually moved forward left close mentioned gap allowed left “claw” fully engagelock bed frame adapter allowed gain control positioner manipulator unlocked taken bed issue returned workstation planning red reference markers placed bodybase pedicle plan uploaded workstation red reference markers positions planned bodies registration moving planning notice red reference markers place plan laptop unk mazor x manipulator | NAV3260 | NAV4103 | NAV3255 | NAV3259 | NAV4105 | NAV3260NAV4103NAV3255NAV3259NAV4105 | True
2177 | 703228485 | X - 3 Define Issue | 2019-06-26T15:06:38.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3 Define Issue During 3Define scan, had to make second attempt because an error message popped up saying it couldn't calculate work volume. 1300 Arm mounted to bed 1325 Arm Draped 1359 BMU Mounted to patient 1340 3Define Scan -First attempt caused error messaging stating it couldn't calculate working volume. 1342 2nd 3Define Scan Successful 1345 Draw Spine. Labeled off L5 1346 AP 3 Images 1349 OBL 4 Images 1353 Registration Achieved on Advanced with green values 1354 Visual Verification 1355 L4 Right 1356 L5 Right 1358 S1 Right 1400 S1 Left 1402 L5 Left 1405 L4 Left 1407 Arm Cleared Right. UNK MAZOR X 3DEFINE CAMERA | During 3Define scan, had to make second attempt because an error message popped up saying it couldn't calculate work volume. 1300 Arm mounted to bed 1325 Arm Draped 1359 BMU Mounted to patient 1340 3Define Scan -First attempt caused error messaging stating it couldn't calculate working volume. 1342 2nd 3Define Scan Successful 1345 Draw Spine. Labeled off L5 1346 AP 3 Images 1349 OBL 4 Images 1353 Registration Achieved on Advanced with green values 1354 Visual Verification 1355 L4 Right 1356 L5 Right 1358 S1 Right 1400 S1 Left 1402 L5 Left 1405 L4 Left 1407 Arm Cleared Right. | x define issue scan make second attempt error message popped saying calculate work volume arm mounted bed arm draped bmu mounted scan attempt caused error messaging stating calculate working volume scan successful draw spine labeled ap images obl images registration achieved advanced green values visual verification right right right left left left arm cleared right unk mazor x camera | NAV3244 | NAV3246 | NAV4101 | NAV4099 | NAV4095 | NAV3244NAV3246NAV4101NAV4099NAV4095 | True
8333 | 704398753 | XSE - Broken Screwdriver | 2021-05-17T21:01:15.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1650 | KIT SVC MAS1650 UNIVER SCRWDRVR 04S NEW | XSE - Broken Screwdriver 2021-May-17 MPXR 839246 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the green universal screwdriver snapped at the connection point between the driver shank and the handle when putting sets together after they had been through the wash. The driver was not broken during the procedure or immediately after because they used it to dismount and disassemble the robot. It could have broken at some point during the cleaning process. No patient was present. 2021-05-18 IUD x3 (Other): No new information. 2021-07-07 email 2021-07-07 10:45 (Rep): No new information. 2021-07-19 IUD (Other): No new information. 2021-08-12 IUD (Other): No new information. 2021-08-23 IUD (Other): No new information. KIT SVC MAS1650 UNIVER SCRWDRVR 04S NEW | 2021-May-17 MPXR 839246 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the green universal screwdriver snapped at the connection point between the driver shank and the handle when putting sets together after they had been through the wash. The driver was not broken during the procedure or immediately after because they used it to dismount and disassemble the robot. It could have broken at some point during the cleaning process. No patient was present. 2021-05-18 IUD x3 (Other): No new information. 2021-07-07 email 2021-07-07 10:45 (Rep): No new information. 2021-07-19 IUD (Other): No new information. 2021-08-12 IUD (Other): No new information. 2021-08-23 IUD (Other): No new information. | xse broken screwdriver mpxr outside procedure green universal screwdriver snapped connection point driver shank handle putting sets wash driver broken procedure immediately dismount disassemble robot broken point cleaning process present iud new email new iud new iud new iud new kit svc univer scrwdrvr new | NAV3299 | NAV3251 | NAV3295 | NAV2055 | NAV3249 | NAV3299NAV3251NAV3295NAV2055NAV3249 | True
5054 | 703592664 | X- SW defects & freeze | 2020-02-03T00:08:44.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X- SW defects & freeze 2018-1-15 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a software bug and freeze. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SOFTWARE KIT VER-2.3.1 | 2018-1-15 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a software bug and freeze. There was no patient harm and the procedure was not delayed over an hour. | x sw defects freeze ccr spinal procedure software bug freeze procedure delayed hour mazor x software kit | NAV3248 | NAV3246 | NAV4091 | NAV3242 | NAV3247 | NAV3248NAV3246NAV4091NAV3242NAV3247 | True
3569 | 703378361 | XSE - Arm Failed Advanced Accuracy Test | 2019-09-24T16:56:06.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Arm Failed Advanced Accuracy Test 2019-09-24 SCL1902883, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported hat planning for the procedure was completed by the surgeon. The surgeon first completed a L3-S1 OLIF and then flipped the patient to a prone position. The surgical system was mounted to the patient using the Schanz bond and BMB connected to a Schanz pin inserted into the right PSIS. 3Define scan and Draw Spine were successfully completed. Registration was completed on the first attempt using auto registration and the basic algorithm. The surgeon decided to drill all of the trajectories first prior to placing screws. All of the left side trajectories were drilled starting at the top followed by the right side trajectories. Following drilling of the last trajectory, the surgeon began placing screws working backwards in drilling order. When inserting the screw at right L3, the screw appeared to be inferior to the prior drilled hole. Navigation showed the trajectory to be inferior to plan. A platform shift was suspected and a new registration was done with the towers and screws already at right L4-S1. Registration was unsuccessful after multiple attempts so the screws were removed in order to achieve registration. A third registration was performed quickly using the auto registration and basic algorithm. The surgical arm was sent to all trajectories to place k-wires and then screws. After the procedure, the surgical arm failed an advanced accuracy test. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. 2019-09-24 email 2019-09-24 12:30 (Rep): Additional information received from a manufacturer representative reported that during the surgical procedure the surgeon noted the surgical arm wrist was slightly loose. By wrist, the representative was referring to the point where the arm guide attached to the arm. The representative was waiting for post-op CT to confirm accuracy of the screws compared to the plan. Final AP and lateral fluoro images appeared to show L4-S1 screws were accurately placed, but L3 trajectories were slightly inferior to plan. During the procedure, the surgeon's glove got stuck between the tracker and the drill connection on two separate occasions. The representative also mentioned the long cannula was stuck in the arm guide. The surgeon had to apply excessive force to the cannula in order to remove it from the arm guide. These events and the wrist being loose, led the representative to perform a 10 point accuracy test after the procedure. MAZOR X SYSTEM TPL0059 | 2019-09-24 SCL1902883, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported hat planning for the procedure was completed by the surgeon. The surgeon first completed a L3-S1 OLIF and then flipped the patient to a prone position. The surgical system was mounted to the patient using the Schanz bond and BMB connected to a Schanz pin inserted into the right PSIS. 3Define scan and Draw Spine were successfully completed. Registration was completed on the first attempt using auto registration and the basic algorithm. The surgeon decided to drill all of the trajectories first prior to placing screws. All of the left side trajectories were drilled starting at the top followed by the right side trajectories. Following drilling of the last trajectory, the surgeon began placing screws working backwards in drilling order. When inserting the screw at right L3, the screw appeared to be inferior to the prior drilled hole. Navigation showed the trajectory to be inferior to plan. A platform shift was suspected and a new registration was done with the towers and screws already at right L4-S1. Registration was unsuccessful after multiple attempts so the screws were removed in order to achieve registration. A third registration was performed quickly using the auto registration and basic algorithm. The surgical arm was sent to all trajectories to place k-wires and then screws. After the procedure, the surgical arm failed an advanced accuracy test. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. 2019-09-24 email 2019-09-24 12:30 (Rep): Additional information received from a manufacturer representative reported that during the surgical procedure the surgeon noted the surgical arm wrist was slightly loose. By wrist, the representative was referring to the point where the arm guide attached to the arm. The representative was waiting for post-op CT to confirm accuracy of the screws compared to the plan. Final AP and lateral fluoro images appeared to show L4-S1 screws were accurately placed, but L3 trajectories were slightly inferior to plan. During the procedure, the surgeon's glove got stuck between the tracker and the drill connection on two separate occasions. The representative also mentioned the long cannula was stuck in the arm guide. The surgeon had to apply excessive force to the cannula in order to remove it from the arm guide. These events and the wrist being loose, led the representative to perform a 10 point accuracy test after the procedure. | xse arm failed advanced accuracy test ccr spinal procedure treat degenerative disc fixation hat planning procedure completed surgeon surgeon completed olif flipped prone position surgical mounted using schanz bond bmb connected schanz pin inserted right psis scan draw spine successfully completed registration completed attempt using auto registration basic algorithm surgeon decided drill trajectories prior placing screws left trajectories drilled starting followed right trajectories following drilling trajectory surgeon began placing screws working backwards drilling order inserting screw right screw appeared inferior prior drilled hole navigation showed trajectory inferior plan platform shift suspected new registration towers screws right registration unsuccessful multiple attempts screws removed order achieve registration registration performed quickly using auto registration basic algorithm surgical arm sent trajectories place kwires screws procedure surgical arm failed advanced accuracy test procedure delayed hour case completed email additional surgical procedure surgeon noted surgical arm wrist slightly loose wrist referring point arm guide attached arm waiting postop ct confirm accuracy screws compared plan final ap lateral fluoro images appeared screws accurately placed trajectories slightly inferior plan procedure surgeons glove got stuck tracker drill connection separate occasions mentioned long cannula stuck arm guide surgeon apply excessive force cannula order remove arm guide events wrist loose led perform point accuracy test procedure mazor x | NAV5000 | NAV4093 | NAV3299 | NAV3243 | NAV4097 | NAV5000NAV4093NAV3299NAV3243NAV4097 | False
850 | 703178491 | X - Surgical System Issue | 2019-05-29T15:34:27.000000+0000 | MazorX | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X - Surgical System Issue Separated from Comp-0000284. This complaint will handle the positioner not engaging to the bed. Issues continued during mounting the system to the bed frame adaptor. Both LEDs on top of the positioner were illuminated, NRC pressed the lock button. Error appeared on the screen saying only one cleat was locked. CE verified this by looking under the table, the right cleat was fully engaged, and the left cleat was completely back within the positioner plate. Pressing on the shoulder and pressing on the positioner bottom portion were tried as is the fix for a halfway engaged cleat. At this point CE called CSC, no one picked up (it was past 5pm EST). Hard re-start was attempted, this did not fix issue. CE called a couple FSEs directly until someone picked up. The override window was opened under the direction of the FSE. Surgeon and clinical engineer attempted to press the sensor on the left cleat in order to be able to use the override. Since the cleat was within the positioner it was difficult to do, surgeon was able to get the cleat to extend out to press the sensor. Many attempts to use the override to engage and release the cleat were tried. CE went to look at system in dicators, the left positioner sensor kept flashing back and forth between green and red, even though the left LED was stable. Surgeon decided to abort use of the robot. In order to get the system off the bed system plus bedframe adaptor had to be removed together as a unit. After the system was taken off the bed, the pressure dropped significantly down to 4.625. Surgeon wanted it out of the room, system was shut down without placing the system back into the workstation. In the storage closet CE attempted to recreate the failure, however it was unsuccessful. When pressing on all sensors individually they appeared to function when watching the system indications. Then while attempting to put the system back in the workstation, the positioner dropped in the Z axis a couple inches. Because of that the main cover could not close, the sensors weren’t engaged. Once mounted onto the bed frame adaptor again, system was able to be put back in home and returned to its resting place in the workstation. Main cover closed. Export is available, no CCR written as it was a scan and plan case and we didn’t make it into operation. Photos of error messages available. UNK MAZOR X POSITIONER | Separated from Comp-0000284. This complaint will handle the positioner not engaging to the bed. Issues continued during mounting the system to the bed frame adaptor. Both LEDs on top of the positioner were illuminated, NRC pressed the lock button. Error appeared on the screen saying only one cleat was locked. CE verified this by looking under the table, the right cleat was fully engaged, and the left cleat was completely back within the positioner plate. Pressing on the shoulder and pressing on the positioner bottom portion were tried as is the fix for a halfway engaged cleat. At this point CE called CSC, no one picked up (it was past 5pm EST). Hard re-start was attempted, this did not fix issue. CE called a couple FSEs directly until someone picked up. The override window was opened under the direction of the FSE. Surgeon and clinical engineer attempted to press the sensor on the left cleat in order to be able to use the override. Since the cleat was within the positioner it was difficult to do, surgeon was able to get the cleat to extend out to press the sensor. Many attempts to use the override to engage and release the cleat were tried. CE went to look at system in dicators, the left positioner sensor kept flashing back and forth between green and red, even though the left LED was stable. Surgeon decided to abort use of the robot. In order to get the system off the bed system plus bedframe adaptor had to be removed together as a unit. After the system was taken off the bed, the pressure dropped significantly down to 4.625. Surgeon wanted it out of the room, system was shut down without placing the system back into the workstation. In the storage closet CE attempted to recreate the failure, however it was unsuccessful. When pressing on all sensors individually they appeared to function when watching the system indications. Then while attempting to put the system back in the workstation, the positioner dropped in the Z axis a couple inches. Because of that the main cover could not close, the sensors weren’t engaged. Once mounted onto the bed frame adaptor again, system was able to be put back in home and returned to its resting place in the workstation. Main cover closed. Export is available, no CCR written as it was a scan and plan case and we didn’t make it into operation. Photos of error messages available. | x surgical issue separated complaint handle positioner engaging bed issues continued mounting bed frame adaptor leds positioner illuminated nrc pressed lock button error appeared screen saying cleat locked ce verified looking table right cleat fully engaged left cleat completely positioner plate pressing shoulder pressing positioner portion tried fix halfway engaged cleat point ce called csc picked hard restart attempted did fix issue ce called couple fses directly picked override window opened direction fse surgeon clinical engineer attempted press sensor left cleat order able use override cleat positioner difficult surgeon able cleat extend press sensor attempts use override engage release cleat tried ce went look dicators left positioner sensor kept flashing forth green red left led stable surgeon decided abort use robot order bed plus bedframe adaptor removed unit taken bed pressure dropped significantly surgeon wanted room shut placing workstation storage closet ce attempted recreate failure unsuccessful pressing sensors individually appeared function watching indications attempting workstation positioner dropped z axis couple inches main cover close sensors weren’t engaged mounted bed frame adaptor able home returned resting place workstation main cover closed export available ccr written scan plan case didn’t make operation photos error messages available unk mazor x positioner | NAV3250 | NAV4095 | NAV4105 | NAV3255 | NAV3259 | NAV3250NAV4095NAV4105NAV3255NAV3259 | False
1704 | 703209617 | X- SW Issues | 2019-06-15T17:19:11.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X- SW Issues Software crash after 1st Registration – soft restart resolved issue - screen brought back to segmentation and registration had to be completed second time. Interface with Dual Clamp for several trajectories in thoracic – T5 unable to be lateralized further in planning safely to avoid clamp – surgeons choose to freehand MAZOR X SOFTWARE KIT VER-3.0 | Software crash after 1st Registration – soft restart resolved issue - screen brought back to segmentation and registration had to be completed second time. Interface with Dual Clamp for several trajectories in thoracic – T5 unable to be lateralized further in planning safely to avoid clamp – surgeons choose to freehand | x sw issues software crash registration – soft restart resolved issue screen brought segmentation registration completed second time interface dual clamp trajectories thoracic – unable lateralized planning safely avoid clamp – surgeons choose freehand mazor x software kit | NAV3246 | NAV3247 | NAV3242 | NAV4112 | NAV3248 | NAV3246NAV3247NAV3242NAV4112NAV3248 | True
5027 | 703591670 | X - 3Define/Shoulder Shift/Reg issues | 2020-01-31T18:32:30.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - 3Define/Shoulder Shift/Reg issues 2018-08-08 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case plan was reviewed by representative and surgeon the morning of the procedure. There were no major concerns discussed. L4-S1 were planned for an open procedure with the intention of utilizing the disposable clamp. The disposable clamp was mounted to L4 with no complications. 3Define failed on the first attempt. The arm was sent to the up position and then 3Define was attempted again. 3Define was successful on the second attempt. Registration was extremely difficult and took multiple attempts, new images and platform adjustments. After multiple attempts of wagging and trying to get a clearer image, L4 was finally registered which was the only issue in the construct. L5 and S1 would register. When lowering the table to get the procedure started, the mayo was caught under the arm where the tablet attaches and applied a lot of pressure to the system, causing a shoulder shift. There was need to re-register. The surgeon switched from the clamp to the a single schanz in the left PSIS. Registration was achieved on the first attempt. There were no complications throughout the procedure. All wires and screws were placed accurately. There was no patient harm. MAZOR X SYSTEM TPL0059 | 2018-08-08 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case plan was reviewed by representative and surgeon the morning of the procedure. There were no major concerns discussed. L4-S1 were planned for an open procedure with the intention of utilizing the disposable clamp. The disposable clamp was mounted to L4 with no complications. 3Define failed on the first attempt. The arm was sent to the up position and then 3Define was attempted again. 3Define was successful on the second attempt. Registration was extremely difficult and took multiple attempts, new images and platform adjustments. After multiple attempts of wagging and trying to get a clearer image, L4 was finally registered which was the only issue in the construct. L5 and S1 would register. When lowering the table to get the procedure started, the mayo was caught under the arm where the tablet attaches and applied a lot of pressure to the system, causing a shoulder shift. There was need to re-register. The surgeon switched from the clamp to the a single schanz in the left PSIS. Registration was achieved on the first attempt. There were no complications throughout the procedure. All wires and screws were placed accurately. There was no patient harm. | x shiftreg issues ccr spinal procedure case plan reviewed surgeon morning procedure major concerns discussed planned open procedure intention utilizing disposable clamp disposable clamp mounted complications failed attempt arm sent position attempted successful second attempt registration extremely difficult took multiple attempts new images platform adjustments multiple attempts wagging trying clearer image finally registered issue construct register lowering table procedure started mayo caught arm tablet attaches applied lot pressure causing shoulder shift need reregister surgeon switched clamp single schanz left psis registration achieved attempt complications procedure wires screws placed accurately mazor x | NAV3242 | NAV4112 | NAV3299 | NAV1012 | NAV4108 | NAV3242NAV4112NAV3299NAV1012NAV4108 | True
9801 | 704723233 | XSE - Broken Screwdriver | 2021-11-18T16:23:08.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1650 | MAS1650 UNIVERSAL SCREWDRIVER | XSE - Broken Screwdriver 2021-11-14 mPXR 897056 (HCP, Rep): Medtronic received information regarding a guidance system instrument being used intra-operatively of a spinal procedure. It was reported that the tip of the driver sheared off while trying to tighten the schanz arm while the surgeon was adjusting the screw height. This happened on the first use of the instrument. The issue extended the surgical time by about 5 minutes. There was no impact on the patient impact. 2021-11-30 IUD (Other): No new information. 2021-12-03 email 2021-12-03 15:16 (Rep): No new information. 2021-12-08 email 2021-12-08 11:31, email 2021-12-09 14:02 (Rep): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-13 IUD (Other): No new information. MAS1650 UNIVERSAL SCREWDRIVER | 2021-11-14 mPXR 897056 (HCP, Rep): Medtronic received information regarding a guidance system instrument being used intra-operatively of a spinal procedure. It was reported that the tip of the driver sheared off while trying to tighten the schanz arm while the surgeon was adjusting the screw height. This happened on the first use of the instrument. The issue extended the surgical time by about 5 minutes. There was no impact on the patient impact. 2021-11-30 IUD (Other): No new information. 2021-12-03 email 2021-12-03 15:16 (Rep): No new information. 2021-12-08 email 2021-12-08 11:31, email 2021-12-09 14:02 (Rep): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-13 IUD (Other): No new information. | xse broken screwdriver mpxr instrument intraoperatively spinal procedure tip driver sheared trying tighten schanz arm surgeon adjusting screw height happened use instrument issue extended surgical time minutes impact impact iud new email new email email new iud new iud new universal screwdriver | NAV3299 | NAV3295 | NAV3246 | NAV5000 | NAV3243 | NAV3299NAV3295NAV3246NAV5000NAV3243 | True
813 | 703176373 | X - Software issue | 2019-05-28T20:29:03.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - Software issue This complaint was separated from Comp-0000454: During movement of the arm (3Define scan ) we did not receive any movement sound. We stop the movement and re scan and the sound returned MAZOR X SOFTWARE KIT VER-4.0.0 | This complaint was separated from Comp-0000454: During movement of the arm (3Define scan ) we did not receive any movement sound. We stop the movement and re scan and the sound returned | x software issue complaint separated movement arm did receive movement sound stop movement scan sound returned mazor x software kit | NAV3246 | NAV3247 | NAV4091 | NAV3248 | NAV3249 | NAV3246NAV3247NAV4091NAV3248NAV3249 | True
5824 | 703815316 | XSE - OOBF/SRA CANBUS Error | 2020-06-22T23:58:09.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - OOBF/SRA CANBUS Error 2020-06-22 00691309 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the surgical arm caused a CANBUS error message 295 at noted 2, 3, 4 to be displayed. The manufacturer representative noted that calibration was not found on the driver error. There was no patient involvement. 2020-06-25 IUD x4 (Other): No new information. 2020-07-07 IUD (Other): No new information. 2020-07-22 IUD x3 (Other): No new information. 2020-08-05 IUD (Other): No new information. 2020-09-03 IUD x2 (Other): No new information. 2020-09-08 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-06-22 00691309 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the surgical arm caused a CANBUS error message 295 at noted 2, 3, 4 to be displayed. The manufacturer representative noted that calibration was not found on the driver error. There was no patient involvement. 2020-06-25 IUD x4 (Other): No new information. 2020-07-07 IUD (Other): No new information. 2020-07-22 IUD x3 (Other): No new information. 2020-08-05 IUD (Other): No new information. 2020-09-03 IUD x2 (Other): No new information. 2020-09-08 IUD (Other): No new information. | xse oobfsra canbus error prior procedure surgical arm caused canbus error message noted displayed noted calibration driver error iud new iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV2047 | NAV2054 | NAV3249 | NAV4095 | NAV4098 | NAV2047NAV2054NAV3249NAV4095NAV4098 | False
9610 | 704682644 | XSE - Damaged Positioner Handle | 2021-10-26T18:24:08.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 30 | ASM0364-01S | KIT SVC ASM0364-01S RIGHT HANDLE | XSE - Damaged Positioner Handle 2021-Oct-26 Salesforce C00236829 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the left handle of the positioner was damaged. No patient was present. 2021-10-27 IUD (Other): No new information. 2021-11-03 IUD x3 (Other): No new information. 2021-11-12 IUD (Other): No new information. 2022-01-28 IUD (Other): No new information. 2022-02-22 IUD (Other): No new information. KIT SVC ASM0364-01S RIGHT HANDLE | 2021-Oct-26 Salesforce C00236829 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the left handle of the positioner was damaged. No patient was present. 2021-10-27 IUD (Other): No new information. 2021-11-03 IUD x3 (Other): No new information. 2021-11-12 IUD (Other): No new information. 2022-01-28 IUD (Other): No new information. 2022-02-22 IUD (Other): No new information. | xse damaged positioner handle salesforce outside procedure left handle positioner damaged present iud new iud new iud new iud new iud new kit svc right handle | NAV3251 | NAV2058 | NAV3252 | NAV2055 | NAV3299 | NAV3251NAV2058NAV3252NAV2055NAV3299 | True
317 | 703078080 | XSE - Schanz Bond/BM Screws Stripped | 2019-03-28T13:39:28.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MEC5196-01 | RNX-ACC-063 | XSE - Schanz Bond/BM Screws Stripped 2019-Mar-26 SCL1900738 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Schanz Bond and Bone Mount Bridge screws were stripped out. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. RNX-ACC-063 | 2019-Mar-26 SCL1900738 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Schanz Bond and Bone Mount Bridge screws were stripped out. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse schanz bondbm screws stripped spinal procedure schanz bond bone mount bridge screws stripped noted additional complications anticipated | NAV3299 | NAV3295 | NAV4100 | NAV5000 | NAV3261 | NAV3299NAV3295NAV4100NAV5000NAV3261 | True
5610 | 703764783 | XSE - L2 Deviated Lateral | 2020-05-21T18:14:44.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - L2 Deviated Lateral 2020-05-04 00689326, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a revision case from L2-S1 was being done. Hardware was previously placed at L3 to L1. A Schanz pin was placed in the PSIS and connected to the bone mount bridge to mount the surgical system to the patient. The surgeon instrumented L1, L2 and then S1 right. All k-wires were placed first and then screws were placed. Right L2 was found to have a lateral breach. The plan was adjusted and the surgical arm was sent to the trajectory again to place the screw. The manufacturer representative noted that registration was completed with green values and no shifts were seen. There was no patient harm and the procedure was delayed less than an hour. 2020-06-01, email 2020-06-01 08:00 IUD (Rep): Additional information received from a manufacturer representative reported that the t rajectories were deviated less than 3.5 mm. There was no patient harm and the surgeon decided to leave the screws out. A cause or suspected cause was not determined. MAZOR X SYSTEM TPL0059 | 2020-05-04 00689326, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a revision case from L2-S1 was being done. Hardware was previously placed at L3 to L1. A Schanz pin was placed in the PSIS and connected to the bone mount bridge to mount the surgical system to the patient. The surgeon instrumented L1, L2 and then S1 right. All k-wires were placed first and then screws were placed. Right L2 was found to have a lateral breach. The plan was adjusted and the surgical arm was sent to the trajectory again to place the screw. The manufacturer representative noted that registration was completed with green values and no shifts were seen. There was no patient harm and the procedure was delayed less than an hour. 2020-06-01, email 2020-06-01 08:00 IUD (Rep): Additional information received from a manufacturer representative reported that the t rajectories were deviated less than 3.5 mm. There was no patient harm and the surgeon decided to leave the screws out. A cause or suspected cause was not determined. | xse deviated lateral iud spinal procedure revision case hardware previously placed schanz pin placed psis connected bone mount bridge mount surgical surgeon instrumented right kwires placed screws placed right lateral breach plan adjusted surgical arm sent trajectory place screw noted registration completed green values shifts seen procedure delayed hour email iud additional t rajectories deviated mm surgeon decided leave screws cause suspected cause determined mazor x | NAV5000 | NAV1012 | NAV3246 | NAV4112 | NAV2054 | NAV5000NAV1012NAV3246NAV4112NAV2054 | True
1141 | 703191182 | X-Registration issues & Deviations | 2019-06-05T18:00:55.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X-Registration issues & Deviations 2018-6-18 SCL1802264(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the initial registration attempt failed due to the beads of the marker not being clear in the lateral scout image. The second spin with the O-arm was successful after lowering the arm and 3D marker. The manufacturer representative planned screws and the surgeon reviewed each trajectory. The left T12 trajectory was sent, and drilled. Left T12 was significantly medial to plan. Left T12 was resent, but still produced similar results. L1 trajectories were sent, but were both medial to the plan. The surgeon decided to freehand the rest of the procedure. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-6-18 SCL1802264(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the initial registration attempt failed due to the beads of the marker not being clear in the lateral scout image. The second spin with the O-arm was successful after lowering the arm and 3D marker. The manufacturer representative planned screws and the surgeon reviewed each trajectory. The left T12 trajectory was sent, and drilled. Left T12 was significantly medial to plan. Left T12 was resent, but still produced similar results. L1 trajectories were sent, but were both medial to the plan. The surgeon decided to freehand the rest of the procedure. There was no patient harm and the procedure was not delayed over an hour. | xregistration issues deviations spinal procedure initial registration attempt failed beads marker clear lateral scout image second spin oarm successful lowering arm marker planned screws surgeon reviewed trajectory left trajectory sent drilled left significantly medial plan left resent produced similar results trajectories sent medial plan surgeon decided freehand rest procedure procedure delayed hour mazor x | NAV3242 | NAV3244 | NAV4091 | NAV1012 | NAV5000 | NAV3242NAV3244NAV4091NAV1012NAV5000 | True
7980 | 704325002 | XSE - Wireless Tablet Not Connecting | 2021-04-06T16:10:26.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0337-01 | SURGEON&APOS;S SCREEN | XSE - Wireless Tablet Not Connecting 2021-Apr-06 Salesforce C00223890 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon screen was not connecting to the system. The manufacturer representative was able to get into the twomon, but they were not able to log into the option that is provided. There was no patient involvement. 2021-04-13 IUD (Other): Additional information was received stating that it was confirmed he was able to get the tablet to connect. 2021-05-06 email 2021-05-06 15:31 (Rep): Additional information was received stating that the screen was connected and it was user error the manufacturer representative's part. SURGEON&APOS;S SCREEN | 2021-Apr-06 Salesforce C00223890 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon screen was not connecting to the system. The manufacturer representative was able to get into the twomon, but they were not able to log into the option that is provided. There was no patient involvement. 2021-04-13 IUD (Other): Additional information was received stating that it was confirmed he was able to get the tablet to connect. 2021-05-06 email 2021-05-06 15:31 (Rep): Additional information was received stating that the screen was connected and it was user error the manufacturer representative's part. | xse wireless tablet connecting salesforce outside procedure surgeon screen connecting able twomon able log option provided iud additional stating confirmed able tablet connect email additional stating screen connected user error representatives surgeonaposs screen | NAV3253 | NAV2047 | NAV2054 | NAV3246 | NAV3262 | NAV3253NAV2047NAV2054NAV3246NAV3262 | True
6859 | 704051245 | XSE - Damaged Accuracy Pointer | 2020-11-04T14:43:07.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 20 | MAS2285-01S | KIT SVC MAS2285-01S ACCURACY POINTER | XSE - Damaged Accuracy Pointer 2020-Nov-04 00700194 (Rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the accuracy pointer seemed bent. When placing it in the divot, it would seem inaccurate until he rotated. There was no patient involved or additional complications reported or anticipated as a result of the event. 2020-11-06 IUD (Other): No new information. KIT SVC MAS2285-01S ACCURACY POINTER | 2020-Nov-04 00700194 (Rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the accuracy pointer seemed bent. When placing it in the divot, it would seem inaccurate until he rotated. There was no patient involved or additional complications reported or anticipated as a result of the event. 2020-11-06 IUD (Other): No new information. | xse damaged accuracy pointer outside procedure accuracy pointer bent placing divot inaccurate rotated involved additional complications anticipated iud new kit svc accuracy pointer | NAV4092 | NAV3295 | NAV3299 | NAV2054 | NAV3264 | NAV4092NAV3295NAV3299NAV2054NAV3264 | True
1567 | 703205540 | X- Shoulder Issues | 2019-06-13T17:08:16.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X- Shoulder Issues LED lights in shoulder. When shoulder was all the way to the left the LED would not lock the shoulder it would continue to blink. This was true in certain ranges. Surgeon reviewed plan prior to procedure with RTC and discussed axial angles and skive potential. Surgeon chose to use the single schanz in the right psis. Automatic registration was achieved using the basic algorithm with good values off of L5. 3Define surface scanner was not found/recognized RTC attempted 4 scan attempts after the first to RTC restarted realsense dcm application after that attempted soft restart. Advanced using generic work volume 30 inflate 30 coronal shift. Drawspine stylus was very challenging to pick up. S2 left was a red screw RTC adjusted the plan and surgeon confirmed that the adjustments were okay.1130 10 pt 1547 3pt 1700 mounted 1723 draped 1839 attached 1841-1844 3Define attempts, real sense dcm restarted, soft restart 1845 drawspine 1855 registered S1 left 1918 S2R 1922 S2L 1933 S1 R 1938. Overall surgeon was happy with the accuracy. UNK MAZOR X SHOULDER | LED lights in shoulder. When shoulder was all the way to the left the LED would not lock the shoulder it would continue to blink. This was true in certain ranges. Surgeon reviewed plan prior to procedure with RTC and discussed axial angles and skive potential. Surgeon chose to use the single schanz in the right psis. Automatic registration was achieved using the basic algorithm with good values off of L5. 3Define surface scanner was not found/recognized RTC attempted 4 scan attempts after the first to RTC restarted realsense dcm application after that attempted soft restart. Advanced using generic work volume 30 inflate 30 coronal shift. Drawspine stylus was very challenging to pick up. S2 left was a red screw RTC adjusted the plan and surgeon confirmed that the adjustments were okay.1130 10 pt 1547 3pt 1700 mounted 1723 draped 1839 attached 1841-1844 3Define attempts, real sense dcm restarted, soft restart 1845 drawspine 1855 registered S1 left 1918 S2R 1922 S2L 1933 S1 R 1938. Overall surgeon was happy with the accuracy. | x shoulder issues led lights shoulder shoulder way left led lock shoulder continue blink true certain ranges surgeon reviewed plan prior procedure rtc discussed axial angles skive potential surgeon chose use single schanz right psis automatic registration achieved using basic algorithm good values surface scanner foundrecognized rtc attempted scan attempts rtc restarted realsense dcm application attempted soft restart advanced using generic work volume inflate coronal shift drawspine stylus challenging pick left red screw rtc adjusted plan surgeon confirmed adjustments pt mounted draped attached attempts real sense dcm restarted soft restart drawspine registered left r overall surgeon happy accuracy unk mazor x shoulder | NAV4108 | NAV3259 | NAV2058 | NAV3244 | NAV4106 | NAV4108NAV3259NAV2058NAV3244NAV4106 | False
7078 | 704115056 | XSE - Damaged Surgeon Monitor | 2020-12-09T19:04:27.000000+0000 | MazorX Stealth Edition | NAV2055 | PRODUCTS DAMAGED DURING SHIPPING | 20 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE - Damaged Surgeon Monitor 2020-12-09 00702382 (Rep): Medtronic received information regarding a guidance system being used during a demo procedure. It was reported that the surgeon monitor was broken and hanging by parts when the system was being uncrated. The system was able to be used. There was no patient involvement. 2020-12-18 IUD x4 (Other): No new information. 2021-Jan-13 IUDx3 (other): No new information. 2021-Feb-04 IUDx4 (other): No new information. 2021-02-20 IUD x2 (Other): No new information. 2021-03-09 IUD (Other): No new information. 2021-05-31 IUD (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2020-12-09 00702382 (Rep): Medtronic received information regarding a guidance system being used during a demo procedure. It was reported that the surgeon monitor was broken and hanging by parts when the system was being uncrated. The system was able to be used. There was no patient involvement. 2020-12-18 IUD x4 (Other): No new information. 2021-Jan-13 IUDx3 (other): No new information. 2021-Feb-04 IUDx4 (other): No new information. 2021-02-20 IUD x2 (Other): No new information. 2021-03-09 IUD (Other): No new information. 2021-05-31 IUD (Other): No new information. | xse damaged surgeon monitor demo procedure surgeon monitor broken hanging parts uncrated able iud new new new iud new iud new iud new kit svc srgn screen arm | NAV3253 | NAV2054 | NAV2058 | NAV3251 | NAV3249 | NAV3253NAV2054NAV2058NAV3251NAV3249 | False
1541 | 703205196 | X - Software Crash | 2019-06-13T13:07:39.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X - Software Crash This complaint will be separated to 4 complaints. This complaint will handle the SW crash (see bolted lines). 3Define cover came off - SCL1801461 X eye broken - SCL1801467 SW bugs - noise when no movement, CCR depth - SCL1801462 software crashed after system was placed onto bed. kicked out of operate mode and had to re-do accuracy test while arm was on bed. software crashed ~15 times after wires were completed and arm was still on bed. no bovie use was taking place on multiple crashes. system went black and kicked the software all the way back to the login screen. Dr. Hrynkiw Grandview Hospital 4.26.18 L4-L5 MIS PSF -planning done in clinic a few days prior to surgery-rt shanz pin platform-registration fast and accurate.-surgeon placed 4 wires Multiple issues encountered below: 1. software crashed after system was placed onto bed. kicked out of operate mode and had to re-do accuracy test while arm was on bed. 2. drawspine camera turned completely backwards. arm was properly calibrated prior to case. 3. software crashed ~15 times after wires were completed and arm was still on bed. no bovie use was taking place on multiple crashes. system went black and kicked the software all the way back to the login screen. 4. upon taking off drape, 3 define camera cover became dislodged extremely easy. needs to be looked at to make sure it's not broken. 5. trajectory depth on CCR says L3 rt depth was edited, this is not correct. 6. On Monday, 4/23, RTC reported during CS training at hospital that when demo case was being performed, arm was lowered to simulate moving the arm closer to the patient and right when "lower arm" was clicked, the software displayed "trajectory no longer accurate" and when we clicked lower the arm again and then stopped, the arm was no longer moving but the workstation was making the audible sound like the arm was still moving. Have not heard back from service team about this issue. MAZOR X SOFTWARE KIT VER-2.3 | This complaint will be separated to 4 complaints. This complaint will handle the SW crash (see bolted lines). 3Define cover came off - SCL1801461 X eye broken - SCL1801467 SW bugs - noise when no movement, CCR depth - SCL1801462 software crashed after system was placed onto bed. kicked out of operate mode and had to re-do accuracy test while arm was on bed. software crashed ~15 times after wires were completed and arm was still on bed. no bovie use was taking place on multiple crashes. system went black and kicked the software all the way back to the login screen. Dr. Hrynkiw Grandview Hospital 4.26.18 L4-L5 MIS PSF -planning done in clinic a few days prior to surgery-rt shanz pin platform-registration fast and accurate.-surgeon placed 4 wires Multiple issues encountered below: 1. software crashed after system was placed onto bed. kicked out of operate mode and had to re-do accuracy test while arm was on bed. 2. drawspine camera turned completely backwards. arm was properly calibrated prior to case. 3. software crashed ~15 times after wires were completed and arm was still on bed. no bovie use was taking place on multiple crashes. system went black and kicked the software all the way back to the login screen. 4. upon taking off drape, 3 define camera cover became dislodged extremely easy. needs to be looked at to make sure it's not broken. 5. trajectory depth on CCR says L3 rt depth was edited, this is not correct. 6. On Monday, 4/23, RTC reported during CS training at hospital that when demo case was being performed, arm was lowered to simulate moving the arm closer to the patient and right when "lower arm" was clicked, the software displayed "trajectory no longer accurate" and when we clicked lower the arm again and then stopped, the arm was no longer moving but the workstation was making the audible sound like the arm was still moving. Have not heard back from service team about this issue. | x software crash complaint separated complaints complaint handle sw crash cover came x eye broken sw bugs noise movement ccr depth software crashed placed bed kicked operate mode redo accuracy test arm bed software crashed times wires completed arm bed bovie use taking place multiple crashes went black kicked software way login screen dr hrynkiw grandview hospital mis psf planning clinic days prior surgeryrt shanz pin platformregistration fast accuratesurgeon placed wires multiple issues encountered software crashed placed bed kicked operate mode redo accuracy test arm bed drawspine camera turned completely backwards arm properly calibrated prior case software crashed times wires completed arm bed bovie use taking place multiple crashes went black kicked software way login screen taking drape define camera cover dislodged extremely easy needs looked make sure broken trajectory depth ccr says rt depth edited correct monday rtc cs training hospital demo case performed arm lowered simulate moving arm closer right lower arm clicked software displayed trajectory longer accurate clicked lower arm stopped arm longer moving workstation making audible sound like arm moving heard service team issue mazor x software kit | NAV3247 | NAV3248 | NAV1012 | NAV4099 | NAV4095 | NAV3247NAV3248NAV1012NAV4099NAV4095 | True
8284 | 704388361 | XSE ¦ Navigated Tools Database Miss | 2021-05-11T15:49:30.000000+0000 | MazorX Stealth Edition | NAV2056 | SOFTWARE ISSUES | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE ¦ Navigated Tools Database Miss 2021-May-11 Salesforce C00226218 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the manufacturer representative booted the system on this morning, he received the following three error messages - Error 9254 - Failed to load navigation module nav mode will be disabled Error 10145 - MDT DB installation failed MDT implants will not be supported Database version mismatch He tried to hard reboot the system waiting 5 minutes in between each time with no resolution. No patient was present. 2021-05-12 IUD x3 (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-May-11 Salesforce C00226218 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the manufacturer representative booted the system on this morning, he received the following three error messages - Error 9254 - Failed to load navigation module nav mode will be disabled Error 10145 - MDT DB installation failed MDT implants will not be supported Database version mismatch He tried to hard reboot the system waiting 5 minutes in between each time with no resolution. No patient was present. 2021-05-12 IUD x3 (Other): No new information. | xse navigated tools database miss salesforce outside procedure booted morning following error messages error failed load navigation module nav mode disabled error mdt db installation failed mdt implants supported database version mismatch tried hard reboot waiting minutes time resolution present iud new mazorx sw installation kit | NAV3246 | NAV2047 | NAV3247 | NAV2056 | NAV3249 | NAV3246NAV2047NAV3247NAV2056NAV3249 | True
579 | 703152732 | X - shoulder shift | 2019-05-14T14:18:16.000000+0000 | MazorX | NAV4108 | SHOULDER SHIFTS | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - shoulder shift 2019-May-09 SCL1901160 intake form, CCR (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that a shoulder shift occurred after the last trajectory was sent and the drill guide was being removed on two cases in a row. Nothing needed to be done since the wires were already in and accurate. Shoulder values and an accuracy test were performed before the case. The rep felt there could be a joint issue. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X SHOULDER | 2019-May-09 SCL1901160 intake form, CCR (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that a shoulder shift occurred after the last trajectory was sent and the drill guide was being removed on two cases in a row. Nothing needed to be done since the wires were already in and accurate. Shoulder values and an accuracy test were performed before the case. The rep felt there could be a joint issue. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x shoulder shift ccr spinal procedure shoulder shift occurred trajectory sent drill guide removed cases row needed wires accurate shoulder values accuracy test performed case felt joint issue additional complications anticipated unk mazor x shoulder | NAV4108 | NAV4106 | NAV4107 | NAV3261 | NAV3252 | NAV4108NAV4106NAV4107NAV3261NAV3252 | True
8177 | 704366334 | XSE - Manipulator Will Not Release | 2021-04-28T16:24:08.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 30 | KIT0925-01 | SVC KIT0925-01 MANIPULATOR SOLENOID | XSE - Manipulator Will Not Release 2021-Apr-28 Salesforce C00225440 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manipulator would not release from the positioner during a software upgrade of the guidance system. There was no patient involvement. 2021-04-30 IUD x4 (Other): No new information. 2021-06-17 IUDx3 (other): No new information. 2021-12-01 302651448, IUD x5 (Other): No new information. 2021-12-02 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. SVC KIT0925-01 MANIPULATOR SOLENOID | 2021-Apr-28 Salesforce C00225440 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manipulator would not release from the positioner during a software upgrade of the guidance system. There was no patient involvement. 2021-04-30 IUD x4 (Other): No new information. 2021-06-17 IUDx3 (other): No new information. 2021-12-01 302651448, IUD x5 (Other): No new information. 2021-12-02 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse manipulator release salesforce outside procedure manipulator release positioner software upgrade iud new new iud new iud new iud new svc manipulator solenoid | NAV3260 | NAV3249 | NAV2047 | NAV3251 | NAV3264 | NAV3260NAV3249NAV2047NAV3251NAV3264 | True
145 | 703014953 | XSE-tablet not charging/accuracy pointer | 2019-02-19T16:47:57.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | MAS2285-01 | ACCURACY POINTER | XSE-tablet not charging/accuracy pointer 2019-Feb-12 SCL1900367, intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon tablet would not hold a charge and the accuracy pointer was bent. It was noted there was no patient harm and no delay in the surgery. There were no additional complications reported or anticipated as a result of the event. ACCURACY POINTER | 2019-Feb-12 SCL1900367, intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon tablet would not hold a charge and the accuracy pointer was bent. It was noted there was no patient harm and no delay in the surgery. There were no additional complications reported or anticipated as a result of the event. | xsetablet chargingaccuracy pointer spinal procedure surgeon tablet hold charge accuracy pointer bent noted delay surgery additional complications anticipated accuracy pointer | NAV3299 | NAV3295 | NAV3253 | NAV1006 | NAV2054 | NAV3299NAV3295NAV3253NAV1006NAV2054 | True
8436 | 704414162 | XSE - System Did Not Work | 2021-05-25T21:24:48.000000+0000 | MazorX Stealth Edition | NAV1013 | MAZOR X/STEALTH EDITION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - System Did Not Work 2021-04-27 Email (HCP): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the system did not work. There was no patient reported present at the time. 2021-05-25 email 2021-05-25 16:46 (Other): No new information. 2021-06-16 email 2021-06-16 10:31, email 2021-06-26 10:31, email 2021-07-05 14:31, email 2021-07-15 20:15 (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-04-27 Email (HCP): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the system did not work. There was no patient reported present at the time. 2021-05-25 email 2021-05-25 16:46 (Other): No new information. 2021-06-16 email 2021-06-16 10:31, email 2021-06-26 10:31, email 2021-07-05 14:31, email 2021-07-15 20:15 (Other): No new information. | xse did work email outside procedure did work present time email new email email email email new mazor x | NAV3246 | NAV5000 | NAV2047 | NAV3242 | NAV1012 | NAV3246NAV5000NAV2047NAV3242NAV1012 | False
1836 | 703218593 | X- SW Issues | 2019-06-20T19:41:00.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X- SW Issues Registration issues - will be more specific and add details after receiving IF and CCR From CCR: 1st registration-Automatic registration was achieved using the basic algorithm off of L5 with good values. L5 and L4 bilateral trajectories were omitted due to existing implanted hardware. 2nd registration-automatic registration failed, attempting semi-automatic registration did not allow RTC to complete registration. Manual registrationwas also attempted without success. Labeling of the vertebral levels showed different then intended on registration screen leading to a mismatch when switching over to registration from segmentation. Registration was abandoned, RTC started a completely new registration including new 3Define and Draw Spine. Once AP and OBL locations were sent images were obtained off T12 body and new registration passed using automatic registration algorithm. MAZOR X SOFTWARE KIT VER-2.3.1 | Registration issues - will be more specific and add details after receiving IF and CCR From CCR: 1st registration-Automatic registration was achieved using the basic algorithm off of L5 with good values. L5 and L4 bilateral trajectories were omitted due to existing implanted hardware. 2nd registration-automatic registration failed, attempting semi-automatic registration did not allow RTC to complete registration. Manual registrationwas also attempted without success. Labeling of the vertebral levels showed different then intended on registration screen leading to a mismatch when switching over to registration from segmentation. Registration was abandoned, RTC started a completely new registration including new 3Define and Draw Spine. Once AP and OBL locations were sent images were obtained off T12 body and new registration passed using automatic registration algorithm. | x sw issues registration issues specific add details receiving ccr ccr registrationautomatic registration achieved using basic algorithm good values bilateral trajectories omitted existing implanted hardware registrationautomatic registration failed attempting semiautomatic registration did allow rtc complete registration manual registrationwas attempted success labeling vertebral levels showed different intended registration screen leading mismatch switching registration segmentation registration abandoned rtc started completely new registration including new draw spine ap obl locations sent images obtained body new registration passed using automatic registration algorithm mazor x software kit | NAV3242 | NAV4112 | NAV4091 | NAV3247 | NAV3240 | NAV3242NAV4112NAV4091NAV3247NAV3240 | False
8413 | 704409097 | XSE - System Unresponsive | 2021-05-21T23:56:11.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 30 | KIT0830 | USA REGIONAL KIT KIT0830 | XSE - System Unresponsive 2021-05-21 C00226882 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the system was unresponsive. When setting up for a case the manufacturer representative got an error for the X-eye camera not being plugged in. She went to clear the error message but the system got stuck with processing on the monitor. She let the system sit for 10 minutes to make sure it was not just operating slowly. However the processing error did not get cleared. She then did a soft shutdown on the system and was able to get the error to clear. No patient was present. USA REGIONAL KIT KIT0830 | 2021-05-21 C00226882 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the system was unresponsive. When setting up for a case the manufacturer representative got an error for the X-eye camera not being plugged in. She went to clear the error message but the system got stuck with processing on the monitor. She let the system sit for 10 minutes to make sure it was not just operating slowly. However the processing error did not get cleared. She then did a soft shutdown on the system and was able to get the error to clear. No patient was present. | xse unresponsive outside procedure unresponsive setting case got error xeye camera plugged went clear error message got stuck processing monitor let sit minutes make sure just operating slowly processing error did cleared did soft shutdown able error clear present usa regional kit | NAV3248 | NAV3262 | NAV3249 | NAV4110 | NAV2047 | NAV3248NAV3262NAV3249NAV4110NAV2047 | True
4360 | 703537939 | X - 3Define Failure, Left L5 Deviation | 2019-12-30T14:54:49.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - 3Define Failure, Left L5 Deviation 2018-09-11 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed for a PSF at L2-L5 with a revision of pre-existing L4-L5 screws and placement of interbodies at L2-L3 and L3-L4. The 3Define did not execute and the generic work volume had to be used. Guidewires were placed bilaterally at L2 and L3. Existing hardware at L4-L5 was removes and guidewires were placed at new trajectories. Left L5 stimulated low. The pilot hole was probed but no breach was found. The surgeon deem the screw to be slightly medial. The level was fused from a previous surgery so the surgeon opted to leave that screw out the construct. Other trajectories were placed accurately. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2018-09-11 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed for a PSF at L2-L5 with a revision of pre-existing L4-L5 screws and placement of interbodies at L2-L3 and L3-L4. The 3Define did not execute and the generic work volume had to be used. Guidewires were placed bilaterally at L2 and L3. Existing hardware at L4-L5 was removes and guidewires were placed at new trajectories. Left L5 stimulated low. The pilot hole was probed but no breach was found. The surgeon deem the screw to be slightly medial. The level was fused from a previous surgery so the surgeon opted to leave that screw out the construct. Other trajectories were placed accurately. No patient harm was reported and surgery was delayed less than an hour. | x failure left deviation ccr spinal procedure planning completed psf revision preexisting screws placement interbodies did execute generic work volume guidewires placed bilaterally existing hardware removes guidewires placed new trajectories left stimulated low pilot hole probed breach surgeon deem screw slightly medial level fused previous surgery surgeon opted leave screw construct trajectories placed accurately surgery delayed hour mazor x | NAV4091 | NAV5000 | NAV3242 | NAV3243 | NAV2047 | NAV4091NAV5000NAV3242NAV3243NAV2047 | True
6416 | 703937207 | XSE - Missing components | 2020-09-02T21:47:47.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | XSE - Missing components 2020-08-31 00695404, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the universal screw driver and accuracy pointer was missing. There was no patient involvement. 2020-09-09 IUD (Other): No new information. 2020-09-15 IUD (Other): No new information. 2020-09-15 e1 (Rep): No new information. 2020-10-06 IUD (Other): No new information. 2020-10-13, IUD e1 (Other): No new information. 2020-11-20 IUD x2 (Other): No new information. 2021-03-16 IUD (Other): No new information. UNK MAZOR X ACCESSORIES | 2020-08-31 00695404, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the universal screw driver and accuracy pointer was missing. There was no patient involvement. 2020-09-09 IUD (Other): No new information. 2020-09-15 IUD (Other): No new information. 2020-09-15 e1 (Rep): No new information. 2020-10-06 IUD (Other): No new information. 2020-10-13, IUD e1 (Other): No new information. 2020-11-20 IUD x2 (Other): No new information. 2021-03-16 IUD (Other): No new information. | xse missing components iud outside procedure universal screw driver accuracy pointer missing iud new iud new new iud new iud new iud new iud new unk mazor x accessories | NAV2053 | NAV3299 | NAV2054 | NAV1006 | NAV3262 | NAV2053NAV3299NAV2054NAV1006NAV3262 | True
10270 | 704843462 | XSE - Unable to Connect to C-Arm | 2022-02-02T21:05:40.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Unable to Connect to C-Arm 2022-Feb-02 Salesforce C00242613 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was unable to be connected to the site's new C-arm. A flat panel adapter kit was received, but the manufacturer representative still had an issue connecting the system. There was no patient involvement. 2022-02-08 IUD x3 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved s etting a static IP on the C-arm, added a new DICOM device to the system using the setup page, use DICOM storage server setup on the C-arm, assigned an IP address on the guidance system, and then trying to push an exam to the system, but that failed. The representative noted that they were unaware that the C-arm had been previously using DHCP to connect to the hospital network and the previous DICOM device had been set up, but filled out correctly. The representative suspected the cause was having two DICOM devices on the system. 2022-02-22 IUD x4 (Other): No new information. UNK MAZOR X WORKSTATION | 2022-Feb-02 Salesforce C00242613 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was unable to be connected to the site's new C-arm. A flat panel adapter kit was received, but the manufacturer representative still had an issue connecting the system. There was no patient involvement. 2022-02-08 IUD x3 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved s etting a static IP on the C-arm, added a new DICOM device to the system using the setup page, use DICOM storage server setup on the C-arm, assigned an IP address on the guidance system, and then trying to push an exam to the system, but that failed. The representative noted that they were unaware that the C-arm had been previously using DHCP to connect to the hospital network and the previous DICOM device had been set up, but filled out correctly. The representative suspected the cause was having two DICOM devices on the system. 2022-02-22 IUD x4 (Other): No new information. | xse unable connect carm salesforce outside procedure unable connected sites new carm flat panel adapter kit issue connecting iud additional troubleshooting involved s etting static ip carm added new dicom using setup page use dicom storage server setup carm assigned ip address trying push exam failed noted unaware carm previously using dhcp connect hospital network previous dicom set filled correctly suspected cause having dicom devices iud new unk mazor x workstation | NAV2058 | NAV3258 | NAV4108 | NAV3264 | NAV2047 | NAV2058NAV3258NAV4108NAV3264NAV2047 | True
564 | 703150807 | XSE - stripped screw/screwdriver loose | 2019-05-13T18:35:40.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | XSE - stripped screw/screwdriver loose 2019-May-09 SCL1901147 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the screws stripped on the Rear Panel Adapter for Tablet and the Universal Screwdriver handle was loose and would not tighten. There was no patient harm. There were no additional complications reported or a nticipated as a result of the event. UNK MAZOR X ACCESSORIES | 2019-May-09 SCL1901147 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the screws stripped on the Rear Panel Adapter for Tablet and the Universal Screwdriver handle was loose and would not tighten. There was no patient harm. There were no additional complications reported or a nticipated as a result of the event. | xse stripped screwscrewdriver loose spinal procedure screws stripped rear panel adapter tablet universal screwdriver handle loose tighten additional complications nticipated unk mazor x accessories | NAV3299 | NAV3295 | NAV2054 | NAV3251 | NAV4097 | NAV3299NAV3295NAV2054NAV3251NAV4097 | True
3664 | 703409144 | X - paint/shoulder button not working | 2019-10-11T16:02:36.000000+0000 | MazorX | NAV3252 | SHOULDER ISSUES | 20 | ASM0207-02 | SHOULDER | X - paint/shoulder button not working 2019-Oct-10 SCL1903049, SCL1903050 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that there was a paint problem and the shoulder handle button was not working. There was no patient harm or additional complications reported or anticipated as a result of the event. SHOULDER | 2019-Oct-10 SCL1903049, SCL1903050 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that there was a paint problem and the shoulder handle button was not working. There was no patient harm or additional complications reported or anticipated as a result of the event. | x paintshoulder button working outside procedure paint problem shoulder handle button working additional complications anticipated shoulder | NAV2054 | NAV3249 | NAV3252 | NAV3251 | NAV4107 | NAV2054NAV3249NAV3252NAV3251NAV4107 | True
7098 | 704116681 | XSE - SRA Not on Trajectory | 2020-12-10T07:04:19.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 30 | ASM0206-01 | SURGICAL ARM | XSE - SRA Not on Trajectory 2020-12-06 00702425 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon placed the first screw successfully during a T8-L3 spinal fusion. During placement of the second screw, a "Surgical Arm not on trajectory" error message was displayed. The error message occurred when the surgeon was using very delicately using the dilator and the use would not typically cause the error to appear. The surgical arm was resent to the trajectory, but a error 8798 at Axis A4 was displayed. The surgical arm was then sent to the clear up position, but a 8798 (Max Closing Tolerance) error was displayed. A Fatal Error 295 was displayed when the surgical arm was sent attempted to be sent to the clear up position again. The surgical could not be moved. The emergency stop button was pressed on the surgical arm and all lights remained green on the arm. The surgical arm was manually moved into the clear up position and a hard restart of the system was done. The surgical arm passed an advanced accuracy test. During troubleshooting, the surgeon had placed three screws manually. The guidance system was used to place the remaining screws. There was no patient harm and the procedure was delayed less than an hour. 2020-12-10 email 2020-12-10 16:45 (Rep, For): Additional information received from a manufacturer representative reported that J4 of the surgical arm did not achieve the target location and the cause was unknown. 2021-02-01 IUD (Other, For): No new information. 2021-01-06 email 2021-01-06 16:45 (Other, For): No new information. 2021-02-02 email 2021-02-02 18:45, IUD (Rep, For): No new information. 2021-Mar-12 email 3/12 (Rep, For): No new information. 2022-01-18 IUD (Other, For): No new information. 2022-02-02 302712302 (Other, For): No new information. 2022-02-14 IUD x4 (Other): No new information. SURGICAL ARM | 2020-12-06 00702425 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon placed the first screw successfully during a T8-L3 spinal fusion. During placement of the second screw, a "Surgical Arm not on trajectory" error message was displayed. The error message occurred when the surgeon was using very delicately using the dilator and the use would not typically cause the error to appear. The surgical arm was resent to the trajectory, but a error 8798 at Axis A4 was displayed. The surgical arm was then sent to the clear up position, but a 8798 (Max Closing Tolerance) error was displayed. A Fatal Error 295 was displayed when the surgical arm was sent attempted to be sent to the clear up position again. The surgical could not be moved. The emergency stop button was pressed on the surgical arm and all lights remained green on the arm. The surgical arm was manually moved into the clear up position and a hard restart of the system was done. The surgical arm passed an advanced accuracy test. During troubleshooting, the surgeon had placed three screws manually. The guidance system was used to place the remaining screws. There was no patient harm and the procedure was delayed less than an hour. 2020-12-10 email 2020-12-10 16:45 (Rep, For): Additional information received from a manufacturer representative reported that J4 of the surgical arm did not achieve the target location and the cause was unknown. 2021-02-01 IUD (Other, For): No new information. 2021-01-06 email 2021-01-06 16:45 (Other, For): No new information. 2021-02-02 email 2021-02-02 18:45, IUD (Rep, For): No new information. 2021-Mar-12 email 3/12 (Rep, For): No new information. 2022-01-18 IUD (Other, For): No new information. 2022-02-02 302712302 (Other, For): No new information. 2022-02-14 IUD x4 (Other): No new information. | xse sra trajectory spinal procedure surgeon placed screw successfully spinal fusion placement second screw surgical arm trajectory error message displayed error message occurred surgeon using delicately using dilator use typically cause error appear surgical arm resent trajectory error axis displayed surgical arm sent clear position error displayed fatal error displayed surgical arm sent attempted sent clear position surgical moved emergency stop button pressed surgical arm lights remained green arm surgical arm manually moved clear position hard restart surgical arm passed advanced accuracy test troubleshooting surgeon placed screws manually place remaining screws procedure delayed hour email additional surgical arm did achieve target location cause unknown iud new email new email iud new email new iud new new iud new surgical arm | NAV4101 | NAV2047 | NAV3246 | NAV3249 | NAV4095 | NAV4101NAV2047NAV3246NAV3249NAV4095 | True
5220 | 703619852 | X - BMB Plate Stripped | 2020-02-17T20:06:47.000000+0000 | MazorX | NAV4100 | STRIPPED THREADS | 20 | MAS2188-01 | BM BRIDGE | X - BMB Plate Stripped 2020-02-17 00684254 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the BMB plate is stripped and needs to be replaced. There was no patient involvement. 2020-02-25 IUD x2: No new information. 2020-03-12 IUD x3: No new information. 2020-03-25 Email 2020-03-25 11:45, IUD x2: No new information. 2020-04-14 IUD x3: No new information. 2020-04-29 IUD x3 (Other): No new information. BM BRIDGE | 2020-02-17 00684254 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the BMB plate is stripped and needs to be replaced. There was no patient involvement. 2020-02-25 IUD x2: No new information. 2020-03-12 IUD x3: No new information. 2020-03-25 Email 2020-03-25 11:45, IUD x2: No new information. 2020-04-14 IUD x3: No new information. 2020-04-29 IUD x3 (Other): No new information. | x bmb plate stripped outside procedure bmb plate stripped needs replaced iud new iud new email iud new iud new iud new bm bridge | NAV4100 | NAV3299 | NAV3295 | NAV1006 | NAV3251 | NAV4100NAV3299NAV3295NAV1006NAV3251 | True
6786 | 704031337 | X - 3Define Not Functioning | 2020-10-23T18:36:45.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | KIT0749-01 | KIT SPARE PART FOR THE 3D CAMERA | X - 3Define Not Functioning 2020-10-23 00699400 (Rep): Medtronic received information regarding a guidance system being during a spinal procedure. It was reported that the 3Define scan failed on the first attempt. A soft restart was done and the second scan worked. The case was able to be continued. there was no patient harm and the procedure was delayed less than an hour. KIT SPARE PART FOR THE 3D CAMERA | 2020-10-23 00699400 (Rep): Medtronic received information regarding a guidance system being during a spinal procedure. It was reported that the 3Define scan failed on the first attempt. A soft restart was done and the second scan worked. The case was able to be continued. there was no patient harm and the procedure was delayed less than an hour. | x functioning spinal procedure scan failed attempt soft restart second scan worked case able continued procedure delayed hour kit spare camera | NAV3244 | NAV3246 | NAV3299 | NAV3247 | NAV3249 | NAV3244NAV3246NAV3299NAV3247NAV3249 | True
4686 | 703562338 | XSE - Joint 5 Not Locking | 2020-01-15T00:03:25.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Joint 5 Not Locking 2020-01-14 00682076 (REP, HCP): Medtronic received information regarding a guidance system in use. It was reported that the system will not lock J5. When the representative goes to move the arm, joint five will not lock and slips. There was no patient involved. 2020-01-17 email 2020-01-17 1306 (Rep): Additional information received from the rep reporting that no troubleshooting was completed at the time of the event, only that tech support was called. The rep felt certain that they would not be able to fix the issue within the manipulator sensor. 2020-01-16 IUD x3 (Other): Additional information received from a field service engineer reporting that the system was checked on-site and the complaint was confirmed. It was determined that the pneumatic line was pinched at the manipulator D5. This was repaired on-site and the system was fully functional. SURGICAL ARM | 2020-01-14 00682076 (REP, HCP): Medtronic received information regarding a guidance system in use. It was reported that the system will not lock J5. When the representative goes to move the arm, joint five will not lock and slips. There was no patient involved. 2020-01-17 email 2020-01-17 1306 (Rep): Additional information received from the rep reporting that no troubleshooting was completed at the time of the event, only that tech support was called. The rep felt certain that they would not be able to fix the issue within the manipulator sensor. 2020-01-16 IUD x3 (Other): Additional information received from a field service engineer reporting that the system was checked on-site and the complaint was confirmed. It was determined that the pneumatic line was pinched at the manipulator D5. This was repaired on-site and the system was fully functional. | xse joint locking use lock goes arm joint lock slips involved email additional reporting troubleshooting completed time tech support called felt certain able fix issue manipulator sensor iud additional field service engineer reporting checked onsite complaint confirmed determined pneumatic line pinched manipulator repaired onsite fully functional surgical arm | NAV3260 | NAV4098 | NAV3264 | NAV3249 | NAV2044 | NAV3260NAV4098NAV3264NAV3249NAV2044 | True
3275 | 703301946 | XSE - Surgical Arm Joint Issue/OOBF | 2019-08-08T19:20:36.000000+0000 | MazorX Stealth Edition | NAV3294 | ARM GUIDE ISSUES | 30 | KIT0897 | 5.5/6.0 ARM GUIDE | XSE - Surgical Arm Joint Issue/OOBF 2019-08-05 SCL1902302, PDF (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm had a questionable joint. There was no patient involved. 2019-08-08 SCL1902347, PDF (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when conducting an accuracy test, the pointer was not aligning correctly with the reference point. Troubleshooting revealed an issue with a defective arm guide. Two other arm guides were tested with the system and there were no errors with the accuracy tests. An arm guide from a different system was also tested and there were no accuracy issues. A case scheduled later that week was completed without any issues. Information unrelated to this event was included in a separate PE; See PE 703306859 for No Registration. 5.5/6.0 ARM GUIDE | 2019-08-05 SCL1902302, PDF (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm had a questionable joint. There was no patient involved. 2019-08-08 SCL1902347, PDF (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when conducting an accuracy test, the pointer was not aligning correctly with the reference point. Troubleshooting revealed an issue with a defective arm guide. Two other arm guides were tested with the system and there were no errors with the accuracy tests. An arm guide from a different system was also tested and there were no accuracy issues. A case scheduled later that week was completed without any issues. Information unrelated to this event was included in a separate PE; See PE 703306859 for No Registration. | xse surgical arm joint issueoobf pdf outside procedure surgical arm questionable joint involved pdf outside procedure conducting accuracy test pointer aligning correctly reference point troubleshooting revealed issue defective arm guide arm guides tested errors accuracy tests arm guide different tested accuracy issues case scheduled later week completed issues unrelated included separate pe pe registration arm guide | NAV3294 | NAV3249 | NAV4092 | NAV4098 | NAV4097 | NAV3294NAV3249NAV4092NAV4098NAV4097 | True
7125 | 704123095 | XSE ¦ Surgeon Monitor Not Powering | 2020-12-14T17:20:27.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE ¦ Surgeon Monitor Not Powering 2020-Dec-14 00702607 IUD (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used prior to a procedure. It was reported that when powering on the system, the surgeon monitor was not receiving power. He did a soft restart and the surgeon monitor was still not receiving power. He did a hard restart and then was receiving power to the surgeon monitor. The surgeon screen did cut out one more time while they were draping. He didn't reboot at that time. There was no patient harm or additional complications reported or anticipated as a result of the event. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2020-Dec-14 00702607 IUD (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used prior to a procedure. It was reported that when powering on the system, the surgeon monitor was not receiving power. He did a soft restart and the surgeon monitor was still not receiving power. He did a hard restart and then was receiving power to the surgeon monitor. The surgeon screen did cut out one more time while they were draping. He didn't reboot at that time. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse surgeon monitor powering iud prior procedure powering surgeon monitor receiving power did soft restart surgeon monitor receiving power did hard restart receiving power surgeon monitor surgeon screen did cut time draping didnt reboot time additional complications anticipated kit svc srgn screen arm | NAV3253 | NAV3263 | NAV3261 | NAV2047 | NAV5000 | NAV3253NAV3263NAV3261NAV2047NAV5000 | True
114 | 703000386 | X - Drape Issue | 2019-02-11T11:56:51.000000+0000 | MazorX | NAV2048 | DOCUMENT ISSUES | 20 | KIT0574 | MAZOR X SPINE DISPOSABLE KIT KIT0574 | X - Drape Issue 2019-01-31 SCL1900323, Intake Form 2019-02-11 06:20, Image 2019-02-11 06:20 x3, e1, e2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were plastic "hairs" on the welding line of the new drape for the surgical system and no indication for use (IFUs) were included in the box. There was no patient harm and surgery was no delayed more than an hour. The case was successfully completed. 2019-02-11 email 2019-02-26 14:34 (Rep): Additional information received from a manufacturer representative reported that two indications for use (IFUs) were not included in the box. MAZOR X SPINE DISPOSABLE KIT KIT0574 | 2019-01-31 SCL1900323, Intake Form 2019-02-11 06:20, Image 2019-02-11 06:20 x3, e1, e2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were plastic "hairs" on the welding line of the new drape for the surgical system and no indication for use (IFUs) were included in the box. There was no patient harm and surgery was no delayed more than an hour. The case was successfully completed. 2019-02-11 email 2019-02-26 14:34 (Rep): Additional information received from a manufacturer representative reported that two indications for use (IFUs) were not included in the box. | x drape issue image spinal procedure plastic hairs welding line new drape surgical indication use included box surgery delayed hour case successfully completed email additional indications use included box mazor x spine disposable kit | NAV3232 | NAV2054 | NAV3246 | NAV3248 | NAV5000 | NAV3232NAV2054NAV3246NAV3248NAV5000 | False
3673 | 703411717 | X - Communication Cable Malfunction | 2019-10-14T19:31:30.000000+0000 | MazorX | NAV3256 | CABLE ISSUES | 20 | ASM0206-01 | SURGICAL ARM | X - Communication Cable Malfunction 2019-06-02 Attachment (REP): Medtronic received information regarding a guidance system during servicing. It was reported that there was a malfunction with the communication cable. No patient was present. SURGICAL ARM | 2019-06-02 Attachment (REP): Medtronic received information regarding a guidance system during servicing. It was reported that there was a malfunction with the communication cable. No patient was present. | x communication cable malfunction attachment servicing malfunction communication cable present surgical arm | NAV2047 | NAV3256 | NAV3249 | NAV2054 | NAV4091 | NAV2047NAV3256NAV3249NAV2054NAV4091 | True
1597 | 703207751 | X- WS Issues | 2019-06-14T14:40:26.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- WS Issues The D5 manipulator doesn't straighten and makes it difficult to get off the bed. The break near the D5 sensor is broken. The robot loses pressure to 4.7 or lower once the robot comes off bed. Audible hissing (air leak) noise heard once robot is off the bed. Issue being solved on SCL1801028. UNK MAZOR X WORKSTATION | The D5 manipulator doesn't straighten and makes it difficult to get off the bed. The break near the D5 sensor is broken. The robot loses pressure to 4.7 or lower once the robot comes off bed. Audible hissing (air leak) noise heard once robot is off the bed. Issue being solved on SCL1801028. | x ws issues manipulator doesnt straighten makes difficult bed break near sensor broken robot loses pressure lower robot comes bed audible hissing noise heard robot bed issue solved unk mazor x workstation | NAV3264 | NAV3260 | NAV3250 | NAV3255 | NAV3295 | NAV3264NAV3260NAV3250NAV3255NAV3295 | True
6727 | 704015424 | XSE - Damaged Image Adaptor Latch | 2020-10-15T20:15:26.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 20 | KIT0586-03 | REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | XSE - Damaged Image Adaptor Latch 2020-10-15 00698872 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter latch did not pinch down or rotate up and down. There was no patient involvement. 2020-11-16 IUD (Other): No new information. 2020-11-18 email 2020-11-18 07:45, email 2020-11-18 07:46, IUD (Rep): No new information. 2020-12-11 e1, e2 (Other): No new information. 2020-12-16 IUD (Other): No new information. REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | 2020-10-15 00698872 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter latch did not pinch down or rotate up and down. There was no patient involvement. 2020-11-16 IUD (Other): No new information. 2020-11-18 email 2020-11-18 07:45, email 2020-11-18 07:46, IUD (Rep): No new information. 2020-12-11 e1, e2 (Other): No new information. 2020-12-16 IUD (Other): No new information. | xse damaged image adaptor latch outside procedure image adapter latch did pinch rotate iud new email email iud new new iud new replacement grip set image adaptor | NAV3258 | NAV2054 | NAV3299 | NAV3295 | NAV3237 | NAV3258NAV2054NAV3299NAV3295NAV3237 | True
2858 | 703256439 | X - Shoulder Issue / Registration Issue | 2019-07-12T16:23:34.000000+0000 | MazorX | NAV4108 | SHOULDER SHIFTS | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Shoulder Issue / Registration Issue Shoulder shift occurred during instrumentation with little pressure applied to system, unable to reach OBL trajectory , registration issues after initially achieving good values on same segment earlier in lab UNK MAZOR X SHOULDER | Shoulder shift occurred during instrumentation with little pressure applied to system, unable to reach OBL trajectory , registration issues after initially achieving good values on same segment earlier in lab | x shoulder issue registration issue shoulder shift occurred instrumentation little pressure applied unable reach obl trajectory registration issues initially achieving good values segment earlier lab unk mazor x shoulder | NAV4108 | NAV4095 | NAV4107 | NAV3261 | NAV4101 | NAV4108NAV4095NAV4107NAV3261NAV4101 | True
6678 | 703999908 | XSE - Schanz Arm Not Locking | 2020-10-07T20:10:34.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Schanz Arm Not Locking 2020-10-07 00698333 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arm would not lock into place. The issue was discovered outside of a case. There was no patient involvement. 2020-Nov-05 email 11/5 (rep): Additional information indicated that the accounts equipment repair team was able to fix both Schanz arms internally. They will not be sent back or replaced at this time. KIT SVC MAS2351-04S SCHANZ ARM | 2020-10-07 00698333 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arm would not lock into place. The issue was discovered outside of a case. There was no patient involvement. 2020-Nov-05 email 11/5 (rep): Additional information indicated that the accounts equipment repair team was able to fix both Schanz arms internally. They will not be sent back or replaced at this time. | xse schanz arm locking outside procedure schanz arm lock place issue discovered outside case email additional indicated accounts equipment repair team able fix schanz arms internally sent replaced time kit svc schanz arm | NAV3299 | NAV3298 | NAV4093 | NAV2053 | NAV3244 | NAV3299NAV3298NAV4093NAV2053NAV3244 | True
322 | 703078085 | XSE - 3Define failed/arm error | 2019-03-28T14:40:31.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | XSE - 3Define failed/arm error 2019-Mar-26 SCL1900744 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3Define scan failed and the rep had to use a generic work volume. During the last screw placement, the rep received the following error, "Error 8025 Pre-defined Position is Not Reachable. Check Collision with Shoulder/Postioner or WorkVolume". The surgeon requested the rep Clear Up and but the Arm would not clear up. The surgeon then grabbed the arm and manually moved it out of the way. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Mar-26 SCL1900744 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3Define scan failed and the rep had to use a generic work volume. During the last screw placement, the rep received the following error, "Error 8025 Pre-defined Position is Not Reachable. Check Collision with Shoulder/Postioner or WorkVolume". The surgeon requested the rep Clear Up and but the Arm would not clear up. The surgeon then grabbed the arm and manually moved it out of the way. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse failedarm error spinal procedure scan failed use generic work volume screw placement following error error predefined position reachable check collision shoulderpostioner workvolume surgeon requested clear arm clear surgeon grabbed arm manually moved way noted additional complications anticipated surgical arm | NAV3249 | NAV4095 | NAV3244 | NAV4101 | NAV5000 | NAV3249NAV4095NAV3244NAV4101NAV5000 | True
10402 | 704874760 | XSE - OOBF/SRA Issue, Pneumatic Line Cut | 2022-02-21T17:07:29.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 40 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - OOBF/SRA Issue, Pneumatic Line Cut 2022-Feb-21 Salesforce C00243946 (Rep, HCP): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the manufacturer representative was opening the guidance system and after moving the manipulator and heard a rush of air and saw that the line had been cut. It was also reported that the system was failing the BIST test. No patient was present. 2022-02-22 IUD x5 (Other): No new information. 2022-03-09 IUD x2 (Other): No new information. 2022-03-20 e1, e2, e3, (Rep): Additional information received from a manufacturer representative reported that the surgical arm failed the BIST test when preparing the system for a case. The surgical arm got stuck in a position where it was unable to be moved to a predefined location. A hard restart was done, but the same thing occurred until the representative went into the system settings and reset the movements. The surgical arm was able to move, but when doing a BIST test again, it failed. 2022-03-23 e5, IUD x2 (Rep): The first arm that arrived had failure with 3-define. 2022-03-29 e6, e7, e8, IUD (Other): No new information. 2022-03-30 e9, e10, e11, e12 (Rep): No new information. 2022-04-05 302775626, IUD x6 (Rep): No new information. 2022-05-11 302811994 (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2022-Feb-21 Salesforce C00243946 (Rep, HCP): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the manufacturer representative was opening the guidance system and after moving the manipulator and heard a rush of air and saw that the line had been cut. It was also reported that the system was failing the BIST test. No patient was present. 2022-02-22 IUD x5 (Other): No new information. 2022-03-09 IUD x2 (Other): No new information. 2022-03-20 e1, e2, e3, (Rep): Additional information received from a manufacturer representative reported that the surgical arm failed the BIST test when preparing the system for a case. The surgical arm got stuck in a position where it was unable to be moved to a predefined location. A hard restart was done, but the same thing occurred until the representative went into the system settings and reset the movements. The surgical arm was able to move, but when doing a BIST test again, it failed. 2022-03-23 e5, IUD x2 (Rep): The first arm that arrived had failure with 3-define. 2022-03-29 e6, e7, e8, IUD (Other): No new information. 2022-03-30 e9, e10, e11, e12 (Rep): No new information. 2022-04-05 302775626, IUD x6 (Rep): No new information. 2022-05-11 302811994 (Other): No new information. | xse oobfsra issue pneumatic line cut salesforce outside proc edure opening moving manipulator heard rush air saw line cut failing bist test present iud new iud new additional surgical arm failed bist test preparing case surgical arm got stuck position unable moved predefined location hard restart thing occurred went settings reset movements surgical arm able doing bist test failed iud arm arrived failure iud new new iud new new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV3244 | NAV3264 | NAV2054 | NAV3249NAV4095NAV3244NAV3264NAV2054 | True
6386 | 703933477 | X - Registration Issue | 2020-09-01T12:59:28.000000+0000 | MazorX | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issue 2020-08-30 CCR (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation.. It was reported that 2 screws were not able to be executed during a T3-L3 procedure due to registration issues. There was no patient harm and the procedure was delayed less than an hour. 2020-09-01 email 2020-09-01 11:15 (Rep): Additional information received from a manufacturer representative reported that the system could not register T3 during the second segment registration during the procedure. The second segment was four T3-T9 and the first segment was for T10-L3. The first segment was registered successfully. The surgeon did not want to do a third registration for just T3 and they did not think there was a shift. The other representative at the case thought there was a shift in registration. The surgeon decided to freehand T3. Troubleshooting involved trying several new registrations of the segment while excluding top and bottom vertebrae to get better success rates. T2, T9, T8, and T2 were removed from the registration process, leaving T3 and T5/T6 that wa s already executed. Registration could not be completed and the cause was not determined. MAZOR X SYSTEM TPL0059 | 2020-08-30 CCR (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation.. It was reported that 2 screws were not able to be executed during a T3-L3 procedure due to registration issues. There was no patient harm and the procedure was delayed less than an hour. 2020-09-01 email 2020-09-01 11:15 (Rep): Additional information received from a manufacturer representative reported that the system could not register T3 during the second segment registration during the procedure. The second segment was four T3-T9 and the first segment was for T10-L3. The first segment was registered successfully. The surgeon did not want to do a third registration for just T3 and they did not think there was a shift. The other representative at the case thought there was a shift in registration. The surgeon decided to freehand T3. Troubleshooting involved trying several new registrations of the segment while excluding top and bottom vertebrae to get better success rates. T2, T9, T8, and T2 were removed from the registration process, leaving T3 and T5/T6 that wa s already executed. Registration could not be completed and the cause was not determined. | x registration issue ccr spinal procedure treat deformity fixation screws able executed procedure registration issues procedure delayed hour email additional register second segment registration procedure second segment segment segment registered successfully surgeon did want registration just did think shift case thought shift registration surgeon decided freehand troubleshooting involved trying new registrations segment excluding vertebrae better success rates removed registration process leaving wa s executed registration completed cause determined mazor x | NAV4112 | NAV3242 | NAV4098 | NAV5000 | NAV3243 | NAV4112NAV3242NAV4098NAV5000NAV3243 | True
705 | 703172464 | X - 3DEFINE/SOFTWARE BUG ISSUES | 2019-05-24T19:00:22.000000+0000 | MazorX | NAV3239 | INACCURATE WORK VOLUMES | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3DEFINE/SOFTWARE BUG ISSUES Surgical Arm 12 was installed and Site ATP was performed, with the below occurences. M/F Description: Instability in the performance of the 3Define camera. Scan would either timeout or return a blank scan. Reason for M/F: Voltage loss accross USB cable. M/F Description: Main Cover does not close when system is stored. Reason for M/F: Due to a known issue with a sensor on systems 13 and below. After changing the ini file to CheckSysStorage=No. The issue continued, possible bug. UNK MAZOR X 3DEFINE CAMERA | Surgical Arm 12 was installed and Site ATP was performed, with the below occurences. M/F Description: Instability in the performance of the 3Define camera. Scan would either timeout or return a blank scan. Reason for M/F: Voltage loss accross USB cable. M/F Description: Main Cover does not close when system is stored. Reason for M/F: Due to a known issue with a sensor on systems 13 and below. After changing the ini file to CheckSysStorage=No. The issue continued, possible bug. | x bug issues surgical arm installed site atp performed occurences mf description instability performance camera scan timeout return blank scan reason mf voltage loss accross usb cable mf description main cover does close stored reason mf known issue sensor systems changing ini file checksysstorageno issue continued possible bug unk mazor x camera | NAV3244 | NAV2054 | NAV2058 | NAV4099 | NAV4110 | NAV3244NAV2054NAV2058NAV4099NAV4110 | False
7363 | 704188410 | XSE - SSY Unlocked From Bed | 2021-01-22T01:33:31.000000+0000 | MazorX Stealth Edition | NAV2057 | SURGICAL SYSTEM ISSUES | 20 | ASM0214-01 | SURGICAL SYSTEM | XSE - SSY Unlocked From Bed 2021-Jan-21 Salesforce C00219729 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that after the screws were successfully placed during the procedure, the surgical arm was moved to the right to make room to navigate the interbody and the mount unlocked without anyone touching it. The surgeon was unable to navigate the interbody using the previous registration so the C-arm was used to place it. The event occurred during a T5-S1 Scan & Plan procedure using percutaneous voyager system. There was no patient harm and the procedure was delayed less than an hour. 2021-01-29 IUD x3 (Other): No new information. 2021-02-14 email 2021-02-14 11:30 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved going into the software while in Operate mode and trying to determine what happened, but they couldn't determine what happened. The cause was not determined. No field service was completed so the issue was unlikely to be resolved. SURGICAL SYSTEM | 2021-Jan-21 Salesforce C00219729 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that after the screws were successfully placed during the procedure, the surgical arm was moved to the right to make room to navigate the interbody and the mount unlocked without anyone touching it. The surgeon was unable to navigate the interbody using the previous registration so the C-arm was used to place it. The event occurred during a T5-S1 Scan & Plan procedure using percutaneous voyager system. There was no patient harm and the procedure was delayed less than an hour. 2021-01-29 IUD x3 (Other): No new information. 2021-02-14 email 2021-02-14 11:30 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved going into the software while in Operate mode and trying to determine what happened, but they couldn't determine what happened. The cause was not determined. No field service was completed so the issue was unlikely to be resolved. | xse ssy unlocked bed salesforce spinal p rocedure screws successfully placed procedure surgical arm moved right make room navigate interbody mount unlocked touching surgeon unable navigate interbody using previous registration carm place occurred scan plan procedure using percutaneous voyager procedure delayed hour iud new email additional troubleshooting involved going software operate mode trying determine happened determine happened cause determined field service completed issue unlikely resolved surgical | NAV2057 | NAV3249 | NAV3246 | NAV3250 | NAV2058 | NAV2057NAV3249NAV3246NAV3250NAV2058 | True
2255 | 703230987 | XSE - deviation | 2019-06-27T15:50:41.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - deviation 2019-Jun-25 SCL1901844 intake form (HCp,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that a screw deviation was discovered on final trajectory executed for left L5 during a L4-S1 posterior spine fusion procedure with Mazor X Stealth Edition. Surgeon Instrumented left L5 twice with the same result. Surgeon intra-operatively removed and forgo placement of left L5 screw. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-Jun-25 CCR (HCP,rep): Additional information received indicated that patient history included degenerative and a previous L4-L5 & L5-S1 ALIF performed during a prior surgery. Both right L5 & left L5 screws, were planned proud to allow ideal rod alignment and ease of placement. The mount was a 4x120mm schanz pin placed into the right PSIS. The robot was mounted and positioned midline with ~50mm clearance above the patient. The schanz arm was extended more than 70 degrees and tightened with green driver. The 3Define failed once due to the drape covering the camera, and a second attempt with 3Define was executed with success. The Snap shot tracker was captured and verified accuracy with passive planar probe with arm guide divot. The arm was sent to the AP for 1 AP c-arm image. Arm was sent to the oblique and the 3D marker was adjusted to avoid schanz pin in the field of view. 4 oblique c-arm images were taken to achieve registration. Registration was performed with one-dot and passed after the first attempt. The surgeon reviewed and approved L4-S1 reg istration. The arm was sent to each trajectory to mark incision. The surgeon created bilateral partial wiltsie incisions and fascial incisions were made with the long handle blade through the arm guide when it was sent to each trajectory. Surgeon did not use a tap and placed the screw after creating a pilot osteotomy with the drill. Surgeon used drill guide anchor and tamping tool for every trajectory A lateral skive detected with dilator at right L4 and the dilator was repositioned. The arm vibrated twice when attempting to drill. Arm was resent and lowered. The pedicle was drilled and screw placed successfully. For the right S1, L5, left L4, and S1, a normal workflow was executed. The screws were placed successfully. For left L5, a normal workflow was executed. The screw was placed and a C-Arm image was taken to confirm placement for all screws. Left L5 appeared lateral. The screw was removed. A normal workflow was executed for the second time and screw appeared to be lateral again after C-Arm image was taken. Navigation was checked for accuracy with the arm guide divot and navigation accuracy was confirmed with the passive planar prove on the anatomy. The surgeon decided to remove and leave left L5 out and placed rods bilaterally. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jun-25 SCL1901844 intake form (HCp,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that a screw deviation was discovered on final trajectory executed for left L5 during a L4-S1 posterior spine fusion procedure with Mazor X Stealth Edition. Surgeon Instrumented left L5 twice with the same result. Surgeon intra-operatively removed and forgo placement of left L5 screw. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-Jun-25 CCR (HCP,rep): Additional information received indicated that patient history included degenerative and a previous L4-L5 & L5-S1 ALIF performed during a prior surgery. Both right L5 & left L5 screws, were planned proud to allow ideal rod alignment and ease of placement. The mount was a 4x120mm schanz pin placed into the right PSIS. The robot was mounted and positioned midline with ~50mm clearance above the patient. The schanz arm was extended more than 70 degrees and tightened with green driver. The 3Define failed once due to the drape covering the camera, and a second attempt with 3Define was executed with success. The Snap shot tracker was captured and verified accuracy with passive planar probe with arm guide divot. The arm was sent to the AP for 1 AP c-arm image. Arm was sent to the oblique and the 3D marker was adjusted to avoid schanz pin in the field of view. 4 oblique c-arm images were taken to achieve registration. Registration was performed with one-dot and passed after the first attempt. The surgeon reviewed and approved L4-S1 reg istration. The arm was sent to each trajectory to mark incision. The surgeon created bilateral partial wiltsie incisions and fascial incisions were made with the long handle blade through the arm guide when it was sent to each trajectory. Surgeon did not use a tap and placed the screw after creating a pilot osteotomy with the drill. Surgeon used drill guide anchor and tamping tool for every trajectory A lateral skive detected with dilator at right L4 and the dilator was repositioned. The arm vibrated twice when attempting to drill. Arm was resent and lowered. The pedicle was drilled and screw placed successfully. For the right S1, L5, left L4, and S1, a normal workflow was executed. The screws were placed successfully. For left L5, a normal workflow was executed. The screw was placed and a C-Arm image was taken to confirm placement for all screws. Left L5 appeared lateral. The screw was removed. A normal workflow was executed for the second time and screw appeared to be lateral again after C-Arm image was taken. Navigation was checked for accuracy with the arm guide divot and navigation accuracy was confirmed with the passive planar prove on the anatomy. The surgeon decided to remove and leave left L5 out and placed rods bilaterally. | xse deviation spinal procedure screw deviation discovered final trajectory executed left posterior spine fusion procedure mazor x stealth edition surgeon instrumented left twice surgeon intraoperatively removed forgo placement left screw additional complications anticipated ccr additional indicated history included degenerative previous alif performed prior surgery right left screws planned proud allow ideal rod alignment ease placement mount schanz pin placed right psis robot mounted positioned midline clearance schanz arm extended degrees tightened green driver failed drape covering camera second attempt executed success snap shot tracker captured verified accuracy passive planar probe arm guide divot arm sent ap ap carm image arm sent oblique marker adjusted avoid schanz pin field view oblique carm images taken achieve registration registration performed onedot passed attempt surgeon reviewed approved reg istration arm sent trajectory mark incision surgeon created bilateral partial wiltsie incisions fascial incisions long handle blade arm guide sent trajectory surgeon did use tap placed screw creating pilot osteotomy drill surgeon drill guide anchor tamping tool trajectory lateral skive detected dilator right dilator repositioned arm vibrated twice attempting drill arm resent lowered pedicle drilled screw placed successfully right left normal workflow executed screws placed successfully left normal workflow executed screw placed carm image taken confirm placement screws left appeared lateral screw removed normal workflow executed second time screw appeared lateral carm image taken navigation checked accuracy arm guide divot navigation accuracy confirmed passive planar prove anatomy surgeon decided remove leave left placed rods bilaterally mazor x software installation kit | NAV5000 | NAV3299 | NAV3243 | NAV4091 | NAV3237 | NAV5000NAV3299NAV3243NAV4091NAV3237 | True
51 | 702978994 | X- WS Issue | 2019-01-28T16:19:44.000000+0000 | MazorX | NAV3255 | BED FRAME ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- WS Issue Case aborted due to system functionality. Trauma case, surgeon had preop CT done, trays open, kit open. Patient presented PJK and surgeon said he needed the flat top spine table for patient indications. Robot was not able to be used and aborted because bed frame adaptor does not fit on flat top Jackson. UNK MAZOR X WORKSTATION | Case aborted due to system functionality. Trauma case, surgeon had preop CT done, trays open, kit open. Patient presented PJK and surgeon said he needed the flat top spine table for patient indications. Robot was not able to be used and aborted because bed frame adaptor does not fit on flat top Jackson. | x ws issue case aborted functionality trauma case surgeon preop ct trays open kit open presented pjk surgeon said needed flat spine table indications robot able aborted bed frame adaptor does fit flat jackson unk mazor x workstation | NAV3264 | NAV3255 | NAV2058 | NAV3244 | NAV3250 | NAV3264NAV3255NAV2058NAV3244NAV3250 | True
1684 | 703209556 | X - XEYE ISSUES | 2019-06-15T14:17:49.000000+0000 | MazorX | NAV3247 | EXITS | 30 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - XEYE ISSUES 2018-01-29 CCR, SCL1800319 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat kyphosis with fixation at T9-S2. It was reported that the surgeon and manufacturer representative completed planning for the procedure to take place in two segments prior to the case. The dual clamp was used and attached to the patient at T9 and T11 for T9-T11 and a schanz arm with pins inserted in the patient's bilateral iliac crests for L5-S2. X Eyes not working – bovie related issue – slow software. 2 registrations were necessary. For the first registration, images looked good but wouldn't register. A new OBL image was taken and registration was obtained. The second registration was fast and accurate. Draw spine was not utilized because of an X eyes error. The surface scan was acceptable for both registrations. L5 was deviated inferior to plan, thought to be due to screw following old trajectory. All other trajectories were accurate. Bovie caused the system to restart at the beginning of the ca se. The X eyes were unplugged which would not allow the system to restart. The X eyes were not working again after plugging them back in. The software was also very slow throughout the procedure. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SOFTWARE KIT VER-2.3.1 | 2018-01-29 CCR, SCL1800319 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat kyphosis with fixation at T9-S2. It was reported that the surgeon and manufacturer representative completed planning for the procedure to take place in two segments prior to the case. The dual clamp was used and attached to the patient at T9 and T11 for T9-T11 and a schanz arm with pins inserted in the patient's bilateral iliac crests for L5-S2. X Eyes not working – bovie related issue – slow software. 2 registrations were necessary. For the first registration, images looked good but wouldn't register. A new OBL image was taken and registration was obtained. The second registration was fast and accurate. Draw spine was not utilized because of an X eyes error. The surface scan was acceptable for both registrations. L5 was deviated inferior to plan, thought to be due to screw following old trajectory. All other trajectories were accurate. Bovie caused the system to restart at the beginning of the ca se. The X eyes were unplugged which would not allow the system to restart. The X eyes were not working again after plugging them back in. The software was also very slow throughout the procedure. No patient harm was reported and surgery was delayed less than an hour. | x xeye issues ccr spinal procedure treat kyphosis fixation surgeon completed planning procedure place segments prior case dual clamp attached schanz arm pins inserted patients bilateral iliac crests x eyes working – bovie related issue – slow software registrations necessary registration images looked good wouldnt register new obl image taken registration obtained second registration fast accurate draw spine utilized x eyes error surface scan acceptable registrations deviated inferior plan thought screw following old trajectory trajectories accurate bovie caused restart beginning ca se x eyes unplugged allow restart x eyes working plugging software slow procedure surgery delayed hour mazor x software kit | NAV4091 | NAV3248 | NAV3242 | NAV3247 | NAV4110 | NAV4091NAV3248NAV3242NAV3247NAV4110 | True
1761 | 703213556 | X- X-eye issues | 2019-06-18T19:38:19.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X- X-eye issues Grounding issues with bovie – crashing the system. Crossview error during registration. Lateral skive due to unilateral retraction UNK MAZOR X X-EYE CAMERA | Grounding issues with bovie – crashing the system. Crossview error during registration. Lateral skive due to unilateral retraction | x xeye issues grounding issues bovie – crashing crossview error registration lateral skive unilateral retraction unk mazor x xeye camera | NAV4110 | NAV3254 | NAV2054 | NAV4109 | NAV4095 | NAV4110NAV3254NAV2054NAV4109NAV4095 | True
215 | 703041178 | X - deviation | 2019-03-06T19:39:20.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - deviation 2019-Mar-04 SCL1900514 msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the right L4 screw stimulated low, and the HCP removed the screw and probed the pedicle and it seemed the screw was placed slightly inferior to the plan. The screw was replaced by hand and it stimulated slightly higher. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM TPL0059 | 2019-Mar-04 SCL1900514 msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the right L4 screw stimulated low, and the HCP removed the screw and probed the pedicle and it seemed the screw was placed slightly inferior to the plan. The screw was replaced by hand and it stimulated slightly higher. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x deviation msg spinal procedure right screw stimulated low hcp removed screw probed pedicle screw placed slightly inferior plan screw replaced hand stimulated slightly higher noted additional complications anticipated mazor x | NAV5000 | NAV4091 | NAV2054 | NAV4112 | NAV1006 | NAV5000NAV4091NAV2054NAV4112NAV1006 | True
7991 | 704325814 | XSE - Instrument Verification Lost | 2021-04-07T00:10:32.000000+0000 | MazorX Stealth Edition | NAV2056 | SOFTWARE ISSUES | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Instrument Verification Lost 2021-Apr-03 Salesforce C00224130 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the system would drop verified instruments randomly within the procedure. The team would have to reverify all instruments through-out the procedure. There was no impact on the patient outcome. 2021-04-07 email 2021-04-07 14:14, IUD (Rep): Additional information received from a manufacturer representative reported a demo case was being completed on a cadaver. One procedure was using Scan & Plan and a bug was found. The other case was a CT to Fluoro fusion. The cause of the loss verification was not determined. The demo case was delayed less than an hour. See 704325815 for the other case. 2021-04-16 email 2021-04-16 12:31 (Rep): No new information. 2021-04-20 IUD x2 (Other): No new information. 2021-05-17 email 2021-05-17 12:45, e1 (Rep): No new information. 2021-06-03 email 2021-06-03 09:31, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the event occurred during training and has since been used without issue. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Apr-03 Salesforce C00224130 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the system would drop verified instruments randomly within the procedure. The team would have to reverify all instruments through-out the procedure. There was no impact on the patient outcome. 2021-04-07 email 2021-04-07 14:14, IUD (Rep): Additional information received from a manufacturer representative reported a demo case was being completed on a cadaver. One procedure was using Scan & Plan and a bug was found. The other case was a CT to Fluoro fusion. The cause of the loss verification was not determined. The demo case was delayed less than an hour. See 704325815 for the other case. 2021-04-16 email 2021-04-16 12:31 (Rep): No new information. 2021-04-20 IUD x2 (Other): No new information. 2021-05-17 email 2021-05-17 12:45, e1 (Rep): No new information. 2021-06-03 email 2021-06-03 09:31, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the event occurred during training and has since been used without issue. | xse instrument verification lost salesforce drop verified instruments randomly procedure team reverify instruments procedure impact outcome email iud additional demo case completed cadaver procedure using scan plan bug case ct fluoro fusion cause loss verification determined demo case delayed hour case email new iud new email new email additional occurred training issue mazorx sw installation kit | NAV3246 | NAV2056 | NAV3299 | NAV3235 | NAV2058 | NAV3246NAV2056NAV3299NAV3235NAV2058 | True
4365 | 703538323 | X - Deviation/Skive Right L4 | 2019-12-30T19:24:25.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation/Skive Right L4 2018-08-16 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the procedure was planned for a L4-L5 fusion. The surgery went well except for a lateral skive on right L4. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2018-08-16 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the procedure was planned for a L4-L5 fusion. The surgery went well except for a lateral skive on right L4. No patient harm was reported and surgery was delayed less than an hour. | x deviationskive right ccr spinal procedure procedure planned fusion surgery went lateral skive right surgery delayed hour mazor x | NAV4091 | NAV3242 | NAV2047 | NAV4095 | NAV5000 | NAV4091NAV3242NAV2047NAV4095NAV5000 | True
5869 | 703824131 | X - CANBUS Error Node 3 | 2020-06-26T20:01:01.000000+0000 | MazorX | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | X - CANBUS Error Node 3 2020-06-26 00691626, CCR (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was powered up and during an advanced accuracy test, a CANBUS error at node 3 was displayed prior to the case. The manufacturer representative did a hard restart of the system to clear the error. After the case, the representative pressed the back button and the screen went black. A soft restart of the system was done to get the software to pop up again to the approve registration page. There was no patient involvement. 2020-06-29 IUD x3 (Other): No new information. 2020-07-02 IUD x7 (Other): No new information. 2020-07-17 IUD (Other): No new information. 2020-07-28 IUD x5 (Other): No new information. 2021-09-09 302555445, IUD x4 (Other): No new information. 2021-12-08 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-06-26 00691626, CCR (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was powered up and during an advanced accuracy test, a CANBUS error at node 3 was displayed prior to the case. The manufacturer representative did a hard restart of the system to clear the error. After the case, the representative pressed the back button and the screen went black. A soft restart of the system was done to get the software to pop up again to the approve registration page. There was no patient involvement. 2020-06-29 IUD x3 (Other): No new information. 2020-07-02 IUD x7 (Other): No new information. 2020-07-17 IUD (Other): No new information. 2020-07-28 IUD x5 (Other): No new information. 2021-09-09 302555445, IUD x4 (Other): No new information. 2021-12-08 IUD (Other): No new information. | x canbus error node ccr outside procedure powered advanced accuracy test canbus error node displayed prior case did hard restart clear error case pressed button screen went black soft restart software pop approve registration page iud new iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV2047 | NAV3247 | NAV3261 | NAV4095 | NAV3249 | NAV2047NAV3247NAV3261NAV4095NAV3249 | False
2423 | 703235139 | X - SOFTWARE ISSUE | 2019-06-29T18:19:54.000000+0000 | MazorX | NAV3247 | EXITS | 20 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X - SOFTWARE ISSUE 1. Software froze and crashed during case. Software also performed very slow. 2. After SW crash, received X Eye communication error after checking camera function in utilities (both cameras worked prior to crash). Powering system down for ~1min was not successful. Post op, system was shut down for >5min, after which, camera appeared to function normally. UNKNOWN MAZOR X SW | 1. Software froze and crashed during case. Software also performed very slow. 2. After SW crash, received X Eye communication error after checking camera function in utilities (both cameras worked prior to crash). Powering system down for ~1min was not successful. Post op, system was shut down for >5min, after which, camera appeared to function normally. | x software issue software froze crashed case software performed slow sw crash x eye communication error checking camera function utilities powering successful post op shut camera appeared function normally unknown mazor x sw | NAV3248 | NAV3247 | NAV3246 | NAV4091 | NAV1012 | NAV3248NAV3247NAV3246NAV4091NAV1012 | True
3331 | 703315121 | XSE - registration/hit WV/navlock issues | 2019-08-16T15:34:07.000000+0000 | MazorX Stealth Edition | NAV4096 | ARM HIT WORK VOLUME | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - registration/hit WV/navlock issues 2019-08-13 SCL1902373, Intake Form, CCR (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat deformity with fixation. It was reported that a T2-L2 open procedure was planned with the surgeon. Due to some of the pedicles being too small to comfortably fit a sizable screw, some pedicles were left out of the plan. For the first segment, a dual spinous process clamp was attached to T10 and T12 and connected to the surgical system via a bone mounting bridge. A 3Define scan was performed and Draw Spine was completed with a navigated passive planar prove. AP and oblique trajectories were sent and both images were accepted; however, when trying to register a warning that the AP image had been auto rotated appeared. The representative went back in the software and a new AP image was taken. Registration was successful from L2-T8 with green values at each level. All trajectories were sent, drilled and screws were placed. When sending the arm to the L1 and L2 trajectories, the surgical arm ran into the bone mounting bridge. L1 and L2 on the right side were obstructed by the clamp and the bone mounting bridge so the surgeon decided that they would move the clamp and try those trajectories at the end of the case. For the second segment, the dual spinous process clam was moved to the T4 and T6 ve rtebrae. A 3Define scan was performed and Draw Spine was completed with a passive planar prove. AP and oblique images were taken. The AP image was accepted on the first try, but several oblique images had to be taken to best visualize the anatomy of interest. Registration was successful when laveling off of T6 using the basic algorithm. All values were green except for T3 which had red values. The surgeon decided to go ahead and execute all trajectories except for T3. The surgeon planned to try and register T3 by itself. After placing all of the screws for the second segment, registration was completed for T3. Green values were achieved so the trajectory was drilled and screws placed. For segment 3, the surgeon tried to go back to L1 and L2 on the right side. The mount was unlocked and the surgeon used a single clamp on the L4 spinous process. A 3Define scan was performed and Draw Spine was completed using a passive planar probe. AP and oblique images were taken and both were accepted on the first try. Registration was suc cessful off of L3 with all green values. Both trajectories were sent, drill and had screws placed. All screws were placed and looked acceptable in fluoroscopic images. Neuro monitoring was used during the procedure and all screws were stimulated with good values. The representative noted that the drill bit bound and locked in the NavLock tracker and the surgeon's glove got stuck in between the tracker and the drill bit at least two times and tore the surgeons clove causing them to re-glove. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-08-22 email 2019-08-22 18:45 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2019-08-13 SCL1902373, Intake Form, CCR (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat deformity with fixation. It was reported that a T2-L2 open procedure was planned with the surgeon. Due to some of the pedicles being too small to comfortably fit a sizable screw, some pedicles were left out of the plan. For the first segment, a dual spinous process clamp was attached to T10 and T12 and connected to the surgical system via a bone mounting bridge. A 3Define scan was performed and Draw Spine was completed with a navigated passive planar prove. AP and oblique trajectories were sent and both images were accepted; however, when trying to register a warning that the AP image had been auto rotated appeared. The representative went back in the software and a new AP image was taken. Registration was successful from L2-T8 with green values at each level. All trajectories were sent, drilled and screws were placed. When sending the arm to the L1 and L2 trajectories, the surgical arm ran into the bone mounting bridge. L1 and L2 on the right side were obstructed by the clamp and the bone mounting bridge so the surgeon decided that they would move the clamp and try those trajectories at the end of the case. For the second segment, the dual spinous process clam was moved to the T4 and T6 ve rtebrae. A 3Define scan was performed and Draw Spine was completed with a passive planar prove. AP and oblique images were taken. The AP image was accepted on the first try, but several oblique images had to be taken to best visualize the anatomy of interest. Registration was successful when laveling off of T6 using the basic algorithm. All values were green except for T3 which had red values. The surgeon decided to go ahead and execute all trajectories except for T3. The surgeon planned to try and register T3 by itself. After placing all of the screws for the second segment, registration was completed for T3. Green values were achieved so the trajectory was drilled and screws placed. For segment 3, the surgeon tried to go back to L1 and L2 on the right side. The mount was unlocked and the surgeon used a single clamp on the L4 spinous process. A 3Define scan was performed and Draw Spine was completed using a passive planar probe. AP and oblique images were taken and both were accepted on the first try. Registration was suc cessful off of L3 with all green values. Both trajectories were sent, drill and had screws placed. All screws were placed and looked acceptable in fluoroscopic images. Neuro monitoring was used during the procedure and all screws were stimulated with good values. The representative noted that the drill bit bound and locked in the NavLock tracker and the surgeon's glove got stuck in between the tracker and the drill bit at least two times and tore the surgeons clove causing them to re-glove. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-08-22 email 2019-08-22 18:45 (Rep): No new information. | xse registrationhit wvnavlock issues ccr spinal procedure treat deformity fixation open procedure planned surgeon pedicles small comfortably fit sizable screw pedicles left plan segment dual spinous process clamp attached connected surgical bone mounting bridge scan performed draw spine completed navigated passive planar prove ap oblique trajectories sent images accepted trying register warning ap image auto rotated appeared went software new ap image taken registration successful green values level trajectories sent drilled screws placed sending arm trajectories surgical arm ran bone mounting bridge right obstructed clamp bone mounting bridge surgeon decided clamp try trajectories end case second segment dual spinous process clam moved ve rtebrae scan performed draw spine completed passive planar prove ap oblique images taken ap image accepted try oblique images taken best visualize anatomy registration successful laveling using basic algorithm values green red values surgeon decided ahead execute trajectories surgeon planned try register placing screws second segment registration completed green values achieved trajectory drilled screws placed segment surgeon tried right mount unlocked surgeon single clamp spinous process scan performed draw spine completed using passive planar probe ap oblique images taken accepted try registration suc cessful green values trajectories sent drill screws placed screws placed looked acceptable fluoroscopic images neuro monitoring procedure screws stimulated good values noted drill bit bound locked navlock tracker surgeons glove got stuck tracker drill bit times tore surgeons clove causing reglove additional complications anticipated email new mazor x | NAV3243 | NAV3242 | NAV5000 | NAV3246 | NAV3239 | NAV3243NAV3242NAV5000NAV3246NAV3239 | False
4549 | 703550662 | X- 3Define & arm issue | 2020-01-08T13:55:15.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 30 | ASM0206-01 | SURGICAL ARM | X- 3Define & arm issue 2018-8-6 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done by the surgeon prior to the procedure. The clamp was chosen as the platform. The 3Define scan completely failed, so the surgeon used the default work volume. The surgical arm had difficulty reaching trajectories, as it would stop in the middle of a movement and say that the arm failed to reach its destination. The surgical arm would have to be sent multiple times to get to the trajectories. There was no patient harm and the procedure was not delayed over an hour. SURGICAL ARM | 2018-8-6 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done by the surgeon prior to the procedure. The clamp was chosen as the platform. The 3Define scan completely failed, so the surgeon used the default work volume. The surgical arm had difficulty reaching trajectories, as it would stop in the middle of a movement and say that the arm failed to reach its destination. The surgical arm would have to be sent multiple times to get to the trajectories. There was no patient harm and the procedure was not delayed over an hour. | x arm issue ccr spinal procedure planning surgeon prior procedure clamp chosen platform scan completely failed surgeon default work volume surgical arm difficulty reaching trajectories stop middle movement say arm failed reach destination surgical arm sent multiple times trajectories procedure delayed hour surgical arm | NAV4095 | NAV3244 | NAV4101 | NAV4096 | NAV3242 | NAV4095NAV3244NAV4101NAV4096NAV3242 | True
9757 | 704716147 | XSE - Screw Deviation at T3 | 2021-11-15T20:09:12.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Screw Deviation at T3 2021-11-12 00726082 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the left T3 was found to be deviated during a T3-pelvis fusion procedure. The procedure was completed in three registrations. The frame was clamped on T4 and T6 when the screw was placed at T3. The surgeon worked up the spine and used the Midas drill during the procedure. The screw was found to be deviated 3-4 mm during a confirmation spin. The use of the guidance system was aborted and the screw was repositioned freehand. The manufacturer representative noted there was a potential for soft tissue pressure due to the proximity of the screw to the end of the incision. None of the other screws were inaccurate. There was no patient harm and the procedure was delayed less than an hour. 2021-11-22 IUD (Other): No new information. 2021-12-02 IUD (Other): No new information. 2021-12-13 IUD, email 2021-12-15 01:15 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2021-11-12 00726082 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the left T3 was found to be deviated during a T3-pelvis fusion procedure. The procedure was completed in three registrations. The frame was clamped on T4 and T6 when the screw was placed at T3. The surgeon worked up the spine and used the Midas drill during the procedure. The screw was found to be deviated 3-4 mm during a confirmation spin. The use of the guidance system was aborted and the screw was repositioned freehand. The manufacturer representative noted there was a potential for soft tissue pressure due to the proximity of the screw to the end of the incision. None of the other screws were inaccurate. There was no patient harm and the procedure was delayed less than an hour. 2021-11-22 IUD (Other): No new information. 2021-12-02 IUD (Other): No new information. 2021-12-13 IUD, email 2021-12-15 01:15 (Rep): No new information. | xse screw deviation spinal procedure left deviated fusion procedure procedure completed registrations frame clamped screw placed surgeon worked spine midas drill procedure screw deviated mm confirmation spin use aborted screw repositioned freehand noted potential soft tissue pressure proximity screw end incision screws inaccurate procedure delayed hour iud new iud new iud email new mazor x | NAV5000 | NAV4112 | NAV1012 | NAV4090 | NAV3243 | NAV5000NAV4112NAV1012NAV4090NAV3243 | True
3339 | 703319821 | X - Positioner Stuck in Y Axis | 2019-08-20T14:13:28.000000+0000 | MazorX | NAV4104 | UNABLE TO MOVE | 20 | ASM0214-01 | SURGICAL SYSTEM | X - Positioner Stuck in Y Axis 2019-08-13 SCL1902380, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the position got stuck about halfway in the Y-axis. The representative was able to free the position, but it continued to stick when moving up or down. The issue had occurred in the past, but was unable to be resolved today. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. 2019-08-20 email 2019-08-20 19:08 (Rep): Additional information received from a manufacturer representative reported the cause of the positioner getting stuck was a broken spring. SURGICAL SYSTEM | 2019-08-13 SCL1902380, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the position got stuck about halfway in the Y-axis. The representative was able to free the position, but it continued to stick when moving up or down. The issue had occurred in the past, but was unable to be resolved today. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. 2019-08-20 email 2019-08-20 19:08 (Rep): Additional information received from a manufacturer representative reported the cause of the positioner getting stuck was a broken spring. | x positioner stuck y axis spinal procedure position got stuck halfway yaxis able free position continued stick moving issue occurred past unable resolved today procedure delayed hour case completed email additional cause positioner getting stuck broken spring surgical | NAV4104 | NAV3299 | NAV4095 | NAV3251 | NAV3244 | NAV4104NAV3299NAV4095NAV3251NAV3244 | True
3253 | 703298950 | XSE - Registration Issues/Aborted | 2019-08-07T12:17:06.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059-05 | MAZOR X SYSTEM (REV 5) | XSE - Registration Issues/Aborted 2019-07-30 SCL1902271, Intake Form (HCP, Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a patient was present with Magic Grow rods in from T4-L2 and they were planning a correction from T2-L2. After the exposure they pulled out the Magic rods and left the screws in place for registration. Every time they would go to register they would be able to achieve segmentation but would always get extremely poor values at all levels. There would be no apparent shift at the levels but they still were "red" and would not let them continue. This happened multiple times. After multiple attempts to register the same results were achieved. After approximately 15 minutes the surgeon said to get the system off the bed and out of the way, he would put lead on and freehand the screws. The surgeon ended up free-handing all of the screws. After the case the reps went through the case data and tried to troubleshoot the registration process. It was discovered that if the region of inter est was modified they were able to register the bodies that did not have the Magic rod in the CT. The vertebraes that had some semblance of the rod near them would have had bad values but no apparent shift. The issue extended the surgical time by less than one hour. There was no impact on the patient outcome. MAZOR X SYSTEM (REV 5) | 2019-07-30 SCL1902271, Intake Form (HCP, Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a patient was present with Magic Grow rods in from T4-L2 and they were planning a correction from T2-L2. After the exposure they pulled out the Magic rods and left the screws in place for registration. Every time they would go to register they would be able to achieve segmentation but would always get extremely poor values at all levels. There would be no apparent shift at the levels but they still were "red" and would not let them continue. This happened multiple times. After multiple attempts to register the same results were achieved. After approximately 15 minutes the surgeon said to get the system off the bed and out of the way, he would put lead on and freehand the screws. The surgeon ended up free-handing all of the screws. After the case the reps went through the case data and tried to troubleshoot the registration process. It was discovered that if the region of inter est was modified they were able to register the bodies that did not have the Magic rod in the CT. The vertebraes that had some semblance of the rod near them would have had bad values but no apparent shift. The issue extended the surgical time by less than one hour. There was no impact on the patient outcome. | xse registration issuesaborted spinal procedure present magic grow rods planning correction exposure pulled magic rods left screws place registration time register able achieve segmentation extremely poor values levels apparent shift levels red let continue happened multiple times multiple attempts register results achieved approximately minutes surgeon said bed way lead freehand screws surgeon ended freehanding screws case reps went case data tried troubleshoot registration process discovered region inter est modified able register bodies did magic rod ct vertebraes semblance rod near bad values apparent shift issue extended surgical time hour impact outcome mazor x | NAV4112 | NAV3242 | NAV3260 | NAV2058 | NAV3253 | NAV4112NAV3242NAV3260NAV2058NAV3253 | True
543 | 703143650 | X - low pressure | 2019-05-08T15:43:15.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | ASM0214-02 | SSY | X - low pressure 2019-May-06 SCL1901099 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the system would not build pressure. After multiple hard restarts and tapping on the pneumatic valve, the system finally build pressure up to 6.9 psi. The system was mounted to the bed and when leaving the room, the rep came back and the system dropped to 0.5 psi and would not rebuild pressure. It was noted that it seemed as if there was a leak in the cleats but they were unable to confirm at the moment. Once the teeth were opened and closed, the air leak closed. There was no patient harm but the case was aborted. There were no additional complications reported or anticipated as a result of the event. SSY | 2019-May-06 SCL1901099 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the system would not build pressure. After multiple hard restarts and tapping on the pneumatic valve, the system finally build pressure up to 6.9 psi. The system was mounted to the bed and when leaving the room, the rep came back and the system dropped to 0.5 psi and would not rebuild pressure. It was noted that it seemed as if there was a leak in the cleats but they were unable to confirm at the moment. Once the teeth were opened and closed, the air leak closed. There was no patient harm but the case was aborted. There were no additional complications reported or anticipated as a result of the event. | x low pressure spinal procedure build pressure multiple hard restarts tapping pneumatic valve finally build pressure psi mounted bed leaving room came dropped psi rebuild pressure noted leak cleats unable confirm moment teeth opened closed air leak closed case aborted additional complications anticipated ssy | NAV3264 | NAV4112 | NAV3250 | NAV4095 | NAV4105 | NAV3264NAV4112NAV3250NAV4095NAV4105 | True
205 | 703038322 | XSE - trajectories appeared deep | 2019-03-05T16:05:01.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - trajectories appeared deep 2018-Mar-01 SCL1900498 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the CAD model would change from the dilator to the drill when navigating the dilator during a procedure. The dilator, drill, and screw appeared 5-10 mm deeper than expected on a couple trajectories. It was noted that possible explanations for this could be “implant planned proud”, or a facectomy during the transforaminal lumbar interbody fusion (TLIF) on the left side. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Information unrelated to this event was included in a separate PE; See PE 703038325-Navlock won't verify. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2018-Mar-01 SCL1900498 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the CAD model would change from the dilator to the drill when navigating the dilator during a procedure. The dilator, drill, and screw appeared 5-10 mm deeper than expected on a couple trajectories. It was noted that possible explanations for this could be “implant planned proud”, or a facectomy during the transforaminal lumbar interbody fusion (TLIF) on the left side. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Information unrelated to this event was included in a separate PE; See PE 703038325-Navlock won't verify. | xse trajectories appeared deep spinal procedure cad model change dilator drill navigating dilator procedure dilator drill screw appeared mm deeper expected couple trajectories noted possible explanations “implant planned proud” facectomy transforaminal lumbar interbody fusion left noted additional complications anticipated unrelated included separate pe pe wont verify mazor x software installation kit | NAV3246 | NAV5000 | NAV2054 | NAV3247 | NAV3243 | NAV3246NAV5000NAV2054NAV3247NAV3243 | True
2355 | 703233671 | X - Surgical arm/software issues | 2019-06-28T16:51:55.000000+0000 | MazorX | NAV3246 | DEFECTS | 30 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X - Surgical arm/software issues It was reported that the software auto-corrected mislabeling attempt software not prioritizing solutions appropriately so surgical arm making unnecessary movements to execute trajectories. MAZOR X SOFTWARE KIT VER-3.0 | It was reported that the software auto-corrected mislabeling attempt software not prioritizing solutions appropriately so surgical arm making unnecessary movements to execute trajectories. | x surgical armsoftware issues software autocorrected mislabeling attempt software prioritizing solutions appropriately surgical arm making unnecessary movements execute trajectories mazor x software kit | NAV3246 | NAV3247 | NAV3248 | NAV4091 | NAV3242 | NAV3246NAV3247NAV3248NAV4091NAV3242 | True
321 | 703078085 | XSE - 3Define failed/arm error | 2019-03-28T14:40:31.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 30 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - 3Define failed/arm error 2019-Mar-26 SCL1900744 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3Define scan failed and the rep had to use a generic work volume. During the last screw placement, the rep received the following error, "Error 8025 Pre-defined Position is Not Reachable. Check Collision with Shoulder/Postioner or WorkVolume". The surgeon requested the rep Clear Up and but the Arm would not clear up. The surgeon then grabbed the arm and manually moved it out of the way. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X 3DEFINE CAMERA | 2019-Mar-26 SCL1900744 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3Define scan failed and the rep had to use a generic work volume. During the last screw placement, the rep received the following error, "Error 8025 Pre-defined Position is Not Reachable. Check Collision with Shoulder/Postioner or WorkVolume". The surgeon requested the rep Clear Up and but the Arm would not clear up. The surgeon then grabbed the arm and manually moved it out of the way. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse failedarm error spinal procedure scan failed use generic work volume screw placement following error error predefined position reachable check collision shoulderpostioner workvolume surgeon requested clear arm clear surgeon grabbed arm manually moved way noted additional complications anticipated unk mazor x camera | NAV3244 | NAV2047 | NAV4095 | NAV2054 | NAV4110 | NAV3244NAV2047NAV4095NAV2054NAV4110 | True
3757 | 703435812 | XSE - Failed Accuracy Test | 2019-10-28T12:00:47.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | XSE - Failed Accuracy Test 2019-Oct-28 00676934 (for,HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during the arm set up process the system failed the accuracy check where the test pointer land on the side of testing hole. There was no residue of drape in all 3 groove and screw hole and the pointer tip was not bent. Repeating the step with Mazor core sterile arm guide returned same result. They resent the arm and restart the system. After Mazor was aborted, a closer look at the end effector revealed that there was movement on the joint where it is mobile deviated from original position when tightening arm guide to end effector. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-10-29 IUD x 8, e1, e2, e3, e4, e5, e6, e7, e8, e9, e10 (Rep, For): No new information. 2020-02-26 IUD (Other, For): No new information. 2020-07-27 IUD (Other, For): No new information. 2021-12-07 IUD (Other, For): No new information. SURGICAL ARM | 2019-Oct-28 00676934 (for,HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during the arm set up process the system failed the accuracy check where the test pointer land on the side of testing hole. There was no residue of drape in all 3 groove and screw hole and the pointer tip was not bent. Repeating the step with Mazor core sterile arm guide returned same result. They resent the arm and restart the system. After Mazor was aborted, a closer look at the end effector revealed that there was movement on the joint where it is mobile deviated from original position when tightening arm guide to end effector. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-10-29 IUD x 8, e1, e2, e3, e4, e5, e6, e7, e8, e9, e10 (Rep, For): No new information. 2020-02-26 IUD (Other, For): No new information. 2020-07-27 IUD (Other, For): No new information. 2021-12-07 IUD (Other, For): No new information. | xse failed accuracy test spinal procedure arm set process failed accuracy check test pointer land testing hole residue drape groove screw hole pointer tip bent repeating step mazor core sterile arm guide returned resent arm restart mazor aborted closer look end effector revealed movement joint mobile deviated original position tightening arm guide end effector additional complications anticipated iud x new iud new iud new iud new surgical arm | NAV1012 | NAV5000 | NAV4112 | NAV3264 | NAV1006 | NAV1012NAV5000NAV4112NAV3264NAV1006 | False
1658 | 703208433 | X - Registration issue | 2019-06-14T20:38:28.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Registration issue It was reported that the flouro seg w/basic 1st attempt red value at L5, tried manual seg w/ poor results took new OBL, flouro seg w/basic green values at all surgical levels surgeon visually verified matching accuracies at all surgical level MAZOR X SOFTWARE KIT VER-3.0.2 | It was reported that the flouro seg w/basic 1st attempt red value at L5, tried manual seg w/ poor results took new OBL, flouro seg w/basic green values at all surgical levels surgeon visually verified matching accuracies at all surgical level | x registration issue flouro seg wbasic attempt red value tried manual seg w poor results took new obl flouro seg wbasic green values surgical levels surgeon visually verified matching accuracies surgical level mazor x software kit | NAV3246 | NAV3242 | NAV4091 | NAV2047 | NAV2054 | NAV3246NAV3242NAV4091NAV2047NAV2054 | True
9485 | 704651155 | XSE - Registration Difficulties | 2021-10-07T18:11:40.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Difficulties 2021-Oct-07 Salesforce C00235943 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a SI fusion case, the manufacturer representative was having registration issues. The CT was segmented into L5 and S1 with plans only on S1 for the fusion. After acquiring the AP/OBL, he went to register off L5 and received a lower quality image error for the OBL. The manufacturer representative reacquired the OBL, but the issue persisted. He noticed that the red dot had shifted form the CT, so he went back and added a trajectory to L5 which resolved the red dot moving. However, he was still getting errors when trying to register. He reacquired both AP and OBL and they were good quality. He tried manual registration as well, but could not get it. At this point, they aborted the use of the guidance system and switched over to navigation/imaging systems. This caused a 30-40 minute delay to the case but no other patient impact. 2021-10-11 IUD x2 (Other): No new information. 2021-10-25 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-Oct-07 Salesforce C00235943 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a SI fusion case, the manufacturer representative was having registration issues. The CT was segmented into L5 and S1 with plans only on S1 for the fusion. After acquiring the AP/OBL, he went to register off L5 and received a lower quality image error for the OBL. The manufacturer representative reacquired the OBL, but the issue persisted. He noticed that the red dot had shifted form the CT, so he went back and added a trajectory to L5 which resolved the red dot moving. However, he was still getting errors when trying to register. He reacquired both AP and OBL and they were good quality. He tried manual registration as well, but could not get it. At this point, they aborted the use of the guidance system and switched over to navigation/imaging systems. This caused a 30-40 minute delay to the case but no other patient impact. 2021-10-11 IUD x2 (Other): No new information. 2021-10-25 IUD (Other): No new information. | xse registration difficulties salesforce spinal p rocedure si fusion case having registration issues ct segmented plans fusion acquiring apobl went register lower quality image error obl reacquired obl issue persisted noticed red dot shifted ct went added trajectory resolved red dot moving getting errors trying register reacquired ap obl good quality tried manual registration point aborted use switched navigationimaging systems caused minute delay case impact iud new iud new mazor x | NAV4112 | NAV3242 | NAV3243 | NAV5000 | NAV3248 | NAV4112NAV3242NAV3243NAV5000NAV3248 | True
8123 | 704356797 | XSE ¦ No Video on Surgeon Screen | 2021-04-23T00:19:54.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 30 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE ¦ No Video on Surgeon Screen 2021-Apr-22 Salesforce C00225157 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the system was powered on this morning, the surgeon tablet was not functioning. No patient was present. 2021-04-27 IUD x4 (Other): No new information. 2021-05-21 IUD x2 (Other): No new information. 2021-06-24 IUD x4 (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2021-Apr-22 Salesforce C00225157 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the system was powered on this morning, the surgeon tablet was not functioning. No patient was present. 2021-04-27 IUD x4 (Other): No new information. 2021-05-21 IUD x2 (Other): No new information. 2021-06-24 IUD x4 (Other): No new information. | xse video surgeon screen salesforce outside procedure powered morning surgeon tablet functioning present iud new iud new iud new kit svc srgn screen arm | NAV3253 | NAV3249 | NAV2058 | NAV2047 | NAV3299 | NAV3253NAV3249NAV2058NAV2047NAV3299 | True
9176 | 704566904 | XSE - Deviation/SRA Failed Accuracy Test | 2021-08-19T21:44:21.000000+0000 | MazorX Stealth Edition | NAV3239 | INACCURATE WORK VOLUMES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Deviation/SRA Failed Accuracy Test 2021-Aug-19 Salesforce C00233208, IUD x4 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm passed an accuracy test prior to the procedure. During the procedure, the 3Define scan did not register correctly and the surgical arm collided with the patient when being sent for C-arm registration shots. After this, the case was restarted and the surgeon proceeded with placing screws. The screws were way off from the plan. The surgical arm was tested after the procedure and the arm failed the 10 point accuracy test. There was no patient harm and the procedure was delayed less than an hour. 2021-09-16 email 2021-09-16 11:16 (Rep): Additional information received from a manufacturer representative reported that a minimally invasive L4-L5 TLIF was done. The L4 and L5 screws on the left side deviated. It was unknown how much the screws were deviated. Troubleshooting involved restarting the whole case, reattaching the 3Define lens, taking new images and re-registering the patient. The use of the guidance system was aborted and the procedure was completed with navigation. 2021-09-21 IUD x2 (Other): No new information. 2021-12-01 302652181, IUD x3 (Other): No new information. 2021-12-05 IUD x2 (Other): No new information. 2021-12-07 IUD x4 (Other): No new information. SURGICAL ARM | 2021-Aug-19 Salesforce C00233208, IUD x4 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm passed an accuracy test prior to the procedure. During the procedure, the 3Define scan did not register correctly and the surgical arm collided with the patient when being sent for C-arm registration shots. After this, the case was restarted and the surgeon proceeded with placing screws. The screws were way off from the plan. The surgical arm was tested after the procedure and the arm failed the 10 point accuracy test. There was no patient harm and the procedure was delayed less than an hour. 2021-09-16 email 2021-09-16 11:16 (Rep): Additional information received from a manufacturer representative reported that a minimally invasive L4-L5 TLIF was done. The L4 and L5 screws on the left side deviated. It was unknown how much the screws were deviated. Troubleshooting involved restarting the whole case, reattaching the 3Define lens, taking new images and re-registering the patient. The use of the guidance system was aborted and the procedure was completed with navigation. 2021-09-21 IUD x2 (Other): No new information. 2021-12-01 302652181, IUD x3 (Other): No new information. 2021-12-05 IUD x2 (Other): No new information. 2021-12-07 IUD x4 (Other): No new information. | xse deviationsra failed accuracy test salesforce iud spinal procedure surgical arm passed accuracy test prior procedure procedure scan did register correctly surgical arm collided sent carm registration shots case restarted surgeon proceeded placing screws screws way plan surgical arm tested procedure arm failed point accuracy test procedure delayed hour email additional minimally invasive tlif screws left deviated unknown screws deviated troubleshooting involved restarting case reattaching lens taking new images reregistering use aborted procedure completed navigation iud new iud new iud new iud new surgical arm | NAV4093 | NAV5000 | NAV1012 | NAV3249 | NAV4092 | NAV4093NAV5000NAV1012NAV3249NAV4092 | False
1248 | 703193549 | X - Software Issue | 2019-06-06T19:22:39.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Software Issue No OBL found -screen froze -soft restart -Generic OBL attempted -processing time only graduating from 10%-25% -over 2min *Image of surgical arm showed to be well under surgical volume -Surgical arm showed collision with body -New 3-Define New Draw Spine attempted -Reference frame not found -no changes were made to mount -plastic was cleared -Proceeded with prior draw spine -OBL trajectory resolved MAZOR X SOFTWARE KIT VER-3.0.2 | No OBL found -screen froze -soft restart -Generic OBL attempted -processing time only graduating from 10%-25% -over 2min *Image of surgical arm showed to be well under surgical volume -Surgical arm showed collision with body -New 3-Define New Draw Spine attempted -Reference frame not found -no changes were made to mount -plastic was cleared -Proceeded with prior draw spine -OBL trajectory resolved | x software issue obl screen froze soft restart generic obl attempted processing time graduating image surgical arm showed surgical volume surgical arm showed collision body new new draw spine attempted reference frame changes mount plastic cleared proceeded prior draw spine obl trajectory resolved mazor x software kit | NAV3248 | NAV4091 | NAV3246 | NAV2047 | NAV4099 | NAV3248NAV4091NAV3246NAV2047NAV4099 | True
3994 | 703483138 | XSE - SRA Vibrating/Making Noise | 2019-11-22T20:13:32.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - SRA Vibrating/Making Noise 2019-11-22 00678980 (Rep): Medtronic received information regarding a guidance system being used during repair. It was reported that the surgical arm was noisy and vibrating. No patient involvement. 2019-12-06 IUD x6 (Other): No new information. 2020-03-31 IUD x3 (Other): No new information. 2020-04-23 IUD (Other): No new information. 2021-04-25 IUD (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2019-11-22 00678980 (Rep): Medtronic received information regarding a guidance system being used during repair. It was reported that the surgical arm was noisy and vibrating. No patient involvement. 2019-12-06 IUD x6 (Other): No new information. 2020-03-31 IUD x3 (Other): No new information. 2020-04-23 IUD (Other): No new information. 2021-04-25 IUD (Other): No new information. | xse sra vibratingmaking noise repair surgical arm noisy vibrating iud new iud new iud new iud new kit svc surgical arm | NAV3249 | NAV4101 | NAV2054 | NAV2047 | NAV4095 | NAV3249NAV4101NAV2054NAV2047NAV4095 | True
1632 | 703208215 | X - Registration Issue | 2019-06-14T18:23:19.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X - Registration Issue post case it was noticed registration cross view errors of T11 and L4 were greater than 12. No known reasoning intraoperatively for this observed by RTC. MAZOR X SOFTWARE KIT VER-2.3 | post case it was noticed registration cross view errors of T11 and L4 were greater than 12. No known reasoning intraoperatively for this observed by RTC. | x registration issue post case noticed registration cross view errors greater known reasoning intraoperatively observed rtc mazor x software kit | NAV3242 | NAV3246 | NAV3247 | NAV4091 | NAV4112 | NAV3242NAV3246NAV3247NAV4091NAV4112 | True
7106 | 704117730 | X - SRA Cannot Reach Final Destination | 2020-12-10T19:05:02.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | X - SRA Cannot Reach Final Destination 2020-12-10 00702467, CCR (Rep): Medtronic received information regarding a guidance system being used after a spinal procedure. It was reported that after the surgical system was removed from the bed, the surgical arm was sent to the clear left position. The manufacturer representative then sent the surgical arm to the drape position to remove the drape, but an error was displayed that the arm could not reach the destination. The representative tried sending the surgical arm to the other predefined positions , but the same error was displayed. A stress test was started and the surgical arm started to move. After the surgical arm moved enough, the stress test was stopped and the arm was able to be sent to the drape position. The system was able stored successfully. There was no patient involvement. 2020-12-10 IUD x3 (Other): No new information. 2020-12-21 IUD x2 (Other): No new information. 2021-Jan-20 IUD (other): No new information. 2021-03-03 302354044, IUD x2 (Other): No new information. 2021-03-18 IUD x5 (Other): No new information. Analysis information included in PE 704059236. 2021-06-15 IUD x4 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-12-10 00702467, CCR (Rep): Medtronic received information regarding a guidance system being used after a spinal procedure. It was reported that after the surgical system was removed from the bed, the surgical arm was sent to the clear left position. The manufacturer representative then sent the surgical arm to the drape position to remove the drape, but an error was displayed that the arm could not reach the destination. The representative tried sending the surgical arm to the other predefined positions , but the same error was displayed. A stress test was started and the surgical arm started to move. After the surgical arm moved enough, the stress test was stopped and the arm was able to be sent to the drape position. The system was able stored successfully. There was no patient involvement. 2020-12-10 IUD x3 (Other): No new information. 2020-12-21 IUD x2 (Other): No new information. 2021-Jan-20 IUD (other): No new information. 2021-03-03 302354044, IUD x2 (Other): No new information. 2021-03-18 IUD x5 (Other): No new information. Analysis information included in PE 704059236. 2021-06-15 IUD x4 (Other): No new information. | x sra reach final destination ccr spinal procedure surgical removed bed surgical arm sent clear left position sent surgical arm drape position remove drape error displayed arm reach destination tried sending surgical arm predefined positions error displayed stress test started surgical arm started surgical arm moved stress test stopped arm able sent drape position able stored successfully iud new iud new iud new iud new iud new analysis included pe iud new kit svc surgical arm refurb | NAV4095 | NAV3249 | NAV1006 | NAV4101 | NAV4107 | NAV4095NAV3249NAV1006NAV4101NAV4107 | True
9703 | 704701622 | XSE - Damaged Workstation Handle | 2021-11-05T22:31:41.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | MEC5416-01S | KIT SVC MEC5416-01S CART MAIN HANDLE | XSE - Damaged Workstation Handle 2021-Nov-05 Salesforce C00237459 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the main workstation handle was found to be damaged during servicing. There was no patient involvement. Information unrelated to this event was excluded from the PE; See PE 704682644 for positioner handle. 2021-11-12 IUD (Other): No new information. 2021-11-24 IUD (Other): No new information. 2021-12-07 IUD (Other): No new information. 2021-12-17 IUD (Other): No new information. 2022-01-14 email 2022-01-14 15:01 (Rep): Additional information received from a manufacturer representative reported that the handle was broken by customer. KIT SVC MEC5416-01S CART MAIN HANDLE | 2021-Nov-05 Salesforce C00237459 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the main workstation handle was found to be damaged during servicing. There was no patient involvement. Information unrelated to this event was excluded from the PE; See PE 704682644 for positioner handle. 2021-11-12 IUD (Other): No new information. 2021-11-24 IUD (Other): No new information. 2021-12-07 IUD (Other): No new information. 2021-12-17 IUD (Other): No new information. 2022-01-14 email 2022-01-14 15:01 (Rep): Additional information received from a manufacturer representative reported that the handle was broken by customer. | xse damaged workstation handle salesforce outside procedure main workstation handle damaged servicing unrelated excluded pe pe positioner handle iud new iud new iud new iud new email additional handle broken customer kit svc cart main handle | NAV2058 | NAV3251 | NAV2055 | NAV3247 | NAV3256 | NAV2058NAV3251NAV2055NAV3247NAV3256 | True
3243 | 703294879 | X - Drill Doesn't Fit in Drill Guide | 2019-08-05T19:30:44.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | KIT0737 | SHORT DRILL GUIDE AND CANNULA KIT0737 | X - Drill Doesn't Fit in Drill Guide 2019-08-01 SCL1902285, Intake Form (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the short drill bit did not fit through the short drill guide. There was no patient involvement. 2019-09-18 email 2019-09-18 09:01 (Rep): No new information. 2019-09-20 email 2019-09-20 11:30, email 2019-09-20 12:15 x2 (Rep): No new information. SHORT DRILL GUIDE AND CANNULA KIT0737 | 2019-08-01 SCL1902285, Intake Form (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the short drill bit did not fit through the short drill guide. There was no patient involvement. 2019-09-18 email 2019-09-18 09:01 (Rep): No new information. 2019-09-20 email 2019-09-20 11:30, email 2019-09-20 12:15 x2 (Rep): No new information. | x drill doesnt fit drill guide outside procedure short drill bit did fit short drill guide email new email email new short drill guide cannula | NAV3299 | NAV3295 | NAV1006 | NAV3264 | NAV3231 | NAV3299NAV3295NAV1006NAV3264NAV3231 | True
10378 | 704867778 | XSE - Difficulties Building Pressure | 2022-02-16T21:04:10.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Difficulties Building Pressure 2022-Feb-16 Salesforce C00243767 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system is having difficulties building pressure. No patient was present. 2022-02-18 IUD x4 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-04-10 302781931, IUD (Other): No new information. 2022-04-11 IUD x2 (Other): No new information. 2022-05-08 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2022-Feb-16 Salesforce C00243767 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system is having difficulties building pressure. No patient was present. 2022-02-18 IUD x4 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-04-10 302781931, IUD (Other): No new information. 2022-04-11 IUD x2 (Other): No new information. 2022-05-08 IUD (Other): No new information. | xse difficulties building pressure salesforce outside procedure having difficulties building pressure present iud new iud new iud new iud new iud new kit svc cart elec panel assy | NAV3264 | NAV2058 | NAV2047 | NAV2054 | NAV3249 | NAV3264NAV2058NAV2047NAV2054NAV3249 | True
179 | 703027410 | X - pedal broken | 2019-02-26T18:47:36.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - pedal broken 2019-Feb-20 SCL1900425, jpg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the plastic part on the black pedal was not attached to the pedal like the red one. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM TPL0059 | 2019-Feb-20 SCL1900425, jpg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the plastic part on the black pedal was not attached to the pedal like the red one. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. | x pedal broken jpg spinal procedure plastic black pedal attached pedal like red additional complications anticipated mazor x | NAV5000 | NAV3295 | NAV2054 | NAV4112 | NAV3242 | NAV5000NAV3295NAV2054NAV4112NAV3242 | False
10004 | 704768976 | XSE ¦ Pneumatic Leak | 2021-12-16T20:30:59.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE ¦ Pneumatic Leak 2021-Dec-16 Salesforce C00239868 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during preventative maintenance the field service engineer discovered that there is a leak in the system and believed it was the pneumatic panel that was leaking. It was noted that the replacement compressor had an out of box failure. There was no patient involved. 2022-01-05 email 2022-01-05 12:46 (Rep): The out of box failure was that the new pneumatic panel would not lift the arm up or down. 2022-01-10 IUD x2 (Other): No new information. 2022-01-14 email 2022-01-14 14:31 (Rep): The manufacturer representative reseated all the cables and it seemed to work better then it did when he decided to change it. The original part was still in use. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2021-Dec-16 Salesforce C00239868 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during preventative maintenance the field service engineer discovered that there is a leak in the system and believed it was the pneumatic panel that was leaking. It was noted that the replacement compressor had an out of box failure. There was no patient involved. 2022-01-05 email 2022-01-05 12:46 (Rep): The out of box failure was that the new pneumatic panel would not lift the arm up or down. 2022-01-10 IUD x2 (Other): No new information. 2022-01-14 email 2022-01-14 14:31 (Rep): The manufacturer representative reseated all the cables and it seemed to work better then it did when he decided to change it. The original part was still in use. | xse pneumatic leak salesforce outside procedure preventative maintenance field service engineer discovered leak believed pneumatic panel leaking noted replacement compressor box failure involved email box failure new pneumatic panel lift arm iud new email reseated cables work better did decided change original use kit svc pneumatic panel | NAV3264 | NAV2054 | NAV2058 | NAV2057 | NAV3254 | NAV3264NAV2054NAV2058NAV2057NAV3254 | True
10063 | 704792589 | XSE - No Display to Surgeon Tablet | 2022-01-05T20:34:22.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - No Display to Surgeon Tablet 2021-Dec-30 Salesforce C00240309, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that there was no display on the surgeon tablet when turning the guidance system on. There was no patient involvement. 2022-01-12 IUD (Other): No new information. SURGEON SCREEN | 2021-Dec-30 Salesforce C00240309, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that there was no display on the surgeon tablet when turning the guidance system on. There was no patient involvement. 2022-01-12 IUD (Other): No new information. | xse display surgeon tablet salesforce iud outside proc edure display surgeon tablet turning iud new surgeon screen | NAV3253 | NAV3264 | NAV3249 | NAV2054 | NAV3299 | NAV3253NAV3264NAV3249NAV2054NAV3299 | True
5808 | 703809837 | XSE - Positioner Dropping in Z Axis | 2020-06-18T19:56:06.000000+0000 | MazorX Stealth Edition | NAV4105 | UNEXPECTED POSITIONER MOVEMENT | 20 | ASM0214-02 | SSY | XSE - Positioner Dropping in Z Axis 2020-06-18 00691126 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative mounted the surgical system to the bed and verified the pressure was at 6.1 or 6.2. Once mounted, the position would drop in the Z-axis. The manual lock of the surgical system was in the unlocked position. There was no patient involvement. 2020-06-30 IUD (Other): No new information. 2020-07-06 IUD x3 (Other): No new information. SSY | 2020-06-18 00691126 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative mounted the surgical system to the bed and verified the pressure was at 6.1 or 6.2. Once mounted, the position would drop in the Z-axis. The manual lock of the surgical system was in the unlocked position. There was no patient involvement. 2020-06-30 IUD (Other): No new information. 2020-07-06 IUD x3 (Other): No new information. | xse positioner dropping z axis spinal procedure mounted surgical bed verified pressure mounted position drop zaxis manual lock surgical unlocked position iud new iud new ssy | NAV4105 | NAV4101 | NAV2057 | NAV3264 | NAV4104 | NAV4105NAV4101NAV2057NAV3264NAV4104 | True
3228 | 703285749 | XSE - arm failed accuracy test OOB | 2019-07-30T19:50:52.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | XSE - arm failed accuracy test OOB 2019-Jul-29 SCL1902244 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the rep received a surgical arm that failed advanced accuracy test on both the left and right side. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Jul-29 SCL1902244 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the rep received a surgical arm that failed advanced accuracy test on both the left and right side. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. | xse arm failed accuracy test oob outside procedure surgical arm failed advanced accuracy test left right additional complications anticipated surgical arm | NAV2054 | NAV4093 | NAV4092 | NAV3249 | NAV3264 | NAV2054NAV4093NAV4092NAV3249NAV3264 | True
7046 | 704107085 | XSE ¦ Replacement Accuracy Pointer | 2020-12-04T23:01:34.000000+0000 | MazorX Stealth Edition | NAV3294 | ARM GUIDE ISSUES | 30 | MAS2330-03S | KIT SVC MAS2330-03S ARM GUIDE | XSE ¦ Replacement Accuracy Pointer 2020-Dec-04 00702107 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the accuracy pointer and arm guide needed to be replaced in the setup drawer. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-14 IUD x3 (Other): No new information. 2021-01-07 email 2021-01-07 06:45 (Rep): Additional information was received stating that the accuracy pointer was bent and there was slight play in the arm guide. 2021-02-02 email 2021-02-02 05:01: (Rep): The cause of the accuracy pointer being bent and the slight play in arm guide was due to the bent arm guide at the tip where the pointer enters the divot and wanted to replace arm guide as a second layover of defense but it had slight play. KIT SVC MAS2330-03S ARM GUIDE | 2020-Dec-04 00702107 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the accuracy pointer and arm guide needed to be replaced in the setup drawer. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-14 IUD x3 (Other): No new information. 2021-01-07 email 2021-01-07 06:45 (Rep): Additional information was received stating that the accuracy pointer was bent and there was slight play in the arm guide. 2021-02-02 email 2021-02-02 05:01: (Rep): The cause of the accuracy pointer being bent and the slight play in arm guide was due to the bent arm guide at the tip where the pointer enters the divot and wanted to replace arm guide as a second layover of defense but it had slight play. | xse replacement accuracy pointer outside procedure accuracy pointer arm guide needed replaced setup drawer additional complications anticipated iud new email additional stating accuracy pointer bent slight play arm guide email cause accuracy pointer bent slight play arm guide bent arm guide tip pointer enters divot wanted replace arm guide second layover defense slight play kit svc arm guide | NAV3294 | NAV3299 | NAV3295 | NAV2054 | NAV2046 | NAV3294NAV3299NAV3295NAV2054NAV2046 | True
3516 | 703366363 | XSE - Arm Moved to Close to Work Volume | 2019-09-17T17:12:58.000000+0000 | MazorX Stealth Edition | NAV3239 | INACCURATE WORK VOLUMES | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - Arm Moved to Close to Work Volume 2019-09-10 SCL1902777, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after reinstalling the software on the system for the external monitor, the safety margin on the no fly zone of the work volume needs to be adjusted. The 3Define scan looked fairly accurate, but the surgical arm moved too close to the Schanz pin while moving to the AP and oblique trajectories. Three was no patient harm and the procedure was not delayed more than an hour. UNK MAZOR X 3DEFINE CAMERA | 2019-09-10 SCL1902777, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after reinstalling the software on the system for the external monitor, the safety margin on the no fly zone of the work volume needs to be adjusted. The 3Define scan looked fairly accurate, but the surgical arm moved too close to the Schanz pin while moving to the AP and oblique trajectories. Three was no patient harm and the procedure was not delayed more than an hour. | xse arm moved close work volume spinal procedure reinstalling software external monitor safety margin fly zone work volume needs adjusted scan looked fairly accurate surgical arm moved close schanz pin moving ap oblique trajectories procedure delayed hour unk mazor x camera | NAV4096 | NAV3244 | NAV3239 | NAV4101 | NAV3248 | NAV4096NAV3244NAV3239NAV4101NAV3248 | True
9508 | 704660601 | XSE - Lit - Morse - Malfuntion | 2021-10-13T16:39:46.000000+0000 | MazorX Stealth Edition | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Lit - Morse - Malfuntion Morse K, Heath M, Avrumova F, Defrancesco C, Fabricant P, Lebl D, Widmann R. Comprehensive Error Analysis for Robotic-assisted Placement of Pedicle Screws in Pediatric Spinal Deformity: The Initial Learning Curve. Journal of Pediatric Orthopaedics. 2021. DOI: 10.1 097/BPO.0000000000001842. Abstract: Surgical navigation improves pedicle screw insertion accuracy and reliability. Robotic-assisted spinal surgery and screw placement has not been fully assessed in pediatric patients with spine deformity undergoing posterior spinal fusion. The purpose of this study was to describe the learning curve for robotically assisted pedicle screw placement in pediatric patients. A retrospective review on a consecutive series of the first 19 pediatric patients who underwent posterior spinal fusion by a single surgeon using robotic navigation was performed. Demographics, curve parameters, pedicle diameter, vertebral rotation, and additional outcome measures were recorded. Screw position was assessed with calibrated intraoperative 3-dimensional fluoroscopic images. All complications of planned and placed robotically placed screws were recorded. A total of 194 left-sided screws were planned as robot assisted. One hundred sixty-eight of the robotically planned screws (86.6%) were placed with robot assistance; 29 robotically planned screws (15.0%) were abandoned or converted to freehand. The mean time per robotically placed screw was 3.6 ±2.4 minutes. Fifteen breaches (8.9%) and 1 anterior perforation occurred with 2 critical (> 2 to 4 mm) breaches, 1 was associated with a durotomy, and both occurred in the first case. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. Six breaches occurred in the first case. The odds ratio of obtaining a breach in screws with a matched trajectory was 0.275 (95% confidence interval: 0.089-0.848). Screw time and accuracy improved and the number of breaches decreased after 10 cases. This series had 2 critical breaches (between 2 and 4 mm) on the first case. Overall, excluding the 2 critical breaches, 98.8% of robotically executed screws were placed without a critical breach, which is comparable to previous pediatric deformity studies. Caution should be exercised during the initial training period to avoid complications as experience and training lead to an improved understanding of surgical planning, skive, and soft-tissue pressure on the end-effector all of which can impact accuracy of robotically assisted pedicle screw placement. Reported Events: 1. After use of the guidance system during spinal procedures to place screws, 15 medial breaches occurred. One 1mm anterior perforation occurred during the second case at the T4 level. 11 unplanned lateral breaches occurred; however, none of them were clinically or radiographically significant with all screw tips engaging the vertebral body. Two medial breaches occurred in the first case, one of which was associated with a durotomy. Critically breached screws were removed and replaced with either a pedicle screw using freehand technique or a pedicle hook. The durotomy occurred during drilling due to operator error from inadequate tightening of the screw between the bone mount bridge and dual spinous process clamp. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. The second critical breach occurred due to medial skive from improper seating of the anti-skive pin. There were no breaches >4 mm. 2. After use of the guidance system during spinal procedures to place screws, 29 screws were abandoned. 14 screws were abandoned due to soft tissue contact with the robotic end effector resulting in end effector translation relative to the fixed position of the spine. Increasing positive axial rotation was associated with increased rate of abandoned screws. 7 screws were abandoned due to unknown reasons, 3 screws were abandoned after being breached and removed, 2 screws were abandoned due to surgical decision, 1 screw was abandoned due to a clamp/bridge interfering with planned screw trajectory, 1 screw was abandoned due to incorrect position and 1 screw was abandoned due to poor purchase. The abandoned screws were placed freehand. MAZOR X SYSTEM TPL0059 | Morse K, Heath M, Avrumova F, Defrancesco C, Fabricant P, Lebl D, Widmann R. Comprehensive Error Analysis for Robotic-assisted Placement of Pedicle Screws in Pediatric Spinal Deformity: The Initial Learning Curve. Journal of Pediatric Orthopaedics. 2021. DOI: 10.1 097/BPO.0000000000001842. Abstract: Surgical navigation improves pedicle screw insertion accuracy and reliability. Robotic-assisted spinal surgery and screw placement has not been fully assessed in pediatric patients with spine deformity undergoing posterior spinal fusion. The purpose of this study was to describe the learning curve for robotically assisted pedicle screw placement in pediatric patients. A retrospective review on a consecutive series of the first 19 pediatric patients who underwent posterior spinal fusion by a single surgeon using robotic navigation was performed. Demographics, curve parameters, pedicle diameter, vertebral rotation, and additional outcome measures were recorded. Screw position was assessed with calibrated intraoperative 3-dimensional fluoroscopic images. All complications of planned and placed robotically placed screws were recorded. A total of 194 left-sided screws were planned as robot assisted. One hundred sixty-eight of the robotically planned screws (86.6%) were placed with robot assistance; 29 robotically planned screws (15.0%) were abandoned or converted to freehand. The mean time per robotically placed screw was 3.6 ±2.4 minutes. Fifteen breaches (8.9%) and 1 anterior perforation occurred with 2 critical (> 2 to 4 mm) breaches, 1 was associated with a durotomy, and both occurred in the first case. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. Six breaches occurred in the first case. The odds ratio of obtaining a breach in screws with a matched trajectory was 0.275 (95% confidence interval: 0.089-0.848). Screw time and accuracy improved and the number of breaches decreased after 10 cases. This series had 2 critical breaches (between 2 and 4 mm) on the first case. Overall, excluding the 2 critical breaches, 98.8% of robotically executed screws were placed without a critical breach, which is comparable to previous pediatric deformity studies. Caution should be exercised during the initial training period to avoid complications as experience and training lead to an improved understanding of surgical planning, skive, and soft-tissue pressure on the end-effector all of which can impact accuracy of robotically assisted pedicle screw placement. Reported Events: 1. After use of the guidance system during spinal procedures to place screws, 15 medial breaches occurred. One 1mm anterior perforation occurred during the second case at the T4 level. 11 unplanned lateral breaches occurred; however, none of them were clinically or radiographically significant with all screw tips engaging the vertebral body. Two medial breaches occurred in the first case, one of which was associated with a durotomy. Critically breached screws were removed and replaced with either a pedicle screw using freehand technique or a pedicle hook. The durotomy occurred during drilling due to operator error from inadequate tightening of the screw between the bone mount bridge and dual spinous process clamp. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. The second critical breach occurred due to medial skive from improper seating of the anti-skive pin. There were no breaches >4 mm. 2. After use of the guidance system during spinal procedures to place screws, 29 screws were abandoned. 14 screws were abandoned due to soft tissue contact with the robotic end effector resulting in end effector translation relative to the fixed position of the spine. Increasing positive axial rotation was associated with increased rate of abandoned screws. 7 screws were abandoned due to unknown reasons, 3 screws were abandoned after being breached and removed, 2 screws were abandoned due to surgical decision, 1 screw was abandoned due to a clamp/bridge interfering with planned screw trajectory, 1 screw was abandoned due to incorrect position and 1 screw was abandoned due to poor purchase. The abandoned screws were placed freehand. | xse lit morse malfuntion morse k heath m avrumova f defrancesco c fabricant p lebl d widmann r comprehensive error analysis roboticassisted placement pedicle screws pediatric spinal deformity initial learning curve journal pediatric orthopaedics doi abstract surgical navigation improves pedicle screw insertion accuracy reliability roboticassisted spinal surgery screw placement fully assessed pediatric patients spine deformity undergoing posterior spinal fusion purpose study learning curve robotically assisted pedicle screw placement pediatric patients retrospective review consecutive series pediatric patients underwent posterior spinal fusion single surgeon using robotic navigation performed demographics curve parameters pedicle diameter vertebral rotation additional outcome measures recorded screw position assessed calibrated intraoperative fluoroscopic images complications planned placed robotically placed screws recorded total leftsided screws planned robot assisted sixtyeight robotically planned screws placed robot assistance robotically planned screws abandoned converted freehand mean time robotically placed screw ± minutes breaches anterior perforation occurred critical breaches associated durotomy occurred case intraoperativepostoperative neuromonitoring changes sequela durotomy breaches occurred case odds ratio obtaining breach screws matched trajectory screw time accuracy improved number breaches decreased cases series critical breaches case overall excluding critical breaches robotically executed screws placed critical breach comparable previous pediatric deformity studies caution exercised initial training period avoid complications experience training lead improved understanding surgical planning skive softtissue pressure endeffector impact accuracy robotically assisted pedicle screw placement events use spinal procedures place screws medial breaches occurred anterior perforation occurred second case level unplanned lateral breaches occurred clinically radiographically significant screw tips engaging vertebral body medial breaches occurred case associated durotomy critically breached screws removed replaced pedicle screw using freehand technique pedicle hook durotomy occurred drilling operator error inadequate tightening screw bone mount bridge dual spinous process clamp intraoperativepostoperative neuromonitoring changes sequela durotomy second critical breach occurred medial skive improper seating antiskive pin breaches mm use spinal procedures place screws screws abandoned screws abandoned soft tissue contact robotic end effector resulting end effector translation relative fixed position spine increasing positive axial rotation associated increased rate abandoned screws screws abandoned unknown reasons screws abandoned breached removed screws abandoned surgical decision screw abandoned clampbridge interfering planned screw trajectory screw abandoned incorrect position screw abandoned poor purchase abandoned screws placed freehand mazor x | NAV4098 | NAV3261 | NAV5000 | NAV3243 | NAV4090 | NAV4098NAV3261NAV5000NAV3243NAV4090 | True
9510 | 704660601 | XSE - Lit - Morse - Malfuntion | 2021-10-13T16:39:46.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Lit - Morse - Malfuntion Morse K, Heath M, Avrumova F, Defrancesco C, Fabricant P, Lebl D, Widmann R. Comprehensive Error Analysis for Robotic-assisted Placement of Pedicle Screws in Pediatric Spinal Deformity: The Initial Learning Curve. Journal of Pediatric Orthopaedics. 2021. DOI: 10.1 097/BPO.0000000000001842. Abstract: Surgical navigation improves pedicle screw insertion accuracy and reliability. Robotic-assisted spinal surgery and screw placement has not been fully assessed in pediatric patients with spine deformity undergoing posterior spinal fusion. The purpose of this study was to describe the learning curve for robotically assisted pedicle screw placement in pediatric patients. A retrospective review on a consecutive series of the first 19 pediatric patients who underwent posterior spinal fusion by a single surgeon using robotic navigation was performed. Demographics, curve parameters, pedicle diameter, vertebral rotation, and additional outcome measures were recorded. Screw position was assessed with calibrated intraoperative 3-dimensional fluoroscopic images. All complications of planned and placed robotically placed screws were recorded. A total of 194 left-sided screws were planned as robot assisted. One hundred sixty-eight of the robotically planned screws (86.6%) were placed with robot assistance; 29 robotically planned screws (15.0%) were abandoned or converted to freehand. The mean time per robotically placed screw was 3.6 ±2.4 minutes. Fifteen breaches (8.9%) and 1 anterior perforation occurred with 2 critical (> 2 to 4 mm) breaches, 1 was associated with a durotomy, and both occurred in the first case. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. Six breaches occurred in the first case. The odds ratio of obtaining a breach in screws with a matched trajectory was 0.275 (95% confidence interval: 0.089-0.848). Screw time and accuracy improved and the number of breaches decreased after 10 cases. This series had 2 critical breaches (between 2 and 4 mm) on the first case. Overall, excluding the 2 critical breaches, 98.8% of robotically executed screws were placed without a critical breach, which is comparable to previous pediatric deformity studies. Caution should be exercised during the initial training period to avoid complications as experience and training lead to an improved understanding of surgical planning, skive, and soft-tissue pressure on the end-effector all of which can impact accuracy of robotically assisted pedicle screw placement. Reported Events: 1. After use of the guidance system during spinal procedures to place screws, 15 medial breaches occurred. One 1mm anterior perforation occurred during the second case at the T4 level. 11 unplanned lateral breaches occurred; however, none of them were clinically or radiographically significant with all screw tips engaging the vertebral body. Two medial breaches occurred in the first case, one of which was associated with a durotomy. Critically breached screws were removed and replaced with either a pedicle screw using freehand technique or a pedicle hook. The durotomy occurred during drilling due to operator error from inadequate tightening of the screw between the bone mount bridge and dual spinous process clamp. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. The second critical breach occurred due to medial skive from improper seating of the anti-skive pin. There were no breaches >4 mm. 2. After use of the guidance system during spinal procedures to place screws, 29 screws were abandoned. 14 screws were abandoned due to soft tissue contact with the robotic end effector resulting in end effector translation relative to the fixed position of the spine. Increasing positive axial rotation was associated with increased rate of abandoned screws. 7 screws were abandoned due to unknown reasons, 3 screws were abandoned after being breached and removed, 2 screws were abandoned due to surgical decision, 1 screw was abandoned due to a clamp/bridge interfering with planned screw trajectory, 1 screw was abandoned due to incorrect position and 1 screw was abandoned due to poor purchase. The abandoned screws were placed freehand. MAZOR X SYSTEM TPL0059 | Morse K, Heath M, Avrumova F, Defrancesco C, Fabricant P, Lebl D, Widmann R. Comprehensive Error Analysis for Robotic-assisted Placement of Pedicle Screws in Pediatric Spinal Deformity: The Initial Learning Curve. Journal of Pediatric Orthopaedics. 2021. DOI: 10.1 097/BPO.0000000000001842. Abstract: Surgical navigation improves pedicle screw insertion accuracy and reliability. Robotic-assisted spinal surgery and screw placement has not been fully assessed in pediatric patients with spine deformity undergoing posterior spinal fusion. The purpose of this study was to describe the learning curve for robotically assisted pedicle screw placement in pediatric patients. A retrospective review on a consecutive series of the first 19 pediatric patients who underwent posterior spinal fusion by a single surgeon using robotic navigation was performed. Demographics, curve parameters, pedicle diameter, vertebral rotation, and additional outcome measures were recorded. Screw position was assessed with calibrated intraoperative 3-dimensional fluoroscopic images. All complications of planned and placed robotically placed screws were recorded. A total of 194 left-sided screws were planned as robot assisted. One hundred sixty-eight of the robotically planned screws (86.6%) were placed with robot assistance; 29 robotically planned screws (15.0%) were abandoned or converted to freehand. The mean time per robotically placed screw was 3.6 ±2.4 minutes. Fifteen breaches (8.9%) and 1 anterior perforation occurred with 2 critical (> 2 to 4 mm) breaches, 1 was associated with a durotomy, and both occurred in the first case. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. Six breaches occurred in the first case. The odds ratio of obtaining a breach in screws with a matched trajectory was 0.275 (95% confidence interval: 0.089-0.848). Screw time and accuracy improved and the number of breaches decreased after 10 cases. This series had 2 critical breaches (between 2 and 4 mm) on the first case. Overall, excluding the 2 critical breaches, 98.8% of robotically executed screws were placed without a critical breach, which is comparable to previous pediatric deformity studies. Caution should be exercised during the initial training period to avoid complications as experience and training lead to an improved understanding of surgical planning, skive, and soft-tissue pressure on the end-effector all of which can impact accuracy of robotically assisted pedicle screw placement. Reported Events: 1. After use of the guidance system during spinal procedures to place screws, 15 medial breaches occurred. One 1mm anterior perforation occurred during the second case at the T4 level. 11 unplanned lateral breaches occurred; however, none of them were clinically or radiographically significant with all screw tips engaging the vertebral body. Two medial breaches occurred in the first case, one of which was associated with a durotomy. Critically breached screws were removed and replaced with either a pedicle screw using freehand technique or a pedicle hook. The durotomy occurred during drilling due to operator error from inadequate tightening of the screw between the bone mount bridge and dual spinous process clamp. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. The second critical breach occurred due to medial skive from improper seating of the anti-skive pin. There were no breaches >4 mm. 2. After use of the guidance system during spinal procedures to place screws, 29 screws were abandoned. 14 screws were abandoned due to soft tissue contact with the robotic end effector resulting in end effector translation relative to the fixed position of the spine. Increasing positive axial rotation was associated with increased rate of abandoned screws. 7 screws were abandoned due to unknown reasons, 3 screws were abandoned after being breached and removed, 2 screws were abandoned due to surgical decision, 1 screw was abandoned due to a clamp/bridge interfering with planned screw trajectory, 1 screw was abandoned due to incorrect position and 1 screw was abandoned due to poor purchase. The abandoned screws were placed freehand. | xse lit morse malfuntion morse k heath m avrumova f defrancesco c fabricant p lebl d widmann r comprehensive error analysis roboticassisted placement pedicle screws pediatric spinal deformity initial learning curve journal pediatric orthopaedics doi abstract surgical navigation improves pedicle screw insertion accuracy reliability roboticassisted spinal surgery screw placement fully assessed pediatric patients spine deformity undergoing posterior spinal fusion purpose study learning curve robotically assisted pedicle screw placement pediatric patients retrospective review consecutive series pediatric patients underwent posterior spinal fusion single surgeon using robotic navigation performed demographics curve parameters pedicle diameter vertebral rotation additional outcome measures recorded screw position assessed calibrated intraoperative fluoroscopic images complications planned placed robotically placed screws recorded total leftsided screws planned robot assisted sixtyeight robotically planned screws placed robot assistance robotically planned screws abandoned converted freehand mean time robotically placed screw ± minutes breaches anterior perforation occurred critical breaches associated durotomy occurred case intraoperativepostoperative neuromonitoring changes sequela durotomy breaches occurred case odds ratio obtaining breach screws matched trajectory screw time accuracy improved number breaches decreased cases series critical breaches case overall excluding critical breaches robotically executed screws placed critical breach comparable previous pediatric deformity studies caution exercised initial training period avoid complications experience training lead improved understanding surgical planning skive softtissue pressure endeffector impact accuracy robotically assisted pedicle screw placement events use spinal procedures place screws medial breaches occurred anterior perforation occurred second case level unplanned lateral breaches occurred clinically radiographically significant screw tips engaging vertebral body medial breaches occurred case associated durotomy critically breached screws removed replaced pedicle screw using freehand technique pedicle hook durotomy occurred drilling operator error inadequate tightening screw bone mount bridge dual spinous process clamp intraoperativepostoperative neuromonitoring changes sequela durotomy second critical breach occurred medial skive improper seating antiskive pin breaches mm use spinal procedures place screws screws abandoned screws abandoned soft tissue contact robotic end effector resulting end effector translation relative fixed position spine increasing positive axial rotation associated increased rate abandoned screws screws abandoned unknown reasons screws abandoned breached removed screws abandoned surgical decision screw abandoned clampbridge interfering planned screw trajectory screw abandoned incorrect position screw abandoned poor purchase abandoned screws placed freehand mazor x | NAV4098 | NAV3261 | NAV5000 | NAV3243 | NAV4090 | NAV4098NAV3261NAV5000NAV3243NAV4090 | True
4366 | 703538526 | X - Schanz Pin Won't Fit Schanz Arm | 2019-12-30T19:57:04.000000+0000 | MazorX | NAV2054 | OUT OF BOX FAILURES | 30 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | X - Schanz Pin Won't Fit Schanz Arm 2019-12-30 00681187 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the 80mm Schanz pin where it tapers from 4mm to 3mm was not tapered correctly. This caused the Schanz arm to not go over the Schanz pin. No patient was present. 2020-03-30 IUD x3 (Other): No new information. UNK MAZOR X ACCESSORIES | 2019-12-30 00681187 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the 80mm Schanz pin where it tapers from 4mm to 3mm was not tapered correctly. This caused the Schanz arm to not go over the Schanz pin. No patient was present. 2020-03-30 IUD x3 (Other): No new information. | x schanz pin wont fit schanz arm outside procedure schanz pin tapers tapered correctly caused schanz arm schanz pin present iud new unk mazor x accessories | NAV3299 | NAV2053 | NAV3295 | NAV2054 | NAV3298 | NAV3299NAV2053NAV3295NAV2054NAV3298 | True
5930 | 703839562 | XSE - Stripped Dual Clamp Screw | 2020-07-07T19:10:10.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1355-05S | KIT SVC MAS1355-05S DUAL CLAMP | XSE - Stripped Dual Clamp Screw 2020-07-07 00692171 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws used to tighten down the clamp onto the spinous process of the dual clamp was stripped. There was no patient involvement. 2020-07-17 IUD x2 (Other): No new information. 2020-07-28 e1 (Rep): No new information. 2020-08-03 IUD x2 (Other): No new information. 2020-08-19 IUD (Other): No new information. 2020-08-25 email 2020-08-25 12:30 (Rep): Additional information received from a manufacturer representative reported that the site decided to keep the clamp in case the issue occurs again. KIT SVC MAS1355-05S DUAL CLAMP | 2020-07-07 00692171 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws used to tighten down the clamp onto the spinous process of the dual clamp was stripped. There was no patient involvement. 2020-07-17 IUD x2 (Other): No new information. 2020-07-28 e1 (Rep): No new information. 2020-08-03 IUD x2 (Other): No new information. 2020-08-19 IUD (Other): No new information. 2020-08-25 email 2020-08-25 12:30 (Rep): Additional information received from a manufacturer representative reported that the site decided to keep the clamp in case the issue occurs again. | xse stripped dual clamp screw outside procedure screws tighten clamp spinous process dual clamp stripped iud new new iud new iud new email additional site decided clamp case issue occurs kit svc dual clamp | NAV3299 | NAV3295 | NAV4090 | NAV2054 | NAV3261 | NAV3299NAV3295NAV4090NAV2054NAV3261 | True
9193 | 704571258 | XSE - Software Crash | 2021-08-23T17:45:59.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Software Crash 2021-Aug-23 Salesforce C00233327 (rep, hcp, for): Medtronic received information regarding a guidance system. It was reported that after the installation of software version 5.0 on surgeons planning computer, the software crashes during planning. There was no patient involved. 2021-08-31 e1 (Rep, HCP, For): No new information. 2021-10-01 email 2021-10-01 04:01 (Rep, For): No new information. 2021-10-26 email 2021-10-26 20:05 (Rep, For): No new information. 2021-10-29 IUD (Other, For): No new information. 2021-11-10 IUD (Other, For): No new information. 2021-11-16 email 2021-11-16 10:31 (Rep, For): Additional information received from a manufacturer representative reported that the cause was not determined. 2021-11-24 IUD (Other, For): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Aug-23 Salesforce C00233327 (rep, hcp, for): Medtronic received information regarding a guidance system. It was reported that after the installation of software version 5.0 on surgeons planning computer, the software crashes during planning. There was no patient involved. 2021-08-31 e1 (Rep, HCP, For): No new information. 2021-10-01 email 2021-10-01 04:01 (Rep, For): No new information. 2021-10-26 email 2021-10-26 20:05 (Rep, For): No new information. 2021-10-29 IUD (Other, For): No new information. 2021-11-10 IUD (Other, For): No new information. 2021-11-16 email 2021-11-16 10:31 (Rep, For): Additional information received from a manufacturer representative reported that the cause was not determined. 2021-11-24 IUD (Other, For): No new information. | xse software crash salesforce installation software version surgeons planning computer software crashes planning involved new email new email new iud new iud new email additional cause determined iud new mazorx sw installation kit | NAV3246 | NAV3247 | NAV2056 | NAV3248 | NAV2054 | NAV3246NAV3247NAV2056NAV3248NAV2054 | True
3352 | 703320786 | XSE - Dilator Showed Lateral in Arm Guid | 2019-08-21T01:25:21.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Dilator Showed Lateral in Arm Guid 2019-08-16 SCL1902396, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc using fixation. It was reported that preoperative planning was completed and reviewed with the surgeon to pay close special attention to skin incisions, rod alignment and screw sizes. Two Schanz arms were used as the mount due to clinical indication and surgeon preference. Two AP and four oblique images were needed for registration. Automatic segmentation was successful at L3, but a red value was seen at S1. Manual manipulation of the segmentation line at S1 resulted in green values at each level. The surgical arm was sent to the L2 trajectory. Once the dilator and cannula were placed, navigation showed the dilator was lateral. Another snapshot was done, but the dilator looked a bit lateral. The dilator was tested in the divot of the arm guide and it was accurate. The dilator was further tested on the spinous process and it was accurate. Once the dilator was placed in the arm guide, the trajectory was lateral. The short NavLock worked with the other tools and there was only an issue with the dilator. The surgeon decided to proceed with the case. The issue with the dilator was seen on the whole left side. On the right side, the dilator looked to be in the correct position. After inserting all of the screws, confirmation image with the C-arm showed the left L2 and left L3 screws to be lateral to the planned position. The screws were removed and placed freehand. The rest of the screws looked to be in a good position. The surgeon believed the screws skived off of the planned trajectory. The surgeon further noted that as the screw was being inserted that it did not feel correct for those two trajectories. There was no patient harm and the procedure was not delayed more than an hour. 2019-09-17 email 2019-09-17 22:45 (Rep): Additional information received from a manufacturer representative reported the cause was not determined. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-08-16 SCL1902396, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc using fixation. It was reported that preoperative planning was completed and reviewed with the surgeon to pay close special attention to skin incisions, rod alignment and screw sizes. Two Schanz arms were used as the mount due to clinical indication and surgeon preference. Two AP and four oblique images were needed for registration. Automatic segmentation was successful at L3, but a red value was seen at S1. Manual manipulation of the segmentation line at S1 resulted in green values at each level. The surgical arm was sent to the L2 trajectory. Once the dilator and cannula were placed, navigation showed the dilator was lateral. Another snapshot was done, but the dilator looked a bit lateral. The dilator was tested in the divot of the arm guide and it was accurate. The dilator was further tested on the spinous process and it was accurate. Once the dilator was placed in the arm guide, the trajectory was lateral. The short NavLock worked with the other tools and there was only an issue with the dilator. The surgeon decided to proceed with the case. The issue with the dilator was seen on the whole left side. On the right side, the dilator looked to be in the correct position. After inserting all of the screws, confirmation image with the C-arm showed the left L2 and left L3 screws to be lateral to the planned position. The screws were removed and placed freehand. The rest of the screws looked to be in a good position. The surgeon believed the screws skived off of the planned trajectory. The surgeon further noted that as the screw was being inserted that it did not feel correct for those two trajectories. There was no patient harm and the procedure was not delayed more than an hour. 2019-09-17 email 2019-09-17 22:45 (Rep): Additional information received from a manufacturer representative reported the cause was not determined. | xse dilator showed lateral arm guid ccr spinal procedure treat degenerative disc using fixation preoperative planning completed reviewed surgeon pay close special attention skin incisions rod alignment screw sizes schanz arms mount clinical indication surgeon preference ap oblique images needed registration automatic segmentation successful red value seen manual manipulation segmentation line resulted green values level surgical arm sent trajectory dilator cannula placed navigation showed dilator lateral snapshot dilator looked bit lateral dilator tested divot arm guide accurate dilator tested spinous process accurate dilator placed arm guide trajectory lateral short navlock worked tools issue dilator surgeon decided proceed case issue dilator seen left right dilator looked correct position inserting screws confirmation image carm showed left left screws lateral planned position screws removed placed freehand rest screws looked good position surgeon believed screws skived planned trajectory surgeon noted screw inserted did feel correct trajectories procedure delayed hour email additional cause determined mazor x software installation kit | NAV3246 | NAV5000 | NAV4095 | NAV5001 | NAV3244 | NAV3246NAV5000NAV4095NAV5001NAV3244 | True
798 | 703176145 | X-Positioner Issues | 2019-05-28T19:37:59.000000+0000 | MazorX | NAV3251 | POSITIONER ISSUES | 20 | ASM0283-01 | RNX-PO-117 | X-Positioner Issues DURING PM FOUND RIGHT SRS HANDLE BROKEN RNX-PO-117 | DURING PM FOUND RIGHT SRS HANDLE BROKEN | xpositioner issues pm right srs handle broken | NAV3251 | NAV3295 | NAV3249 | NAV4091 | NAV2058 | NAV3251NAV3295NAV3249NAV4091NAV2058 | True
7139 | 704126497 | XSE - Stripped Screws/Rusted Schanz Arm | 2020-12-15T21:14:31.000000+0000 | MazorX Stealth Edition | NAV3296 | DISCOLORATION | 30 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Stripped Screws/Rusted Schanz Arm 2020-12-15 00702713, 00702714 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the clamp link screw was stripped and two Schanz arms were found to be rusty during sterile processing. There was no patient involvement. 2020-12-22 IUD x6 (Other): No new information. 2021-Jan-20 e1 (rep): It was noted that the clamp link screw always strips. The Shanz arms rust because of stainless steel. The rep did not have the lot numbers or products. Hospital will not use them are trying to remove the rust. The rep was told that the SPD sterilization was not done to guidelines, and advised the account but they said it’s made of steel that rust. KIT SVC MAS2351-04S SCHANZ ARM | 2020-12-15 00702713, 00702714 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the clamp link screw was stripped and two Schanz arms were found to be rusty during sterile processing. There was no patient involvement. 2020-12-22 IUD x6 (Other): No new information. 2021-Jan-20 e1 (rep): It was noted that the clamp link screw always strips. The Shanz arms rust because of stainless steel. The rep did not have the lot numbers or products. Hospital will not use them are trying to remove the rust. The rep was told that the SPD sterilization was not done to guidelines, and advised the account but they said it’s made of steel that rust. | xse stripped screwsrusted schanz arm outside procedure clamp link screw stripped schanz arms rusty sterile processing iud new noted clamp link screw strips shanz arms rust stainless steel did lot numbers products hospital use trying remove rust told spd sterilization guidelines advised account said it’s steel rust kit svc schanz arm | NAV3299 | NAV3296 | NAV4100 | NAV4090 | NAV3295 | NAV3299NAV3296NAV4100NAV4090NAV3295 | True
7265 | 704159268 | X - Stripped Schanz Bond Screws | 2021-01-07T05:25:14.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC4530-02S | KIT SVC MEC4530-02S BALL ARM TRGT SCREW | X - Stripped Schanz Bond Screws 2021-01-05 00703816, mPXR 792983, email (Rep): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that the two tightening screws on the Schanz bond were stripped. There was no patient involvement. 2021-01-21 IUD x2 (Other): No new information. 2021-Feb-04 IUDx2 (other): No new information. 2021-03-08 IUD x2 (Other): No new information. 2021-03-04 email x2 (Rep): No new information. 2021-04-05 IUD (Other): No new information. 2021-04-07 email 2021-04-07 14:31 (Rep): Additional information received from a manufacturer representative reported that the screws stripped due to poor quality of the metal. 2021-04-19 email 2021-04-19 10:16, email 2021-04-19 10:31 (Rep): No new information. 2021-04-22 email 2021-04-22 12:31 (Rep): No new information. 2021-04-29 IUD (Other): No new information. 2021-05-06 IUD (Other): No new information. 2021-06-21 IUD (Other): No new information. 2021-06-29 IUD (Other): No new information. KIT SVC MEC4530-02S BALL ARM TRGT SCREW | 2021-01-05 00703816, mPXR 792983, email (Rep): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that the two tightening screws on the Schanz bond were stripped. There was no patient involvement. 2021-01-21 IUD x2 (Other): No new information. 2021-Feb-04 IUDx2 (other): No new information. 2021-03-08 IUD x2 (Other): No new information. 2021-03-04 email x2 (Rep): No new information. 2021-04-05 IUD (Other): No new information. 2021-04-07 email 2021-04-07 14:31 (Rep): Additional information received from a manufacturer representative reported that the screws stripped due to poor quality of the metal. 2021-04-19 email 2021-04-19 10:16, email 2021-04-19 10:31 (Rep): No new information. 2021-04-22 email 2021-04-22 12:31 (Rep): No new information. 2021-04-29 IUD (Other): No new information. 2021-05-06 IUD (Other): No new information. 2021-06-21 IUD (Other): No new information. 2021-06-29 IUD (Other): No new information. | x stripped schanz bond screws mpxr email outside pr ocedure tightening screws schanz bond stripped iud new new iud new email new iud new email additional screws stripped poor quality metal email email new email new iud new iud new iud new iud new kit svc ball arm trgt screw | NAV3299 | NAV3295 | NAV4090 | NAV3296 | NAV4100 | NAV3299NAV3295NAV4090NAV3296NAV4100 | True
4556 | 703550861 | X-Arm, WS, & 3Define issues | 2020-01-08T15:09:53.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 30 | ASM0206-01 | SURGICAL ARM | X-Arm, WS, & 3Define issues 2018-1-10 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative received a canbus error, an arm shift error at joint one, and error #8025. The 3Define camera did not work and the surgeon had to use the generic work volume. There was no patient harm and the procedure was not delayed over an hour. SURGICAL ARM | 2018-1-10 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative received a canbus error, an arm shift error at joint one, and error #8025. The 3Define camera did not work and the surgeon had to use the generic work volume. There was no patient harm and the procedure was not delayed over an hour. | xarm ws issues ccr spinal procedure canbus error arm shift error joint error camera did work surgeon use generic work volume procedure delayed hour surgical arm | NAV2047 | NAV3244 | NAV4098 | NAV4101 | NAV4095 | NAV2047NAV3244NAV4098NAV4101NAV4095 | True
4164 | 703507909 | XSE - Image Adaptor Latch Damaged | 2019-12-09T15:22:08.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 20 | KIT0586-03 | REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | XSE - Image Adaptor Latch Damaged 2019-12-09 00679827 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the image adapter latch is stuck and nonfunctional. It was believed that it was dropped and is now bent which is preventing it from moving. The rest of the image adapter is fine, its simply the latching mechanism that allows you to secure to the C-Arm. No patient was present. 2019-12-26 IUD x2 (Other): No new information. REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | 2019-12-09 00679827 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the image adapter latch is stuck and nonfunctional. It was believed that it was dropped and is now bent which is preventing it from moving. The rest of the image adapter is fine, its simply the latching mechanism that allows you to secure to the C-Arm. No patient was present. 2019-12-26 IUD x2 (Other): No new information. | xse image adaptor latch damaged outside procedure image adapter latch stuck nonfunctional believed dropped bent preventing moving rest image adapter fine simply latching mechanism allows secure carm present iud new replacement grip set image adaptor | NAV3258 | NAV3299 | NAV2054 | NAV3237 | NAV4101 | NAV3258NAV3299NAV2054NAV3237NAV4101 | True
1548 | 703205320 | X - Plastic Cover issue | 2019-06-13T14:11:57.000000+0000 | MazorX | NAV4099 | PLASTIC COVER ISSUES | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - Plastic Cover issue It was reported that the motor cover falls out of the joint (3). UNK MAZOR X ARM | It was reported that the motor cover falls out of the joint (3). | x plastic cover issue motor cover falls joint unk mazor x arm | NAV4099 | NAV2054 | NAV4098 | NAV2058 | NAV3244 | NAV4099NAV2054NAV4098NAV2058NAV3244 | True
9483 | 704650486 | XSE - Instrument Verification Lost | 2021-10-07T10:11:02.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Instrument Verification Lost 2021-Oct-06 Salesforce C00235913 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that navigation instruments were verified during the setup process for a Scan & Plan procedure. When going into the operate process after the O-arm spin and registration, the instruments needed to be verified again. The verification was lost. There was no patient harm and the procedure was delayed less than an hour. 2021-10-07 email 2021-10-07 11:16 (Rep): No new information. 2021-10-18 email 2021-10-18 02:15 (Rep): The cause of the issue was not determined. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Oct-06 Salesforce C00235913 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that navigation instruments were verified during the setup process for a Scan & Plan procedure. When going into the operate process after the O-arm spin and registration, the instruments needed to be verified again. The verification was lost. There was no patient harm and the procedure was delayed less than an hour. 2021-10-07 email 2021-10-07 11:16 (Rep): No new information. 2021-10-18 email 2021-10-18 02:15 (Rep): The cause of the issue was not determined. | xse instrument verification lost salesforce spinal procedure navigation instruments verified setup process scan plan procedure going operate process oarm spin registration instruments needed verified verification lost procedure delayed hour email new email cause issue determined mazorx sw installation kit | NAV3246 | NAV2053 | NAV3235 | NAV2056 | NAV3248 | NAV3246NAV2053NAV3235NAV2056NAV3248 | True
2448 | 703236073 | X - Surgical Arm Issue | 2019-07-01T14:42:20.000000+0000 | MazorX | NAV4097 | ARM JOINT LOOSE | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical Arm Issue The (metal piece) cable duct P/N MEC6270-01 that the end effector attaches to is loose. The issues is not with the attachment of the end effector itself, but an issue with the metal piece which directly attaches to. Video and additional explanation attached to this email. SURGICAL ARM | The (metal piece) cable duct P/N MEC6270-01 that the end effector attaches to is loose. The issues is not with the attachment of the end effector itself, but an issue with the metal piece which directly attaches to. Video and additional explanation attached to this email. | x surgical arm issue cable duct pn end effector attaches loose issues attachment end effector issue metal piece directly attaches video additional explanation attached email surgical arm | NAV3249 | NAV4097 | NAV3299 | NAV3246 | NAV2054 | NAV3249NAV4097NAV3299NAV3246NAV2054 | True
10135 | 704806393 | XSE - SRA Failed BIST/Error Message | 2022-01-12T16:47:00.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 50 | HRS0350-01S | KIT SVC HRS0350-01S NAV CONNECT CABLE | XSE - SRA Failed BIST/Error Message 2022-Jan-12 Salesforce C00241541, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that while setting up for a lab the system failed it's BIST test. They got an error code stating something was out of calibration. They rebooted and BIST test was fine. Then they got new error code: Error 10273 No Image received from 3Define Camera. It was also reported that the cable connection between the guidance system unit and the navigation camera was not working. The manufacturer representative confirmed this by swapping the cable from another guidance system and the problem followed the cable. No patient was present. 2022-01-13 IUD (Other): No new information. 2022-01-21 IUD x2 (Other): No new information. 2022-01-24 IUD x2 (Other): No new information. 2022-01-27 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. 2022-02-22 IUD x2 (Other): No new information 2022-03-02 02740443 (Other): No new information. 2022-03-03 IUD x3 (Other): No new information. 2022-03-28 IUD (Other): No new information. KIT SVC HRS0350-01S NAV CONNECT CABLE | 2022-Jan-12 Salesforce C00241541, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that while setting up for a lab the system failed it's BIST test. They got an error code stating something was out of calibration. They rebooted and BIST test was fine. Then they got new error code: Error 10273 No Image received from 3Define Camera. It was also reported that the cable connection between the guidance system unit and the navigation camera was not working. The manufacturer representative confirmed this by swapping the cable from another guidance system and the problem followed the cable. No patient was present. 2022-01-13 IUD (Other): No new information. 2022-01-21 IUD x2 (Other): No new information. 2022-01-24 IUD x2 (Other): No new information. 2022-01-27 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. 2022-02-22 IUD x2 (Other): No new information 2022-03-02 02740443 (Other): No new information. 2022-03-03 IUD x3 (Other): No new information. 2022-03-28 IUD (Other): No new information. | xse sra failed bisterror message salesforce iud outside proc edure setting lab failed bist test got error code stating calibration rebooted bist test fine got new error code error image camera cable connection unit navigation camera working confirmed swapping cable problem followed cable present iud new iud new iud new iud new iud new iud new new iud new iud new kit svc nav connect cable | NAV3256 | NAV2047 | NAV3249 | NAV3244 | NAV1006 | NAV3256NAV2047NAV3249NAV3244NAV1006 | True
1610 | 703207910 | X- Workstation Issues | 2019-06-14T16:44:54.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- Workstation Issues Plan was done by Dr. Francis the morning of the surgery. L5-S1 were planned for an open procedure. There were no concerns discussed or changes made to the plan. A single schanz pin was utilized and placed in the right PSIS. there were no complications mounting the platform. Although the pictures were poor quality, registration was achieved on the second attempt. the 3d marker had to be adjusted and a new image was necessary for clarity. There were no complications throughout he procedure. all wires and screws were placed accurately andconfirmed with x-ray. after the procedure when removing the system from the bed, there was some difficulty. after attaching the manipulator, unlocking the air pressure and hitting the unlock button, there was no air pressure assisting with removing from the bed. all of the weight was present when removing from the bed. after the patient had left the room, I re attached the robot to the bed, again with no air pressure assisting. after multiple trouble shooting it was noted that thehood was stuck in between the open and closed state due to friction against the side of the workstation. the door where the image adapter lives seemed to be loose and sat in a position that rubbed against the hood when opening and closing. because of this, the hood was not fully open, causing there to be no air pressure distributed to release the robot from the bed. the hood was adjusted and reset to the full open position and this was recognized by the green dot lit up under the sensors. the robot was then removed. the right door on the workstation still remains in an awkward position but doesn't seem to cause any friction anymore when opening and closing the hood. UNK MAZOR X WORKSTATION | Plan was done by Dr. Francis the morning of the surgery. L5-S1 were planned for an open procedure. There were no concerns discussed or changes made to the plan. A single schanz pin was utilized and placed in the right PSIS. there were no complications mounting the platform. Although the pictures were poor quality, registration was achieved on the second attempt. the 3d marker had to be adjusted and a new image was necessary for clarity. There were no complications throughout he procedure. all wires and screws were placed accurately andconfirmed with x-ray. after the procedure when removing the system from the bed, there was some difficulty. after attaching the manipulator, unlocking the air pressure and hitting the unlock button, there was no air pressure assisting with removing from the bed. all of the weight was present when removing from the bed. after the patient had left the room, I re attached the robot to the bed, again with no air pressure assisting. after multiple trouble shooting it was noted that thehood was stuck in between the open and closed state due to friction against the side of the workstation. the door where the image adapter lives seemed to be loose and sat in a position that rubbed against the hood when opening and closing. because of this, the hood was not fully open, causing there to be no air pressure distributed to release the robot from the bed. the hood was adjusted and reset to the full open position and this was recognized by the green dot lit up under the sensors. the robot was then removed. the right door on the workstation still remains in an awkward position but doesn't seem to cause any friction anymore when opening and closing the hood. | x workstation issues plan dr francis morning surgery planned open procedure concerns discussed changes plan single schanz pin utilized placed right psis complications mounting platform pictures poor quality registration achieved second attempt marker adjusted new image necessary clarity complications procedure wires screws placed accurately andconfirmed xray procedure removing bed difficulty attaching manipulator unlocking air pressure hitting unlock button air pressure assisting removing bed weight present removing bed left room attached robot bed air pressure assisting multiple trouble shooting noted thehood stuck open closed state friction workstation door image adapter lives loose sat position rubbed hood opening closing hood fully open causing air pressure distributed release robot bed hood adjusted reset open position recognized green dot lit sensors robot removed right door workstation remains awkward position doesnt cause friction anymore opening closing hood unk mazor x workstation | NAV3264 | NAV2058 | NAV3250 | NAV3243 | NAV3263 | NAV3264NAV2058NAV3250NAV3243NAV3263 | True
5748 | 703793729 | XSE - Bed Frame Adapter Damaged | 2020-06-09T21:14:00.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | MAS1574-01S | KIT SVC MAS1574-01S BED FRAME ADAPTOR | XSE - Bed Frame Adapter Damaged 2020-06-09 00690421 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the bed frame adapter is damaged a needs to be replaced for Canbus errors. No patient was present. 2020-06-11 IUD x2 (Other): No new information. 2020-06-19 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-06-30 IUD (Other): No new information. 2020-07-13 IUD x2 (Other): No new information. KIT SVC MAS1574-01S BED FRAME ADAPTOR | 2020-06-09 00690421 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the bed frame adapter is damaged a needs to be replaced for Canbus errors. No patient was present. 2020-06-11 IUD x2 (Other): No new information. 2020-06-19 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-06-30 IUD (Other): No new information. 2020-07-13 IUD x2 (Other): No new information. | xse bed frame adapter damaged outside procedure bed frame adapter damaged needs replaced canbus errors present iud new iud new iud new iud new iud new kit svc bed frame adaptor | NAV3255 | NAV2047 | NAV3295 | NAV1006 | NAV2054 | NAV3255NAV2047NAV3295NAV1006NAV2054 | True
9377 | 704621632 | XSE - Schanz Bond Screw Damaged | 2021-09-21T14:00:59.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2271-04S | KIT SVC MAS2271-04S SCHANZ BOND | XSE - Schanz Bond Screw Damaged 2021-Sep-21 Salesforce C00234731 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the schanz bond screw head was damaged and it was not possible to remove the pin. There was no delay and no impact on the patient outcome. 2021-10-07 email 2021-10-07 03:16 (Rep): The cause or potential contributing factors of the damage was stated to be potentially due to the material of the screw head being too soft which gets damaged quickly. 2021-10-22 IUD (Other): No new information. 2021-11-01 IUD (Other): No new information. KIT SVC MAS2271-04S SCHANZ BOND | 2021-Sep-21 Salesforce C00234731 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the schanz bond screw head was damaged and it was not possible to remove the pin. There was no delay and no impact on the patient outcome. 2021-10-07 email 2021-10-07 03:16 (Rep): The cause or potential contributing factors of the damage was stated to be potentially due to the material of the screw head being too soft which gets damaged quickly. 2021-10-22 IUD (Other): No new information. 2021-11-01 IUD (Other): No new information. | xse schanz bond screw damaged salesforce spinal procedure schanz bond screw head damaged possible remove pin delay impact outcome email cause potential contributing factors damage stated potentially material screw head soft gets damaged quickly iud new iud new kit svc schanz bond | NAV3299 | NAV4090 | NAV3295 | NAV2053 | NAV3244 | NAV3299NAV4090NAV3295NAV2053NAV3244 | True
5696 | 703783223 | XSE - Snap Shot Tracker Error | 2020-06-03T14:21:04.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT0898 | TRACKER KIT0898 MAZOR X SNAPSHOT TRACKER | XSE - Snap Shot Tracker Error 2020-06-03 00690025 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that planning was reviewed by the surgeon prior to the case. Special attention was paid to the axial angles and docking surfaces due to the open cortical approach. A bone mount bridge was chosen based on clinical indication and sugeon preference. The bridge was connected to a Schanz pin in the left PSIS to mount the surgical system to the patient. Registration was completed with green values and the surgeon verified no shifts between CT and fluoro. A 3Define scan was completed successfully. There were issues with the snap shot tracker during the procedure. The "Acquire" button was not highlighted and could not be clicked. A soft restart of the system was done to resolve the issue. The 3Define scan then failed twice so a generic work volume was used. The procedure was completed without any other issues. There was no patient harm and the procedure was delay ed less than an hour. 2020-06-04 Email (Rep): Additional information was received from the manufacturer representative stating that this occurred after the 3Define scan. The system would not allow the manufacturer representative to click acquire even though both the reference frame and the snapshot tracker were both being seen. A soft restart and another 3Define scan was needed to resolve the issue. They were told this is a known software bug. 2020-06-12 IUD x2 (Other): No new information. 2020-06-19 IUD (Other): No new information. TRACKER KIT0898 MAZOR X SNAPSHOT TRACKER | 2020-06-03 00690025 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that planning was reviewed by the surgeon prior to the case. Special attention was paid to the axial angles and docking surfaces due to the open cortical approach. A bone mount bridge was chosen based on clinical indication and sugeon preference. The bridge was connected to a Schanz pin in the left PSIS to mount the surgical system to the patient. Registration was completed with green values and the surgeon verified no shifts between CT and fluoro. A 3Define scan was completed successfully. There were issues with the snap shot tracker during the procedure. The "Acquire" button was not highlighted and could not be clicked. A soft restart of the system was done to resolve the issue. The 3Define scan then failed twice so a generic work volume was used. The procedure was completed without any other issues. There was no patient harm and the procedure was delay ed less than an hour. 2020-06-04 Email (Rep): Additional information was received from the manufacturer representative stating that this occurred after the 3Define scan. The system would not allow the manufacturer representative to click acquire even though both the reference frame and the snapshot tracker were both being seen. A soft restart and another 3Define scan was needed to resolve the issue. They were told this is a known software bug. 2020-06-12 IUD x2 (Other): No new information. 2020-06-19 IUD (Other): No new information. | xse snap shot tracker error spinal procedure treat degenerative discs fixation planning reviewed surgeon prior case special attention paid axial angles docking surfaces open cortical approach bone mount bridge chosen based clinical indication sugeon preference bridge connected schanz pin left psis mount surgical registration completed green values surgeon verified shifts ct fluoro scan completed successfully issues snap shot tracker procedure acquire button highlighted clicked soft restart resolve issue scan failed twice generic work volume procedure completed issues procedure delay ed hour email additional stating occurred scan allow click acquire reference frame snapshot tracker seen soft restart scan needed resolve issue told known software bug iud new iud new tracker mazor x snapshot tracker | NAV5000 | NAV3246 | NAV3242 | NAV3244 | NAV3299 | NAV5000NAV3246NAV3242NAV3244NAV3299 | True
192 | 703032188 | XSE - Deviation/binding | 2019-02-28T20:56:26.000000+0000 | MazorX Stealth Edition | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Deviation/binding 2019-Feb-26 SCL1900470, _e1 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a L2-S2 procedure, the left L2 skived and the wire was placed lateral that was noted on a poor quality A/P image. The surgeon manually re-positioned. The arm audibly alarmed as being off trajectory on at least 5 trajectories upon malleting down the drill guide. They cleared up and resent the trajectory. On one trajectory that possibly was the right L4, this happened 3 times. The HCP was instructed to irrigate the arm guide and instruments and re-iterated co-linear placement of the cannulas along the trajectory. Additional information indicated there were six times during two separate cases where the arm moved out of trajectory when the surgeon malleted instruments and did not happen during the previous 15 cases. The HCP made good sized incisions and used very little force. No prompt would appear and the system made an alert noise and the system would read out of trajectory. It was noted that the case was delayed at least 20 minutes but there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Additional information reported by the rep indicated that the Powerease binding was a major issue. They used the powerease adaptor and it still continued to bind. The HCP struggled with long tools and the Powerease to get the drill bit into the drill guide without it getting stuck. The rep said the HCP was about 5’10 and this was a constant struggle. The software was very slow during segmentation and lagging during registration and moving between trajectories. It was also noted that the wireless tablet was too small to visually see anything during the procedure and they had to move the workstation closer so the HCP could see. They consistently had been getting error messages that the arm is out of trajectory with the upgraded software. SURGICAL ARM | 2019-Feb-26 SCL1900470, _e1 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a L2-S2 procedure, the left L2 skived and the wire was placed lateral that was noted on a poor quality A/P image. The surgeon manually re-positioned. The arm audibly alarmed as being off trajectory on at least 5 trajectories upon malleting down the drill guide. They cleared up and resent the trajectory. On one trajectory that possibly was the right L4, this happened 3 times. The HCP was instructed to irrigate the arm guide and instruments and re-iterated co-linear placement of the cannulas along the trajectory. Additional information indicated there were six times during two separate cases where the arm moved out of trajectory when the surgeon malleted instruments and did not happen during the previous 15 cases. The HCP made good sized incisions and used very little force. No prompt would appear and the system made an alert noise and the system would read out of trajectory. It was noted that the case was delayed at least 20 minutes but there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Additional information reported by the rep indicated that the Powerease binding was a major issue. They used the powerease adaptor and it still continued to bind. The HCP struggled with long tools and the Powerease to get the drill bit into the drill guide without it getting stuck. The rep said the HCP was about 5’10 and this was a constant struggle. The software was very slow during segmentation and lagging during registration and moving between trajectories. It was also noted that the wireless tablet was too small to visually see anything during the procedure and they had to move the workstation closer so the HCP could see. They consistently had been getting error messages that the arm is out of trajectory with the upgraded software. | xse deviationbinding spinal procedure procedure left skived wire placed lateral noted poor quality ap image surgeon manually repositioned arm audibly alarmed trajectory trajectories malleting drill guide cleared resent trajectory trajectory possibly right happened times hcp instructed irrigate arm guide instruments reiterated colinear placement cannulas trajectory additional indicated times separate cases arm moved trajectory surgeon malleted instruments did happen previous cases hcp good sized incisions little force prompt appear alert noise read trajectory noted case delayed minutes additional complications anticipated additional indicated powerease binding major issue powerease adaptor continued bind hcp struggled long tools powerease drill bit drill guide getting stuck said hcp ’ constant struggle software slow segmentation lagging registration moving trajectories noted wireless tablet small visually procedure workstation closer hcp consistently getting error messages arm trajectory upgraded software surgical arm | NAV3299 | NAV1007 | NAV5000 | NAV4101 | NAV3243 | NAV3299NAV1007NAV5000NAV4101NAV3243 | True
2747 | 703249196 | X - Tablet communication issue | 2019-07-09T15:41:22.000000+0000 | MazorX | NAV3253 | TABLET ISSUES | 20 | UNK_MAZORX_TBLT | UNK MAZOR X TABLET | X - Tablet communication issue When tablet 15 was placed on the docking station on the surgical system of system 14 a breaker was tripped causing a loss in communication with the shoulder. When tablet 14 was placed on the docking system on the surgical system of system 15 a breaker was tripped causing a loss in communication with the shoulder UNK MAZOR X TABLET | When tablet 15 was placed on the docking station on the surgical system of system 14 a breaker was tripped causing a loss in communication with the shoulder. When tablet 14 was placed on the docking system on the surgical system of system 15 a breaker was tripped causing a loss in communication with the shoulder | x tablet communication issue tablet placed docking station surgical breaker tripped causing loss communication shoulder tablet placed docking surgical breaker tripped causing loss communication shoulder unk mazor x tablet | NAV3253 | NAV2047 | NAV4107 | NAV3259 | NAV4108 | NAV3253NAV2047NAV4107NAV3259NAV4108 | True
3213 | 703285217 | XSE - Surgeon Screen Goes Black | 2019-07-30T14:38:15.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - Surgeon Screen Goes Black 2019-07-23 SCL1902204 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon screen continued to go black and was unable to be used during the case. There was no patient harm reported. 2019-10-07 email 2019-10-07 11:30 (Rep): Additional information received from a manufacturer representative reported the issue was resolved. 2020-01-16 email 2020-01-16 09:30 (Rep): No new information. SURGEON SCREEN | 2019-07-23 SCL1902204 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon screen continued to go black and was unable to be used during the case. There was no patient harm reported. 2019-10-07 email 2019-10-07 11:30 (Rep): Additional information received from a manufacturer representative reported the issue was resolved. 2020-01-16 email 2020-01-16 09:30 (Rep): No new information. | xse surgeon screen goes black spinal procedure surgeon screen continued black unable case email additional issue resolved email new surgeon screen | NAV3253 | NAV2054 | NAV4112 | NAV4101 | NAV2047 | NAV3253NAV2054NAV4112NAV4101NAV2047 | True
7208 | 704142462 | XSE ¦ Target Extender replacement | 2020-12-24T01:19:47.000000+0000 | MazorX Stealth Edition | NAV2044 | ACCESSORY KIT ISSUES | 20 | MAS2310-04S | KIT SVC MAS2310-04S TARGET EXTENDER | XSE ¦ Target Extender replacement 2020-Dec-23 00703319 (rep): Medtronic received information from a manufacturer representative (rep) regarding a device being used outside a procedure. It was reported that the site needs a quote for a new target extender. There was no patient harm or reported as a result of the event. 2020-Dec-29 IUDx2 (other): No new information. KIT SVC MAS2310-04S TARGET EXTENDER | 2020-Dec-23 00703319 (rep): Medtronic received information from a manufacturer representative (rep) regarding a device being used outside a procedure. It was reported that the site needs a quote for a new target extender. There was no patient harm or reported as a result of the event. 2020-Dec-29 IUDx2 (other): No new information. | xse target extender replacement outside procedure site needs quote new target extender new kit svc target extender | NAV3299 | NAV3295 | NAV2054 | NAV2053 | NAV2047 | NAV3299NAV3295NAV2054NAV2053NAV2047 | False
5504 | 703740845 | XSE - Missing Instruments | 2020-05-06T19:44:43.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 50 | MAS1355-05S | KIT SVC MAS1355-05S DUAL CLAMP | XSE - Missing Instruments 2020-05-06 00688410 (Rep): Medtronic received information regarding a guidance system outside of procedure. It was reported that the following instruments were damaged; Bone Mount Bridge, Reference Marker, Arm Guide and Dual Clamp. No patient was present. 2020-05-07 email 2020-05-07 13:00, IUD x2 (Rep): Additional information received from a manufacturer representative reported that the instruments were not damaged, but they were accidentally thrown away by the facility. 2020-05-26 IUD (Other): No new information. KIT SVC MAS1355-05S DUAL CLAMP | 2020-05-06 00688410 (Rep): Medtronic received information regarding a guidance system outside of procedure. It was reported that the following instruments were damaged; Bone Mount Bridge, Reference Marker, Arm Guide and Dual Clamp. No patient was present. 2020-05-07 email 2020-05-07 13:00, IUD x2 (Rep): Additional information received from a manufacturer representative reported that the instruments were not damaged, but they were accidentally thrown away by the facility. 2020-05-26 IUD (Other): No new information. | xse missing instruments outside procedure following instruments damaged bone mount bridge reference marker arm guide dual clamp present email iud additional instruments damaged accidentally thrown away facility iud new kit svc dual clamp | NAV2053 | NAV3299 | NAV3295 | NAV3256 | NAV2054 | NAV2053NAV3299NAV3295NAV3256NAV2054 | True
3190 | 703271181 | XSE - registration issues | 2019-07-22T18:10:47.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - registration issues 2019-Jul-19 SCL1902145 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they were doing a T2-T9 procedure with multiple fractures. The plan was screws at T2-T4 with fractures at T5-T7 and screws at T7(only right side screw due to fracture) -T9. They initially attempted to use clamp on T4 for upper screws, but due to stability, the surgeon chose to use the misty bridge with schanz screws. They were not able to initially get registration due to difficult anatomy so they moved to the lower screws. They were able to get registration and successfully placed screws. They then moved back up to the superior screws and were able to get segmentation but, due to the angulation, were not able to have reachable trajectories. They tried moving the system up higher so that it could reach but with the new pos itioning, in addition to the difficult anatomy, it was not able to get the oblique angle for registration and the surgeon chose to free hand final screws. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jul-19 SCL1902145 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they were doing a T2-T9 procedure with multiple fractures. The plan was screws at T2-T4 with fractures at T5-T7 and screws at T7(only right side screw due to fracture) -T9. They initially attempted to use clamp on T4 for upper screws, but due to stability, the surgeon chose to use the misty bridge with schanz screws. They were not able to initially get registration due to difficult anatomy so they moved to the lower screws. They were able to get registration and successfully placed screws. They then moved back up to the superior screws and were able to get segmentation but, due to the angulation, were not able to have reachable trajectories. They tried moving the system up higher so that it could reach but with the new pos itioning, in addition to the difficult anatomy, it was not able to get the oblique angle for registration and the surgeon chose to free hand final screws. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse registration issues spinal procedure doing procedure multiple fractures plan screws fractures screws initially attempted use clamp upper screws stability surgeon chose use misty bridge schanz screws able initially registration difficult anatomy moved lower screws able registration successfully placed screws moved superior screws able segmentation angulation able reachable trajectories tried moving higher reach new pos itioning addition difficult anatomy able oblique angle registration surgeon chose free hand final screws additional complications anticipated mazor x software installation kit | NAV3243 | NAV3242 | NAV4112 | NAV5000 | NAV3299 | NAV3243NAV3242NAV4112NAV5000NAV3299 | True
163 | 703021733 | X - Loose Connector | 2019-02-22T15:15:35.000000+0000 | MazorX | NAV4109 | LOOSE | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Loose Connector 2019-02-12 SCL1900394, Image 2019-02-22 09:18 (Other): Medtronic received information regarding a guidance system. It was reported that a Node 7 error was displayed and the connector to the X-eye motor was loose during install. There was no patient involvement. MAZOR X SYSTEM TPL0059 | 2019-02-12 SCL1900394, Image 2019-02-22 09:18 (Other): Medtronic received information regarding a guidance system. It was reported that a Node 7 error was displayed and the connector to the X-eye motor was loose during install. There was no patient involvement. | x loose connector image node error displayed connector xeye motor loose install mazor x | NAV4109 | NAV2054 | NAV4091 | NAV3242 | NAV3246 | NAV4109NAV2054NAV4091NAV3242NAV3246 | True
1848 | 703218839 | X - Tablet issue | 2019-06-20T21:58:17.000000+0000 | MazorX | NAV3253 | TABLET ISSUES | 20 | ASM0337-01 | SURGEON&APOS;S SCREEN | X - Tablet issue 2017-11-14 SC17001565(Rep): Medtronic received information regarding a guidance system being used during preventative maintenance. It was reported that the tablet was broken. No patient involvement. SURGEON&APOS;S SCREEN | 2017-11-14 SC17001565(Rep): Medtronic received information regarding a guidance system being used during preventative maintenance. It was reported that the tablet was broken. No patient involvement. | x tablet issue preventative maintenance tablet broken surgeonaposs screen | NAV3253 | NAV3295 | NAV3249 | NAV2055 | NAV2054 | NAV3253NAV3295NAV3249NAV2055NAV2054 | True
8707 | 704466023 | XSE - Deviation/SRA Hit BMB | 2021-06-23T23:55:03.000000+0000 | MazorX Stealth Edition | NAV1012 | SYSTEM REPORTED | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation/SRA Hit BMB 2021-Jun-22 Salesforce C00228986, C00228987, IUD x14, email 2021-06-24 11:01 x2, e1, e2, e3, e4 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a deviation during an open T10 to pelvis procedure. When moving the surgical arm to a different trajectory during the case, the arm hit the bone mount bridge. The surgeon did an accuracy test and everything was accurate. A new snapshot was done as a precaution and navigation was accurate. A 3Define scan was completed to define the work volume during the procedure. The left L1 screw was noted to be a little medial, but the surgeon thought the position was ok. After the procedure, the patient woke up with left side weakness and no movement in their left leg. The surgeon indicated there was significant neurological injury. A CT scan was done and the T10 - L1 screws on the left side were all medial. The manufacturer representative noted that the deviations were not seen on intra op imaging. The amount of deviat ion was unknown. All of the right side screws and L2-pelvis on the left side were placed accurately. The surgeon did not feel the system was inaccurate during the case. The surgeon was standing on the left side of the patient during the procedure. The cause of the deviation was not determined. A revision to reposition the screws was planned. There procedure was delayed less than an hour. The guidance system was tested and the surgical arm failed the 3 and 10 point accuracy tests. 2021-07-14 IUD (Other): No new information. 2021-07-21 IUD x5 (Other): No new information. 2021-09-01 IUD x3 (Other): No new information. 2021-09-16 e5 (Other): No new information. 2021-10-06 e6 (Rep): No new information. 2021-10-07 IUD (Other): No new information. 2021-10-14 e7 (Rep): Additional information received from a manufacturer representative reported that the surgical arm hit the bone mount bridge during registration. An AP image was taken and the bridge was hit when moving the surgical arm to the oblique position. The 3D marker was removed from the arm guide prior to the surgical arm being sent to the oblique trajectory. A navigation accuracy test was done after the anatomical accuracy test. The passive planar probe was used and placed on the patient's anatomy and then the arm guide. Additional anatomical checks were done throughout the case since an instrument was shown to be slightly medial during the procedure. Navigation accuracy and guidance system accuracy were discussed with the surgeon. The surgeon decided to do additional snapshots and accuracy tests were performed afterwards to avoid re-registering the patient. The additional anatomical accuracy checks were accurate. The revision procedure was completed. After the procedure, the patient still had weakness and immobility on their left side. 2021-10-15 IUD x2 (Other): No new information. 2021-11-05 IUD (Other): No new information. 2021-11-11 IUD x2 (Other): No new information. 2021-11-23 IUD x2 (Other): No new information. 2021-12-13 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. 2022-01-27 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-Jun-22 Salesforce C00228986, C00228987, IUD x14, email 2021-06-24 11:01 x2, e1, e2, e3, e4 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a deviation during an open T10 to pelvis procedure. When moving the surgical arm to a different trajectory during the case, the arm hit the bone mount bridge. The surgeon did an accuracy test and everything was accurate. A new snapshot was done as a precaution and navigation was accurate. A 3Define scan was completed to define the work volume during the procedure. The left L1 screw was noted to be a little medial, but the surgeon thought the position was ok. After the procedure, the patient woke up with left side weakness and no movement in their left leg. The surgeon indicated there was significant neurological injury. A CT scan was done and the T10 - L1 screws on the left side were all medial. The manufacturer representative noted that the deviations were not seen on intra op imaging. The amount of deviat ion was unknown. All of the right side screws and L2-pelvis on the left side were placed accurately. The surgeon did not feel the system was inaccurate during the case. The surgeon was standing on the left side of the patient during the procedure. The cause of the deviation was not determined. A revision to reposition the screws was planned. There procedure was delayed less than an hour. The guidance system was tested and the surgical arm failed the 3 and 10 point accuracy tests. 2021-07-14 IUD (Other): No new information. 2021-07-21 IUD x5 (Other): No new information. 2021-09-01 IUD x3 (Other): No new information. 2021-09-16 e5 (Other): No new information. 2021-10-06 e6 (Rep): No new information. 2021-10-07 IUD (Other): No new information. 2021-10-14 e7 (Rep): Additional information received from a manufacturer representative reported that the surgical arm hit the bone mount bridge during registration. An AP image was taken and the bridge was hit when moving the surgical arm to the oblique position. The 3D marker was removed from the arm guide prior to the surgical arm being sent to the oblique trajectory. A navigation accuracy test was done after the anatomical accuracy test. The passive planar probe was used and placed on the patient's anatomy and then the arm guide. Additional anatomical checks were done throughout the case since an instrument was shown to be slightly medial during the procedure. Navigation accuracy and guidance system accuracy were discussed with the surgeon. The surgeon decided to do additional snapshots and accuracy tests were performed afterwards to avoid re-registering the patient. The additional anatomical accuracy checks were accurate. The revision procedure was completed. After the procedure, the patient still had weakness and immobility on their left side. 2021-10-15 IUD x2 (Other): No new information. 2021-11-05 IUD (Other): No new information. 2021-11-11 IUD x2 (Other): No new information. 2021-11-23 IUD x2 (Other): No new information. 2021-12-13 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. 2022-01-27 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. | xse deviationsra hit bmb salesforce iud email spinal procedure deviation open pelvis procedure moving surgical arm different trajectory case arm hit bone mount bridge surgeon did accuracy test accurate new snapshot precaution navigation accurate scan completed define work volume procedure left screw noted little medial surgeon thought position ok procedure woke left weakness movement left leg surgeon indicated significant neurological injury ct scan screws left medial noted deviations seen intra op imaging deviat ion unknown right screws left placed accurately surgeon did feel inaccurate case surgeon standing left procedure cause deviation determined revision reposition screws planned procedure delayed hour tested surgical arm failed point accuracy tests iud new iud new iud new new new iud new additional surgical arm hit bone mount bridge registration ap image taken bridge hit moving surgical arm oblique position marker removed arm guide prior surgical arm sent oblique trajectory navigation accuracy test anatomical accuracy test passive planar probe placed patients anatomy arm guide additional anatomical checks case instrument shown slightly medial procedure navigation accuracy accuracy discussed surgeon surgeon decided additional snapshots accuracy tests performed avoid reregistering additional anatomical accuracy checks accurate revision procedure completed procedure weakness immobility left iud new iud new iud new iud new iud new iud new iud new iud new mazor x | NAV5000 | NAV1012 | NAV4096 | NAV4093 | NAV5001 | NAV5000NAV1012NAV4096NAV4093NAV5001 | True
7811 | 704280748 | XSE - Intermittent BNC Connection | 2021-03-13T01:17:44.000000+0000 | MazorX Stealth Edition | NAV3237 | FLUORO IMAGES | 20 | KIT1218-01 | SVC KIT1218-01 CBL BNC VIDEO | XSE - Intermittent BNC Connection 2021-Mar-12 Salesforce C00222631 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a procedure they were having difficulties transferring images from their C-Arm. They tried multiple cables and ended up have to contort/"wiggle" the cable near the BNC port to get it to send over. There was no patient harm reported and the case was delayed by less than one hour. 2021-03-17 IUD x3 (Other): No new information. 2021-03-31 IUD x3 (Other): No new information. 2021-04-29 IUD (Other): No new information. 2021-06-02 IUD (Other): No new information. 2021-06-14 IUD (Other): No new information. 2022-01-05 IUD (Other): No new information. SVC KIT1218-01 CBL BNC VIDEO | 2021-Mar-12 Salesforce C00222631 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a procedure they were having difficulties transferring images from their C-Arm. They tried multiple cables and ended up have to contort/"wiggle" the cable near the BNC port to get it to send over. There was no patient harm reported and the case was delayed by less than one hour. 2021-03-17 IUD x3 (Other): No new information. 2021-03-31 IUD x3 (Other): No new information. 2021-04-29 IUD (Other): No new information. 2021-06-02 IUD (Other): No new information. 2021-06-14 IUD (Other): No new information. 2022-01-05 IUD (Other): No new information. | xse intermittent bnc connection salesforce spinal p rocedure procedure having difficulties transferring images carm tried multiple cables ended contortwiggle cable near bnc port send case delayed hour iud new iud new iud new iud new iud new iud new svc cbl bnc video | NAV3256 | NAV3237 | NAV2047 | NAV2058 | NAV3253 | NAV3256NAV3237NAV2047NAV2058NAV3253 | True
7858 | 704293422 | XSE - SRA Joint 6 Overcurrent | 2021-03-19T16:02:01.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Joint 6 Overcurrent 2021-Mar-19 Salesforce C00222945 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when attempting to send the surgical arm to the trajectories for the advanced accuracy test, an error message stating the arm failed to reach destination and joint 6 overcurrent was displayed. There was no patient involvement. 2021-03-24 IUD x2 (Other): No new information. 2021-05-18 email 2021-05-18 15:46 (Rep): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Mar-19 Salesforce C00222945 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when attempting to send the surgical arm to the trajectories for the advanced accuracy test, an error message stating the arm failed to reach destination and joint 6 overcurrent was displayed. There was no patient involvement. 2021-03-24 IUD x2 (Other): No new information. 2021-05-18 email 2021-05-18 15:46 (Rep): No new information. | xse sra joint overcurrent salesforce outside procedure attempting send surgical arm trajectories advanced accuracy test error message stating arm failed reach destination joint overcurrent displayed iud new email new kit svc surgical arm refurb | NAV4095 | NAV3249 | NAV4093 | NAV4092 | NAV4098 | NAV4095NAV3249NAV4093NAV4092NAV4098 | True
3693 | 703416696 | X - Surgical Arm Failed Accuracy Test | 2019-10-16T19:55:25.000000+0000 | MazorX | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical Arm Failed Accuracy Test 2019-04-25 SCL1901029 (REP): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the arm was hit by a door being opened as it was being carried into the site. It then failed accuracy test. No patient was present. SURGICAL ARM | 2019-04-25 SCL1901029 (REP): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the arm was hit by a door being opened as it was being carried into the site. It then failed accuracy test. No patient was present. | x surgical arm failed accuracy test outside procedure arm hit door opened carried site failed accuracy test present surgical arm | NAV3249 | NAV4092 | NAV4093 | NAV2054 | NAV4096 | NAV3249NAV4092NAV4093NAV2054NAV4096 | True
8390 | 704405834 | XSE - SRA E-Stop Error/SRA Froze | 2021-05-20T15:54:50.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA E-Stop Error/SRA Froze 2021-May-20 Salesforce C00226749, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgical arm froze while doing the built in self test (BIST). The manufacturer representative restarted the system, but when the system was powered back on an error message was displayed indicating the emergency stop button was activated when it was not. The representative attempted to activate and de-activate the emergency stop button, but there was no change. There was no patient involvement. 2021-05-26 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-May-20 Salesforce C00226749, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgical arm froze while doing the built in self test (BIST). The manufacturer representative restarted the system, but when the system was powered back on an error message was displayed indicating the emergency stop button was activated when it was not. The representative attempted to activate and de-activate the emergency stop button, but there was no change. There was no patient involvement. 2021-05-26 IUD (Other): No new information. | xse sra estop errorsra froze salesforce iud outside proc edure surgical arm froze doing built self test restarted powered error message displayed indicating emergency stop button activated attempted activate deactivate emergency stop button change iud new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV1006 | NAV2047 | NAV4099 | NAV3249NAV4095NAV1006NAV2047NAV4099 | True
7679 | 704247623 | XSE - Medial Screw Placement | 2021-02-24T21:08:37.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Medial Screw Placement 2021-02-24 00707210 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the screw was placed medial during the case. The deviation occurred after taking AP images. The surgeon was performing a lateral retraction and pushing too much which caused the implant to be medial to the pedicle. There was no patient harm and the procedure was delayed less than an hour. 2021-02-26 email 2021-02-26 05:46 (Rep): Additional information was received stating that the screw was deviated between 3.5 and 10mm. 2021-03-03 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-02-24 00707210 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the screw was placed medial during the case. The deviation occurred after taking AP images. The surgeon was performing a lateral retraction and pushing too much which caused the implant to be medial to the pedicle. There was no patient harm and the procedure was delayed less than an hour. 2021-02-26 email 2021-02-26 05:46 (Rep): Additional information was received stating that the screw was deviated between 3.5 and 10mm. 2021-03-03 IUD (Other): No new information. | xse medial screw placement spinal procedure screw placed medial case deviation occurred taking ap images surgeon performing lateral retraction pushing caused implant medial pedicle procedure delayed hour email additional stating screw deviated iud new mazor x | NAV5000 | NAV3242 | NAV3244 | NAV1012 | NAV2047 | NAV5000NAV3242NAV3244NAV1012NAV2047 | True
8773 | 704479859 | XSE - SRA Failed Stress Test | 2021-07-01T11:59:49.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Failed Stress Test 2021-Jun-28 Salesforce C00229385 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed a stress test during installation. There was no patient involvement. 2021-07-17 IUD x2 (Other): Troubleshooting involved completing two field calibrations and four stress tests, the R and D team reviewed logs and deemed system suitable for installation. The installation was completed. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Jun-28 Salesforce C00229385 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed a stress test during installation. There was no patient involvement. 2021-07-17 IUD x2 (Other): Troubleshooting involved completing two field calibrations and four stress tests, the R and D team reviewed logs and deemed system suitable for installation. The installation was completed. | xse sra failed stress test salesforce outside procedure surgical arm failed stress test installation iud troubleshooting involved completing field calibrations stress tests r d team reviewed logs deemed suitable installation installation completed kit svc surgical arm refurb | NAV3249 | NAV2054 | NAV4092 | NAV4096 | NAV2056 | NAV3249NAV2054NAV4092NAV4096NAV2056 | True
10207 | 704821224 | XSE - OOBF/SRAs Failed BIST Test | 2022-01-21T04:54:17.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - OOBF/SRAs Failed BIST Test 2022-Jan-20 Salesforce C00242021 (Rep, for): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the J5 axis of the surgical arm failed during a built in self test (BIST). There was no patient involvement. 2022-01-21 e1, e2, e3, e4 ,e5 (Rep, For): Additional information received from a manufacturer representative reported that the the s urgical arm displayed a Drive 4 overspeed error and the BIST test failed due to joint 5 errors. The test sometimes passed after the software and firmware was updated. A second arm failed out of box. The BIST test failed on joint 3 and a Drive 6 over current error was displayed during a stress test. A field motor calibration did not resolve the issue. When tested on an older version of software, a stress test could not be passed and an error 3016 was displayed. 2022-01-23 IUD (Other, For): No new information. 2022-01-25 IUD (Other, For): No new information. 2022-01-28 Update (Other, For): No new information. 2022-04-04 e6, e7, e8, e9, e10 (Rep, For): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2022-Jan-20 Salesforce C00242021 (Rep, for): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the J5 axis of the surgical arm failed during a built in self test (BIST). There was no patient involvement. 2022-01-21 e1, e2, e3, e4 ,e5 (Rep, For): Additional information received from a manufacturer representative reported that the the s urgical arm displayed a Drive 4 overspeed error and the BIST test failed due to joint 5 errors. The test sometimes passed after the software and firmware was updated. A second arm failed out of box. The BIST test failed on joint 3 and a Drive 6 over current error was displayed during a stress test. A field motor calibration did not resolve the issue. When tested on an older version of software, a stress test could not be passed and an error 3016 was displayed. 2022-01-23 IUD (Other, For): No new information. 2022-01-25 IUD (Other, For): No new information. 2022-01-28 Update (Other, For): No new information. 2022-04-04 e6, e7, e8, e9, e10 (Rep, For): No new information. | xse oobfsras failed bist test salesforce outside proc edure axis surgical arm failed built self test additional s urgical arm displayed drive overspeed error bist test failed joint errors test passed software firmware updated second arm failed box bist test failed joint drive current error displayed stress test field motor calibration did resolve issue tested older version software stress test passed error displayed iud new iud new update new new kit svc surgical arm | NAV3249 | NAV2054 | NAV4106 | NAV3246 | NAV4098 | NAV3249NAV2054NAV4106NAV3246NAV4098 | True
940 | 703181374 | X-Deviation & Accessory Issues | 2019-05-30T19:08:50.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | KIT0737-02 | SHORT DRILL GUIDE AND SHORT CANNULA | X-Deviation & Accessory Issues 2018-10-15 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and the surgical arm was setup. An O-arm spin was completed, and the surgeon planned L3-L5. The surgeon placed screws at L3 and right L4. The surgeon decided not to execute left L4 and L5. Right L3 had a medial skive due to soft tissue pressure, and the k-wire stimulated at nine. The retractors were repositioned and right L3 was resent. The second attempt was placed per the plan and stimulated to 30. It was also noted that the collar of the short drill guide separated from the cannula at the weld. All screws were verified with a lateral image and stimulated to acceptable values. There was no patient harm and the procedure was not delayed over an hour. SHORT DRILL GUIDE AND SHORT CANNULA | 2018-10-15 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and the surgical arm was setup. An O-arm spin was completed, and the surgeon planned L3-L5. The surgeon placed screws at L3 and right L4. The surgeon decided not to execute left L4 and L5. Right L3 had a medial skive due to soft tissue pressure, and the k-wire stimulated at nine. The retractors were repositioned and right L3 was resent. The second attempt was placed per the plan and stimulated to 30. It was also noted that the collar of the short drill guide separated from the cannula at the weld. All screws were verified with a lateral image and stimulated to acceptable values. There was no patient harm and the procedure was not delayed over an hour. | xdeviation accessory issues ccr spinal procedure shoulder calibration checked surgical arm setup oarm spin completed surgeon planned surgeon placed screws right surgeon decided execute left right medial skive soft tissue pressure kwire stimulated retractors repositioned right resent second attempt placed plan stimulated noted collar short drill guide separated cannula weld screws verified lateral image stimulated acceptable values procedure delayed hour short drill guide short cannula | NAV3299 | NAV3295 | NAV4091 | NAV3242 | NAV3244 | NAV3299NAV3295NAV4091NAV3242NAV3244 | True
5830 | 703816534 | XSE - SW Error (No Instrument Missing) | 2020-06-23T15:59:03.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - SW Error (No Instrument Missing) 2020-06-23 00691347, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that after turning the system on during setup for a case, the instruments were all gone when going into the surgeon profiles. The guidance system was rebooted and the manufacturer representative went into a few other profiles, but they were still unable to find the instruments. The system was not used for the case and navigation was used. There was no patient involvement. 2020-06-30 IUD (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-06-23 00691347, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that after turning the system on during setup for a case, the instruments were all gone when going into the surgeon profiles. The guidance system was rebooted and the manufacturer representative went into a few other profiles, but they were still unable to find the instruments. The system was not used for the case and navigation was used. There was no patient involvement. 2020-06-30 IUD (Other): No new information. | xse sw error iud outside procedure turning setup case instruments gone going surgeon profiles rebooted went profiles unable instruments case navigation iud new mazor x software installation kit | NAV3246 | NAV2056 | NAV2047 | NAV3247 | NAV3248 | NAV3246NAV2056NAV2047NAV3247NAV3248 | True
9120 | 704556234 | XSE - Left L4 Deviation | 2021-08-13T13:40:27.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Left L4 Deviation 2021-Aug-13 Salesforce C00232901, email 2021-08-17 10:31, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a deviation during a L3-L4 single position percutaneous case. A Schanz screw was placed in the left PSIS to mount the surgical system to the patient. The first three screws at right L4, right L3 and left L3 were accurately placed. The left L4 screw was deviated. The surgeon tapped the trajectory and that was fine, but the screw had resistance when placed and navigation was slightly lateral to plan (1-2 mm). The drill and tap appeared to be on trajectory for navigation. The trajectory was re-drilled and re-tapped larger, but there was still resistance. Imaging was done and the screw was lateral and superior in the disk space. The screw was removed and the surgeon decided to add an additional level at L5. During the procedure, the surgeon used the drill for the pilot hole and then a separate drill to tap and place the screws. The cause of the deviation was not confirmed, but the representative suspected that the tapped hole was on trajectory and the screw did not follow the hole. There was no patient harm and the procedure was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2021-Aug-13 Salesforce C00232901, email 2021-08-17 10:31, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a deviation during a L3-L4 single position percutaneous case. A Schanz screw was placed in the left PSIS to mount the surgical system to the patient. The first three screws at right L4, right L3 and left L3 were accurately placed. The left L4 screw was deviated. The surgeon tapped the trajectory and that was fine, but the screw had resistance when placed and navigation was slightly lateral to plan (1-2 mm). The drill and tap appeared to be on trajectory for navigation. The trajectory was re-drilled and re-tapped larger, but there was still resistance. Imaging was done and the screw was lateral and superior in the disk space. The screw was removed and the surgeon decided to add an additional level at L5. During the procedure, the surgeon used the drill for the pilot hole and then a separate drill to tap and place the screws. The cause of the deviation was not confirmed, but the representative suspected that the tapped hole was on trajectory and the screw did not follow the hole. There was no patient harm and the procedure was delayed less than an hour. | xse left deviation salesforce email iud spinal procedure deviation single position percutaneous case schanz screw placed left psis mount surgical screws right right left accurately placed left screw deviated surgeon tapped trajectory fine screw resistance placed navigation slightly lateral plan drill tap appeared trajectory navigation trajectory redrilled retapped larger resistance imaging screw lateral superior disk space screw removed surgeon decided add additional level procedure surgeon drill pilot hole separate drill tap place screws cause deviation confirmed suspected tapped hole trajectory screw did follow hole procedure delayed hour mazor x | NAV5000 | NAV3299 | NAV5001 | NAV1012 | NAV4091 | NAV5000NAV3299NAV5001NAV1012NAV4091 | True
1832 | 703218527 | XSE - iron man missing | 2019-06-20T18:58:46.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MEC0625-01 | IRON MAN. | XSE - iron man missing 2019-Jun-18 SCL1901792 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the Iron Man was missing from system drawer. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. IRON MAN. | 2019-Jun-18 SCL1901792 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the Iron Man was missing from system drawer. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. | xse iron man missing outside procedure iron man missing drawer additional complications anticipated iron man | NAV2053 | NAV2054 | NAV3249 | NAV3295 | NAV4092 | NAV2053NAV2054NAV3249NAV3295NAV4092 | True
4125 | 703502512 | XSE - SRA Shaking/System Crash | 2019-12-05T11:20:43.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 40 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - SRA Shaking/System Crash 2019-12-05 00679631 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the arm was shaking and the system crashed during the procedure. During the procedure, the surgical arm was sent to the left L1 trajectory. The dilator was placed in the arm guide followed by drilling and tapping with no issues. When the driver was placed in the arm guide, the surgical arm started shaking. The arm stopped shaking after the emergency button was pressed. After the emergency button was released, the surgical arm worked normally. The representative also reported that when going from the operation to planning in the software, the system crashed. After 2 minutes, the system came back to the same window. There was no patient harm and the procedure was delayed less than one hour. 2019-12-09 IUD x2 (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-12-05 00679631 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the arm was shaking and the system crashed during the procedure. During the procedure, the surgical arm was sent to the left L1 trajectory. The dilator was placed in the arm guide followed by drilling and tapping with no issues. When the driver was placed in the arm guide, the surgical arm started shaking. The arm stopped shaking after the emergency button was pressed. After the emergency button was released, the surgical arm worked normally. The representative also reported that when going from the operation to planning in the software, the system crashed. After 2 minutes, the system came back to the same window. There was no patient harm and the procedure was delayed less than one hour. 2019-12-09 IUD x2 (Other): No new information. | xse sra shakingsystem crash spinal procedure arm shaking crashed procedure procedure surgical arm sent left trajectory dilator placed arm guide followed drilling tapping issues driver placed arm guide surgical arm started shaking arm stopped shaking emergency button pressed emergency button released surgical arm worked normally going operation planning software crashed minutes came window procedure delayed hour iud new mazor x software installation kit | NAV3247 | NAV4101 | NAV3246 | NAV4096 | NAV2047 | NAV3247NAV4101NAV3246NAV4096NAV2047 | True
9729 | 704707955 | XSE - System Needs Repair | 2021-11-10T16:05:08.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 60 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - System Needs Repair 2021-Nov-10 Salesforce C00237642, 00725556 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system needed repair. There was no patient involvement. 2021-11-14 email 2021-11-14 03:00, email 2021-11-14 03:15, email 2021-11-14 03:30 (Rep, For): No new information. 2021-11-18 email 2021-11-18 00:15 (Rep): No new information. 2021-11-21 IUD x2 (Other): No new information. 2021-12-03 email 2021-12-02 02:46 (Rep, For): Additional information received from a manufacturer representative reported the PCs were replaced due to being faulty and a PC fan error. The electrical unit was replaced because of a communication error in the system. 2021-Dec-26 IUD, SAP 302673385 (other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-Nov-10 Salesforce C00237642, 00725556 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system needed repair. There was no patient involvement. 2021-11-14 email 2021-11-14 03:00, email 2021-11-14 03:15, email 2021-11-14 03:30 (Rep, For): No new information. 2021-11-18 email 2021-11-18 00:15 (Rep): No new information. 2021-11-21 IUD x2 (Other): No new information. 2021-12-03 email 2021-12-02 02:46 (Rep, For): Additional information received from a manufacturer representative reported the PCs were replaced due to being faulty and a PC fan error. The electrical unit was replaced because of a communication error in the system. 2021-Dec-26 IUD, SAP 302673385 (other): No new information. | xse needs repair salesforce outside procedure needed repair email email email new email new iud new email additional pcs replaced faulty pc fan error electrical unit replaced communication error iud sap new kit svc cart elec panel assy | NAV2058 | NAV2047 | NAV3249 | NAV3262 | NAV3264 | NAV2058NAV2047NAV3249NAV3262NAV3264 | True
7913 | 704305285 | XSE - No Video on Surgeon Screen | 2021-03-26T00:04:38.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - No Video on Surgeon Screen 2021-Mar-25 Salesforce C00223302 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the system was powered on this morning, the surgeon tablet was not functioning. It was coming on and the manufacturer representative was able to move her mouse to that screen, but the software was not properly extending onto that screen. This was before the system was plugged into the booms. The system was then plugged into the boom monitors and able to use them just fine. No patient was present. 2021-04-09 IUD x3 (Rep): Additional information received from a manufacturer representative reported that the issue would be resolved and then come back. SURGEON SCREEN | 2021-Mar-25 Salesforce C00223302 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the system was powered on this morning, the surgeon tablet was not functioning. It was coming on and the manufacturer representative was able to move her mouse to that screen, but the software was not properly extending onto that screen. This was before the system was plugged into the booms. The system was then plugged into the boom monitors and able to use them just fine. No patient was present. 2021-04-09 IUD x3 (Rep): Additional information received from a manufacturer representative reported that the issue would be resolved and then come back. | xse video surgeon screen salesforce outside procedure powered morning surgeon tablet functioning coming able mouse screen software properly extending screen plugged booms plugged boom monitors able use just fine present iud additional issue resolved come surgeon screen | NAV3253 | NAV3262 | NAV3246 | NAV2058 | NAV3249 | NAV3253NAV3262NAV3246NAV2058NAV3249 | True
1057 | 703186506 | X - Software issue | 2019-06-03T20:25:09.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Software issue 2018-07-17 CCR, SCL1802576 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system performed as intended for k-wire placement. 4 k-wires were placed without issue. The monitor went black for approximately 30 seconds. CTRL,ALT,R was hit and screen came back up. There were no other issues. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SOFTWARE KIT VER-3.0.2 | 2018-07-17 CCR, SCL1802576 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system performed as intended for k-wire placement. 4 k-wires were placed without issue. The monitor went black for approximately 30 seconds. CTRL,ALT,R was hit and screen came back up. There were no other issues. No patient harm was reported and surgery was delayed less than an hour. | x software issue ccr spinal procedure performed intended kwire placement kwires placed issue monitor went black approximately seconds ctrlaltr hit screen came issues surgery delayed hour mazor x software kit | NAV3247 | NAV3248 | NAV3242 | NAV4091 | NAV3261 | NAV3247NAV3248NAV3242NAV4091NAV3261 | True
5016 | 703590349 | XSE - SW Issue w/ Orientation of Scan | 2020-01-31T00:10:55.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - SW Issue w/ Orientation of Scan 2020-01-27 00683132 ID (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for a scan and plan procedure. The portable 3D CT spin was completed and the study was transferred to the system for the manufacturer representative to start the planning process. The representative noted immediately that the orientation of the scan as it was imported to the system software was not correct as the anterior/posterior were switched. The representative as the radiology technician present for the case if the parameters had been set correct and the rad tech assured that they had been. After checking the scan on the CT's screen, it was noted that the patient orientation was not correct for the guidance system planning process; the patient orientation was set to supine and the guidance system required prone. The representatives immediately began trying to adjust the orientation on the software, attempting to utilize the tools in the software to reverse the orientation of the scan. By that time, the surgeon was scrubbed in and became impatient. The surgeon did not understand the workflow that required completing the entire planning process before surgery could being as the surgeon was used to navigating with a different system. After numerous attempts to correct the orientation of the scan in the guidance system software, the scan orientation could not be corrected. The surgeon was adamant about moving on with the case so another spine was taken with the correct patient orientation. The surgeon aborted use of the robot and switched to another navigation system. No patient harm was reported and surgery was delayed less than an hour. 2020-02-03 IUD x2 (Rep): Patient information updated. 2020-02-20 email 2020-02-20 16:45 (Rep): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-01-27 00683132 ID (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for a scan and plan procedure. The portable 3D CT spin was completed and the study was transferred to the system for the manufacturer representative to start the planning process. The representative noted immediately that the orientation of the scan as it was imported to the system software was not correct as the anterior/posterior were switched. The representative as the radiology technician present for the case if the parameters had been set correct and the rad tech assured that they had been. After checking the scan on the CT's screen, it was noted that the patient orientation was not correct for the guidance system planning process; the patient orientation was set to supine and the guidance system required prone. The representatives immediately began trying to adjust the orientation on the software, attempting to utilize the tools in the software to reverse the orientation of the scan. By that time, the surgeon was scrubbed in and became impatient. The surgeon did not understand the workflow that required completing the entire planning process before surgery could being as the surgeon was used to navigating with a different system. After numerous attempts to correct the orientation of the scan in the guidance system software, the scan orientation could not be corrected. The surgeon was adamant about moving on with the case so another spine was taken with the correct patient orientation. The surgeon aborted use of the robot and switched to another navigation system. No patient harm was reported and surgery was delayed less than an hour. 2020-02-03 IUD x2 (Rep): Patient information updated. 2020-02-20 email 2020-02-20 16:45 (Rep): No new information. | xse sw issue w orientation scan id spinal procedure case scan plan procedure portable ct spin completed study transferred start planning process noted immediately orientation scan imported software correct anteriorposterior switched radiology technician present case parameters set correct rad tech assured checking scan cts screen noted orientation correct planning process orientation set supine required prone representatives immediately began trying adjust orientation software attempting utilize tools software reverse orientation scan time surgeon scrubbed impatient surgeon did understand workflow required completing entire planning process surgery surgeon navigating different numerous attempts correct orientation scan software scan orientation corrected surgeon adamant moving case spine taken correct orientation surgeon aborted use robot switched navigation surgery delayed hour iud updated email new mazor x software installation kit | NAV3246 | NAV3235 | NAV3248 | NAV2056 | NAV4095 | NAV3246NAV3235NAV3248NAV2056NAV4095 | True
8497 | 704427978 | XBU ¦ Inaccurate Trajectory/Driver Issue | 2021-06-02T23:01:38.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 50 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XBU ¦ Inaccurate Trajectory/Driver Issue 2021-06-02 00714320 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that there was a variance showing a full screw width off in the axial and lateral view. All four trajectories were not aligning with the instruments. There was also a driver was getting stuck on the Navlock. And, the main screen was being slow to respond and jumping. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-06-21 IUD (Other): No new information. 2021-07-07 email 2021-07-07 18:46 (Rep): No new information. 2021-07-21 IUD (Other): No new information. 2021-08-03 IUD (Other): No new information. 2021-08-19 IUD x2 (Other): No new information. 2021-09-02 email 2021-09-30 11:50 (Rep): Trajectories were deviated about 5mm. Also, it was stated that the driver being stuck in the Navlock was due to a Navlock issue. It was also stated that the issue of the screen blacking out is due to a known issue with software. 2021-10-11 IUD (Other): No new information. 2021-10-26 Attachment (Other): No new information. UNK MAZOR X MONITOR | 2021-06-02 00714320 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that there was a variance showing a full screw width off in the axial and lateral view. All four trajectories were not aligning with the instruments. There was also a driver was getting stuck on the Navlock. And, the main screen was being slow to respond and jumping. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-06-21 IUD (Other): No new information. 2021-07-07 email 2021-07-07 18:46 (Rep): No new information. 2021-07-21 IUD (Other): No new information. 2021-08-03 IUD (Other): No new information. 2021-08-19 IUD x2 (Other): No new information. 2021-09-02 email 2021-09-30 11:50 (Rep): Trajectories were deviated about 5mm. Also, it was stated that the driver being stuck in the Navlock was due to a Navlock issue. It was also stated that the issue of the screen blacking out is due to a known issue with software. 2021-10-11 IUD (Other): No new information. 2021-10-26 Attachment (Other): No new information. | xbu inaccurate trajectorydriver issue intraoperatively spinal procedure variance showing screw width axial lateral view trajectories aligning instruments driver getting stuck navlock main screen slow respond jumping issue extended surgical time hour impact outcome iud new email new iud new iud new iud new email trajectories deviated stated driver stuck navlock navlock issue stated issue screen blacking known issue software iud new attachment new unk mazor x monitor | NAV3261 | NAV5002 | NAV3248 | NAV3244 | NAV1012 | NAV3261NAV5002NAV3248NAV3244NAV1012 | True
9914 | 704745309 | XSE - Shoulder Lock & X-Eye Issues | 2021-12-02T23:11:55.000000+0000 | MazorX Stealth Edition | NAV3254 | X-EYE ISSUES | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | XSE - Shoulder Lock & X-Eye Issues 2021-Dec-02 Salesforce C00239087 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative had difficulty getting the unit to recognize that the shoulder was locked on the surgical arm setup procedure. He moved the shoulder several times before he was able to get the system to recognize the shoulder was locked. The shoulder did physically lock every time they pressed the shoulder lock. It was also noticed that the X-eye camera was not rigidly attached to the post. No patient was present. 2021-12-14 IUD (Other): No new information. 2021-12-21 IUD (Other): No new information. 2022-01-06 IUD x2 (Other): No new information. 2022-02-09 IUD (Other): The X-eye bracket was completely off and the screw securing it was inside the positioner (the camera was dangling on arrival). The Firewire cable at the base of the positioner was unplugged. Also the cover on the bottom rear of the positioner is damaged and was off. The shoulder was overtightened resulting in the piston scraping on the side wall of the shoulder and throwing off the home position sensor. The front shoulder was assembled incorrectly. The cover was not flush and therefore the encoder that gives real time feedback of the shoulder position could not communicate. 2022-02-21 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. UNK MAZOR X X-EYE CAMERA | 2021-Dec-02 Salesforce C00239087 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative had difficulty getting the unit to recognize that the shoulder was locked on the surgical arm setup procedure. He moved the shoulder several times before he was able to get the system to recognize the shoulder was locked. The shoulder did physically lock every time they pressed the shoulder lock. It was also noticed that the X-eye camera was not rigidly attached to the post. No patient was present. 2021-12-14 IUD (Other): No new information. 2021-12-21 IUD (Other): No new information. 2022-01-06 IUD x2 (Other): No new information. 2022-02-09 IUD (Other): The X-eye bracket was completely off and the screw securing it was inside the positioner (the camera was dangling on arrival). The Firewire cable at the base of the positioner was unplugged. Also the cover on the bottom rear of the positioner is damaged and was off. The shoulder was overtightened resulting in the piston scraping on the side wall of the shoulder and throwing off the home position sensor. The front shoulder was assembled incorrectly. The cover was not flush and therefore the encoder that gives real time feedback of the shoulder position could not communicate. 2022-02-21 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. | xse shoulder lock xeye issues salesforce outside procedure difficulty getting unit recognize shoulder locked surgical arm setup procedure moved shoulder times able recognize shoulder locked shoulder did physically lock time pressed shoulder lock noticed xeye camera rigidly attached post present iud new iud new iud new iud xeye bracket completely screw securing inside positioner firewire cable base positioner unplugged cover rear positioner damaged shoulder overtightened resulting piston scraping wall shoulder throwing home position sensor shoulder assembled incorrectly cover flush encoder gives real time feedback shoulder position communicate iud new iud new unk mazor x xeye camera | NAV4110 | NAV3254 | NAV4107 | NAV4106 | NAV4103 | NAV4110NAV3254NAV4107NAV4106NAV4103 | True
6406 | 703935990 | XSE - Deviated Trajectory | 2020-09-02T13:47:08.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Trajectory 2020-09-01 00695703, Intake Form, e1, IUD x3, email 2020-09-21 13:45 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon mentioned that right L5 trajectory felt weird while executing the trajectory. The plan was adjusted and a new trajectory was executed. Confirmation shots showed that the right L5 trajectory was deviated laterally so the surgeon free-handed the trajectory. The cause of the deviation was not determined. An advance accuracy test and setup test was done and the surgical arm passed the tests. All screws were able to be accurately placed. There was no patient harm and the procedure was delayed less than an hour. 2020-10-16 email 2020-10-16 06:16 (Rep): Additional information received from a manufacturer representative reported that the procedure was not related to the surgical arm. 2020-11-12 email 2020-11-12 10:01 (Rep): Additional information received from a manufacturer representative reported the deviation was between 3.5-10 mm. The cause of the deviation was not determined. MAZOR X SYSTEM TPL0059 | 2020-09-01 00695703, Intake Form, e1, IUD x3, email 2020-09-21 13:45 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon mentioned that right L5 trajectory felt weird while executing the trajectory. The plan was adjusted and a new trajectory was executed. Confirmation shots showed that the right L5 trajectory was deviated laterally so the surgeon free-handed the trajectory. The cause of the deviation was not determined. An advance accuracy test and setup test was done and the surgical arm passed the tests. All screws were able to be accurately placed. There was no patient harm and the procedure was delayed less than an hour. 2020-10-16 email 2020-10-16 06:16 (Rep): Additional information received from a manufacturer representative reported that the procedure was not related to the surgical arm. 2020-11-12 email 2020-11-12 10:01 (Rep): Additional information received from a manufacturer representative reported the deviation was between 3.5-10 mm. The cause of the deviation was not determined. | xse deviated trajectory iud email spinal procedure surgeon mentioned right trajectory felt weird executing trajectory plan adjusted new trajectory executed confirmation shots showed right trajectory deviated laterally surgeon freehanded trajectory cause deviation determined advance accuracy test setup test surgical arm passed tests screws able accurately placed procedure delayed hour email additional procedure related surgical arm email additional deviation mm cause deviation determined mazor x | NAV5000 | NAV3243 | NAV4096 | NAV2051 | NAV3249 | NAV5000NAV3243NAV4096NAV2051NAV3249 | True
6443 | 703941946 | XSE - Damaged Screen/SRA Failing | 2020-09-04T21:48:51.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Damaged Screen/SRA Failing 2020-09-04 00695925 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that after doing a preventative maintenance the system was noted to have a damaged surgeon screen and the surgical arm was failing at points 3 and 6 and also reaching its trajectories. No patient was present. 2020-09-09 IUD x5 (Rep): No new information. 2020-10-06 IUD x2 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-09-04 00695925 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that after doing a preventative maintenance the system was noted to have a damaged surgeon screen and the surgical arm was failing at points 3 and 6 and also reaching its trajectories. No patient was present. 2020-09-09 IUD x5 (Rep): No new information. 2020-10-06 IUD x2 (Other): No new information. | xse damaged screensra failing outside procedure doing preventative maintenance noted damaged surgeon screen surgical arm failing points reaching trajectories present iud new iud new kit svc surgical arm refurb | NAV3249 | NAV3244 | NAV1006 | NAV2054 | NAV4101 | NAV3249NAV3244NAV1006NAV2054NAV4101 | True
5936 | 703841340 | XSE - Damaged Schanz Arm | 2020-07-08T15:10:13.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | KIT0284 | 4MM X 120MM SCHANZ SCREW KIT0284 | XSE - Damaged Schanz Arm 2020-07-08 00692246 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arm seized and would not function as intended. There was no patient involvement. 2020-07-14 IUD x3 (Other): No new information. 2020-07-29 IUD x4 (Other): No new information. 2020-08-19 IUD (Other): No new information. 2020-08-27 IUD (Other): No new information. 2020-09-15 IUD x3 (Other): No new information. 4MM X 120MM SCHANZ SCREW KIT0284 | 2020-07-08 00692246 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arm seized and would not function as intended. There was no patient involvement. 2020-07-14 IUD x3 (Other): No new information. 2020-07-29 IUD x4 (Other): No new information. 2020-08-19 IUD (Other): No new information. 2020-08-27 IUD (Other): No new information. 2020-09-15 IUD x3 (Other): No new information. | xse damaged schanz arm outside procedure schanz arm seized function intended iud new iud new iud new iud new iud new x schanz screw | NAV3299 | NAV3244 | NAV3295 | NAV2054 | NAV3249 | NAV3299NAV3244NAV3295NAV2054NAV3249 | True
9099 | 704550395 | XSE ¦ Issues with Instruments | 2021-08-11T01:38:39.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 40 | MAS2271 | KIT SVC MAS2271 SCHANZ BOND 05S NEW | XSE ¦ Issues with Instruments 2021-Aug-10 Salesforce C00232656 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the auxiliary drill guide's teeth were damaged and that the clamp link screw was stripped. In addition, the schanz bond was missing. This was discovered outside of surgery with no patient present. 2021-Aug-19 IUDx4 (other): No new information. KIT SVC MAS2271 SCHANZ BOND 05S NEW | 2021-Aug-10 Salesforce C00232656 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the auxiliary drill guide's teeth were damaged and that the clamp link screw was stripped. In addition, the schanz bond was missing. This was discovered outside of surgery with no patient present. 2021-Aug-19 IUDx4 (other): No new information. | xse issues instruments salesforce outside procedure auxiliary drill guides teeth damaged clamp link screw stripped addition schanz bond missing discovered outside surgery present new kit svc schanz bond new | NAV3299 | NAV2053 | NAV4103 | NAV3251 | NAV4090 | NAV3299NAV2053NAV4103NAV3251NAV4090 | True
5064 | 703593979 | XSE - Schanz Arm Won't Loosen | 2020-02-03T16:11:35.000000+0000 | MazorX Stealth Edition | NAV3298 | PLATFORM ISSUES | 20 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Schanz Arm Won't Loosen 2020-02-03 00683230 ID (Rep) Medtronic received information regarding a guidance system being used at a medical facility. It was reported that the schanz arm would not loosen around the schanz pin and they were not tightened properly. No patient was reported to be involved. 2020-02-12 IUD x3, Email 2020-03-09 17:00 (Other): Additional information received from a manufacturer representative reported that the representative worked with the account and the filed service engineer to resolve this issue. KIT SVC MAS2351-04S SCHANZ ARM | 2020-02-03 00683230 ID (Rep) Medtronic received information regarding a guidance system being used at a medical facility. It was reported that the schanz arm would not loosen around the schanz pin and they were not tightened properly. No patient was reported to be involved. 2020-02-12 IUD x3, Email 2020-03-09 17:00 (Other): Additional information received from a manufacturer representative reported that the representative worked with the account and the filed service engineer to resolve this issue. | xse schanz arm wont loosen id medical facility schanz arm loosen schanz pin tightened properly involved iud email additional worked account filed service engineer resolve issue kit svc schanz arm | NAV3299 | NAV3298 | NAV3295 | NAV2054 | NAV3244 | NAV3299NAV3298NAV3295NAV2054NAV3244 | True
6138 | 703872119 | XSE - Tablet Going Blank | 2020-07-27T16:36:25.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - Tablet Going Blank 2020-03-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon screen blacked out during the procedure. A hard restart of the system was done and the issue was resolved. The screws were placed accurately. There was no patient harm and the procedure was delayed less than an hour. SURGEON SCREEN | 2020-03-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon screen blacked out during the procedure. A hard restart of the system was done and the issue was resolved. The screws were placed accurately. There was no patient harm and the procedure was delayed less than an hour. | xse tablet going blank ccr spinal procedure surgeon screen blacked procedure hard restart issue resolved screws placed accurately procedure delayed hour surgeon screen | NAV3253 | NAV2047 | NAV4112 | NAV3243 | NAV3246 | NAV3253NAV2047NAV4112NAV3243NAV3246 | True
618 | 703161461 | XSE - deviation | 2019-05-19T14:47:45.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - deviation 2019-May-15 SCL1901246 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the star marker was registered, planned and sent to L4-5. It was noted that these screws were very high skive potentials. They were "fighting" different tool options and even the bone anchor tool kept falling off the landing zone. Navigation was showing they were 15-20mm into pedicle instantly. The screws were eventually placed with the arm and AP showed lateral placement. They re-planned the screws to find different options and attempting freehand with no success. The HCP had to leave the OR with no posterior instruments placed. The HCP commented this was due to patient anatomy and the current drilling process. It was noted that without a high speed drill the pilot hole was being created somewhere else as planned. The procedure was delayed for 50 minutes and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-May-15 SCL1901246 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the star marker was registered, planned and sent to L4-5. It was noted that these screws were very high skive potentials. They were "fighting" different tool options and even the bone anchor tool kept falling off the landing zone. Navigation was showing they were 15-20mm into pedicle instantly. The screws were eventually placed with the arm and AP showed lateral placement. They re-planned the screws to find different options and attempting freehand with no success. The HCP had to leave the OR with no posterior instruments placed. The HCP commented this was due to patient anatomy and the current drilling process. It was noted that without a high speed drill the pilot hole was being created somewhere else as planned. The procedure was delayed for 50 minutes and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse deviation spinal procedure star marker registered planned sent noted screws high skive potentials fighting different tool options bone anchor tool kept falling landing zone navigation showing pedicle instantly screws eventually placed arm ap showed lateral placement replanned screws different options attempting freehand success hcp leave posterior instruments placed hcp commented anatomy current drilling process noted high speed drill pilot hole created planned procedure delayed minutes additional complications anticipated mazor x software installation kit | NAV5000 | NAV2054 | NAV1006 | NAV3235 | NAV5001 | NAV5000NAV2054NAV1006NAV3235NAV5001 | True
5719 | 703788945 | XSE - Unreachable Trajectories | 2020-06-05T22:22:35.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Unreachable Trajectories 2020-06-05 00690232, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disks with fixation. It was reported that during a Scan & Plan case, the manufacturer representative was able to upload the scan to the guidance system with a USB and the star marker was able to be located. Segmentation was able to be completed, but no green values were received during registration. The only time green values were received, occurred when placing the screw tip up toward the ceiling. Use of the guidance system was aborted and another guidance system was used. There was no patient harm and the procedure was delayed less than an hour. 2020-06-11 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-06-05 00690232, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disks with fixation. It was reported that during a Scan & Plan case, the manufacturer representative was able to upload the scan to the guidance system with a USB and the star marker was able to be located. Segmentation was able to be completed, but no green values were received during registration. The only time green values were received, occurred when placing the screw tip up toward the ceiling. Use of the guidance system was aborted and another guidance system was used. There was no patient harm and the procedure was delayed less than an hour. 2020-06-11 IUD (Other): No new information. | xse unreachable trajectories ccr spinal procedure treat degenerative disks fixation scan plan case able upload scan usb star marker able located segmentation able completed green values registration time green values occurred placing screw tip ceiling use aborted procedure delayed hour iud new mazor x | NAV3243 | NAV4112 | NAV4095 | NAV5000 | NAV3242 | NAV3243NAV4112NAV4095NAV5000NAV3242 | True
4529 | 703549096 | X - Stripped Schanz Arm & Clamp Link | 2020-01-07T20:00:08.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 30 | MAS2202-04S | KIT SVC MAS2202-04S CLAMP LINK | X - Stripped Schanz Arm & Clamp Link 2020-01-07 00681494 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the schanz arm and clamp link are stripped and need to be replaced. This was discovered in sterilization. No patient was present. 2020-01-15 IUD x3 (Other): No new information. 2020-02-07 Email 2020-02-07 08:45, IUD x3 (Other): No new information. 2020-02-26 IUD x2 (Other): No new information. KIT SVC MAS2202-04S CLAMP LINK | 2020-01-07 00681494 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the schanz arm and clamp link are stripped and need to be replaced. This was discovered in sterilization. No patient was present. 2020-01-15 IUD x3 (Other): No new information. 2020-02-07 Email 2020-02-07 08:45, IUD x3 (Other): No new information. 2020-02-26 IUD x2 (Other): No new information. | x stripped schanz arm clamp link outside procedure schanz arm clamp link stripped need replaced discovered sterilization present iud new email iud new iud new kit svc clamp link | NAV3299 | NAV3295 | NAV4100 | NAV4090 | NAV4095 | NAV3299NAV3295NAV4100NAV4090NAV4095 | True
1878 | 703220494 | X - Surgical System Issue | 2019-06-21T14:10:26.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Surgical System Issue Shoulder failure to lock errors. This happened when shoulder was locked manually and digitally. UNK MAZOR X SHOULDER | Shoulder failure to lock errors. This happened when shoulder was locked manually and digitally. | x surgical issue shoulder failure lock errors happened shoulder locked manually digitally unk mazor x shoulder | NAV4107 | NAV3244 | NAV2054 | NAV4106 | NAV3261 | NAV4107NAV3244NAV2054NAV4106NAV3261 | True
4084 | 703498067 | XSE - Damaged Dual Clamp | 2019-12-03T19:19:38.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1355-03 | DUAL CLAMP | XSE - Damaged Dual Clamp 2019-12-03 00679468, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the dual bridge screws were stripped. The clamp was unable to be used and the representative could not get the screw out. Another clamp was used to proceed with the case. There was no patient harm and the procedure was completed. 2020-01-17 email 2020-01-17 1216 (Rep): Additional information received from the representative reporting that the broken clamp would not be returned due to site refusal. The lot number was not available. DUAL CLAMP | 2019-12-03 00679468, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the dual bridge screws were stripped. The clamp was unable to be used and the representative could not get the screw out. Another clamp was used to proceed with the case. There was no patient harm and the procedure was completed. 2020-01-17 email 2020-01-17 1216 (Rep): Additional information received from the representative reporting that the broken clamp would not be returned due to site refusal. The lot number was not available. | xse damaged dual clamp spinal procedure dual bridge screws stripped clamp unable screw clamp proceed case procedure completed email additional reporting broken clamp returned site refusal lot number available dual clamp | NAV3299 | NAV3295 | NAV4090 | NAV4096 | NAV3261 | NAV3299NAV3295NAV4090NAV4096NAV3261 | True
655 | 703169730 | X- Software/Bedframe/Surgical arm issues | 2019-05-23T16:57:30.000000+0000 | MazorX | NAV3235 | 3D SCANS | 40 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X- Software/Bedframe/Surgical arm issues Platform First segment: RNS metal mini clamp and pinball. Second segment: Headpin in SP of L2 and pinball EBL: 400mL. Planning 14:31 -14:35 Surgeon and CSR reviewed plan prior to case. Initially, surgeon wished to execute S2AI screws, but opted to execute iliac bolts to due lack of purchase at S2. Reviewed plan to complete in two registrations starting with the iliac screws. Discussed approach and platform selection as well as plan to connect to existing hardware. Prep Arm attach: 14:16 (~60 sec) Draping: 14:26 (~45 sec) 14:49 Surgeon made small hole in drape for arm guide attachment. Arm guide attached by surgeon while maintaining clockwise pressure. Platform Mounting (bottom segment) 15:00-15:05 Surgeon attached metal RNS mini clamp to rod at right L5-S1 and applied pinball. When surgeon attached tracking reference marker plate to side of arm, it put a significant amount of tension on the drape to try to place the green arrows over the correct spot and made a small hole in drape (see image). Surgeon suggests a corresponding arrow mark on arm to line up with arrowon plate for easy alignment. Suggests same arrows for BMB and mounting plate as well. 15:05-15:07 Hole made in drape for hawkeye and BMB connected to pinball at 30 degrees. It was determined that despite attaching bed frame 10" below S1, surgeon could not move arm far enough caudal to attach BMB at 15 degrees. 3D Scan (bottom segment) 15:08-15:10 Recording; blue towel utilized over field. Resulting scan clear and detailed. Images and Registration (bottom segment) 15:10 Draw spine; reference stylus recognized without issue. Surgeon labeled L5. 15:11 Sent to AP: 8 images (received error message "3D marker not recognized" x 2) 15:14 Sent to OBL: 6 images (received error message "3D marker in low accuracy) 15:20-15:22 First attempt at labeling L3 failed. Achieved success off L4 with complex algorithm. 15:23 CSR noted that S2 didn't auto fluoro segment despite appearing to be entirely in images (however, images were not of ideal clarity). CSR attempted to manually place a segmentation marker at S2, but upon doing so the SW crashed. CSR waited approx 45 seconds before attempting a soft SW restart, however, this did not work. After waiting several more seconds, surgeon decided to freehand iliac bolts. 15:27 CSR had to restart WS. 15:32 Mazor bedframe obstructing c-arm images for iliac bolts. Platform mounting (top segment) 16:33 Surgeon inserted headpin into SP of L2 and applied pinball. CSR recommended two headpins and connection with the link bridge. Surgeon felt that because patient was fused from L1-S1, one point of fixation was enough stability. He remarked that the headpin in L2 SP had excellent purchase and if it did not, he would have utilized two pins. 16:39-16:40 BMB attached to pinball and locked at 15 degrees. 3D Scan (top segment) 16:40-16:43 Recording; surgeon utilized blue towel over field. Resulting scan was clear and detailed. Images and Registration (top segment) 16:44 Draw spine; reference stylus recognized without issue. Surgeon labeled T11. 16:45 Sent to AP position: 3 images 16:47 Sent to OBL position: 4 images 16:50-16:53 First attempt at labeling T12 failed. Second attempt labeling L1 was successful on complex algorithm. Operation 16:55 Right T10 120mm 16:57 Right T11 128mm 17:00 Right T12 173mm; lowered 2 clicks to 153mm 17:02 Left T12 118mm 17:04 Left T11 127mm 17:05 Left T10 125mm 17:08 Clear left; confirmatory images 17:20 BMU unlocked and arm moved out of field. 18:10 Arm off table (~30 sec) All solution times were very slow to calculate between levels (>10 sec). Attempts to run simulations were also extremely slow. UNKNOWN MAZOR X SW | Platform First segment: RNS metal mini clamp and pinball. Second segment: Headpin in SP of L2 and pinball EBL: 400mL. Planning 14:31 -14:35 Surgeon and CSR reviewed plan prior to case. Initially, surgeon wished to execute S2AI screws, but opted to execute iliac bolts to due lack of purchase at S2. Reviewed plan to complete in two registrations starting with the iliac screws. Discussed approach and platform selection as well as plan to connect to existing hardware. Prep Arm attach: 14:16 (~60 sec) Draping: 14:26 (~45 sec) 14:49 Surgeon made small hole in drape for arm guide attachment. Arm guide attached by surgeon while maintaining clockwise pressure. Platform Mounting (bottom segment) 15:00-15:05 Surgeon attached metal RNS mini clamp to rod at right L5-S1 and applied pinball. When surgeon attached tracking reference marker plate to side of arm, it put a significant amount of tension on the drape to try to place the green arrows over the correct spot and made a small hole in drape (see image). Surgeon suggests a corresponding arrow mark on arm to line up with arrowon plate for easy alignment. Suggests same arrows for BMB and mounting plate as well. 15:05-15:07 Hole made in drape for hawkeye and BMB connected to pinball at 30 degrees. It was determined that despite attaching bed frame 10" below S1, surgeon could not move arm far enough caudal to attach BMB at 15 degrees. 3D Scan (bottom segment) 15:08-15:10 Recording; blue towel utilized over field. Resulting scan clear and detailed. Images and Registration (bottom segment) 15:10 Draw spine; reference stylus recognized without issue. Surgeon labeled L5. 15:11 Sent to AP: 8 images (received error message "3D marker not recognized" x 2) 15:14 Sent to OBL: 6 images (received error message "3D marker in low accuracy) 15:20-15:22 First attempt at labeling L3 failed. Achieved success off L4 with complex algorithm. 15:23 CSR noted that S2 didn't auto fluoro segment despite appearing to be entirely in images (however, images were not of ideal clarity). CSR attempted to manually place a segmentation marker at S2, but upon doing so the SW crashed. CSR waited approx 45 seconds before attempting a soft SW restart, however, this did not work. After waiting several more seconds, surgeon decided to freehand iliac bolts. 15:27 CSR had to restart WS. 15:32 Mazor bedframe obstructing c-arm images for iliac bolts. Platform mounting (top segment) 16:33 Surgeon inserted headpin into SP of L2 and applied pinball. CSR recommended two headpins and connection with the link bridge. Surgeon felt that because patient was fused from L1-S1, one point of fixation was enough stability. He remarked that the headpin in L2 SP had excellent purchase and if it did not, he would have utilized two pins. 16:39-16:40 BMB attached to pinball and locked at 15 degrees. 3D Scan (top segment) 16:40-16:43 Recording; surgeon utilized blue towel over field. Resulting scan was clear and detailed. Images and Registration (top segment) 16:44 Draw spine; reference stylus recognized without issue. Surgeon labeled T11. 16:45 Sent to AP position: 3 images 16:47 Sent to OBL position: 4 images 16:50-16:53 First attempt at labeling T12 failed. Second attempt labeling L1 was successful on complex algorithm. Operation 16:55 Right T10 120mm 16:57 Right T11 128mm 17:00 Right T12 173mm; lowered 2 clicks to 153mm 17:02 Left T12 118mm 17:04 Left T11 127mm 17:05 Left T10 125mm 17:08 Clear left; confirmatory images 17:20 BMU unlocked and arm moved out of field. 18:10 Arm off table (~30 sec) All solution times were very slow to calculate between levels (>10 sec). Attempts to run simulations were also extremely slow. | x softwarebedframesurgical arm issues platform segment rns metal mini clamp pinball second segment headpin sp pinball ebl planning surgeon csr reviewed plan prior case initially surgeon wished execute screws opted execute iliac bolts lack purchase reviewed plan complete registrations starting iliac screws discussed approach platform selection plan connect existing hardware prep arm attach draping surgeon small hole drape arm guide attachment arm guide attached surgeon maintaining clockwise pressure platform mounting surgeon attached metal rns mini clamp rod right applied pinball surgeon attached tracking reference marker plate arm significant tension drape try place green arrows correct spot small hole drape surgeon suggests corresponding arrow mark arm line arrowon plate easy alignment suggests arrows bmb mounting plate hole drape hawkeye bmb connected pinball degrees determined despite attaching bed frame surgeon arm far caudal attach bmb degrees scan recording blue towel utilized field resulting scan clear detailed images registration draw spine reference stylus recognized issue surgeon labeled sent ap images sent obl images attempt labeling failed achieved success complex algorithm csr noted didnt auto fluoro segment despite appearing entirely images csr attempted manually place segmentation marker doing sw crashed csr waited approx seconds attempting soft sw restart did work waiting seconds surgeon decided freehand iliac bolts csr restart ws mazor bedframe obstructing carm images iliac bolts platform mounting surgeon inserted headpin sp applied pinball csr recommended headpins connection link bridge surgeon felt fused point fixation stability remarked headpin sp excellent purchase did utilized pins bmb attached pinball locked degrees scan recording surgeon utilized blue towel field resulting scan clear detailed images registration draw spine reference stylus recognized issue surgeon labeled sent ap position images sent obl position images attempt labeling failed second attempt labeling successful complex algorithm operation right right right lowered clicks left left left clear left confirmatory images bmu unlocked arm moved field arm table solution times slow calculate levels attempts run simulations extremely slow unknown mazor x sw | NAV3242 | NAV3247 | NAV3249 | NAV3255 | NAV3246 | NAV3242NAV3247NAV3249NAV3255NAV3246 | False
2583 | 703242429 | X - Registration issue | 2019-07-04T13:26:08.000000+0000 | MazorX | NAV4112 | NO REGISTRATION | 20 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X - Registration issue Case aborted due to registration failure. UNKNOWN MAZOR X SW | Case aborted due to registration failure. | x registration issue case aborted registration failure unknown mazor x sw | NAV4112 | NAV3246 | NAV3242 | NAV4091 | NAV2054 | NAV4112NAV3246NAV3242NAV4091NAV2054 | True
2333 | 703233206 | X- Software Issues | 2019-06-28T13:31:42.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X- Software Issues Plan was done by me and reviewed by Dr. Gebauer the morning of the surgery. L3-L5 were planned for apercutaneous surgery with a medial to lateral approach for cortical screws. After review, Dr. Gebauer requested that the screws take a more inferior to superior approach in the lateral which was adjusted at each level. There were no other concerns discussed or changes made to the plan. upon initial set up, there were no complications including advanced accuracy test and a check of the cameras before the robot was mounted to the table, we noticed that the screen was stuck in the "processing" state and was frozen. a soft restart was done which corrected the problem once the robot was mounted to the bed, a single schantz pin was utilized and placed into the right psis. there were no complications mounting the platform. During the 3define scan there was some difficulty after 3 failed attempts. We then selected our own parameters to move forward with the surface scan and get to the draw spine process. After attempting to move forward, the work station was stuck in the "processing" state for the second time. another soft reset was conducted but then we had to manually selectour parameters again to get through the surface scan again. There were no complications during the draw spine process. After our AP and oblique images, a body was selected to register off of and registration was successful on the first attempt.After the X's were placed for visual confirmation and attempting to move forward to the send surgical arm screen, the workstation was stuck in the "processing" state for a third time. Before attempting another soft restart, the c-arm tech hit the robot with the c-arm which caused a shoulder shift. we then had to go back and re do our manual surface parameters and re do the draw spine process. The second registration was achieved with no issues. There were no complications throughout instrumentation. all wires and screws were placed accurately and confirmed with x-ray. MAZOR X SOFTWARE KIT VER-2.3 | Plan was done by me and reviewed by Dr. Gebauer the morning of the surgery. L3-L5 were planned for apercutaneous surgery with a medial to lateral approach for cortical screws. After review, Dr. Gebauer requested that the screws take a more inferior to superior approach in the lateral which was adjusted at each level. There were no other concerns discussed or changes made to the plan. upon initial set up, there were no complications including advanced accuracy test and a check of the cameras before the robot was mounted to the table, we noticed that the screen was stuck in the "processing" state and was frozen. a soft restart was done which corrected the problem once the robot was mounted to the bed, a single schantz pin was utilized and placed into the right psis. there were no complications mounting the platform. During the 3define scan there was some difficulty after 3 failed attempts. We then selected our own parameters to move forward with the surface scan and get to the draw spine process. After attempting to move forward, the work station was stuck in the "processing" state for the second time. another soft reset was conducted but then we had to manually selectour parameters again to get through the surface scan again. There were no complications during the draw spine process. After our AP and oblique images, a body was selected to register off of and registration was successful on the first attempt.After the X's were placed for visual confirmation and attempting to move forward to the send surgical arm screen, the workstation was stuck in the "processing" state for a third time. Before attempting another soft restart, the c-arm tech hit the robot with the c-arm which caused a shoulder shift. we then had to go back and re do our manual surface parameters and re do the draw spine process. The second registration was achieved with no issues. There were no complications throughout instrumentation. all wires and screws were placed accurately and confirmed with x-ray. | x software issues plan reviewed dr gebauer morning surgery planned apercutaneous surgery medial lateral approach cortical screws review dr gebauer requested screws inferior superior approach lateral adjusted level concerns discussed changes plan initial set complications including advanced accuracy test check cameras robot mounted table noticed screen stuck processing state frozen soft restart corrected problem robot mounted bed single schantz pin utilized placed right psis complications mounting platform scan difficulty failed attempts selected parameters forward surface scan draw spine process attempting forward work station stuck processing state second time soft reset conducted manually selectour parameters surface scan complications draw spine process ap oblique images body selected register registration successful attemptafter xs placed visual confirmation attempting forward send surgical arm screen workstation stuck processing state time attempting soft restart carm tech hit robot carm caused shoulder shift manual surface parameters draw spine process second registration achieved issues complications instrumentation wires screws placed accurately confirmed xray mazor x software kit | NAV3242 | NAV3247 | NAV4091 | NAV2051 | NAV3249 | NAV3242NAV3247NAV4091NAV2051NAV3249 | False
115 | 703000386 | X - Drape Issue | 2019-02-11T11:56:51.000000+0000 | MazorX | NAV2054 | OUT OF BOX FAILURES | 20 | KIT0574 | MAZOR X SPINE DISPOSABLE KIT KIT0574 | X - Drape Issue 2019-01-31 SCL1900323, Intake Form 2019-02-11 06:20, Image 2019-02-11 06:20 x3, e1, e2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were plastic "hairs" on the welding line of the new drape for the surgical system and no indication for use (IFUs) were included in the box. There was no patient harm and surgery was no delayed more than an hour. The case was successfully completed. 2019-02-11 email 2019-02-26 14:34 (Rep): Additional information received from a manufacturer representative reported that two indications for use (IFUs) were not included in the box. MAZOR X SPINE DISPOSABLE KIT KIT0574 | 2019-01-31 SCL1900323, Intake Form 2019-02-11 06:20, Image 2019-02-11 06:20 x3, e1, e2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were plastic "hairs" on the welding line of the new drape for the surgical system and no indication for use (IFUs) were included in the box. There was no patient harm and surgery was no delayed more than an hour. The case was successfully completed. 2019-02-11 email 2019-02-26 14:34 (Rep): Additional information received from a manufacturer representative reported that two indications for use (IFUs) were not included in the box. | x drape issue image spinal procedure plastic hairs welding line new drape surgical indication use included box surgery delayed hour case successfully completed email additional indications use included box mazor x spine disposable kit | NAV3232 | NAV2054 | NAV3246 | NAV3248 | NAV5000 | NAV3232NAV2054NAV3246NAV3248NAV5000 | True
4381 | 703538883 | XSE - No Video Feed on X-eye | 2019-12-30T23:57:05.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 30 | ELB0057-01 | MAZOR X FIRE WIRE CARD | XSE - No Video Feed on X-eye 2019-12-30 00681203 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that there was no video feed for the X-eye camera. The camera turns toward the reference marker but there is no video. The manufacturer representative verified in the camera tabs that the X-eye was turned on but did not check the connections in the cart. No patient was present. 2020-01-02 IUD x2 (Other): Additional information was received stating that the cable was not plugged into the cart and once that was secured the X-eye functioned properly. 2020-01-09 IUD (Other): No new information. MAZOR X FIRE WIRE CARD | 2019-12-30 00681203 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that there was no video feed for the X-eye camera. The camera turns toward the reference marker but there is no video. The manufacturer representative verified in the camera tabs that the X-eye was turned on but did not check the connections in the cart. No patient was present. 2020-01-02 IUD x2 (Other): Additional information was received stating that the cable was not plugged into the cart and once that was secured the X-eye functioned properly. 2020-01-09 IUD (Other): No new information. | xse video feed xeye outside procedure video feed xeye camera camera turns reference marker video verified camera tabs xeye turned did check connections cart present iud additional stating cable plugged cart secured xeye functioned properly iud new mazor x wire card | NAV4110 | NAV2058 | NAV4109 | NAV2054 | NAV3254 | NAV4110NAV2058NAV4109NAV2054NAV3254 | False
8110 | 704354993 | XSE - Unable to Load C-arm Images | 2021-04-22T04:19:19.000000+0000 | MazorX Stealth Edition | NAV3237 | FLUORO IMAGES | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Unable to Load C-arm Images 2021-Apr-21 Salesforce C00225099 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that images from the C-arm were unable to be loaded onto the guidance system during the procedure. The manufacturer representative indicated the images appeared to load, but they were off-set and the guidance system displayed an error that the images were unrecognizable. Registration could not be completed without images and the use of the guidance system was aborted. There was no patient harm and the procedure was delayed less than an hour. 2021-04-30 IUD x2, email 2021-04-30 09:15 (Rep): Additional information received from a manufacturer representative reported that the C-arm was broken and the guidance system was functioning as intended. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Apr-21 Salesforce C00225099 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that images from the C-arm were unable to be loaded onto the guidance system during the procedure. The manufacturer representative indicated the images appeared to load, but they were off-set and the guidance system displayed an error that the images were unrecognizable. Registration could not be completed without images and the use of the guidance system was aborted. There was no patient harm and the procedure was delayed less than an hour. 2021-04-30 IUD x2, email 2021-04-30 09:15 (Rep): Additional information received from a manufacturer representative reported that the C-arm was broken and the guidance system was functioning as intended. | xse unable load carm images salesforce spinal p rocedure images carm unable loaded procedure indicated images appeared load offset displayed error images unrecognizable registration completed images use aborted procedure delayed hour iud email additional carm broken functioning intended mazorx sw installation kit | NAV3237 | NAV3246 | NAV4112 | NAV3235 | NAV3236 | NAV3237NAV3246NAV4112NAV3235NAV3236 | True
5009 | 703589683 | X- Deviation, 3Define & X-eye issues | 2020-01-30T19:57:36.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X- Deviation, 3Define & X-eye issues 2018-2-6 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done by the surgeon and the manufacturer representative prior to the procedure. A schanz pin was inserted into the right iliac and connected to a schanz arm. No 3Define scan or draw spine were completed as they are both not working. The 3Define camera created lines on the images and draw spine did not work due to the lighting. Registration was completed off of L5 on automatic with the basic algorithm. L4 and L2 had to be manually adjusted because they had red values. All trajectories were executed accurately, except for right L2. Right L2 was medial because the surgeon pulled the zelphi slightly towards the right side while passing the tools. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X 3DEFINE CAMERA | 2018-2-6 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done by the surgeon and the manufacturer representative prior to the procedure. A schanz pin was inserted into the right iliac and connected to a schanz arm. No 3Define scan or draw spine were completed as they are both not working. The 3Define camera created lines on the images and draw spine did not work due to the lighting. Registration was completed off of L5 on automatic with the basic algorithm. L4 and L2 had to be manually adjusted because they had red values. All trajectories were executed accurately, except for right L2. Right L2 was medial because the surgeon pulled the zelphi slightly towards the right side while passing the tools. There was no patient harm and the procedure was not delayed over an hour. | x deviation xeye issues ccr spinal procedure planning surgeon prior procedure schanz pin inserted right iliac connected schanz arm scan draw spine completed working camera created lines images draw spine did work lighting registration completed automatic basic algorithm manually adjusted red values trajectories executed accurately right right medial surgeon pulled zelphi slightly right passing tools procedure delayed hour unk mazor x camera | NAV4110 | NAV3244 | NAV3249 | NAV4091 | NAV3254 | NAV4110NAV3244NAV3249NAV4091NAV3254 | True
9887 | 704742993 | XSE - CANBUS Error Messages | 2021-12-01T23:10:51.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error Messages 2021-Dec-01 Salesforce C00239013, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that a CANBUS error and an error that the emergency stop would not clear was displayed. The manufacturer representative noted that the 24V was missing from the electrical panel. There was no patient involvement. 2021-12-01 IUD x2 (Other): No new information. 2021-12-03 IUD x5 (Other): No new information. 2021-12-20 IUD (Other): No new information. 2022-01-10 IUD x2 (Other): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-14 e3, IUD (Rep): Additional information received from a manufacturer representative reported that they thought the electrical panel was bad since the 24 Vdc was missing. The original electrical panel was installed as they believed the missing 24 Vdc was from a bad main cable. 2022-02-17 IUD x2 (Other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2021-Dec-01 Salesforce C00239013, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that a CANBUS error and an error that the emergency stop would not clear was displayed. The manufacturer representative noted that the 24V was missing from the electrical panel. There was no patient involvement. 2021-12-01 IUD x2 (Other): No new information. 2021-12-03 IUD x5 (Other): No new information. 2021-12-20 IUD (Other): No new information. 2022-01-10 IUD x2 (Other): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-14 e3, IUD (Rep): Additional information received from a manufacturer representative reported that they thought the electrical panel was bad since the 24 Vdc was missing. The original electrical panel was installed as they believed the missing 24 Vdc was from a bad main cable. 2022-02-17 IUD x2 (Other): No new information. | xse canbus error messages salesforce iud outside proc edure canbus error error emergency stop clear displayed noted missing electrical panel iud new iud new iud new iud new iud new iud additional thought electrical panel bad vdc missing original electrical panel installed believed missing vdc bad main cable iud new svc main cable | NAV2047 | NAV3263 | NAV2054 | NAV2053 | NAV1006 | NAV2047NAV3263NAV2054NAV2053NAV1006 | True
6281 | 703903011 | XSE - Screens Going Black/Froze System | 2020-08-13T19:31:24.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 30 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE - Screens Going Black/Froze System 2020-08-12 00694563 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system was booted up with nothing plugged in. When the HDMI was plugged in, the screen did not flicker, but the picture did come up. The manufacturer representative noted that everything was working fine for the first spin of the Scan & Plan case. After a second spin, the system froze and the surgeon and monitor screens went black. The screen on the boom monitor was frozen. The representative tried pushing Ctrl-Alt-R, but nothing would respond. A hard restart of the system was done and the case was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-14 email 2020-08-14 09:01, IUD x3 (Rep): Additional information received from a manufacturer representative reported that the cause of the screens going black and the system freezing was not determined. 2020-08-21 IUD x4 (Other): No new information. 2020-09-03 IUD x3 (Other): No new information. 2020-10-05 IUD (Other): No new information. 2020-10-30 IUD (Other): No new information. 2021-06-20 IUD (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2020-08-12 00694563 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system was booted up with nothing plugged in. When the HDMI was plugged in, the screen did not flicker, but the picture did come up. The manufacturer representative noted that everything was working fine for the first spin of the Scan & Plan case. After a second spin, the system froze and the surgeon and monitor screens went black. The screen on the boom monitor was frozen. The representative tried pushing Ctrl-Alt-R, but nothing would respond. A hard restart of the system was done and the case was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-14 email 2020-08-14 09:01, IUD x3 (Rep): Additional information received from a manufacturer representative reported that the cause of the screens going black and the system freezing was not determined. 2020-08-21 IUD x4 (Other): No new information. 2020-09-03 IUD x3 (Other): No new information. 2020-10-05 IUD (Other): No new information. 2020-10-30 IUD (Other): No new information. 2021-06-20 IUD (Other): No new information. | xse screens going blackfroze spinal procedure booted plugged hdmi plugged screen did flicker picture did come noted working fine spin scan plan case second spin froze surgeon monitor screens went black screen boom monitor frozen tried pushing ctrlaltr respond hard restart case able completed procedure delayed hour email iud additional cause screens going black freezing determined iud new iud new iud new iud new iud new kit svc srgn screen arm | NAV3253 | NAV3261 | NAV3247 | NAV3248 | NAV2054 | NAV3253NAV3261NAV3247NAV3248NAV2054 | True
3679 | 703414247 | X - Schanz arm damaged | 2019-10-15T17:35:21.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2351-04 | SCHANZ ARM | X - Schanz arm damaged 2019-Oct-11 SCL1903065 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 4 schanz arms were damaged at the site. There was no patient harm or additional complications reported or anticipated as a result of the event. SCHANZ ARM | 2019-Oct-11 SCL1903065 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 4 schanz arms were damaged at the site. There was no patient harm or additional complications reported or anticipated as a result of the event. | x schanz arm damaged spinal procedure schanz arms damaged site additional complications anticipated schanz arm | NAV3299 | NAV3295 | NAV3296 | NAV3249 | NAV2053 | NAV3299NAV3295NAV3296NAV3249NAV2053 | True
10460 | 704884489 | XSE - Damaged Positioner | 2022-02-25T17:13:35.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 20 | HRS0327-01S | KIT SVC HRS0327-01S TEETH LOCKING SENSOR | XSE - Damaged Positioner 2022-Feb-25 Salesforce C00244254 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site's system has a broken sensor on the positioner. No patient was present. 2022-03-07 IUD (Other): No new information. 2022-03-11 IUD x2 (Other): No new information. KIT SVC HRS0327-01S TEETH LOCKING SENSOR | 2022-Feb-25 Salesforce C00244254 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site's system has a broken sensor on the positioner. No patient was present. 2022-03-07 IUD (Other): No new information. 2022-03-11 IUD x2 (Other): No new information. | xse damaged positioner salesforce outside procedure sites broken sensor positioner present iud new iud new kit svc teeth locking sensor | NAV3251 | NAV4103 | NAV3249 | NAV4104 | NAV2058 | NAV3251NAV4103NAV3249NAV4104NAV2058 | True
6171 | 703879610 | XSE - Black Screen | 2020-07-30T19:23:11.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XSE - Black Screen 2020-07-30 00693745 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while adjusting the brightness and contrast, the monitor went black briefly. The image appeared shortly after and the manufacturer representative was able to continue with preventative maintenance. The issue was unable to be replicated. There was no patient involvement. UNK MAZOR X MONITOR | 2020-07-30 00693745 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while adjusting the brightness and contrast, the monitor went black briefly. The image appeared shortly after and the manufacturer representative was able to continue with preventative maintenance. The issue was unable to be replicated. There was no patient involvement. | xse black screen outside procedure adjusting brightness contrast monitor went black briefly image appeared shortly able continue preventative maintenance issue unable replicated unk mazor x monitor | NAV3261 | NAV3244 | NAV3253 | NAV2054 | NAV3250 | NAV3261NAV3244NAV3253NAV2054NAV3250 | True
9337 | 704612369 | XSE - Amber Blinking Light on PC | 2021-09-15T17:58:31.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215 | ASM0215 MAZOR X COMPUTER KIT | XSE - Amber Blinking Light on PC 2021-Sep-15 Salesforce C00234494 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the system was powered on this morning, the system was not fully coming up. When she took left side cover off, she could see that the power button on the computer was flashing amber. She hard rebooted and that seemed resolve for a little bit and the computer powered on normally, but then it happened again and after hard rebooting, she could not get to resolve. This is for a training system so no patient present. 2021-09-17 IUD x3 (Other): No new information. 2021-09-30 IUD x3 (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-10-26 IUD (Other): No new information. 2021-10-28 IUD x2 (Other): No new information. 2021-12-20 IUD (Other): No new information. ASM0215 MAZOR X COMPUTER KIT | 2021-Sep-15 Salesforce C00234494 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the system was powered on this morning, the system was not fully coming up. When she took left side cover off, she could see that the power button on the computer was flashing amber. She hard rebooted and that seemed resolve for a little bit and the computer powered on normally, but then it happened again and after hard rebooting, she could not get to resolve. This is for a training system so no patient present. 2021-09-17 IUD x3 (Other): No new information. 2021-09-30 IUD x3 (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-10-26 IUD (Other): No new information. 2021-10-28 IUD x2 (Other): No new information. 2021-12-20 IUD (Other): No new information. | xse amber blinking light pc salesforce outside procedure powered morning fully coming took left cover power button computer flashing amber hard rebooted resolve little bit computer powered normally happened hard rebooting resolve training present iud new iud new iud new iud new iud new iud new mazor x computer kit | NAV3262 | NAV3261 | NAV3248 | NAV2047 | NAV3264 | NAV3262NAV3261NAV3248NAV2047NAV3264 | True
4184 | 703511017 | XSE - 3Define Scan Froze | 2019-12-10T19:23:35.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 30 | ASM0206-01 | SURGICAL ARM | XSE - 3Define Scan Froze 2019-12-10 00679983 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that during a 3Define scan on the first attempt the arm went to the first image and froze. The manufacturer representative performed a soft restart and performed the 3Define scan again and it passed. No patient was present. 2019-12-12 IUD x3 (Other): No new information. SURGICAL ARM | 2019-12-10 00679983 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that during a 3Define scan on the first attempt the arm went to the first image and froze. The manufacturer representative performed a soft restart and performed the 3Define scan again and it passed. No patient was present. 2019-12-12 IUD x3 (Other): No new information. | xse scan froze outside procedure scan attempt arm went image froze performed soft restart performed scan passed present iud new surgical arm | NAV3244 | NAV3248 | NAV2058 | NAV3249 | NAV4095 | NAV3244NAV3248NAV2058NAV3249NAV4095 | True
1361 | 703197554 | X - Eye Functionality Issue | 2019-06-10T18:48:54.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Eye Functionality Issue Separated from SCL1802199 During case on 6/13, x eye camera was not functioning. System was working in pre op check and then intra op showed a red X. CS unplugged and plugged back in and still had no connection UNK MAZOR X X-EYE CAMERA | Separated from SCL1802199 During case on 6/13, x eye camera was not functioning. System was working in pre op check and then intra op showed a red X. CS unplugged and plugged back in and still had no connection | x eye functionality issue separated case x eye camera functioning working pre op check intra op showed red x cs unplugged plugged connection unk mazor x xeye camera | NAV4110 | NAV3244 | NAV3299 | NAV3250 | NAV3295 | NAV4110NAV3244NAV3299NAV3250NAV3295 | True
9273 | 704597410 | XSE ¦ Leak in Manipulator | 2021-09-07T22:12:48.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | KIT1259 | SVC KIT1259 PNM PUSH-IN CON QSM3 02 NEW | XSE ¦ Leak in Manipulator 2021-Sep-07 Salesforce C00234055 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a leak in the manipulator line. The manufacturer representative could hear the leak when the manipulator was angled in extreme positions. The system was still functioning as designed. No patient was present. SVC KIT1259 PNM PUSH-IN CON QSM3 02 NEW | 2021-Sep-07 Salesforce C00234055 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a leak in the manipulator line. The manufacturer representative could hear the leak when the manipulator was angled in extreme positions. The system was still functioning as designed. No patient was present. | xse leak manipulator salesforce outside procedure leak manipulator line hear leak manipulator angled extreme positions functioning designed present svc pnm pushin new | NAV3260 | NAV3264 | NAV3249 | NAV2057 | NAV3253 | NAV3260NAV3264NAV3249NAV2057NAV3253 | True
8231 | 704375759 | XSE - Damaged Green Screwdriver | 2021-05-04T15:45:41.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1650-03S | KIT SVC MAS1650-03S UNIV SCREWDRIVER | XSE - Damaged Green Screwdriver 2021-May-04 Salesforce C00225727 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative dropped the green driver and broke it. No patient was present. 2021-05-11 IUD x2 (Other): No new information. 2021-05-24 email 2021-05-24 16:01, IUD x4 (Rep): No new information. 2021-06-18 IUD x4 (Other): No new information. 2021-07-15 IUD x3 (Other): No new information. 2021-08-04 IUD x4 (Other): No new information. 2021-Sep-01 IUD x3 (other): No new information. 2021-09-22 IUD x2 (Other): No new information. 2021-10-04 IUD (Other): No new information. 2021-10-13 IUD (Other): No new information. 2021-11-04 IUD (Other): No new information. KIT SVC MAS1650-03S UNIV SCREWDRIVER | 2021-May-04 Salesforce C00225727 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative dropped the green driver and broke it. No patient was present. 2021-05-11 IUD x2 (Other): No new information. 2021-05-24 email 2021-05-24 16:01, IUD x4 (Rep): No new information. 2021-06-18 IUD x4 (Other): No new information. 2021-07-15 IUD x3 (Other): No new information. 2021-08-04 IUD x4 (Other): No new information. 2021-Sep-01 IUD x3 (other): No new information. 2021-09-22 IUD x2 (Other): No new information. 2021-10-04 IUD (Other): No new information. 2021-10-13 IUD (Other): No new information. 2021-11-04 IUD (Other): No new information. | xse damaged green screwdriver salesforce outside procedure dropped green driver broke present iud new email iud new iud new iud new iud new iud new iud new iud new iud new iud new kit svc univ screwdriver | NAV3299 | NAV3295 | NAV2047 | NAV2058 | NAV3249 | NAV3299NAV3295NAV2047NAV2058NAV3249 | True
4070 | 703495653 | XSE - Software Message and Pneumatic | 2019-12-02T23:19:01.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE - Software Message and Pneumatic 2019-12-02 00679390 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that at every stage of the 3Define scan a software message was received stating that the movement was completed successfully. This message was also received when making small movements with the star marker. When removing the system from the bed it appeared the pneuamatics for the manipulator would go in and out along with the LEDs not functioning. Also, the Left L4 was displaying correctly as Left L4 #6. No patient was present. 2019-12-13 IUD x3 (Other): No new information. UNK MAZOR X MANIPULATOR | 2019-12-02 00679390 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that at every stage of the 3Define scan a software message was received stating that the movement was completed successfully. This message was also received when making small movements with the star marker. When removing the system from the bed it appeared the pneuamatics for the manipulator would go in and out along with the LEDs not functioning. Also, the Left L4 was displaying correctly as Left L4 #6. No patient was present. 2019-12-13 IUD x3 (Other): No new information. | xse software message pneumatic outside procedure stage scan software message stating movement completed successfully message making small movements star marker removing bed appeared pneuamatics manipulator leds functioning left displaying correctly left present iud new unk mazor x manipulator | NAV3260 | NAV3250 | NAV3264 | NAV3259 | NAV4110 | NAV3260NAV3250NAV3264NAV3259NAV4110 | True
6130 | 703871029 | X - Registration Difficulties | 2020-07-26T23:40:00.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Difficulties 2020-03-13 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that there were registration difficulties during the case. Registration was able to be successfully completed using manual registration. All screws were able to be placed successfully. There was no patient harm and the procedure was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2020-03-13 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that there were registration difficulties during the case. Registration was able to be successfully completed using manual registration. All screws were able to be placed successfully. There was no patient harm and the procedure was delayed less than an hour. | x registration difficulties ccr spinal procedure treat degenerative disc fixation registration difficulties case registration able successfully completed using manual registration screws able placed successfully procedure delayed hour mazor x | NAV3242 | NAV4112 | NAV4091 | NAV3243 | NAV1012 | NAV3242NAV4112NAV4091NAV3243NAV1012 | True
4712 | 703568482 | XSE - CT Scans not Loading | 2020-01-17T16:04:34.000000+0000 | MazorX Stealth Edition | NAV3236 | CT IMAGES | 30 | DSD0038-01 | DVD, PORTABLE DVD REWRITER, GP70NS50 | XSE - CT Scans not Loading 2020-01-16 00682251 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that while loading the CT the representative was receiving import errors from the CDs. The representative attempted to use different CDs and sometimes it worked and other times it did not. There was no patient involvement. 2020-01-21 IUD x2 (Rep, Other): Additional information received from the representative and field service engineer reporting that the system was checked in the field and it was confirmed that the problem was not with the system's CD drive but due to protocol for imaging not being followed. DVD, PORTABLE DVD REWRITER, GP70NS50 | 2020-01-16 00682251 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that while loading the CT the representative was receiving import errors from the CDs. The representative attempted to use different CDs and sometimes it worked and other times it did not. There was no patient involvement. 2020-01-21 IUD x2 (Rep, Other): Additional information received from the representative and field service engineer reporting that the system was checked in the field and it was confirmed that the problem was not with the system's CD drive but due to protocol for imaging not being followed. | xse ct scans loading use loading ct receiving import errors cds attempted use different cds worked times did iud additional field service engineer reporting checked field confirmed problem systems cd drive protocol imaging followed dvd portable dvd rewriter | NAV3257 | NAV3236 | NAV3249 | NAV4095 | NAV3295 | NAV3257NAV3236NAV3249NAV4095NAV3295 | True
7667 | 704246547 | XSE - Failed Registration/SRA Low Moveme | 2021-02-24T09:06:55.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 50 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Failed Registration/SRA Low Moveme 2021-02-23 00707162, IUD, CCR, Intake Form, e1, e2, e3, e4 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the patient already had screws and rods placed from L4 to S1. Pre-op CT scans showed that the rods were too long and went up to L3 even though no screws were placed at L3. Planning was done by the surgeon's assistant from L3 to S1. Only screws at L3 would be placed. Prior to mounting the surgical system to the patient, the rods were removed. The Schanz pin was placed in the left PSIS and the surgical system was mounted to the patient using a Schanz arm. Registration failed at L3 and there was a red cross on the verify step. The failure was due to the rods being visible in the pre-op CT and then being removed. The surgeon went back to planning and re-defined the region of interest to exclude the rods from the region and narrowing the blue square. Both implant at L3 were then planned. When going to the Operate screen, the surgeon screen and workstation monitor went black for a minute. When sending the surgical arm to left L3, the surgical arm was really low so the surgeon pressed the emergency stop button. The surgical arm was sent to the clear up position, but the arm went down into the work volume instead of going up. The surgeon decided to abort the sue of the guidance system and place the L3 screws freehand. Before unmounting the surgical system from the patient, the emergency stop button was pressed and the surgical arm elbow was manually moved out of the work volume. The surgical arm was then able to be sent to the clear up position. There was no patient harm and the procedure was delayed less than an hour. 2021-02-25 e5, e6, e7, e8 (Rep, For): Additional information received from a manufacturer representative reported that going back to exclude the rods from the region of interest, made them go back to the first step of planning while the system setup was already done. The representative questioned if this was the cause of the software issue by not doing an unmount when going back that far in the planning. The surgical system was fixed to the patient using a Schanz screw in the left PSIS connected to the bone mount bridge. The surgeon decided to do this after the rods were removed. After the case, the accuracy and 3Define camera were checked and everything was fine. 2021-03-08 e9, e10, e11, e12, e13, e14, e15 (Rep, For): No new information. 2021-05-27 IUD (Rep, For): Additional information received from a manufacturer representative reported that the probable cause was the surgical arm being lowered down too much by the user to reach short tools and the arm entered the work volume. The representative indicated that the system had reversed the sense of up and down from the zero point based on the work volume limit, which made the arm go down even with the clear up command. 2021-06-15 IUD (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2021-02-23 00707162, IUD, CCR, Intake Form, e1, e2, e3, e4 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the patient already had screws and rods placed from L4 to S1. Pre-op CT scans showed that the rods were too long and went up to L3 even though no screws were placed at L3. Planning was done by the surgeon's assistant from L3 to S1. Only screws at L3 would be placed. Prior to mounting the surgical system to the patient, the rods were removed. The Schanz pin was placed in the left PSIS and the surgical system was mounted to the patient using a Schanz arm. Registration failed at L3 and there was a red cross on the verify step. The failure was due to the rods being visible in the pre-op CT and then being removed. The surgeon went back to planning and re-defined the region of interest to exclude the rods from the region and narrowing the blue square. Both implant at L3 were then planned. When going to the Operate screen, the surgeon screen and workstation monitor went black for a minute. When sending the surgical arm to left L3, the surgical arm was really low so the surgeon pressed the emergency stop button. The surgical arm was sent to the clear up position, but the arm went down into the work volume instead of going up. The surgeon decided to abort the sue of the guidance system and place the L3 screws freehand. Before unmounting the surgical system from the patient, the emergency stop button was pressed and the surgical arm elbow was manually moved out of the work volume. The surgical arm was then able to be sent to the clear up position. There was no patient harm and the procedure was delayed less than an hour. 2021-02-25 e5, e6, e7, e8 (Rep, For): Additional information received from a manufacturer representative reported that going back to exclude the rods from the region of interest, made them go back to the first step of planning while the system setup was already done. The representative questioned if this was the cause of the software issue by not doing an unmount when going back that far in the planning. The surgical system was fixed to the patient using a Schanz screw in the left PSIS connected to the bone mount bridge. The surgeon decided to do this after the rods were removed. After the case, the accuracy and 3Define camera were checked and everything was fine. 2021-03-08 e9, e10, e11, e12, e13, e14, e15 (Rep, For): No new information. 2021-05-27 IUD (Rep, For): Additional information received from a manufacturer representative reported that the probable cause was the surgical arm being lowered down too much by the user to reach short tools and the arm entered the work volume. The representative indicated that the system had reversed the sense of up and down from the zero point based on the work volume limit, which made the arm go down even with the clear up command. 2021-06-15 IUD (Other): No new information. | xse failed registrationsra low moveme iud ccr spinal procedure treat degenerative discs fixation screws rods placed preop ct scans showed rods long went screws placed planning surgeons assistant screws placed prior mounting surgical rods removed schanz pin placed left psis surgical mounted using schanz arm registration failed red cross verify step failure rods visible preop ct removed surgeon went planning redefined region exclude rods region narrowing blue square implant planned going operate screen surgeon screen workstation monitor went black minute sending surgical arm left surgical arm really low surgeon pressed emergency stop button surgical arm sent clear position arm went work volume instead going surgeon decided abort sue place screws freehand unmounting surgical emergency stop button pressed surgical arm elbow manually moved work volume surgical arm able sent clear position procedure delayed hour additional going exclude rods region step planning setup questioned cause software issue doing unmount going far planning surgical fixed using schanz screw left psis connected bone mount bridge surgeon decided rods removed case accuracy camera checked fine new iud additional probable cause surgical arm lowered user reach short tools arm entered work volume indicated reversed sense zero point based work volume limit arm clear command iud new mazor x software installation kit | NAV3246 | NAV3247 | NAV4096 | NAV5000 | NAV3261 | NAV3246NAV3247NAV4096NAV5000NAV3261 | True
2048 | 703223954 | XSE - screen flicker/target extenders | 2019-06-24T19:19:23.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 30 | MAS2310-03 | TARGET EXTENDER | XSE - screen flicker/target extenders 2019-Jun-20 SCL1901818 SCL1901821 intake form x2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they first ran into issues with the target extenders. They openied up two sets and were not able to thread the 3d marker onto it. They opened the third set and it was able to thread properly. They also ran into constant issues of delay with the navigation projections throughout the case. They lost complete functionality of the tablet prior to registration, and the screen went completely black. The workstation monitor lagged and the mouse was jumping all over throughout the case. They also noticed the screen would flicker in and out of the actual screen. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. TARGET EXTENDER | 2019-Jun-20 SCL1901818 SCL1901821 intake form x2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they first ran into issues with the target extenders. They openied up two sets and were not able to thread the 3d marker onto it. They opened the third set and it was able to thread properly. They also ran into constant issues of delay with the navigation projections throughout the case. They lost complete functionality of the tablet prior to registration, and the screen went completely black. The workstation monitor lagged and the mouse was jumping all over throughout the case. They also noticed the screen would flicker in and out of the actual screen. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse screen flickertarget extenders spinal procedure ran issues target extenders openied sets able thread marker opened set able thread properly ran constant issues delay navigation projections case lost complete functionality tablet prior registration screen went completely black workstation monitor lagged mouse jumping case noticed screen flicker actual screen additional complications anticipated target extender | NAV3295 | NAV2054 | NAV3261 | NAV3253 | NAV4100 | NAV3295NAV2054NAV3261NAV3253NAV4100 | True
3148 | 703266628 | X - Schanz Arm Stripped Screws | 2019-07-18T20:00:10.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5750-01 | SCH_CON-SCREW | X - Schanz Arm Stripped Screws Terry Tidwell called to report that both Schanz Arms have stripped screws. At this time we are out of stock of the MEC5750-01. SCH_CON-SCREW | Terry Tidwell called to report that both Schanz Arms have stripped screws. At this time we are out of stock of the MEC5750-01. | x schanz arm stripped screws terry tidwell called report schanz arms stripped screws time stock schconscrew | NAV3299 | NAV3295 | NAV3296 | NAV4100 | NAV4091 | NAV3299NAV3295NAV3296NAV4100NAV4091 | True
6475 | 703946896 | X - SRA Didn't Reach Destination Case 1 | 2020-09-09T13:50:36.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 30 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | X - SRA Didn't Reach Destination Case 1 2020-09-08 00696076, 00696081, IUD x2 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after starting the 3Define scan, the surgical arm started to move, but stopped before or when reaching the first point. An error was displayed, which indicated "3Define Scan Failed. Rescan or Delete the Set Default Work Volume." Another 3Define scan was attempted, but the same thing occurred. The surgical arm was sent to the clear up position and a third scan was started. The 3Define scan was able to be successfully completed. At the end of the procedure, there was a pressure issue that caused the manipulator to tilt to the right. This occurred after successfully connecting the manipulator into the positioner and unmounting the surgical system from the bed. The representative had to physically support the manipulator to make sure it did not drop over the patient. After taking the system out, the issue was resolved and the surgical system was able to be stored. There was no patient harm and the procedure was delayed less than an hour. 2021-Feb-10 IUDx2 (other): No new information. UNK MAZOR X MANIPULATOR | 2020-09-08 00696076, 00696081, IUD x2 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after starting the 3Define scan, the surgical arm started to move, but stopped before or when reaching the first point. An error was displayed, which indicated "3Define Scan Failed. Rescan or Delete the Set Default Work Volume." Another 3Define scan was attempted, but the same thing occurred. The surgical arm was sent to the clear up position and a third scan was started. The 3Define scan was able to be successfully completed. At the end of the procedure, there was a pressure issue that caused the manipulator to tilt to the right. This occurred after successfully connecting the manipulator into the positioner and unmounting the surgical system from the bed. The representative had to physically support the manipulator to make sure it did not drop over the patient. After taking the system out, the issue was resolved and the surgical system was able to be stored. There was no patient harm and the procedure was delayed less than an hour. 2021-Feb-10 IUDx2 (other): No new information. | x sra didnt reach destination case iud spinal procedure starting scan surgical arm started stopped reaching point error displayed indicated scan failed rescan delete set default work volume scan attempted thing occurred surgical arm sent clear position scan started scan able successfully completed end procedure pressure issue caused manipulator tilt right occurred successfully connecting manipulator positioner unmounting surgical bed physically support manipulator make sure did drop taking issue resolved surgical able stored procedure delayed hour new unk mazor x manipulator | NAV3260 | NAV3264 | NAV3244 | NAV3239 | NAV3246 | NAV3260NAV3264NAV3244NAV3239NAV3246 | True
10279 | 704843983 | XSE - SSY Fell on Patient | 2022-02-03T01:06:00.000000+0000 | MazorX Stealth Edition | NAV4105 | UNEXPECTED POSITIONER MOVEMENT | 20 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE - SSY Fell on Patient 2022-Feb-02 Salesforce C00242647 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgical system fell and impacted the patient after guidance system was used. After the surgical arm was used during the procedure, the arm was placed into store position and the manipulator was attached to the back of the positioner. Once the manufacturer representative unlocked the teeth, they went to grab the handle underneath the surgeon screen and the manipulator with the positioner fell dropping the surgical arm on the patient. The representative manually lifted the positioner and manipulator off of the patient and the OR staff held the manipulator up so pressure could be manually released from the pneumatic panel in order to get the compressor to turn on until there was enough pressure to store the surgical system safely. There was no patient harm and the procedure was delayed less than an hour. 2022-02-03 IUD (Other): No new information. 2022-02-03 email 2022-02-03 11:16 (Rep): Additional information received from a manufacturer representative reported the surgical arm fell approximately 12-18 inches. There was no harm to the patient. 2022-02-06 e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the manipulator and positioner fell together straight down and forward. The representative noted that they had been having issues with the compressor on the system and they would often get stuck at a low pressure where the system does not know the pressure was too low to move the surgical arm. 2022-02-08 IUD x3 (Other): No new information. 2022-02-11 IUD x3 (Other): No new information. 2022-03-07 IUD x2 (Other): No new information. 2022-04-011 IUD (Other): No new information. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2022-Feb-02 Salesforce C00242647 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgical system fell and impacted the patient after guidance system was used. After the surgical arm was used during the procedure, the arm was placed into store position and the manipulator was attached to the back of the positioner. Once the manufacturer representative unlocked the teeth, they went to grab the handle underneath the surgeon screen and the manipulator with the positioner fell dropping the surgical arm on the patient. The representative manually lifted the positioner and manipulator off of the patient and the OR staff held the manipulator up so pressure could be manually released from the pneumatic panel in order to get the compressor to turn on until there was enough pressure to store the surgical system safely. There was no patient harm and the procedure was delayed less than an hour. 2022-02-03 IUD (Other): No new information. 2022-02-03 email 2022-02-03 11:16 (Rep): Additional information received from a manufacturer representative reported the surgical arm fell approximately 12-18 inches. There was no harm to the patient. 2022-02-06 e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the manipulator and positioner fell together straight down and forward. The representative noted that they had been having issues with the compressor on the system and they would often get stuck at a low pressure where the system does not know the pressure was too low to move the surgical arm. 2022-02-08 IUD x3 (Other): No new information. 2022-02-11 IUD x3 (Other): No new information. 2022-03-07 IUD x2 (Other): No new information. 2022-04-011 IUD (Other): No new information. | xse ssy fell salesforce spinal p rocedure surgical fell impacted surgical arm procedure arm placed store position manipulator attached positioner unlocked teeth went grab handle underneath surgeon screen manipulator positioner fell dropping surgical arm manually lifted positioner manipulator staff held manipulator pressure manually released pneumatic panel order compressor turn pressure store surgical safely procedure delayed hour iud new email additional surgical arm fell approximately inches additional manipulator positioner fell straight forward noted having issues compressor stuck low pressure does know pressure low surgical arm iud new iud new iud new iud new kit svc surgical new | NAV3264 | NAV4105 | NAV4101 | NAV3251 | NAV3260 | NAV3264NAV4105NAV4101NAV3251NAV3260 | True
1131 | 703190883 | X- Registration Issues | 2019-06-05T15:54:50.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X- Registration Issues Mackie was unable to achieve auto-segmentation. L5/S1 looked fused. The bottom half of S1 looked cut off which may have been the issue. Export files uploaded for case analysis since surgeon wanted to Abort. System performed as intended; 4 k-wires placed with no issue MAZOR X SOFTWARE KIT VER-3.0.2 | Mackie was unable to achieve auto-segmentation. L5/S1 looked fused. The bottom half of S1 looked cut off which may have been the issue. Export files uploaded for case analysis since surgeon wanted to Abort. System performed as intended; 4 k-wires placed with no issue | x registration issues mackie unable achieve autosegmentation looked fused half looked cut issue export files uploaded case analysis surgeon wanted abort performed intended kwires placed issue mazor x software kit | NAV3246 | NAV4112 | NAV3242 | NAV4091 | NAV3247 | NAV3246NAV4112NAV3242NAV4091NAV3247 | True
3710 | 703426811 | XSE - SW bug/arm didn't reach/mismatch | 2019-10-22T18:36:05.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 10 | TPL0059-06 | MAZOR X SYSTEM | XSE - SW bug/arm didn't reach/mismatch 2019-Oct-17 SCL1903125, SCL1903126, intake form x2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that upon initiating advanced accuracy check pre-op, the arm extended to full stretch and then reached first position but rather than getting a green confirmation message, error message was given saying “arm failed to reach final destination” (no error number). The rep selected ‘resend’ and trajectory was sent successfully. This happened again when arm proceeded to first point on the opposite side of shoulder with same result. Also, arm was noted to be ‘noisy’ during case and surgeon stated that it was ‘moving funny.’ After registration, upon dropping tools to first trajectory, Nav showed dilator being significantly superior and medial. Accuracy test performed which appeared accurate, but a new snapshot was done anyways, with no change in result. Dilator retracted slightly and still appeared superior and medial, so skiving not suspected. Black tracker was exchanged for the blue tracker, which improved result slightly but still appeared superior and medial to plan. Surgeon and the rep were confident robot was accurate and elected to proceed. Drill bit appeared accurate. There was no patient harm or additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM | 2019-Oct-17 SCL1903125, SCL1903126, intake form x2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that upon initiating advanced accuracy check pre-op, the arm extended to full stretch and then reached first position but rather than getting a green confirmation message, error message was given saying “arm failed to reach final destination” (no error number). The rep selected ‘resend’ and trajectory was sent successfully. This happened again when arm proceeded to first point on the opposite side of shoulder with same result. Also, arm was noted to be ‘noisy’ during case and surgeon stated that it was ‘moving funny.’ After registration, upon dropping tools to first trajectory, Nav showed dilator being significantly superior and medial. Accuracy test performed which appeared accurate, but a new snapshot was done anyways, with no change in result. Dilator retracted slightly and still appeared superior and medial, so skiving not suspected. Black tracker was exchanged for the blue tracker, which improved result slightly but still appeared superior and medial to plan. Surgeon and the rep were confident robot was accurate and elected to proceed. Drill bit appeared accurate. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse sw bugarm didnt reachmismatch spinal procedure initiating advanced accuracy check preop arm extended stretch reached position getting green confirmation message error message given saying “arm failed reach final destination” selected ‘resend’ trajectory sent successfully happened arm proceeded point opposite shoulder arm noted ‘noisy’ case surgeon stated ‘moving funny’ registration dropping tools trajectory nav showed dilator significantly superior medial accuracy test performed appeared accurate new snapshot anyways change dilator retracted slightly appeared superior medial skiving suspected black tracker exchanged blue tracker improved slightly appeared superior medial plan surgeon confident robot accurate elected proceed drill bit appeared accurate additional complications anticipated mazor x | NAV5000 | NAV3246 | NAV4095 | NAV4092 | NAV5001 | NAV5000NAV3246NAV4095NAV4092NAV5001 | True
5045 | 703592259 | XSE - Deviation | 2020-01-31T21:28:29.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2020-1-31 00683197 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on L4-S1. It was reported that while placing left L4, the drill was on the high speed and the surgeon went too fast into the pedicle. The surgeon backed the screw out and then put it back into the pedicle. Confirmation images showed that the screw was not in the pedicle, but superior to the drilled hole. The surgeon removed the screw and sent the arm to the previous trajectory, and the screw was implanted without further issues. The manufacturer representative felt that the drill bumped into the surgical arm throwing the trajectory off. There was no patient harm and the procedure was not delayed over an hour. 2020-02-20 email 2020-02-20 12:45 (Rep): Additional information received from a manufacturer representative reported that the amount of deviation was unknown. The surgical drill was accidentally left at 250 RPM and when the surgeon went to place the screw it slid off of the top of the lamina. The screw was removed and the system was sent back to the same trajectory. This time the drill was set to 100 RPM. The screw followed the original drill and tapped hole and was accurate. Screw placement deviation unknown exactly. Again, the powerease was accidentally left at 250rpm and when he went to place screw and it slid off top of lamina . Removed screw and sent back to same trajectory with powerease on 100 rpm. It followed the original drill and tapped hole which was accurate the entire time. MAZOR X SYSTEM TPL0059 | 2020-1-31 00683197 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on L4-S1. It was reported that while placing left L4, the drill was on the high speed and the surgeon went too fast into the pedicle. The surgeon backed the screw out and then put it back into the pedicle. Confirmation images showed that the screw was not in the pedicle, but superior to the drilled hole. The surgeon removed the screw and sent the arm to the previous trajectory, and the screw was implanted without further issues. The manufacturer representative felt that the drill bumped into the surgical arm throwing the trajectory off. There was no patient harm and the procedure was not delayed over an hour. 2020-02-20 email 2020-02-20 12:45 (Rep): Additional information received from a manufacturer representative reported that the amount of deviation was unknown. The surgical drill was accidentally left at 250 RPM and when the surgeon went to place the screw it slid off of the top of the lamina. The screw was removed and the system was sent back to the same trajectory. This time the drill was set to 100 RPM. The screw followed the original drill and tapped hole and was accurate. Screw placement deviation unknown exactly. Again, the powerease was accidentally left at 250rpm and when he went to place screw and it slid off top of lamina . Removed screw and sent back to same trajectory with powerease on 100 rpm. It followed the original drill and tapped hole which was accurate the entire time. | xse deviation spinal procedure placing left drill high speed surgeon went fast pedicle surgeon backed screw pedicle confirmation images showed screw pedicle superior drilled hole surgeon removed screw sent arm previous trajectory screw implanted issues felt drill bumped surgical arm throwing trajectory procedure delayed hour email additional deviation unknown surgical drill accidentally left rpm surgeon went place screw slid lamina screw removed sent trajectory time drill set rpm screw followed original drill tapped hole accurate screw placement deviation unknown exactly powerease accidentally left went place screw slid lamina removed screw sent trajectory powerease rpm followed original drill tapped hole accurate entire time mazor x | NAV5000 | NAV3299 | NAV3242 | NAV4090 | NAV4101 | NAV5000NAV3299NAV3242NAV4090NAV4101 | True
4819 | 703578710 | X - Registration Failure | 2020-01-23T21:08:23.000000+0000 | MazorX | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Failure 2018-03-06 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed prior to the case. The surgeon completed decompression and placed L4-L5 interbody prior to the portion of the surgery to be completed with the guidance system. The platform was mounted to the patient with a schanz pin inserted in the patient's right PSIS. The 3Define froze on the first attempt. It was restarted and successful. The Draw Spine was completed successfully. AP and OBL images were taken. Automatic segmentation failed with multiple labeling attempts and note of "low matching quality of oblique image." Due to haloing of the the end plates, the manufacturer representative requested new AP and OBL images and the surgeon obliged. New images were obtained and haloing was mitigated. The representative made multiple labeling attempts but automatic segmentation still had the same error. The representative then attempted manual registration but red values were present and registration was not successful. The surgeon requested another C-Arm to be brought in and to remove the robot from the surgical table. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2018-03-06 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed prior to the case. The surgeon completed decompression and placed L4-L5 interbody prior to the portion of the surgery to be completed with the guidance system. The platform was mounted to the patient with a schanz pin inserted in the patient's right PSIS. The 3Define froze on the first attempt. It was restarted and successful. The Draw Spine was completed successfully. AP and OBL images were taken. Automatic segmentation failed with multiple labeling attempts and note of "low matching quality of oblique image." Due to haloing of the the end plates, the manufacturer representative requested new AP and OBL images and the surgeon obliged. New images were obtained and haloing was mitigated. The representative made multiple labeling attempts but automatic segmentation still had the same error. The representative then attempted manual registration but red values were present and registration was not successful. The surgeon requested another C-Arm to be brought in and to remove the robot from the surgical table. No patient harm was reported and surgery was delayed less than an hour. | x registration failure ccr spinal procedure planning completed prior case surgeon completed decompression placed interbody prior portion surgery completed platform mounted schanz pin inserted patients right psis froze attempt restarted successful draw spine completed successfully ap obl images taken automatic segmentation failed multiple labeling attempts note low matching quality oblique image haloing end plates requested new ap obl images surgeon obliged new images obtained haloing mitigated multiple labeling attempts automatic segmentation error attempted manual registration red values present registration successful surgeon requested carm brought remove robot surgical table surgery delayed hour mazor x | NAV3242 | NAV4112 | NAV4091 | NAV2047 | NAV3244 | NAV3242NAV4112NAV4091NAV2047NAV3244 | True
6826 | 704044224 | XSE - Deviation | 2020-10-30T18:41:10.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2020-10-30 00699917, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon placed the left L4 screw and when moving to left L5 they noticed that the screw and instrumentation was about 1.5-2 cm off. A new snapshot was taken and the surgeon tried to re-verify the instrumentations, but the dilator was unable to be verified as it was far from the divot. New instruments were used and they were able to complete verification. The tap being used was accurate and the dilator was able to be successfully verified. The manufacturer representative was able to verify all instruments after the procedure. All screws and There was no patient harm and the procedure was delayed less than an hour. 2020-12-09 email 2020-12-09 16:00 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined. MAZOR X SYSTEM TPL0059 | 2020-10-30 00699917, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon placed the left L4 screw and when moving to left L5 they noticed that the screw and instrumentation was about 1.5-2 cm off. A new snapshot was taken and the surgeon tried to re-verify the instrumentations, but the dilator was unable to be verified as it was far from the divot. New instruments were used and they were able to complete verification. The tap being used was accurate and the dilator was able to be successfully verified. The manufacturer representative was able to verify all instruments after the procedure. All screws and There was no patient harm and the procedure was delayed less than an hour. 2020-12-09 email 2020-12-09 16:00 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined. | xse deviation iud spinal procedure surgeon placed left screw moving left noticed screw instrumentation cm new snapshot taken surgeon tried reverify instrumentations dilator unable verified far divot new instruments able complete verification tap accurate dilator able successfully verified able verify instruments procedure screws procedure delayed hour email additional cause determined mazor x | NAV1012 | NAV5000 | NAV3246 | NAV2054 | NAV2051 | NAV1012NAV5000NAV3246NAV2054NAV2051 | True
5120 | 703603760 | X - Missing Screws | 2020-02-07T20:14:38.000000+0000 | MazorX | NAV2053 | MISSING PARTS | 30 | MEC5192-01S | KIT SVC MEC5192-01S RNX-ACC-055 | X - Missing Screws 2020-02-07 00683677 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the screw on the bone mount bridge that screws into the shoulder and one of the screws on the reference marker were missing. This was discovered before a case with no patient present. There was no patient involvement. 2020-02-20 IUD x2 (Other): No system allegation. KIT SVC MEC5192-01S RNX-ACC-055 | 2020-02-07 00683677 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the screw on the bone mount bridge that screws into the shoulder and one of the screws on the reference marker were missing. This was discovered before a case with no patient present. There was no patient involvement. 2020-02-20 IUD x2 (Other): No system allegation. | x missing screws outside procedure screw bone mount bridge screws shoulder screws reference marker missing discovered case present iud allegation kit svc | NAV2053 | NAV3299 | NAV2054 | NAV4091 | NAV3249 | NAV2053NAV3299NAV2054NAV4091NAV3249 | True
7373 | 704192572 | XSE - SRA Failing Advanced Accuracy Test | 2021-01-25T17:34:38.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 20 | ASM0206-03S | SURGICAL ARM SPARE PART ASM0206-03S | XSE - SRA Failing Advanced Accuracy Test 2021-Jan-25 Salesforce C00219820 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a L4-L5 MIS TLIF, there was a shoulder shift after completing the left side and sending the surgical arm to the right L4 trajectory. The surgical arm was sent to the trajectory and when the surgeon placed the scalpel down the arm guide a shoulder shift occurred. The surgeon was confident that the surgical arm did not move and they did not move it so they continued with the right L4 trajectory. The surgeon decided to freehand right L5 rather than re-image and re-register the patient. The three trajectories executed with the guidance system were accurate. After the procedure, the manufacturer representative removed the system from the bed and reset the movement of the surgical arm. The surgical arm failed an advanced accuracy test. The shoulder was then recalibrated. When having the shoulder straight and in front, the calibration was about 1 in the front and back and was completed correctly. The surgical arm movement was reset and the arm still failed the accuracy test. There was no patient harm and the procedure was delayed less than an hour. 2021-01-27 IUD x2 (Other): No new information. 2021-02-24 IUD x2 (Other): No new information. 2021-04-13 IUD (Other): no new information. 2021-05-02 302418161, IUD (Other): No new information. 2021-05-04 IUD x6 (Other): No new information. 2021-06-27 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. SURGICAL ARM SPARE PART ASM0206-03S | 2021-Jan-25 Salesforce C00219820 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a L4-L5 MIS TLIF, there was a shoulder shift after completing the left side and sending the surgical arm to the right L4 trajectory. The surgical arm was sent to the trajectory and when the surgeon placed the scalpel down the arm guide a shoulder shift occurred. The surgeon was confident that the surgical arm did not move and they did not move it so they continued with the right L4 trajectory. The surgeon decided to freehand right L5 rather than re-image and re-register the patient. The three trajectories executed with the guidance system were accurate. After the procedure, the manufacturer representative removed the system from the bed and reset the movement of the surgical arm. The surgical arm failed an advanced accuracy test. The shoulder was then recalibrated. When having the shoulder straight and in front, the calibration was about 1 in the front and back and was completed correctly. The surgical arm movement was reset and the arm still failed the accuracy test. There was no patient harm and the procedure was delayed less than an hour. 2021-01-27 IUD x2 (Other): No new information. 2021-02-24 IUD x2 (Other): No new information. 2021-04-13 IUD (Other): no new information. 2021-05-02 302418161, IUD (Other): No new information. 2021-05-04 IUD x6 (Other): No new information. 2021-06-27 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. | xse sra failing advanced accuracy test salesforce spinal p rocedure mis tlif shoulder shift completing left sending surgical arm right trajectory surgical arm sent trajectory surgeon placed scalpel arm guide shoulder shift occurred surgeon confident surgical arm did did continued right trajectory surgeon decided freehand right reimage reregister trajectories executed accurate procedure removed bed reset movement surgical arm surgical arm failed advanced accuracy test shoulder recalibrated having shoulder straight calibration completed correctly surgical arm movement reset arm failed accuracy test procedure delayed hour iud new iud new iud new iud new iud new iud new iud new surgical arm spare | NAV4108 | NAV4093 | NAV3249 | NAV4101 | NAV3252 | NAV4108NAV4093NAV3249NAV4101NAV3252 | True
3354 | 703320791 | XSE - Target Extender Not Functioning | 2019-08-21T01:53:30.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2310-03 | TARGET EXTENDER | XSE - Target Extender Not Functioning 2019-08-16 SCL1902397, PDF (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the target extender was not functioning properly and it was very difficult to screw the 3D marker to the target extender. The representative also mentioned the 3Define camera was not functioning properly during a demo case. All of the cables were verified to be connected properly, but the firmware was empty in the software. There was no patient involved. 2019-10-02 e1, e2, email 2019-10-02 18:15 (Rep): Additional information received from a manufacturer representative reported that no troubleshooting was done. TARGET EXTENDER | 2019-08-16 SCL1902397, PDF (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the target extender was not functioning properly and it was very difficult to screw the 3D marker to the target extender. The representative also mentioned the 3Define camera was not functioning properly during a demo case. All of the cables were verified to be connected properly, but the firmware was empty in the software. There was no patient involved. 2019-10-02 e1, e2, email 2019-10-02 18:15 (Rep): Additional information received from a manufacturer representative reported that no troubleshooting was done. | xse target extender functioning pdf spinal procedure target extender functioning properly difficult screw marker target extender mentioned camera functioning properly demo case cables verified connected properly firmware software involved email additional troubleshooting target extender | NAV3299 | NAV3295 | NAV3244 | NAV3246 | NAV3250 | NAV3299NAV3295NAV3244NAV3246NAV3250 | True
9617 | 704683685 | XSE - Multiple Issues | 2021-10-27T02:23:53.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 30 | MEC4955-01S | KIT SVC MEC4955-01S CNCTR PANEL COV | XSE - Multiple Issues 2021-Oct-26 Salesforce C00236858 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the camera handle is loose and is not sturdy enough to move the camera. It was also reported that the system will intermittently display a controller error. Hard reboot will resolve the error but it will reoccur. Also, the front top drawer threading has come out. It is causing issues opening and closing the drawer and will likely break soon. No patient was present. 2021-10-27 IUD x6 (Other): No new information. 2021-11-08 IUD x2 (Other): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-22 IUD (Other): No new information. KIT SVC MEC4955-01S CNCTR PANEL COV | 2021-Oct-26 Salesforce C00236858 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the camera handle is loose and is not sturdy enough to move the camera. It was also reported that the system will intermittently display a controller error. Hard reboot will resolve the error but it will reoccur. Also, the front top drawer threading has come out. It is causing issues opening and closing the drawer and will likely break soon. No patient was present. 2021-10-27 IUD x6 (Other): No new information. 2021-11-08 IUD x2 (Other): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-22 IUD (Other): No new information. | xse multiple issues salesforce outside procedure camera handle loose sturdy camera intermittently display controller error hard reboot resolve error reoccur drawer threading come causing issues opening closing drawer likely break soon present iud new iud new iud new iud new kit svc cnctr panel cov | NAV2047 | NAV3244 | NAV3251 | NAV3253 | NAV2057 | NAV2047NAV3244NAV3251NAV3253NAV2057 | False
7257 | 704158437 | X - Unable To Use Disk Drive | 2021-01-06T21:25:20.000000+0000 | MazorX | NAV3257 | CD DRIVE ISSUES | 30 | KIT1241-01 | SVC KIT1241-01 DVD RW GP70NS50 | X - Unable To Use Disk Drive 2021-01-06 00703792 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the disk drive was unable to be used on the guidance system. The USB drive was used to proceed with the case. There was no patient involvement. 2021-01-13 IUD (Other): No new information. 2021-03-10 IUD x2 (Other): No new information. SVC KIT1241-01 DVD RW GP70NS50 | 2021-01-06 00703792 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the disk drive was unable to be used on the guidance system. The USB drive was used to proceed with the case. There was no patient involvement. 2021-01-13 IUD (Other): No new information. 2021-03-10 IUD x2 (Other): No new information. | x unable use disk drive outside procedure disk drive unable usb drive proceed case iud new iud new svc dvd rw | NAV3257 | NAV3250 | NAV4095 | NAV3249 | NAV3244 | NAV3257NAV3250NAV4095NAV3249NAV3244 | True
3565 | 703378274 | X - Unexpected Arm Movement | 2019-09-24T16:02:14.000000+0000 | MazorX | NAV3246 | DEFECTS | 30 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - Unexpected Arm Movement 2019-09-22 SCL1902882, Intake Form, CCR (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity using fixation. It was reported that the surgical system was docked to the patient and the surgical arm was moved to the drape position. While exposing the anatomy, the surgical arm suddenly descended to a negative degree and the nurse could move the arm without applying any force. No error messages were displayed on the workstation monitor. The representative tried to send the surgical arm to the clear up position. Sound of arm movement was heard, but the arm did not move and the monitor froze. This lasted for a few minutes. A reboot of the system was done. The representative tried to send the arm to a predefined position, but no movement occurred and an error message stating the "Arm failed to reach destination" was displayed. A reset movement was done after rebooting the system, everything was ok. An advanced accuracy test was done and the case was eventually completed successfully. There was no patient harm and the procedure was delayed for 10-15 minutes. 2019-09-29 email 2019-09-29 05:21, image 2019-09-29 05:21 x2, email 2019-10-01 23:45 (Rep): Additional information received from a m anufacturer representative reported that the surgical arm did not hit the patient. MAZOR X SOFTWARE KIT VER-4.0.0 | 2019-09-22 SCL1902882, Intake Form, CCR (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity using fixation. It was reported that the surgical system was docked to the patient and the surgical arm was moved to the drape position. While exposing the anatomy, the surgical arm suddenly descended to a negative degree and the nurse could move the arm without applying any force. No error messages were displayed on the workstation monitor. The representative tried to send the surgical arm to the clear up position. Sound of arm movement was heard, but the arm did not move and the monitor froze. This lasted for a few minutes. A reboot of the system was done. The representative tried to send the arm to a predefined position, but no movement occurred and an error message stating the "Arm failed to reach destination" was displayed. A reset movement was done after rebooting the system, everything was ok. An advanced accuracy test was done and the case was eventually completed successfully. There was no patient harm and the procedure was delayed for 10-15 minutes. 2019-09-29 email 2019-09-29 05:21, image 2019-09-29 05:21 x2, email 2019-10-01 23:45 (Rep): Additional information received from a m anufacturer representative reported that the surgical arm did not hit the patient. | x unexpected arm movement ccr spinal procedure treat deformity using fixation surgical docked surgical arm moved drape position exposing anatomy surgical arm suddenly descended negative degree nurse arm applying force error messages displayed workstation monitor tried send surgical arm clear position sound arm movement heard arm did monitor froze lasted minutes reboot tried send arm predefined position movement occurred error message stating arm failed reach destination displayed reset movement rebooting ok advanced accuracy test case eventually completed successfully procedure delayed minutes email image email additional m anufacturer surgical arm did hit mazor x software kit | NAV4095 | NAV3246 | NAV3248 | NAV4101 | NAV2047 | NAV4095NAV3246NAV3248NAV4101NAV2047 | True
8514 | 704430661 | XSE - Screw Deviation | 2021-06-03T23:02:14.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Screw Deviation 2021-Jun-03 Salesforce C00227803 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there was a deviated screw during the case. There was no patient harm and the procedure was delayed less than an hour. 2021-06-25 email 6/25 (rep): No new information. 2021-08-09 email 2021-08-09 11:16 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2021-Jun-03 Salesforce C00227803 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there was a deviated screw during the case. There was no patient harm and the procedure was delayed less than an hour. 2021-06-25 email 6/25 (rep): No new information. 2021-08-09 email 2021-08-09 11:16 (Rep): No new information. | xse screw deviation salesforce spinal p rocedure deviated screw case procedure delayed hour email new email new mazor x | NAV5000 | NAV3242 | NAV3237 | NAV4112 | NAV4090 | NAV5000NAV3242NAV3237NAV4112NAV4090 | True
2636 | 703245698 | X - Workstation Problems | 2019-07-06T14:37:55.000000+0000 | MazorX | NAV2053 | MISSING PARTS | 20 | MEC0625-01 | IRON MAN. | X - Workstation Problems - Missing Iron men - Monitor elevator causes friction when operating it both up and down. - Joint 3 slips when EMB is pressed and SRA motion is halted. - Clinical team reported 3define camera malfunctioned. Error message: "surface scanner not found" IRON MAN. | - Missing Iron men - Monitor elevator causes friction when operating it both up and down. - Joint 3 slips when EMB is pressed and SRA motion is halted. - Clinical team reported 3define camera malfunctioned. Error message: "surface scanner not found" | x workstation problems missing iron men monitor elevator causes friction operating joint slips emb pressed sra motion halted clinical team camera malfunctioned error message surface scanner iron man | NAV3261 | NAV2054 | NAV4091 | NAV4095 | NAV2058 | NAV3261NAV2054NAV4091NAV4095NAV2058 | False
3036 | 703263642 | X - Arm Issues | 2019-07-17T16:02:29.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | X - Arm Issues Monitor separated during shipping causing touchscreen to be non-functional. The slides were also broken. SURGICAL ARM | Monitor separated during shipping causing touchscreen to be non-functional. The slides were also broken. | x arm issues monitor separated shipping causing touchscreen nonfunctional slides broken surgical arm | NAV3261 | NAV3249 | NAV2055 | NAV4101 | NAV3295 | NAV3261NAV3249NAV2055NAV4101NAV3295 | False
6878 | 704054295 | XSE - OOBF/3Define Camera Issue | 2020-11-05T18:44:04.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - OOBF/3Define Camera Issue 2020-Nov-05 00700304 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the 3Define camera failed on the replacement surgical arm that was shipped for the repair. There were multiple troubleshooting steps taken and the red and green cable within the arm is damaged. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-11-08 IUD x 10 (Other): No new information. 2020-12-04 IUD x6 (Other): No new information 2021-01-04 IUD x4 (Other): No new information. 2021-02-01 IUD (Other): No new information. 2021-02-04 IUD (Other): No new information. 2021-04-09 IUD (Other): No new information. 2021-06-01 IUD x4 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-Nov-05 00700304 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the 3Define camera failed on the replacement surgical arm that was shipped for the repair. There were multiple troubleshooting steps taken and the red and green cable within the arm is damaged. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-11-08 IUD x 10 (Other): No new information. 2020-12-04 IUD x6 (Other): No new information 2021-01-04 IUD x4 (Other): No new information. 2021-02-01 IUD (Other): No new information. 2021-02-04 IUD (Other): No new information. 2021-04-09 IUD (Other): No new information. 2021-06-01 IUD x4 (Other): No new information. | xse camera issue outside procedure camera failed replacement surgical arm shipped repair multiple troubleshooting steps taken red green cable arm damaged additional complications anticipated iud x new iud new iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV2054 | NAV3244 | NAV1006 | NAV3256 | NAV3249NAV2054NAV3244NAV1006NAV3256 | True
3368 | 703327151 | XSE - Failed Accuracy Test | 2019-08-23T21:39:27.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 30 | ASM0206-01 | SURGICAL ARM | XSE - Failed Accuracy Test 2019-08-20 SCL1902423, SCL1902443, Intake Form, video 2019-08-28 08:36 (Rep): Medtronic received information regarding a guidance system being used prior to a surgical procedure. It was reported that the surgical arm failed operational setup for replacement. The surgical arm was replaced, but after installation the arm failed accuracy testing. Joint 6 of the surgical arm was able to be moved causing the arm to be inaccurate. There was no patient harm reported. 2019-09-16 email 2019-09-16 14:00 (Rep): No new information. 2019-09-20 email 2019-09-20 15:30 (Rep): Additional information received from a manufacturer representative reported there was no delay to the procedure and there was no patient harm. The surgery was completed freehand with navigation. SURGICAL ARM | 2019-08-20 SCL1902423, SCL1902443, Intake Form, video 2019-08-28 08:36 (Rep): Medtronic received information regarding a guidance system being used prior to a surgical procedure. It was reported that the surgical arm failed operational setup for replacement. The surgical arm was replaced, but after installation the arm failed accuracy testing. Joint 6 of the surgical arm was able to be moved causing the arm to be inaccurate. There was no patient harm reported. 2019-09-16 email 2019-09-16 14:00 (Rep): No new information. 2019-09-20 email 2019-09-20 15:30 (Rep): Additional information received from a manufacturer representative reported there was no delay to the procedure and there was no patient harm. The surgery was completed freehand with navigation. | xse failed accuracy test video prior surgical procedure surgical arm failed operational setup replacement surgical arm replaced installation arm failed accuracy testing joint surgical arm able moved causing arm inaccurate email new email additional delay procedure surgery completed freehand navigation surgical arm | NAV4092 | NAV2054 | NAV3249 | NAV4097 | NAV4101 | NAV4092NAV2054NAV3249NAV4097NAV4101 | True
3959 | 703475871 | XSE - Pressure Issues and Noise | 2019-11-19T20:11:37.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Pressure Issues and Noise 2019-11-19 00678676, ID (Rep) Medtronic received information regarding a guidance system. It was reported by the manufacturer representative that the was continuously building and then releasing pressure. When the pressure was released the system made a loud noise. However, pneumatics were not lost at any time. No patient was involved. UNK MAZOR X WORKSTATION | 2019-11-19 00678676, ID (Rep) Medtronic received information regarding a guidance system. It was reported by the manufacturer representative that the was continuously building and then releasing pressure. When the pressure was released the system made a loud noise. However, pneumatics were not lost at any time. No patient was involved. | xse pressure issues noise id continuously building releasing pressure pressure released loud noise pneumatics lost time involved unk mazor x workstation | NAV3264 | NAV2053 | NAV2054 | NAV2047 | NAV4095 | NAV3264NAV2053NAV2054NAV2047NAV4095 | True
8905 | 704510008 | XSE - Damaged Instruments | 2021-07-20T00:10:25.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2271 | KIT SVC MAS2271 SCHANZ BOND 05S NEW | XSE - Damaged Instruments 2021-Jul-19 Salesforce C00230423 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw for the alligator clamp became stripped and was attached to the Schanz bond. The manufacturer representative was unable to remove the screw. There was no patient involvement. 2021-Aug-02 IUDx3 (other): No new information. 2021-Aug-30 IUD (other): No new information. KIT SVC MAS2271 SCHANZ BOND 05S NEW | 2021-Jul-19 Salesforce C00230423 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw for the alligator clamp became stripped and was attached to the Schanz bond. The manufacturer representative was unable to remove the screw. There was no patient involvement. 2021-Aug-02 IUDx3 (other): No new information. 2021-Aug-30 IUD (other): No new information. | xse damaged instruments salesforce outside procedure screw alligator clamp stripped attached schanz bond unable remove screw new iud new kit svc schanz bond new | NAV3299 | NAV4090 | NAV3256 | NAV3264 | NAV3251 | NAV3299NAV4090NAV3256NAV3264NAV3251 | True
5399 | 703673185 | XSE - Positioner Plate Handle Button | 2020-03-18T18:21:30.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 20 | KIT0934-01 | SVC KIT0934-01 FRONT PLATE BUTTONS LEFT | XSE - Positioner Plate Handle Button 2020-03-18 00686570 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the field service engineer requested a left positioner handle button. There was no patient involvement. 2020-03-24 IUD x4 (Other): No new information. 2020-04-07 Email 2020-04-07 12:00, Email 2020-04-07 12:00, IUD (Rep, Other): Additional information received from a manufacturer rep resentative reported that the wiring for the button was broken off. It was hanging loose until the engineer could replace it. A new button was ordered to replace the broken one. 2020-04-21 IUD x5 (Other): No new information. SVC KIT0934-01 FRONT PLATE BUTTONS LEFT | 2020-03-18 00686570 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the field service engineer requested a left positioner handle button. There was no patient involvement. 2020-03-24 IUD x4 (Other): No new information. 2020-04-07 Email 2020-04-07 12:00, Email 2020-04-07 12:00, IUD (Rep, Other): Additional information received from a manufacturer rep resentative reported that the wiring for the button was broken off. It was hanging loose until the engineer could replace it. A new button was ordered to replace the broken one. 2020-04-21 IUD x5 (Other): No new information. | xse positioner plate handle button outside procedure field service engineer requested left positioner handle button iud new email email iud additional resentative wiring button broken hanging loose engineer replace new button ordered replace broken iud new svc plate buttons left | NAV3251 | NAV2058 | NAV3252 | NAV2054 | NAV3249 | NAV3251NAV2058NAV3252NAV2054NAV3249 | True
10320 | 704853444 | XSE - SRA Errors | 2022-02-08T21:57:00.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 30 | UK |  | XSE - SRA Errors 2022-Feb-08 Salesforce C00242897 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when sending the surgical arm to the first trajectory during the procedure, an 4059: Joint Shift Error at J5 was displayed followed by Error 4120: Arm Did Not Reach Trajectory error. A soft reboot was done, but the issue was not reso lved. The surgeon decided to abort the use of the guidance system and complete the procedure with navigation. There was no patient harm and the procedure was delayed over an hour. 2022-02-14 IUD x3 (Other): No new information. 2022-03-13 IUD (Other): No new information. 2022-04-10 302780219 (Other): No new information. 2022-04-12 IUD (Other): No new information. | 2022-Feb-08 Salesforce C00242897 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when sending the surgical arm to the first trajectory during the procedure, an 4059: Joint Shift Error at J5 was displayed followed by Error 4120: Arm Did Not Reach Trajectory error. A soft reboot was done, but the issue was not reso lved. The surgeon decided to abort the use of the guidance system and complete the procedure with navigation. There was no patient harm and the procedure was delayed over an hour. 2022-02-14 IUD x3 (Other): No new information. 2022-03-13 IUD (Other): No new information. 2022-04-10 302780219 (Other): No new information. 2022-04-12 IUD (Other): No new information. | xse sra errors salesforce spinal p rocedure sending surgical arm trajectory procedure joint shift error displayed followed error arm did reach trajectory error soft reboot issue reso lved surgeon decided abort use complete procedure navigation procedure delayed hour iud new iud new new iud new | NAV4095 | NAV4098 | NAV2047 | NAV3249 | NAV1006 | NAV4095NAV4098NAV2047NAV3249NAV1006 | True
8542 | 704435899 | XSE - System Unresponsive | 2021-06-07T23:04:37.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - System Unresponsive 2021-Jun-04 Salesforce C00227961 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was shutdown for over an hour between two cases. The system was booted up and prepared for the second case, but a 3Define timed out error was displayed. A hard restart was done to attempt to the clear the error, but the entire system became unresponsive. The manufacturer representative was unable to open or close the workstation hood. There was no patient involvement. UNK MAZOR X WORKSTATION | 2021-Jun-04 Salesforce C00227961 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was shutdown for over an hour between two cases. The system was booted up and prepared for the second case, but a 3Define timed out error was displayed. A hard restart was done to attempt to the clear the error, but the entire system became unresponsive. The manufacturer representative was unable to open or close the workstation hood. There was no patient involvement. | xse unresponsive salesforce outside procedure shutdown hour cases booted prepared second case timed error displayed hard restart attempt clear error entire unresponsive unable open close workstation hood unk mazor x workstation | NAV2058 | NAV2047 | NAV3264 | NAV3262 | NAV4110 | NAV2058NAV2047NAV3264NAV3262NAV4110 | True
2760 | 703249219 | X - Accessories Issue | 2019-07-09T16:13:07.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 70 | MAS1248-03 | SHORT DRILL GUIDE | X - Accessories Issue SC opened following Elad report in attached mail regarding tools related issues: "Marker extender (carbon)- 3D Marker screw was much too tight. Short Drill bit got stuck in the sort drill guide. Long drill bit got stuck in the long drill guide. Both were released using pliers and lots of force. Long drill guide got damaged (bent teeth), as well as both drill bits. RNS arm tube (14-8) was used with RNS tools. Cannulas had too much freedom in the arm tube. New Arm Tube needs to be tested and provided. Accuracy pointer is missing from the surgical set. Needs to be provided. Tablet adaptor (tablet side) too tight on the tablet side while on the drape - hard to tighten the screws. SHORT DRILL GUIDE | SC opened following Elad report in attached mail regarding tools related issues: "Marker extender (carbon)- 3D Marker screw was much too tight. Short Drill bit got stuck in the sort drill guide. Long drill bit got stuck in the long drill guide. Both were released using pliers and lots of force. Long drill guide got damaged (bent teeth), as well as both drill bits. RNS arm tube (14-8) was used with RNS tools. Cannulas had too much freedom in the arm tube. New Arm Tube needs to be tested and provided. Accuracy pointer is missing from the surgical set. Needs to be provided. Tablet adaptor (tablet side) too tight on the tablet side while on the drape - hard to tighten the screws. | x accessories issue sc opened following elad report attached mail tools related issues marker extender marker screw tight short drill bit got stuck sort drill guide long drill bit got stuck long drill guide released using pliers lots force long drill guide got damaged drill bits rns arm tube rns tools cannulas freedom arm tube new arm tube needs tested provided accuracy pointer missing surgical set needs provided tablet adaptor tight tablet drape hard tighten screws short drill guide | NAV3299 | NAV3295 | NAV2053 | NAV3253 | NAV3297 | NAV3299NAV3295NAV2053NAV3253NAV3297 | True
4491 | 703546312 | X - Deviation/Skive Right L4 | 2020-01-06T20:16:25.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation/Skive Right L4 2018-06-04 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed by the manufacturer representative prior to surgery for a TLIF at L4-L5 and the surgeon agreed with the plan. The representative notified the surgeon of several potential skive issues. The representative also noted the patient's high BMI of 56. The system cameras were checked, the advanced accuracy test was done, and the shoulder calibration was checked for arm set-up. It took longer than usual to position the patient and provide sufficient airway clearance. The surgeon and and representative discussed multiple back-up solutions for the schanz pin if the 120mm pin was not long enough. The representative also discussed back-up plan with the surgeon in case registration was not possible. The arm was attached to the bed and locked at 15 degrees. The schanz pin stuck out of the patient's skin just enough to attach the ball joint. While taking the AP image, the radiology technician accidentally manipulated the image digitally. The representative had the rad tech re-calibrate the C-Arm. New AP and OBL images were taken and registration was achieved on the first try with green values at all surgical levels; the surgeon visually verified. The first trajectory sent was right L4 which was a challenging trajectory due to the high skive potential. The surgeon skived multiple times, with the representative adjusting the plan each time to help counteract the skive potential. After several failed attempts to place accurately, the representative recommended moving on to other trajectories and in the meantime would work on a better trajectory for right L4. The surgeon then drilled and placed k-wires accurately at the other three trajectories and returned to right L4 after the representative had significantly altered the plan. The surgeon felt the trajectory looked good and drilled accurately. The surgeon was pleased with the use of the guidance system and the overall speed and accuracy noting that the images were hard to see but that was easily overcome with the robot. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2018-06-04 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed by the manufacturer representative prior to surgery for a TLIF at L4-L5 and the surgeon agreed with the plan. The representative notified the surgeon of several potential skive issues. The representative also noted the patient's high BMI of 56. The system cameras were checked, the advanced accuracy test was done, and the shoulder calibration was checked for arm set-up. It took longer than usual to position the patient and provide sufficient airway clearance. The surgeon and and representative discussed multiple back-up solutions for the schanz pin if the 120mm pin was not long enough. The representative also discussed back-up plan with the surgeon in case registration was not possible. The arm was attached to the bed and locked at 15 degrees. The schanz pin stuck out of the patient's skin just enough to attach the ball joint. While taking the AP image, the radiology technician accidentally manipulated the image digitally. The representative had the rad tech re-calibrate the C-Arm. New AP and OBL images were taken and registration was achieved on the first try with green values at all surgical levels; the surgeon visually verified. The first trajectory sent was right L4 which was a challenging trajectory due to the high skive potential. The surgeon skived multiple times, with the representative adjusting the plan each time to help counteract the skive potential. After several failed attempts to place accurately, the representative recommended moving on to other trajectories and in the meantime would work on a better trajectory for right L4. The surgeon then drilled and placed k-wires accurately at the other three trajectories and returned to right L4 after the representative had significantly altered the plan. The surgeon felt the trajectory looked good and drilled accurately. The surgeon was pleased with the use of the guidance system and the overall speed and accuracy noting that the images were hard to see but that was easily overcome with the robot. No patient harm was reported and surgery was delayed less than an hour. | x deviationskive right ccr spinal procedure planning completed prior surgery tlif surgeon agreed plan notified surgeon potential skive issues noted patients high bmi cameras checked advanced accuracy test shoulder calibration checked arm setup took longer usual position provide sufficient airway clearance surgeon discussed multiple backup solutions schanz pin pin long discussed backup plan surgeon case registration possible arm attached bed locked degrees schanz pin stuck patients skin just attach ball joint taking ap image radiology technician accidentally manipulated image digitally rad tech recalibrate carm new ap obl images taken registration achieved try green values surgical levels surgeon visually verified trajectory sent right challenging trajectory high skive potential surgeon skived multiple times adjusting plan time help counteract skive potential failed attempts place accurately recommended moving trajectories meantime work better trajectory right surgeon drilled placed kwires accurately trajectories returned right significantly altered plan surgeon felt trajectory looked good drilled accurately surgeon pleased use overall speed accuracy noting images hard easily overcome robot surgery delayed hour mazor x | NAV4091 | NAV3242 | NAV3243 | NAV2051 | NAV5001 | NAV4091NAV3242NAV3243NAV2051NAV5001 | True
6989 | 704092606 | XSE - Positioner Issue in Z Direction | 2020-11-26T22:55:18.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 40 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Positioner Issue in Z Direction 2020-11-26 00701610, 00702672, IUD x5 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was shipped and after being unpacked, the positioner was not in the storage position anymore. The over ride panel in the software was used to open the Z axis and put the positioner in the correct storage position. After mounting and unmounting the surgical system to the bed several time, the positioner dropped in the Z direction and couldn't be put back even when using the over ride. The surgical system was not working, the wires in the surgical arm were broken and the main cable was not working. There was no patient involvement. 2021-02-09 email 2021-02-09 06:16 (Rep, For): Additional information received from a manufacturer representative reported that service could not be completed due to COVID-19 restrictions. 2021-05-28 00714003, IUD x4 (Other, For): No new information. 2021-06-08 IUD x3 (Other): No new information. 2021-Aug-16 IUD (other): No new information. 2021-Sep-02 IUD (other): No new information. 2021-09-29 IUD x3 (Other): No new information. 2021-11-11 IUD x2 (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-11-26 00701610, 00702672, IUD x5 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was shipped and after being unpacked, the positioner was not in the storage position anymore. The over ride panel in the software was used to open the Z axis and put the positioner in the correct storage position. After mounting and unmounting the surgical system to the bed several time, the positioner dropped in the Z direction and couldn't be put back even when using the over ride. The surgical system was not working, the wires in the surgical arm were broken and the main cable was not working. There was no patient involvement. 2021-02-09 email 2021-02-09 06:16 (Rep, For): Additional information received from a manufacturer representative reported that service could not be completed due to COVID-19 restrictions. 2021-05-28 00714003, IUD x4 (Other, For): No new information. 2021-06-08 IUD x3 (Other): No new information. 2021-Aug-16 IUD (other): No new information. 2021-Sep-02 IUD (other): No new information. 2021-09-29 IUD x3 (Other): No new information. 2021-11-11 IUD x2 (Other): No new information. | xse positioner issue z direction iud outside procedure shipped unpacked positioner storage position anymore ride panel software open z axis positioner correct storage position mounting unmounting surgical bed time positioner dropped z direction using ride surgical working wires surgical arm broken main cable working email additional service completed restrictions iud new iud new iud new iud new iud new iud new kit svc surgical arm | NAV4104 | NAV4105 | NAV3249 | NAV2055 | NAV2047 | NAV4104NAV4105NAV3249NAV2055NAV2047 | True
6533 | 703969234 | X - Pressure Loss (3 of 3) | 2020-09-22T00:00:08.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 40 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | X - Pressure Loss (3 of 3) 2020-09-21 00697159 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that at the end of the case the system was losing pressure when the manufacturer representative was going to take the arm off of the bed. The pressure would drop enough that she could get the positioner off, but could not close the cabinet. She was able to get it closed and stored eventually. No patient was present at this time. 2020-10-06 IUD x2 (Other): No new information. 2020-10-14 IUD x5 (Other): No new information. 2020-11-02 IUD x3 (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2020-09-21 00697159 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that at the end of the case the system was losing pressure when the manufacturer representative was going to take the arm off of the bed. The pressure would drop enough that she could get the positioner off, but could not close the cabinet. She was able to get it closed and stored eventually. No patient was present at this time. 2020-10-06 IUD x2 (Other): No new information. 2020-10-14 IUD x5 (Other): No new information. 2020-11-02 IUD x3 (Other): No new information. | x pressure loss outside procedure end case losing pressure going arm bed pressure drop positioner close cabinet able closed stored eventually present time iud new iud new iud new kit svc surgical syst type | NAV3264 | NAV4101 | NAV3250 | NAV3251 | NAV2054 | NAV3264NAV4101NAV3250NAV3251NAV2054 | True
4222 | 703515724 | X - Registration Issues, Deviation Right | 2019-12-12T19:08:48.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issues, Deviation Right 2018-05-29 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat trauma with fixation at T12-L3. It was reported that setup included: shoulder calibration check, advanced accuracy test, and surgical arm setup. A Schanz screw was placed in the right PSIS along with a BMB and Schanz bond. A 3Define scan and DrawSpine were completed. AP and multiple OBL images were taken. Fluoro segmentation was done using the basic algorithm, it did not include L3. A new OBL image was obtained. Fluoro imaging was done again with the basic algorithm and revalues were seen at L1. The monitoring wires were moved and OBL images were taken again. Fluoro segmentation was done again with the basic algorithm and green values were seen at all levels. The surgeon visually verified matching accuracies at all surgical levels. Bilateral L1 and left L3 were placed accurately. Right L3 was placed and it was deviated superior due to soft tissue pressure. Right L2 was sent to determine if there was a shift if the patient shifted. It was determined that there was a change in the patient positioning, so use of the guidance system was aborted. Right L3 and bilateral L2 were placed freehand. Upon reflection, the surgeon states that there was some segmental instability and in the future they would use a dual clamp in similar situations. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-05-29 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat trauma with fixation at T12-L3. It was reported that setup included: shoulder calibration check, advanced accuracy test, and surgical arm setup. A Schanz screw was placed in the right PSIS along with a BMB and Schanz bond. A 3Define scan and DrawSpine were completed. AP and multiple OBL images were taken. Fluoro segmentation was done using the basic algorithm, it did not include L3. A new OBL image was obtained. Fluoro imaging was done again with the basic algorithm and revalues were seen at L1. The monitoring wires were moved and OBL images were taken again. Fluoro segmentation was done again with the basic algorithm and green values were seen at all levels. The surgeon visually verified matching accuracies at all surgical levels. Bilateral L1 and left L3 were placed accurately. Right L3 was placed and it was deviated superior due to soft tissue pressure. Right L2 was sent to determine if there was a shift if the patient shifted. It was determined that there was a change in the patient positioning, so use of the guidance system was aborted. Right L3 and bilateral L2 were placed freehand. Upon reflection, the surgeon states that there was some segmental instability and in the future they would use a dual clamp in similar situations. There was no patient harm and the procedure was not delayed over an hour. | x registration issues deviation right ccr spinal procedure treat trauma fixation setup included shoulder calibration check advanced accuracy test surgical arm setup schanz screw placed right psis bmb schanz bond scan drawspine completed ap multiple obl images taken fluoro segmentation using basic algorithm did include new obl image obtained fluoro imaging basic algorithm revalues seen monitoring wires moved obl images taken fluoro segmentation basic algorithm green values seen levels surgeon visually verified matching accuracies surgical levels bilateral left placed accurately right placed deviated superior soft tissue pressure right sent determine shift shifted determined change positioning use aborted right bilateral placed freehand reflection surgeon states segmental instability future use dual clamp similar situations procedure delayed hour mazor x | NAV4091 | NAV3242 | NAV4108 | NAV3251 | NAV2051 | NAV4091NAV3242NAV4108NAV3251NAV2051 | True
6793 | 704032357 | XSE - Image Transfer Issue | 2020-10-24T18:37:03.000000+0000 | MazorX Stealth Edition | NAV3237 | FLUORO IMAGES | 30 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - Image Transfer Issue 2020-10-22 00699449 (Rep, For): Medtronic received information regarding a guidance system outside of procedure, during set up. It was reported that there was an image transfer issue from the C-arm to the guidance system. The guidance system was not able to be integrated with C-Arm system. Troubleshooting involved checking the frame grabber and it was found that live images were not transferring from C-arm to the guidance system. The driver details were then checked and they were not able to find video capture card driver in the device manager. The probable cause was stated to be due to the PC needing to be replaced since there was an issue with the video capture card inside of the PC. No patient was present. 2021-01-16 IUD, email 2021-01-16 00:15 (Rep, For): Additional information received from a manufacturer representative reported that the PC was replaced to resolve the issue. ASM0215-06S MAZOR X 5820 PC STATION | 2020-10-22 00699449 (Rep, For): Medtronic received information regarding a guidance system outside of procedure, during set up. It was reported that there was an image transfer issue from the C-arm to the guidance system. The guidance system was not able to be integrated with C-Arm system. Troubleshooting involved checking the frame grabber and it was found that live images were not transferring from C-arm to the guidance system. The driver details were then checked and they were not able to find video capture card driver in the device manager. The probable cause was stated to be due to the PC needing to be replaced since there was an issue with the video capture card inside of the PC. No patient was present. 2021-01-16 IUD, email 2021-01-16 00:15 (Rep, For): Additional information received from a manufacturer representative reported that the PC was replaced to resolve the issue. | xse image transfer issue outside procedure set image transfer issue carm able integrated carm troubleshooting involved checking frame grabber live images transferring carm driver details checked able video capture card driver manager probable cause stated pc needing replaced issue video capture card inside pc present iud email additional pc replaced resolve issue mazor x pc station | NAV3262 | NAV3237 | NAV3246 | NAV2054 | NAV3295 | NAV3262NAV3237NAV3246NAV2054NAV3295 | True
3386 | 703330081 | XSE - Surgical System Not Locking/OOBF | 2019-08-26T21:58:20.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Surgical System Not Locking/OOBF 2019-08-21 SCL1902428, PDF (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that there was an intermittent issue with the surgical system. When going into operate mode, the system showed a message to "Please Lock the Shoulder", even though the shoulder was locked and the logic sensor was showing the shoulder was also locked. The representative tried moving the shoulder around to see if it happened between a state when they were locking and locking the shoulder, but the issue was no recreated. The system was able to go into operating mode without the issue occurring again. There was no patient involvement. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-08-21 SCL1902428, PDF (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that there was an intermittent issue with the surgical system. When going into operate mode, the system showed a message to "Please Lock the Shoulder", even though the shoulder was locked and the logic sensor was showing the shoulder was also locked. The representative tried moving the shoulder around to see if it happened between a state when they were locking and locking the shoulder, but the issue was no recreated. The system was able to go into operating mode without the issue occurring again. There was no patient involvement. | xse surgical lockingoobf pdf prior procedure intermittent issue surgical going operate mode showed message lock shoulder shoulder locked logic sensor showing shoulder locked tried moving shoulder happened state locking locking shoulder issue recreated able operating mode issue occurring mazor x software installation kit | NAV4107 | NAV2054 | NAV3246 | NAV4106 | NAV4103 | NAV4107NAV2054NAV3246NAV4106NAV4103 | True
2719 | 703247330 | X - Arm joint shifts/CANBUS error | 2019-07-08T20:54:56.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | X - Arm joint shifts/CANBUS error Malfunction occuring on Arm #12 Arm is shaking during several movement, and the issue should be inspected. Arm Shift Detected on two different occasions 1 st time Arm Shift detected in Joint #2 (Happened twice) – 79th Procedure @ SHHEC. 2 nd time Arm Shift detected in joint #4 (once) followed by Fatal Error: CANBUS connection error message – 80th procedure @ SHHEC SURGICAL ARM | Malfunction occuring on Arm #12 Arm is shaking during several movement, and the issue should be inspected. Arm Shift Detected on two different occasions 1 st time Arm Shift detected in Joint #2 (Happened twice) – 79th Procedure @ SHHEC. 2 nd time Arm Shift detected in joint #4 (once) followed by Fatal Error: CANBUS connection error message – 80th procedure @ SHHEC | x arm joint shiftscanbus error malfunction occuring arm arm shaking movement issue inspected arm shift detected different occasions st time arm shift detected joint – procedure shhec nd time arm shift detected joint followed fatal error canbus connection error message – procedure shhec surgical arm | NAV4098 | NAV4101 | NAV4095 | NAV2054 | NAV3242 | NAV4098NAV4101NAV4095NAV2054NAV3242 | True
341 | 703078844 | X - Deviations | 2019-03-28T17:21:50.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviations 2019-03-21 SCL1900734, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that L4-S1 was registered and then the surgeon began drilling and placing guidewires on the patients left side at L4, L5 and S1. All three of the trajectories were off and were observed to be too low (outside of the pedicle) when checked with fluoroscopy. The manufacturer representative indicated there was difficulties with registration and there were minor accuracy/neuro-monitoring issues. Registration was able to be repeated and the trajectories were re-drilled. All trajectories then looked accurate. There was no patient harm and the surgery was not prolonged more than an hour. The case was completed. 2019-04-03 e1 (Rep): Additional information received from a manufacturer representative reported the patient underwent a posterior spinal fusion of L4, L5 and S1. The patient was in the prone position during the procedure. After an x-ray was taken on the left sided trajectories, the surgeon determined that all trajectories were low so another registration was done. The cause was determined to be an arm or patient shift since all trajectories on the right side were accurate after registration was redone. 2019-04-10 e2 (Rep): Additional information received from a manufacturer representative reported the registration difficulties were referring to the need to re-register after the trajectories were found to be inferior. The representative clarified there were no neuro-monitoring issues. MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-03-21 SCL1900734, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that L4-S1 was registered and then the surgeon began drilling and placing guidewires on the patients left side at L4, L5 and S1. All three of the trajectories were off and were observed to be too low (outside of the pedicle) when checked with fluoroscopy. The manufacturer representative indicated there was difficulties with registration and there were minor accuracy/neuro-monitoring issues. Registration was able to be repeated and the trajectories were re-drilled. All trajectories then looked accurate. There was no patient harm and the surgery was not prolonged more than an hour. The case was completed. 2019-04-03 e1 (Rep): Additional information received from a manufacturer representative reported the patient underwent a posterior spinal fusion of L4, L5 and S1. The patient was in the prone position during the procedure. After an x-ray was taken on the left sided trajectories, the surgeon determined that all trajectories were low so another registration was done. The cause was determined to be an arm or patient shift since all trajectories on the right side were accurate after registration was redone. 2019-04-10 e2 (Rep): Additional information received from a manufacturer representative reported the registration difficulties were referring to the need to re-register after the trajectories were found to be inferior. The representative clarified there were no neuro-monitoring issues. | x deviations ccr spinal procedure treat degenerative disc fixation registered surgeon began drilling placing guidewires patients left trajectories observed low checked fluoroscopy indicated difficulties registration minor accuracyneuromonitoring issues registration able repeated trajectories redrilled trajectories looked accurate surgery prolonged hour case completed additional underwent posterior spinal fusion prone position procedure xray taken left sided trajectories surgeon determined trajectories low registration cause determined arm shift trajectories right accurate registration redone additional registration difficulties referring need reregister trajectories inferior clarified neuromonitoring issues mazor x software kit | NAV3242 | NAV5000 | NAV4091 | NAV3243 | NAV3248 | NAV3242NAV5000NAV4091NAV3243NAV3248 | True
1768 | 703213638 | X- Arm issue | 2019-06-18T20:27:40.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X- Arm issue Aborted Case due to J1 error. See steps taken by CSR below: CSR received a joint one error. CSR went in to super user, service tools and reset movement which did not fix the issue. CSR did a soft then hard restart which failed to fix issue. CSR was told by field engineer this is a know issue and that usually fixes the issue, after multiple attempts the surgeon aborted use of the rbt. CSR re calibrated shoulder which didn't help. CSR took the system off the bed but the arm would not go to storage. CSR shut the system down and took it to the hall to continue trying to fix. After restart and two more attempts to reset movement the system worked. UNK MAZOR X ARM | Aborted Case due to J1 error. See steps taken by CSR below: CSR received a joint one error. CSR went in to super user, service tools and reset movement which did not fix the issue. CSR did a soft then hard restart which failed to fix issue. CSR was told by field engineer this is a know issue and that usually fixes the issue, after multiple attempts the surgeon aborted use of the rbt. CSR re calibrated shoulder which didn't help. CSR took the system off the bed but the arm would not go to storage. CSR shut the system down and took it to the hall to continue trying to fix. After restart and two more attempts to reset movement the system worked. | x arm issue aborted case error steps taken csr csr joint error csr went super user service tools reset movement did fix issue csr did soft hard restart failed fix issue csr told field engineer know issue usually fixes issue multiple attempts surgeon aborted use rbt csr calibrated shoulder didnt help csr took bed arm storage csr shut took hall continue trying fix restart attempts reset movement worked unk mazor x arm | NAV3249 | NAV4107 | NAV3263 | NAV4101 | NAV2058 | NAV3249NAV4107NAV3263NAV4101NAV2058 | False
7118 | 704123074 | XSE - Damaged Instrumentation Trays | 2020-12-14T17:13:08.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 90 | KIT0899 | FRAME KIT0899 ROBOTIC FRAME CONNECTION | XSE - Damaged Instrumentation Trays 2020-12-11 00702590, e1, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instrumentation trays were damaged and the instruments failed tracking during preventative maintenance. The trajectory error ranged between 0.554 and 0.976 which was above the maximum amount allowed. Troubleshooting involved trying different spheres, trying different Navlock trackers and snapshot tracker, using long and short tools, adjusting the mounted position of the surgical system, hard restarts of the guidance system and trying different surgical arms. The stealth camera was viewing all of the spheres well during testing. The most probable cause was noted to be the washing and cleaning by the sterile processing department. The instruments were in rough condition and the edges. The edges of the trackers were chipped and the paint was faded. The trackers were difficult to detect during testing. There was no patient involvement. 2021-01-10 IUD (Other): No new information. 2021-01-10 IUD x2 (Other): No new information. 2021-02-02 email 2021-02-02 13:45, IUD (Other): No new information. 2021-02-24 IUD x2, email 2021-02-23 11:31 (Rep): No new information. 2021-03-09 IUD (Rep): Additional information received from a manufacturer representative reported that the issue was determined to be the camera. The camera was replaced and the issue was resolved. 2021-03-16 email 2021-03-16 20:45 (Rep): Additional information received from a manufacturer representative reported that all of the instruments were tested and found to be functioning as designed. FRAME KIT0899 ROBOTIC FRAME CONNECTION | 2020-12-11 00702590, e1, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instrumentation trays were damaged and the instruments failed tracking during preventative maintenance. The trajectory error ranged between 0.554 and 0.976 which was above the maximum amount allowed. Troubleshooting involved trying different spheres, trying different Navlock trackers and snapshot tracker, using long and short tools, adjusting the mounted position of the surgical system, hard restarts of the guidance system and trying different surgical arms. The stealth camera was viewing all of the spheres well during testing. The most probable cause was noted to be the washing and cleaning by the sterile processing department. The instruments were in rough condition and the edges. The edges of the trackers were chipped and the paint was faded. The trackers were difficult to detect during testing. There was no patient involvement. 2021-01-10 IUD (Other): No new information. 2021-01-10 IUD x2 (Other): No new information. 2021-02-02 email 2021-02-02 13:45, IUD (Other): No new information. 2021-02-24 IUD x2, email 2021-02-23 11:31 (Rep): No new information. 2021-03-09 IUD (Rep): Additional information received from a manufacturer representative reported that the issue was determined to be the camera. The camera was replaced and the issue was resolved. 2021-03-16 email 2021-03-16 20:45 (Rep): Additional information received from a manufacturer representative reported that all of the instruments were tested and found to be functioning as designed. | xse damaged instrumentation trays iud outside procedure instrumentation trays damaged instruments failed tracking preventative maintenance trajectory error ranged maximum allowed troubleshooting involved trying different spheres trying different navlock trackers snapshot tracker using long short tools adjusting mounted position surgical hard restarts trying different surgical arms stealth camera viewing spheres testing probable cause noted washing cleaning sterile processing department instruments rough condition edges edges trackers chipped paint faded trackers difficult detect testing iud new iud new email iud new iud email new iud additional issue determined camera camera replaced issue resolved email additional instruments tested functioning designed frame robotic frame connection | NAV3299 | NAV3244 | NAV3295 | NAV3249 | NAV3262 | NAV3299NAV3244NAV3295NAV3249NAV3262 | True
4723 | 703569598 | X-Registration issues/X-eye not tracking | 2020-01-20T23:56:53.000000+0000 | MazorX | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X-Registration issues/X-eye not tracking 2018-03-06 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that surgeon, fellow and representative planned case prior to procedure and approved implant placement and trajectory plan. Single Arm Schanz pin was agreed to be placed in left PSIS for the MIS procedure. Cage was placed in the lateral position at L3 - L4. Surgeon opted to add screws to L1 - L2 and L5 - S1 for stability. System was mounted and placed onto bed without any issue. Arm was draped and pin placed in Left PSIS. The guidance system was mounted to patient without a problem. While draping, scrubbed tech wrapped red rubber bands for X Eye cameras around the neck of the camera, causing inability for the camera to move properly. 3Define was executed properly as planned. X Eyes were unable to consistently pick up the reference marker and stylus. Representative opted to move forward to generic AP and Oblique location for 3D Marker acquisition. Significant manual adjustments were needed in order to place 3D marker in proper location. L1 - S1 was unable to be achieved in the imaging due to larger patient. Surgeon requested to register twice to achieve L5 - S1. First before registering L1 - L2, Surgeon also requested to take the images for the upper portion of the registration ahead of time and save in case we needed them later to which the representative explained it was not possible regardless. L5 - S1 screws were placed with accuracy bilaterally and checked with neuromonitoring successfully. Surgeon requested to bring up the L1 - L2 images to try for registration. Representative explained why it was not possible and proceeded to accept the images to walk the surgeon through the process of why it will not work. (After accepting saved images and moving to segmentation) When placing reference marker in saved oblique Image, the cross-view line was vertical to the vertebral bodies instead of horizontal as normal. Representative explained the cross- view problems and why registration could not be achieved. Surgeon agreed to take additional AP and OB image to capture L1 - L2. After achieving RBCs, representative placed reference marker in oblique image to achieve registration and again saw the cross-view lines running vertical to the spine as opposed to horizontal across a vertebral body. Surgeon opted to freehand the top screws due to lengthy duration of the procedure. No patient harm was reported. MAZOR X SYSTEM TPL0059 | 2018-03-06 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that surgeon, fellow and representative planned case prior to procedure and approved implant placement and trajectory plan. Single Arm Schanz pin was agreed to be placed in left PSIS for the MIS procedure. Cage was placed in the lateral position at L3 - L4. Surgeon opted to add screws to L1 - L2 and L5 - S1 for stability. System was mounted and placed onto bed without any issue. Arm was draped and pin placed in Left PSIS. The guidance system was mounted to patient without a problem. While draping, scrubbed tech wrapped red rubber bands for X Eye cameras around the neck of the camera, causing inability for the camera to move properly. 3Define was executed properly as planned. X Eyes were unable to consistently pick up the reference marker and stylus. Representative opted to move forward to generic AP and Oblique location for 3D Marker acquisition. Significant manual adjustments were needed in order to place 3D marker in proper location. L1 - S1 was unable to be achieved in the imaging due to larger patient. Surgeon requested to register twice to achieve L5 - S1. First before registering L1 - L2, Surgeon also requested to take the images for the upper portion of the registration ahead of time and save in case we needed them later to which the representative explained it was not possible regardless. L5 - S1 screws were placed with accuracy bilaterally and checked with neuromonitoring successfully. Surgeon requested to bring up the L1 - L2 images to try for registration. Representative explained why it was not possible and proceeded to accept the images to walk the surgeon through the process of why it will not work. (After accepting saved images and moving to segmentation) When placing reference marker in saved oblique Image, the cross-view line was vertical to the vertebral bodies instead of horizontal as normal. Representative explained the cross- view problems and why registration could not be achieved. Surgeon agreed to take additional AP and OB image to capture L1 - L2. After achieving RBCs, representative placed reference marker in oblique image to achieve registration and again saw the cross-view lines running vertical to the spine as opposed to horizontal across a vertebral body. Surgeon opted to freehand the top screws due to lengthy duration of the procedure. No patient harm was reported. | xregistration issuesxeye tracking ccr spinal procedure surgeon fellow planned case prior procedure approved implant placement trajectory plan single arm schanz pin agreed placed left psis mis procedure cage placed lateral position surgeon opted add screws stability mounted placed bed issue arm draped pin placed left psis mounted problem draping scrubbed tech wrapped red rubber bands x eye cameras neck camera causing inability camera properly executed properly planned x eyes unable consistently pick reference marker stylus opted forward generic ap oblique location marker acquisition significant manual adjustments needed order place marker proper location unable achieved imaging larger surgeon requested register twice achieve registering surgeon requested images upper portion registration ahead time save case needed later explained possible regardless screws placed accuracy bilaterally checked neuromonitoring successfully surgeon requested bring images try registration explained possible proceeded accept images walk surgeon process work placing reference marker saved oblique image crossview line vertical vertebral bodies instead horizontal normal explained cross view problems registration achieved surgeon agreed additional ap ob image capture achieving rbcs placed reference marker oblique image achieve registration saw crossview lines running vertical spine opposed horizontal vertebral body surgeon opted freehand screws lengthy duration procedure mazor x | NAV4110 | NAV4112 | NAV3233 | NAV3244 | NAV2051 | NAV4110NAV4112NAV3233NAV3244NAV2051 | True
4452 | 703543843 | X - Deviation/Registration issues | 2020-01-03T21:18:01.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation/Registration issues 2018-09-24 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for three vertebrae on level L4 - S1. Schanz pin was used and docked into L PSIS along with schanz arm. 3Define was performed along with draw spine successfully. Registration was failed on first attempt due to low quality of oblique image. AP and oblique images were retaken, and registration was successful on second attempt. The only change was including more of the sacrum for S1. Screws were performed MIS on the right side. L4 skived lateral and L5 was low in the pedicle. No patient harm was reported. MAZOR X SYSTEM TPL0059 | 2018-09-24 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for three vertebrae on level L4 - S1. Schanz pin was used and docked into L PSIS along with schanz arm. 3Define was performed along with draw spine successfully. Registration was failed on first attempt due to low quality of oblique image. AP and oblique images were retaken, and registration was successful on second attempt. The only change was including more of the sacrum for S1. Screws were performed MIS on the right side. L4 skived lateral and L5 was low in the pedicle. No patient harm was reported. | x deviationregistration issues ccr spinal procedure case vertebrae level schanz pin docked l psis schanz arm performed draw spine successfully registration failed attempt low quality oblique image ap oblique images retaken registration successful second attempt change including sacrum screws performed mis right skived lateral low pedicle mazor x | NAV3242 | NAV4091 | NAV3246 | NAV4112 | NAV3299 | NAV3242NAV4091NAV3246NAV4112NAV3299 | True
8806 | 704489535 | XSE - Unable to Mount to Bed | 2021-07-08T00:03:20.000000+0000 | MazorX Stealth Edition | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 30 | ASM0263-01 | RNX-WS-1040 | XSE - Unable to Mount to Bed 2021-Jul-06 Salesforce C00229776 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical system showed all of the lights and the teeth came out when going to mount the system to the bed, but the teeth would not engage. The manufacturer representative was unable to mount the surgical system to the bed and an error message was displayed. The representative tried the next day and the same issue occurred. When the guidance system was shut down, a blue screen was displayed. There was no patient involvement. 2021-07-15 IUD x4 (Other): No new information. 2021-07-26 IUD x2 (Other): No new information. 2021-08-09 IUD x5 (Other): No new information. RNX-WS-1040 | 2021-Jul-06 Salesforce C00229776 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical system showed all of the lights and the teeth came out when going to mount the system to the bed, but the teeth would not engage. The manufacturer representative was unable to mount the surgical system to the bed and an error message was displayed. The representative tried the next day and the same issue occurred. When the guidance system was shut down, a blue screen was displayed. There was no patient involvement. 2021-07-15 IUD x4 (Other): No new information. 2021-07-26 IUD x2 (Other): No new information. 2021-08-09 IUD x5 (Other): No new information. | xse unable mount bed salesforce outside procedure surgical showed lights teeth came going mount bed teeth engage unable mount surgical bed error message displayed tried day issue occurred shut blue screen displayed iud new iud new iud new | NAV3250 | NAV4103 | NAV3253 | NAV3246 | NAV2054 | NAV3250NAV4103NAV3253NAV3246NAV2054 | True
4239 | 703517524 | X - Deviation L1/Arm hit pt/3Define fail | 2019-12-13T14:35:47.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 30 | ASM0206-01 | SURGICAL ARM | X - Deviation L1/Arm hit pt/3Define fail 2018-04-12 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a trauma injury with kyphoplasty using cement augmentation using percutaneous approach. It was reported that the surgeon and manufacturer representative completed planning prior to the procedure. The system was mounted to the patient by a single schanz pin with schanz ball inserted in the patient's right iliac crest. The 3Define scan failed and so draw spine was not utilized. Scan and plan was then successful. During the operating the arm ran into the patient and the work volume had to be adjusted multiple times intra-operatively in order to reach different trajectories. Left L1 trajectory had a lateral skive. All other trajectories were accurate. At the end of the surgery, there was a joint to error. The manufacturer representative reset movement and the issue was resolved. No patient harm was reported and surgery was delayed less than an hour. SURGICAL ARM | 2018-04-12 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a trauma injury with kyphoplasty using cement augmentation using percutaneous approach. It was reported that the surgeon and manufacturer representative completed planning prior to the procedure. The system was mounted to the patient by a single schanz pin with schanz ball inserted in the patient's right iliac crest. The 3Define scan failed and so draw spine was not utilized. Scan and plan was then successful. During the operating the arm ran into the patient and the work volume had to be adjusted multiple times intra-operatively in order to reach different trajectories. Left L1 trajectory had a lateral skive. All other trajectories were accurate. At the end of the surgery, there was a joint to error. The manufacturer representative reset movement and the issue was resolved. No patient harm was reported and surgery was delayed less than an hour. | x deviation hit fail ccr spinal procedure treat trauma injury kyphoplasty using cement augmentation using percutaneous approach surgeon completed planning prior procedure mounted single schanz pin schanz ball inserted patients right iliac crest scan failed draw spine utilized scan plan successful operating arm ran work volume adjusted multiple times intraoperatively order reach different trajectories left trajectory lateral skive trajectories accurate end surgery joint error reset movement issue resolved surgery delayed hour surgical arm | NAV4096 | NAV4091 | NAV3244 | NAV3249 | NAV4095 | NAV4096NAV4091NAV3244NAV3249NAV4095 | False
9700 | 704701618 | XSE - Unable To Install Tools | 2021-11-05T22:31:11.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1348 | KIT KIT1348 TOOLS PACKAGE | XSE - Unable To Install Tools 2021-Nov-05 Salesforce C00237453 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while attempting to install the tools update the manufacturer representative getting an error about being unable to update the database. After selecting the package, the software would prompt if the database should be updated and after it was selected the failure message would appear. He tried rebooting multiple times without success. No patient was present. 2021-11-09 IUD x2 (Other): No new information. 2021-11-29 e1, e3, e4, e5, IUD (Rep): Additional information received from a manufacturer representative reported the same USB was used for all installations. The update was successful with a different USB. 2021-12-01 e6, e7, e8, e9, e10, e11, e12, e13, e14 (Rep): No new information. 2021-12-13 IUD (Other): No new information. KIT KIT1348 TOOLS PACKAGE | 2021-Nov-05 Salesforce C00237453 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while attempting to install the tools update the manufacturer representative getting an error about being unable to update the database. After selecting the package, the software would prompt if the database should be updated and after it was selected the failure message would appear. He tried rebooting multiple times without success. No patient was present. 2021-11-09 IUD x2 (Other): No new information. 2021-11-29 e1, e3, e4, e5, IUD (Rep): Additional information received from a manufacturer representative reported the same USB was used for all installations. The update was successful with a different USB. 2021-12-01 e6, e7, e8, e9, e10, e11, e12, e13, e14 (Rep): No new information. 2021-12-13 IUD (Other): No new information. | xse unable install tools salesforce outside procedure attempting install tools update getting error unable update database selecting package software prompt database updated selected failure message appear tried rebooting multiple times success present iud new iud additional usb installations update successful different usb new iud new kit tools package | NAV3246 | NAV2054 | NAV3249 | NAV2047 | NAV3299 | NAV3246NAV2054NAV3249NAV2047NAV3299 | True
1424 | 703200171 | X- Positioner Issues | 2019-06-11T16:49:46.000000+0000 | MazorX | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X- Positioner Issues Robot also got stuck on bed rail adapter, one tooth engaged the mount while the other did not. We had to remove the robot with bed rail mount attached. We then reattached the bedrail mount and were able to successfully attach robot. Once robot was installed on bed registration was quick and efficient,no software issues. UNK MAZOR X POSITIONER | Robot also got stuck on bed rail adapter, one tooth engaged the mount while the other did not. We had to remove the robot with bed rail mount attached. We then reattached the bedrail mount and were able to successfully attach robot. Once robot was installed on bed registration was quick and efficient,no software issues. | x positioner issues robot got stuck bed rail adapter tooth engaged mount did remove robot bed rail mount attached reattached bedrail mount able successfully attach robot robot installed bed registration quick efficientno software issues unk mazor x positioner | NAV3250 | NAV4103 | NAV3244 | NAV4104 | NAV4105 | NAV3250NAV4103NAV3244NAV4104NAV4105 | True
4823 | 703578917 | XSE- X-eye & firewire cable issues | 2020-01-24T00:07:25.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 30 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE- X-eye & firewire cable issues 2020-1-23 00682677(Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the x-eye camera is not functioning and the firewire cable needs to be replaced. No patient involved. 2020-01-24 IUD (Other): Additional information received from the field service engineer reporting that the complaint was confirmed. The system was serviced on-site. The firewire card was replaced and the firewire cabling was re-seated throughout the system. It was identified that a connection at the bottom of the surgical system that was seated enough to provide power to the x-eye but no communication for a video feed was the likely primary cause of the issue. Re-seating this cable and replacing the firewire card fixed the x-eye video feed. The system was tested and was fully functional. 2020-02-07 IUD x2 (Rep): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2020-1-23 00682677(Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the x-eye camera is not functioning and the firewire cable needs to be replaced. No patient involved. 2020-01-24 IUD (Other): Additional information received from the field service engineer reporting that the complaint was confirmed. The system was serviced on-site. The firewire card was replaced and the firewire cabling was re-seated throughout the system. It was identified that a connection at the bottom of the surgical system that was seated enough to provide power to the x-eye but no communication for a video feed was the likely primary cause of the issue. Re-seating this cable and replacing the firewire card fixed the x-eye video feed. The system was tested and was fully functional. 2020-02-07 IUD x2 (Rep): No new information. | xse xeye firewire cable issues spinal procedure xeye camera functioning firewire cable needs replaced involved iud additional field service engineer reporting complaint confirmed serviced onsite firewire card replaced firewire cabling reseated identified connection surgical seated provide power xeye communication video feed likely primary cause issue reseating cable replacing firewire card fixed xeye video feed tested fully functional iud new mazor x pc station | NAV4110 | NAV3262 | NAV3256 | NAV3254 | NAV1006 | NAV4110NAV3262NAV3256NAV3254NAV1006 | True
6971 | 704085613 | XSE - Unable To Use Surgeon Screen | 2020-11-23T18:58:27.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - Unable To Use Surgeon Screen 2020-11-18 00701375 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon screen stopped working during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-04 IUD x3 (Other): No new information. SURGEON SCREEN | 2020-11-18 00701375 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon screen stopped working during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-04 IUD x3 (Other): No new information. | xse unable use surgeon screen spinal procedure surgeon screen stopped working procedure procedure delayed hour iud new surgeon screen | NAV3253 | NAV3246 | NAV1006 | NAV3244 | NAV4112 | NAV3253NAV3246NAV1006NAV3244NAV4112 | True
2171 | 703228215 | X - 3Define Issue | 2019-06-26T12:27:11.000000+0000 | MazorX | NAV3239 | INACCURATE WORK VOLUMES | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define Issue software issue leading to re-registration. 3D scan was inaccurate. location of spine in the VTK was off. UNK MAZOR X 3DEFINE CAMERA | software issue leading to re-registration. 3D scan was inaccurate. location of spine in the VTK was off. | x issue software issue leading reregistration scan inaccurate location spine vtk unk mazor x camera | NAV3244 | NAV4095 | NAV3250 | NAV4110 | NAV2054 | NAV3244NAV4095NAV3250NAV4110NAV2054 | False
3372 | 703329407 | X - Deviation | 2019-08-26T15:03:33.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation 2019-08-12 SCL1902425, Intake Form, email 2019-08-26 10:11 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done prior to the surgery and reviewed by the surgeon. Mounting of the surgical system to the patient was also discussed prior to the surgery. The representative suggested using disposable clamp or single piece of the dual clamp, but the surgeon preferred a less bulky options so they decided to use the MIST bridge. During the procedure, a 3Define scan and Draw spine were completed with no issues. Registration required troubleshooting. Images were eventually accepted and labeling was attempted, but segmentation would not occur due to poor oblique image quality. A new oblique image was taken and registration was accepted with no shifts detected. The surgeon executed T3 right and then followed with the screw. The process was continued for every trajectory in the following order: T4 right, T5 right, T2 left, T4 left, and T5 left. The representative noted that the surgeon appeared to be having issues getting the cannula and drill guide through to the bony surface so they suggested the long knife be used to dissect the muscle and allow easier passing. The surgeon did not want to use the long knife so the representative suggested the blunt trocar could be used with the cannula, but the surgeon did not like that option and preferred to continue with the drill guide and cannula. Due to axial angles of 10 degrees, the surgeon had slight issues fitting the tools inside the incision. The surgeon noted that all drilled trajectories prior to T5 left felt solid and in the bone. After frilling T5 left, the surgeon noted that the reduction tube did not feel "right". The surgeon pulled the wire and dissected down to reveal that the drill had deviated medially, but had not disturbed the spinal cord or caused a dural tear. There were no CSF leaks and neuromonitoring informed the surgeon that all motor functions and leg signals were completely fine and normal. The surgeon continued to move on to the decompression where the plan was to insert a cage at T3/T4 and T4/T5. During the decompression, a posterior dural tear occurred. There was no relation between the deviation and the dural tear. The surgeon stitched the dural tear and continued with the plan to insert the cages. During the insertion, another dural tear occurred. There was no relation between the second dural tear and deviation. The surgeon decided to abandoned the plan to insert the cages and settle for the decompression and fixation. The representative indicated the deviation was likely caused by unilateral retraction and soft tissue pressure. The surgeon eventually placed the screw manually using a Jamshidi needle and C-arm. All other wires and screws were acceptable according to the surgeon. The representative reported that the patient was doing ok and they were able to move their legs a day after the procedure. There was no patient harm, but the procedure was delayed more than an hour due to the deviation. The case was completed. MAZOR X SYSTEM TPL0059 | 2019-08-12 SCL1902425, Intake Form, email 2019-08-26 10:11 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was done prior to the surgery and reviewed by the surgeon. Mounting of the surgical system to the patient was also discussed prior to the surgery. The representative suggested using disposable clamp or single piece of the dual clamp, but the surgeon preferred a less bulky options so they decided to use the MIST bridge. During the procedure, a 3Define scan and Draw spine were completed with no issues. Registration required troubleshooting. Images were eventually accepted and labeling was attempted, but segmentation would not occur due to poor oblique image quality. A new oblique image was taken and registration was accepted with no shifts detected. The surgeon executed T3 right and then followed with the screw. The process was continued for every trajectory in the following order: T4 right, T5 right, T2 left, T4 left, and T5 left. The representative noted that the surgeon appeared to be having issues getting the cannula and drill guide through to the bony surface so they suggested the long knife be used to dissect the muscle and allow easier passing. The surgeon did not want to use the long knife so the representative suggested the blunt trocar could be used with the cannula, but the surgeon did not like that option and preferred to continue with the drill guide and cannula. Due to axial angles of 10 degrees, the surgeon had slight issues fitting the tools inside the incision. The surgeon noted that all drilled trajectories prior to T5 left felt solid and in the bone. After frilling T5 left, the surgeon noted that the reduction tube did not feel "right". The surgeon pulled the wire and dissected down to reveal that the drill had deviated medially, but had not disturbed the spinal cord or caused a dural tear. There were no CSF leaks and neuromonitoring informed the surgeon that all motor functions and leg signals were completely fine and normal. The surgeon continued to move on to the decompression where the plan was to insert a cage at T3/T4 and T4/T5. During the decompression, a posterior dural tear occurred. There was no relation between the deviation and the dural tear. The surgeon stitched the dural tear and continued with the plan to insert the cages. During the insertion, another dural tear occurred. There was no relation between the second dural tear and deviation. The surgeon decided to abandoned the plan to insert the cages and settle for the decompression and fixation. The representative indicated the deviation was likely caused by unilateral retraction and soft tissue pressure. The surgeon eventually placed the screw manually using a Jamshidi needle and C-arm. All other wires and screws were acceptable according to the surgeon. The representative reported that the patient was doing ok and they were able to move their legs a day after the procedure. There was no patient harm, but the procedure was delayed more than an hour due to the deviation. The case was completed. | x deviation email spinal procedure planning prior surgery reviewed surgeon mounting surgical discussed prior surgery suggested using disposable clamp single piece dual clamp surgeon preferred bulky options decided use mist bridge procedure scan draw spine completed issues registration required troubleshooting images eventually accepted labeling attempted segmentation occur poor oblique image quality new oblique image taken registration accepted shifts detected surgeon executed right followed screw process continued trajectory following order right right left left left noted surgeon appeared having issues getting cannula drill guide bony surface suggested long knife dissect muscle allow easier passing surgeon did want use long knife suggested blunt trocar cannula surgeon did like option preferred continue drill guide cannula axial angles degrees surgeon slight issues fitting tools inside incision surgeon noted drilled trajectories prior left felt solid bone frilling left surgeon noted reduction tube did feel right surgeon pulled wire dissected reveal drill deviated medially disturbed spinal cord caused dural tear csf leaks neuromonitoring informed surgeon motor functions leg signals completely fine normal surgeon continued decompression plan insert cage decompression posterior dural tear occurred relation deviation dural tear surgeon stitched dural tear continued plan insert cages insertion dural tear occurred relation second dural tear deviation surgeon decided abandoned plan insert cages settle decompression fixation indicated deviation likely caused unilateral retraction soft tissue pressure surgeon eventually placed screw manually using jamshidi needle carm wires screws acceptable according surgeon doing ok able legs day procedure procedure delayed hour deviation case completed mazor x | NAV3242 | NAV5000 | NAV4091 | NAV3253 | NAV3299 | NAV3242NAV5000NAV4091NAV3253NAV3299 | True
3173 | 703268428 | XSE - Arm Shift | 2019-07-19T15:37:28.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Arm Shift 2019-Jul-18 SCL1902136 intake form (HCP,rep): ***Please see SR/PE 603268430 for complaint against Powerease*** Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that it was the Surgeon's first time using the robot. After exposing and getting the disposable clamp on L4, Registration was achieved after taking a few shots. Registration was not achieved on the first try, so more images were needed. The surgeon was getting anxious about the time of everything. The arm was sent to L2 right and surgeon began placing cannula, and drill guide down to mallet. The surgeon then drilled and then tapped with a 5.5 tap. The surgeon tried placing screw with power but the 6.5 screw would not advance. The surgeon then tapped with 6.5 and tried placing screw with hand held handle. The arm alerted that it was off trajectory. The rep resent arm and surgeon tried putting screw in with power. The Powerease was not working right, and was wobbly and caused the arm to shift a few cm’s. The surgeon was frustrated with how long everything was taking and said they would try another time with the robot, but they had to get these screws in. The robot was aborted and screws were placed free hand. The rep felt the arm should not have moved that easily. An engineer came out the next day to check the arm and could not reproduce the issue. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jul-18 SCL1902136 intake form (HCP,rep): ***Please see SR/PE 603268430 for complaint against Powerease*** Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that it was the Surgeon's first time using the robot. After exposing and getting the disposable clamp on L4, Registration was achieved after taking a few shots. Registration was not achieved on the first try, so more images were needed. The surgeon was getting anxious about the time of everything. The arm was sent to L2 right and surgeon began placing cannula, and drill guide down to mallet. The surgeon then drilled and then tapped with a 5.5 tap. The surgeon tried placing screw with power but the 6.5 screw would not advance. The surgeon then tapped with 6.5 and tried placing screw with hand held handle. The arm alerted that it was off trajectory. The rep resent arm and surgeon tried putting screw in with power. The Powerease was not working right, and was wobbly and caused the arm to shift a few cm’s. The surgeon was frustrated with how long everything was taking and said they would try another time with the robot, but they had to get these screws in. The robot was aborted and screws were placed free hand. The rep felt the arm should not have moved that easily. An engineer came out the next day to check the arm and could not reproduce the issue. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse arm shift srpe complaint powerease spinal procedure surgeons time using robot exposing getting disposable clamp registration achieved taking shots registration achieved try images needed surgeon getting anxious time arm sent right surgeon began placing cannula drill guide mallet surgeon drilled tapped tap surgeon tried placing screw power screw advance surgeon tapped tried placing screw hand held handle arm alerted trajectory resent arm surgeon tried putting screw power powerease working right wobbly caused arm shift cm’s surgeon frustrated long taking said try time robot screws robot aborted screws placed free hand felt arm moved easily engineer came day check arm reproduce issue additional complications anticipated mazor x software installation kit | NAV3242 | NAV5000 | NAV4108 | NAV4096 | NAV4098 | NAV3242NAV5000NAV4108NAV4096NAV4098 | True
1097 | 703188432 | X - Schanz Bond Issue | 2019-06-04T17:58:31.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | MEC4530-01 | SLOTTED BALL ARM TARGET SCREW | X - Schanz Bond Issue Schanz Bond screws are stripped SLOTTED BALL ARM TARGET SCREW | Schanz Bond screws are stripped | x schanz bond issue schanz bond screws stripped slotted ball arm target screw | NAV3295 | NAV3299 | NAV4090 | NAV4091 | NAV4095 | NAV3295NAV3299NAV4090NAV4091NAV4095 | True
9289 | 704602484 | XSE ¦ Positioner and Manipulator Issues | 2021-09-09T22:14:40.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 40 | KIT0759-01 | PNEUMATIC REPAIR KIT | XSE ¦ Positioner and Manipulator Issues 2021-Sep-09 Salesforce C00234211 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that after the case the positioner and manipulator appeared to lose power. The manufacturer representative completed a reboot and found tubing pinched near the manipulator elbow. There was no known impact to the patient outcome. 2021-09-13 IUD x2 (Other): No new information. 2021-10-04 email 2021-10-04 14:16 (Rep): Additional information was received stating that the surgical procedure was a spinal procedure. There was no delay to the procedure. 2021-10-08 IUD (Other): No new information. PNEUMATIC REPAIR KIT | 2021-Sep-09 Salesforce C00234211 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that after the case the positioner and manipulator appeared to lose power. The manufacturer representative completed a reboot and found tubing pinched near the manipulator elbow. There was no known impact to the patient outcome. 2021-09-13 IUD x2 (Other): No new information. 2021-10-04 email 2021-10-04 14:16 (Rep): Additional information was received stating that the surgical procedure was a spinal procedure. There was no delay to the procedure. 2021-10-08 IUD (Other): No new information. | xse positioner manipulator issues salesforce case positioner manipulator appeared lose power completed reboot tubing pinched near manipulator elbow known impact outcome iud new email additional stating surgical procedure spinal procedure delay procedure iud new pneumatic repair kit | NAV3260 | NAV3264 | NAV3251 | NAV4103 | NAV3249 | NAV3260NAV3264NAV3251NAV4103NAV3249 | True
9086 | 704547667 | XSE - OOBF/PC Failed & Screens Flicker | 2021-08-09T21:37:44.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 60 | ASM0215 | ASM0215 MAZOR X COMPUTER KIT | XSE - OOBF/PC Failed & Screens Flicker 2021-Aug-09 Salesforce C00231919 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that both screens flickered to black, the arm worked. The day after install both screens went to black and neither came back on. No patient was present. 2021-08-13 IUD x6 (Rep): Additional information received from a manufacturer representative reported that the replacement PC failed out of the box. 2021-09-05 e1 (Rep): No new information. 2021-09-07 IUD x2, e2 (Rep): Additional information received from a manufacturer representative reported the replacement PC would not provide an image to the screens. 2021-09-22 IUD (Other): No new information. 2021-09-30 e3, e4, e5, IUD x2 (Other): No new information. ASM0215 MAZOR X COMPUTER KIT | 2021-Aug-09 Salesforce C00231919 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that both screens flickered to black, the arm worked. The day after install both screens went to black and neither came back on. No patient was present. 2021-08-13 IUD x6 (Rep): Additional information received from a manufacturer representative reported that the replacement PC failed out of the box. 2021-09-05 e1 (Rep): No new information. 2021-09-07 IUD x2, e2 (Rep): Additional information received from a manufacturer representative reported the replacement PC would not provide an image to the screens. 2021-09-22 IUD (Other): No new information. 2021-09-30 e3, e4, e5, IUD x2 (Other): No new information. | xse oobfpc failed screens flicker salesforce outside procedure screens flickered black arm worked day install screens went black came present iud additional replacement pc failed box new iud additional replacement pc provide image screens iud new iud new mazor x computer kit | NAV3262 | NAV2054 | NAV3247 | NAV3261 | NAV3253 | NAV3262NAV2054NAV3247NAV3261NAV3253 | True
2692 | 703246594 | X - ARM ISSUE | 2019-07-08T14:31:16.000000+0000 | MazorX | NAV3256 | CABLE ISSUES | 20 | ASM0206-01 | SURGICAL ARM | X - ARM ISSUE Loaner arm 63 was attached to system-we found Joint 1 in Arm non-functional-located broken wire inside arm which needs repair. SURGICAL ARM | Loaner arm 63 was attached to system-we found Joint 1 in Arm non-functional-located broken wire inside arm which needs repair. | x arm issue loaner arm attached systemwe joint arm nonfunctionallocated broken wire inside arm needs repair surgical arm | NAV3249 | NAV4098 | NAV4101 | NAV4097 | NAV4095 | NAV3249NAV4098NAV4101NAV4097NAV4095 | False
8521 | 704431184 | XSE - Surgical System Damaged | 2021-06-04T03:02:14.000000+0000 | MazorX Stealth Edition | NAV4105 | UNEXPECTED POSITIONER MOVEMENT | 30 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - Surgical System Damaged 2021-Jun-03 Salesforce C00227815, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the positioner teeth sensor tang had popped out and appeared slightly bent. The manufacturer representative bent the tang back into place. When inserting the tang into the positioner, the teeth sensors were not triggering correctly. The representative adjusted the sensors and the three lights came up, but the other manipulator light did not. Increasing the pressure to high in the override panel of the workstation resulted in a loud noise and the positioner dropping to the floor. The representative then had difficulties moving the positioner. A hard restart of the system was done, the representative waited five minutes, did another hard restart and then moved the system out of the room. The representative then tried to remount the surgical system to the bed and force close the teeth. Green lights were showed on all sensors, but the representative was unable to move the positioner along the Z-axis. When trying to move the positioner, the surgeon screen went black and there was a fata error on node position 7. The representative noted the manipulator and workstation hood would not open. The surgical arm was put in the storage position and the system was manually unmounted from the bed. A different system was used for the case. There was no patient involvement. 2021-06-14 IUD x9 (Other): No new information. 2021-Jul-09 IUDx7 (other): No new information. 2021-09-01 302547474, IUD x5 (Other): No new information. 2021-09-14 302561435 (Other): No new information. 2021-09-21 IUD x2 (Other): No new information. 2021-09-26 IUD x5 (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-Jun-03 Salesforce C00227815, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the positioner teeth sensor tang had popped out and appeared slightly bent. The manufacturer representative bent the tang back into place. When inserting the tang into the positioner, the teeth sensors were not triggering correctly. The representative adjusted the sensors and the three lights came up, but the other manipulator light did not. Increasing the pressure to high in the override panel of the workstation resulted in a loud noise and the positioner dropping to the floor. The representative then had difficulties moving the positioner. A hard restart of the system was done, the representative waited five minutes, did another hard restart and then moved the system out of the room. The representative then tried to remount the surgical system to the bed and force close the teeth. Green lights were showed on all sensors, but the representative was unable to move the positioner along the Z-axis. When trying to move the positioner, the surgeon screen went black and there was a fata error on node position 7. The representative noted the manipulator and workstation hood would not open. The surgical arm was put in the storage position and the system was manually unmounted from the bed. A different system was used for the case. There was no patient involvement. 2021-06-14 IUD x9 (Other): No new information. 2021-Jul-09 IUDx7 (other): No new information. 2021-09-01 302547474, IUD x5 (Other): No new information. 2021-09-14 302561435 (Other): No new information. 2021-09-21 IUD x2 (Other): No new information. 2021-09-26 IUD x5 (Other): No new information. | xse surgical damaged salesforce iud outside proc edure positioner teeth sensor tang popped appeared slightly bent bent tang place inserting tang positioner teeth sensors triggering correctly adjusted sensors lights came manipulator light did increasing pressure high override panel workstation resulted loud noise positioner dropping floor difficulties moving positioner hard restart waited minutes did hard restart moved room tried remount surgical bed force close teeth green lights showed sensors unable positioner zaxis trying positioner surgeon screen went black fata error node position noted manipulator workstation hood open surgical arm storage position manually unmounted bed different case iud new new iud new new iud new iud new kit svc surgical syst type | NAV4104 | NAV4103 | NAV3251 | NAV2047 | NAV3264 | NAV4104NAV4103NAV3251NAV2047NAV3264 | False
9036 | 704534264 | XSE - Black Surgeon Screen | 2021-08-02T16:18:19.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0337-01 | SURGEON&APOS;S SCREEN | XSE - Black Surgeon Screen 2021-Aug-02 Salesforce C00231193, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgeon screen was black. The software was reinstalled and the issue was resolved. There was no patient involvement. 2021-08-04 IUD x2 (Other): No new information. SURGEON&APOS;S SCREEN | 2021-Aug-02 Salesforce C00231193, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgeon screen was black. The software was reinstalled and the issue was resolved. There was no patient involvement. 2021-08-04 IUD x2 (Other): No new information. | xse black surgeon screen salesforce iud outside proc edure surgeon screen black software reinstalled issue resolved iud new surgeonaposs screen | NAV3253 | NAV3261 | NAV3247 | NAV3262 | NAV3249 | NAV3253NAV3261NAV3247NAV3262NAV3249 | True
3323 | 703314971 | XSE - Arm Not Moving Joint 6/OOBF | 2019-08-16T15:03:01.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Arm Not Moving Joint 6/OOBF 2019-Aug-12 SCL1902360 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the engineer replaced the surgical system and after the replacement the surgical arm would not move at joint 6. There was no patient involvement or additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Aug-12 SCL1902360 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the engineer replaced the surgical system and after the replacement the surgical arm would not move at joint 6. There was no patient involvement or additional complications reported or anticipated as a result of the event. | xse arm moving joint spinal procedure engineer replaced surgical replacement surgical arm joint additional complications anticipated surgical arm | NAV4101 | NAV2054 | NAV3249 | NAV4098 | NAV5000 | NAV4101NAV2054NAV3249NAV4098NAV5000 | True
9256 | 704591011 | XSE - Intermittent Black Monitors | 2021-09-02T14:10:20.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - Intermittent Black Monitors 2021-Sep-01 Salesforce C00233868 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the workstation and surgeon screens would intermittently go black and then return after between 10 and 15 seconds. There was no patient harm and the procedure was not delayed. 2021-Sep-03 IUDx5 (other): No new information. 2021-09-21 IUD x2 (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-18 IUD x2 (Other): No new information. 2021-11-21 IUD x4 (Other): No new information. 2021-11-29 IUD x2 (Other): No new information. 2021-12-05 IUD x2 (Other): No new information. 2021-12-09 IUD (Other): No new information. SURGEON SCREEN | 2021-Sep-01 Salesforce C00233868 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the workstation and surgeon screens would intermittently go black and then return after between 10 and 15 seconds. There was no patient harm and the procedure was not delayed. 2021-Sep-03 IUDx5 (other): No new information. 2021-09-21 IUD x2 (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-18 IUD x2 (Other): No new information. 2021-11-21 IUD x4 (Other): No new information. 2021-11-29 IUD x2 (Other): No new information. 2021-12-05 IUD x2 (Other): No new information. 2021-12-09 IUD (Other): No new information. | xse intermittent black monitors salesforce spinal p rocedure workstation surgeon screens intermittently black return seconds procedure delayed new iud new iud new iud new iud new iud new iud new iud new surgeon screen | NAV3253 | NAV3261 | NAV2047 | NAV2058 | NAV3237 | NAV3253NAV3261NAV2047NAV2058NAV3237 | True
5848 | 703819728 | XSE - OOBF/SRA Damaged Shoulder | 2020-06-24T19:59:15.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - OOBF/SRA Damaged Shoulder 2020-06-24 00691458, IUD (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the surgical arm was received damaged. The shoulder handle of the surgical arm was damaged. The manufacturer representative was able to take the shoulder handle off the old arm and use it on the new arm. There was no patient involvement. 2020-07-01 IUD (Other): No new information. 2020-07-14 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-06-24 00691458, IUD (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the surgical arm was received damaged. The shoulder handle of the surgical arm was damaged. The manufacturer representative was able to take the shoulder handle off the old arm and use it on the new arm. There was no patient involvement. 2020-07-01 IUD (Other): No new information. 2020-07-14 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. | xse oobfsra damaged shoulder iud prior procedure surgical arm damaged shoulder handle surgical arm damaged able shoulder handle old arm use new arm iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV3252 | NAV2054 | NAV4107 | NAV3251 | NAV3249NAV3252NAV2054NAV4107NAV3251 | True
6254 | 703898289 | XSE - Inaccuracy | 2020-08-11T19:30:16.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | XSE - Inaccuracy 2020-08-11 00694413 00694413 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a L4-L5 lateral case after checking the anatomical landmarks and accuracy the system was showing inaccurate. The surgeon re-registered the patient before implanting the screws and the system was accurate from that point on and placed the screws without an issue. The manufacturer representative and the surgeon agreed that either the Schanz pin was not in the PSIS or the patient moved at some point. The Schanz arm was not tight when then attached it to the pin. They will be exchanging the arm for a replacement. They used another Schanz arm that was on site for the procedure. The issue extended the surgical time by less than 1 hour. There was no known impact on the patient outcome. 2020-08-24 e1 (Rep): Additional information was received stating that two Schanz arms that would not fully tighten at the joint closest to where they mount to the arm, but neither of these were utilized in this case (they were both removed prior to going into operate). One was fixed afterwards and the other will be replaced. 2020-09-23 email 2020-09-23 11:30 (Rep): Additional information received from a manufacturer representative reported that the Schanz arms were discarded. UNK MAZOR X ACCESSORIES | 2020-08-11 00694413 00694413 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a L4-L5 lateral case after checking the anatomical landmarks and accuracy the system was showing inaccurate. The surgeon re-registered the patient before implanting the screws and the system was accurate from that point on and placed the screws without an issue. The manufacturer representative and the surgeon agreed that either the Schanz pin was not in the PSIS or the patient moved at some point. The Schanz arm was not tight when then attached it to the pin. They will be exchanging the arm for a replacement. They used another Schanz arm that was on site for the procedure. The issue extended the surgical time by less than 1 hour. There was no known impact on the patient outcome. 2020-08-24 e1 (Rep): Additional information was received stating that two Schanz arms that would not fully tighten at the joint closest to where they mount to the arm, but neither of these were utilized in this case (they were both removed prior to going into operate). One was fixed afterwards and the other will be replaced. 2020-09-23 email 2020-09-23 11:30 (Rep): Additional information received from a manufacturer representative reported that the Schanz arms were discarded. | xse inaccuracy intraoperatively spinal procedure lateral case checking anatomical landmarks accuracy showing inaccurate surgeon reregistered implanting screws accurate point placed screws issue surgeon agreed schanz pin psis moved point schanz arm tight attached pin exchanging arm replacement schanz arm site procedure issue extended surgical time hour known impact outcome additional stating schanz arms fully tighten joint closest mount arm utilized case fixed replaced email additional schanz arms discarded unk mazor x accessories | NAV3299 | NAV5000 | NAV3298 | NAV4093 | NAV4101 | NAV3299NAV5000NAV3298NAV4093NAV4101 | True
6453 | 703944262 | XSE - Low Pressure Error | 2020-09-08T13:49:31.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Low Pressure Error 2020-09-08 00695987, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a low pressure error message was displayed during the procedure. There was no patient harm and the procedure was not delayed. 2020-09-09 IUD x2, email 2020-09-10 08:31 (Other): No new information. 2020-09-14 IUD x5 (Other): No new information. 2021-12-08 IUD x2 (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2020-09-08 00695987, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a low pressure error message was displayed during the procedure. There was no patient harm and the procedure was not delayed. 2020-09-09 IUD x2, email 2020-09-10 08:31 (Other): No new information. 2020-09-14 IUD x5 (Other): No new information. 2021-12-08 IUD x2 (Other): No new information. | xse low pressure error iud spinal procedure low pressure error message displayed procedure procedure delayed iud email new iud new iud new kit svc cart elec panel assy | NAV3264 | NAV1006 | NAV2047 | NAV2054 | NAV5000 | NAV3264NAV1006NAV2047NAV2054NAV5000 | True
1364 | 703197626 | X - Workstation issue | 2019-06-10T16:02:55.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Workstation issue Also, the system was having trouble building pressure past 1.0 psi so they tapped on the main valve which worked but when putting the pneumatic panel back in the air lines were pinched. We will need to replace the entire Pneumatic Panel. UNK MAZOR X WORKSTATION | Also, the system was having trouble building pressure past 1.0 psi so they tapped on the main valve which worked but when putting the pneumatic panel back in the air lines were pinched. We will need to replace the entire Pneumatic Panel. | x workstation issue having trouble building pressure past psi tapped main valve worked putting pneumatic panel air lines pinched need replace entire pneumatic panel unk mazor x workstation | NAV3264 | NAV2058 | NAV2047 | NAV2054 | NAV3256 | NAV3264NAV2058NAV2047NAV2054NAV3256 | True
5657 | 703772415 | XSE - OOBF/Shoulder Not Within Spec | 2020-05-27T18:17:21.000000+0000 | MazorX Stealth Edition | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | ASM0207-02S | KIT SVC ASM0207-02S SHOULDER | XSE - OOBF/Shoulder Not Within Spec 2020-05-26 00689612 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that when testing the shoulder tightness at install, the shoulder was not in spec. The field service engineer (FSE) had to tighten the shoulder. The FSE also reported that the five screws at the back shoulder were not tightened and this caused the shoulder to move when it was locked. The screws were tightened to resolve the issue. There was no patient involvement. 2020-06-03 IUD (Other): No new information. KIT SVC ASM0207-02S SHOULDER | 2020-05-26 00689612 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that when testing the shoulder tightness at install, the shoulder was not in spec. The field service engineer (FSE) had to tighten the shoulder. The FSE also reported that the five screws at the back shoulder were not tightened and this caused the shoulder to move when it was locked. The screws were tightened to resolve the issue. There was no patient involvement. 2020-06-03 IUD (Other): No new information. | xse oobfshoulder spec prior procedure testing shoulder tightness install shoulder spec field service engineer tighten shoulder fse screws shoulder tightened caused shoulder locked screws tightened resolve issue iud new kit svc shoulder | NAV4106 | NAV4108 | NAV2054 | NAV4107 | NAV3252 | NAV4106NAV4108NAV2054NAV4107NAV3252 | True
286 | 703066540 | X - 3Define Camera Not Working | 2019-03-21T17:19:36.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define Camera Not Working 2019-03-18 SCL1900680, Intake Form, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define was not visible during pre-operation setup when the software was started. A soft or hard restart did not resolve the issue. After completing the hard restart, a connection error message was displayed. The manufacturer representative checked all of the connections and looked at the task manager to see if the camera was listed. Technical services recommended uninstalling and reinstalling the software, but the manufacturer representative did not have the software to do that. The case was able to be completed using the generic work volume. There was no patient harm and the procedure was not delayed more than an hour. UNK MAZOR X 3DEFINE CAMERA | 2019-03-18 SCL1900680, Intake Form, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define was not visible during pre-operation setup when the software was started. A soft or hard restart did not resolve the issue. After completing the hard restart, a connection error message was displayed. The manufacturer representative checked all of the connections and looked at the task manager to see if the camera was listed. Technical services recommended uninstalling and reinstalling the software, but the manufacturer representative did not have the software to do that. The case was able to be completed using the generic work volume. There was no patient harm and the procedure was not delayed more than an hour. | x camera working spinal procedure visible preoperation setup software started soft hard restart did resolve issue completing hard restart connection error message displayed checked connections looked task manager camera listed technical services recommended uninstalling reinstalling software did software case able completed using generic work volume procedure delayed hour unk mazor x camera | NAV3244 | NAV2047 | NAV3248 | NAV4095 | NAV3246 | NAV3244NAV2047NAV3248NAV4095NAV3246 | True
5701 | 703784471 | XSE - Instruments Missing From SW | 2020-06-03T22:20:52.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Instruments Missing From SW 2020-06-03 00690069, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instruments were not in the software and would not show up. The field service engineer planned to uninstall and reinstall the software. There was no patient involvement. 2020-06-10 IUD (Other): No new information. 2020-07-07 e1, e2 (Rep): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-06-03 00690069, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instruments were not in the software and would not show up. The field service engineer planned to uninstall and reinstall the software. There was no patient involvement. 2020-06-10 IUD (Other): No new information. 2020-07-07 e1, e2 (Rep): No new information. | xse instruments missing sw iud outside procedure instruments software field service engineer planned uninstall reinstall software iud new new mazor x software installation kit | NAV3246 | NAV2056 | NAV3248 | NAV3247 | NAV2054 | NAV3246NAV2056NAV3248NAV3247NAV2054 | True
5063 | 703593796 | X- Communication Issues | 2020-02-03T14:21:47.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0219-01 | WORKSTATION | X- Communication Issues 2018-6-27 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a canbus error on node four appeared prior to the procedure. A hard reset was completed, and the surgical arm was setup with no additional error messages. The case proceeded without further error. There was no patient harm and the procedure was not delayed over an hour. WORKSTATION | 2018-6-27 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a canbus error on node four appeared prior to the procedure. A hard reset was completed, and the surgical arm was setup with no additional error messages. The case proceeded without further error. There was no patient harm and the procedure was not delayed over an hour. | x communication issues ccr spinal procedure canbus error node appeared prior procedure hard reset completed surgical arm setup additional error messages case proceeded error procedure delayed hour workstation | NAV2047 | NAV4095 | NAV3249 | NAV3264 | NAV4098 | NAV2047NAV4095NAV3249NAV3264NAV4098 | True
5150 | 703606121 | XSE - Failed 3Define | 2020-02-10T20:02:57.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Failed 3Define 2020-02-10 00683761 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while performing 3Define, the camera failed at the last section. She received Error 3716 and 3017. The manufacturer representative created a working volume to continue. No patient was present. 2020-02-18 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-02-10 00683761 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while performing 3Define, the camera failed at the last section. She received Error 3716 and 3017. The manufacturer representative created a working volume to continue. No patient was present. 2020-02-18 IUD (Other): No new information. | xse failed outside procedure performing camera failed section error created working volume continue present iud new mazor x | NAV3244 | NAV2047 | NAV2054 | NAV1012 | NAV4112 | NAV3244NAV2047NAV2054NAV1012NAV4112 | True
2154 | 703226556 | X - Software issue | 2019-06-25T19:53:12.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - Software issue 1) 11:23-11:37 Software crash/freeze x4 The first crashed occurred when the surgeon began to use a bovie. The icons froze the mouse was still able to move but screen did not respond. A prompt appeared to either shut down Mazor X or wait until X responds. XCSR chose to wait until X SW responded, after a minute the prompt appeared again and XCSR chose to shutdown SW. X came back up and brought XCSR to the patient folder all 3 check marks were still checked allowing to move forward in operate process Second crash occurred and kicked XCSR to surgeon login in screen 60 seconds after 3rd crash occurred and brought back to same log in screen XCSR entered in the kit code and acquire new image and began 3 point check Fourth crash occurred when the arm was going through its 3point check it brought XCSR to the same page and 3point check was performed to completion this time 2) X eyes was not locating reference plate after several attempts to look for it. Elbow was at 120 degrees, drape was clear and pulled away. Lights were turned away as well which worked in the past. After taken the AP image with generic solution the reference plate was then found. Surgeon was very displeased that the same issue occurred two procedures in a row. MAZOR X SOFTWARE KIT VER-2.3.1 | 1) 11:23-11:37 Software crash/freeze x4 The first crashed occurred when the surgeon began to use a bovie. The icons froze the mouse was still able to move but screen did not respond. A prompt appeared to either shut down Mazor X or wait until X responds. XCSR chose to wait until X SW responded, after a minute the prompt appeared again and XCSR chose to shutdown SW. X came back up and brought XCSR to the patient folder all 3 check marks were still checked allowing to move forward in operate process Second crash occurred and kicked XCSR to surgeon login in screen 60 seconds after 3rd crash occurred and brought back to same log in screen XCSR entered in the kit code and acquire new image and began 3 point check Fourth crash occurred when the arm was going through its 3point check it brought XCSR to the same page and 3point check was performed to completion this time 2) X eyes was not locating reference plate after several attempts to look for it. Elbow was at 120 degrees, drape was clear and pulled away. Lights were turned away as well which worked in the past. After taken the AP image with generic solution the reference plate was then found. Surgeon was very displeased that the same issue occurred two procedures in a row. | x software issue software crashfreeze crashed occurred surgeon began use bovie icons froze mouse able screen did respond prompt appeared shut mazor x wait x responds xcsr chose wait x sw responded minute prompt appeared xcsr chose shutdown sw x came brought xcsr folder check marks checked allowing forward operate process second crash occurred kicked xcsr surgeon login screen seconds crash occurred brought log screen xcsr entered kit code acquire new image began point check fourth crash occurred arm going check brought xcsr page check performed completion time x eyes locating reference plate attempts look elbow degrees drape clear pulled away lights turned away worked past taken ap image generic solution reference plate surgeon displeased issue occurred procedures row mazor x software kit | NAV3247 | NAV3248 | NAV4091 | NAV3246 | NAV4099 | NAV3247NAV3248NAV4091NAV3246NAV4099 | True
2970 | 703261189 | X - Missing Parts/Dead Pixel at Install | 2019-07-16T14:58:16.000000+0000 | MazorX | NAV2054 | OUT OF BOX FAILURES | 40 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | X - Missing Parts/Dead Pixel at Install These issues were found during the install. - There was no power cable - Missing 9 arm screws for installation. - The monitor also has a dead pixel. - System came pre-calibrated and the calibrations should've been deleted. He will need to revisit the site now. We are schedulign Drew for Wednesday, March 14th. UNK MAZOR X ACCESSORIES | These issues were found during the install. - There was no power cable - Missing 9 arm screws for installation. - The monitor also has a dead pixel. - System came pre-calibrated and the calibrations should've been deleted. He will need to revisit the site now. We are schedulign Drew for Wednesday, March 14th. | x missing partsdead pixel install issues install power cable missing arm screws installation monitor dead pixel came precalibrated calibrations shouldve deleted need revisit site schedulign drew wednesday march unk mazor x accessories | NAV2053 | NAV2054 | NAV2044 | NAV3261 | NAV3233 | NAV2053NAV2054NAV2044NAV3261NAV3233 | True
5509 | 703742265 | XSE - Front Plate Damaged | 2020-05-07T15:45:12.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | KIT0838-01 | SVC KIT0838-01 FRONT PLATE UPGRADE | XSE - Front Plate Damaged 2020-05-07 00688453 (Rep): Medtronic received information regarding a guidance system outside of procedure. It was reported that the front plate needs to be replaced. No patient was present. 2020-05-14 IUD (Other): No new information. 2020-05-26 email 2020-05-26 09:00 (Other): No new information. SVC KIT0838-01 FRONT PLATE UPGRADE | 2020-05-07 00688453 (Rep): Medtronic received information regarding a guidance system outside of procedure. It was reported that the front plate needs to be replaced. No patient was present. 2020-05-14 IUD (Other): No new information. 2020-05-26 email 2020-05-26 09:00 (Other): No new information. | xse plate damaged outside procedure plate needs replaced present iud new email new svc plate upgrade | NAV4100 | NAV2058 | NAV3249 | NAV3252 | NAV3295 | NAV4100NAV2058NAV3249NAV3252NAV3295 | True
6102 | 703869964 | XSE - SRA Moves Easily/DVI Not Working | 2020-07-26T16:24:09.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - SRA Moves Easily/DVI Not Working 2020-05-15 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that the surgical arm was easily pushed out of trajectory and the DVI port of the PC was not working. The procedure was able to be successfully completed. There was no patient harm and the procedure was delayed less than an hour. KIT SVC ASM0206-01S SURGICAL ARM | 2020-05-15 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that the surgical arm was easily pushed out of trajectory and the DVI port of the PC was not working. The procedure was able to be successfully completed. There was no patient harm and the procedure was delayed less than an hour. | xse sra moves easilydvi working ccr spinal procedure treat degenerative disc fixation surgical arm easily pushed trajectory dvi port pc working procedure able successfully completed procedure delayed hour kit svc surgical arm | NAV3249 | NAV4095 | NAV2047 | NAV2058 | NAV3264 | NAV3249NAV4095NAV2047NAV2058NAV3264 | False
2384 | 703234014 | X - Drawspine issue | 2019-06-28T20:14:01.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Drawspine issue After doing the first drawspine, we received an error #295 “A position is unreachable. Consider re-drawing spine using a different vertebral body”. Following the problems we had already experienced, we decided use the generic AP position. UNK MAZOR X X-EYE CAMERA | After doing the first drawspine, we received an error #295 “A position is unreachable. Consider re-drawing spine using a different vertebral body”. Following the problems we had already experienced, we decided use the generic AP position. | x drawspine issue doing drawspine error “a position unreachable consider redrawing spine using different vertebral body” following problems experienced decided use generic ap position unk mazor x xeye camera | NAV4110 | NAV4095 | NAV3244 | NAV2047 | NAV2058 | NAV4110NAV4095NAV3244NAV2047NAV2058 | True
6624 | 703991565 | XSE - Software Bug | 2020-10-02T19:17:06.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Software Bug 2020-10-01 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that there was a software bug during the procedure. There was no patient harm. 2020-10-05 email 2020-10-05 10:45, email 2020-10-05 11:15 (Rep): Additional information received from a manufacturer representative reported that the 3Define scan failed during the procedure. The surgical arm moved to the first position, but no image was displayed on the screen. The manufacturer representative was not sure if the issue was due to hardware or software. Three 3Define scans were attempted without success. A soft restart of the system was done and the issue was resolved. The cause was not determined and there was no patient harm. The procedure was delayed less than an hour. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-10-01 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that there was a software bug during the procedure. There was no patient harm. 2020-10-05 email 2020-10-05 10:45, email 2020-10-05 11:15 (Rep): Additional information received from a manufacturer representative reported that the 3Define scan failed during the procedure. The surgical arm moved to the first position, but no image was displayed on the screen. The manufacturer representative was not sure if the issue was due to hardware or software. Three 3Define scans were attempted without success. A soft restart of the system was done and the issue was resolved. The cause was not determined and there was no patient harm. The procedure was delayed less than an hour. | xse software bug ccr spinal procedure treat degenerative disc fixation software bug procedure email email additional scan failed procedure surgical arm moved position image displayed screen sure issue hardware software scans attempted success soft restart issue resolved cause determined procedure delayed hour mazor x software installation kit | NAV3246 | NAV3247 | NAV3248 | NAV3244 | NAV3242 | NAV3246NAV3247NAV3248NAV3244NAV3242 | True
6547 | 703973140 | XSE - SRA Failed Advanced Accuracy Test | 2020-09-23T16:01:05.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - SRA Failed Advanced Accuracy Test 2020-09-23 00697293 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the advanced accuracy test prior to a case. The surgical arm failed at the first point. There was no patient involvement. 2020-09-25 x5 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-09-23 00697293 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the advanced accuracy test prior to a case. The surgical arm failed at the first point. There was no patient involvement. 2020-09-25 x5 (Other): No new information. | xse sra failed advanced accuracy test outside procedure surgical arm failed advanced accuracy test prior case surgical arm failed point new kit svc surgical arm refurb | NAV4093 | NAV3249 | NAV2054 | NAV4092 | NAV4095 | NAV4093NAV3249NAV2054NAV4092NAV4095 | True
8648 | 704457516 | XSE - Not Booting Up/Error Messages | 2021-06-18T15:52:04.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 60 | CBL0256-01 | CABLE, SH, MFM, OTHER, L=5M, MAIN SYSTEM | XSE - Not Booting Up/Error Messages 2021-Jun-18 Salesforce C00228729, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the guidance system would not boot up past the splash screen. The manufacturer representative powered down the system for five minutes and then power the system back up, but the issue was not resolved. The PC power was cycled and Error Message 4496 Robot Calibration Not Found was displayed. The representative was able to log into the software, but Error Message 3344 Emergency Stop Button Pressed was displayed. The system would not allow the representative to reset the emergency stop button. Another hard restart was done and the issue was not resolved. There was no patient involvement. 2021-06-25 IUD x6 (Other): No new information. 2021-07-18 IUD x5 (Other): No new information. 2021-08-02 IUD (Other): No new information. 2021-08-02 IUD x6 (Other): No new information. 2021-08-13 IUD x11 (Other): No new information. 2021-09-12 302559211, IUD x2 (Other): No new information. 2021-09-20 IUD (Other): No new information. 2021-10-04 IUD x2 (Other): No new information. 2021-10-24 IUD (Other): No new information. 2021-10-31 IUD x3 (Other): No new information. 2021-11-08 IUD (Other): No new information. 2021-11-08 IUD x5, 302621970 (Other): No new information. 2021-12-08 IUD (Other): No new information. CABLE, SH, MFM, OTHER, L=5M, MAIN SYSTEM | 2021-Jun-18 Salesforce C00228729, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the guidance system would not boot up past the splash screen. The manufacturer representative powered down the system for five minutes and then power the system back up, but the issue was not resolved. The PC power was cycled and Error Message 4496 Robot Calibration Not Found was displayed. The representative was able to log into the software, but Error Message 3344 Emergency Stop Button Pressed was displayed. The system would not allow the representative to reset the emergency stop button. Another hard restart was done and the issue was not resolved. There was no patient involvement. 2021-06-25 IUD x6 (Other): No new information. 2021-07-18 IUD x5 (Other): No new information. 2021-08-02 IUD (Other): No new information. 2021-08-02 IUD x6 (Other): No new information. 2021-08-13 IUD x11 (Other): No new information. 2021-09-12 302559211, IUD x2 (Other): No new information. 2021-09-20 IUD (Other): No new information. 2021-10-04 IUD x2 (Other): No new information. 2021-10-24 IUD (Other): No new information. 2021-10-31 IUD x3 (Other): No new information. 2021-11-08 IUD (Other): No new information. 2021-11-08 IUD x5, 302621970 (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse booting uperror messages salesforce iud outside p rocedure boot past splash screen powered minutes power issue resolved pc power cycled error message robot calibration displayed able log software error message emergency stop button pressed displayed allow reset emergency stop button hard restart issue resolved iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new cable sh mfm main | NAV2047 | NAV3246 | NAV3248 | NAV3256 | NAV3262 | NAV2047NAV3246NAV3248NAV3256NAV3262 | True
2044 | 703223944 | X - SOFTWARE ISSUE | 2019-06-24T19:48:33.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X - SOFTWARE ISSUE When going over the plan clicking from one vertebral body to the next the system froze the only way to overcome it is to exit and go back./ Dr. Devito T11-L5 CHOA 8.28.17 SYSTEM #25 mazor xccr preoperative setup : pressures were low had to wait a while when starting advanced accuracy test arm wouldnt unlock we had to lock and unlock numerous times. plan: Dr Devito came to the workstation and planned himself We discussed the levels and he was hesitant from the beginning it may not be a good robot case due to anatomy when clicking from one vertebrae to the other the system froze it was fixed by exiting and getting back in Mounted bed 8:10 draped 8:54 incision 8:58 mounted pt 9:54 3 define 9:55 successful first attempt draw spine 9:56 successful first attempt labeled off L2 registration ap 4 attempts obl 4 attempts successful 1st attempt for first segment when attempting to register T11 and T12 the obl image was upside down and we could not change it the system and the software wouldnt allow we were waiting for c-arm tech to obtain more images and Dr Devito decide to disconnect robot and freehand. operation: Dr Devito decided to operate only T11-L4 no longer L5 we successfully sent to L4-L1trajectories with the exception of the HEMI there was a trajectory we added at L3 it came in contact with the bridge so Dr Devito decided to proceed and disregard this. Also when lowering the arm down at the first trajectory it did a complete 360 degree rotation and didnt even move down and replicated this 3x and did not show in the simulation Post op convo: Dr Devito was happy with the results of what we did execute He was also a little annoyed intraoperatively the way the robot was moving. MAZOR X SOFTWARE KIT VER-3.0 | When going over the plan clicking from one vertebral body to the next the system froze the only way to overcome it is to exit and go back./ Dr. Devito T11-L5 CHOA 8.28.17 SYSTEM #25 mazor xccr preoperative setup : pressures were low had to wait a while when starting advanced accuracy test arm wouldnt unlock we had to lock and unlock numerous times. plan: Dr Devito came to the workstation and planned himself We discussed the levels and he was hesitant from the beginning it may not be a good robot case due to anatomy when clicking from one vertebrae to the other the system froze it was fixed by exiting and getting back in Mounted bed 8:10 draped 8:54 incision 8:58 mounted pt 9:54 3 define 9:55 successful first attempt draw spine 9:56 successful first attempt labeled off L2 registration ap 4 attempts obl 4 attempts successful 1st attempt for first segment when attempting to register T11 and T12 the obl image was upside down and we could not change it the system and the software wouldnt allow we were waiting for c-arm tech to obtain more images and Dr Devito decide to disconnect robot and freehand. operation: Dr Devito decided to operate only T11-L4 no longer L5 we successfully sent to L4-L1trajectories with the exception of the HEMI there was a trajectory we added at L3 it came in contact with the bridge so Dr Devito decided to proceed and disregard this. Also when lowering the arm down at the first trajectory it did a complete 360 degree rotation and didnt even move down and replicated this 3x and did not show in the simulation Post op convo: Dr Devito was happy with the results of what we did execute He was also a little annoyed intraoperatively the way the robot was moving. | x software issue going plan clicking vertebral body froze way overcome exit dr devito choa mazor xccr preoperative setup pressures low wait starting advanced accuracy test arm wouldnt unlock lock unlock numerous times plan dr devito came workstation planned discussed levels hesitant beginning good robot case anatomy clicking vertebrae froze fixed exiting getting mounted bed draped incision mounted pt define successful attempt draw spine successful attempt labeled registration ap attempts obl attempts successful attempt segment attempting register obl image upside change software wouldnt allow waiting carm tech obtain images dr devito decide disconnect robot freehand operation dr devito decided operate longer successfully sent exception hemi trajectory added came contact bridge dr devito decided proceed disregard lowering arm trajectory did complete degree rotation didnt replicated did simulation post op convo dr devito happy results did execute little annoyed intraoperatively way robot moving mazor x software kit | NAV3246 | NAV3248 | NAV4091 | NAV4112 | NAV3247 | NAV3246NAV3248NAV4091NAV4112NAV3247 | True
9920 | 704747740 | XSE - Screw Stripped | 2021-12-04T03:11:52.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2202-04S | KIT SVC MAS2202-04S CLAMP LINK | XSE - Screw Stripped 2021-Dec-03 Salesforce C00239157 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw was stripped. There was no patient involvement. 2021-12-16 IUD (Other): No new information. 2021-12-23 email 2021-12-23 15:15 (Rep): Additional information received from a manufacturer representative reported that the screw head was stripped. 2022-01-10 email 2022-01-10 15:16 (Rep): No new information. 2022-02-01 email 2022-02-01 15:38 (Rep): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-21 IUD (Other): No new information. KIT SVC MAS2202-04S CLAMP LINK | 2021-Dec-03 Salesforce C00239157 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw was stripped. There was no patient involvement. 2021-12-16 IUD (Other): No new information. 2021-12-23 email 2021-12-23 15:15 (Rep): Additional information received from a manufacturer representative reported that the screw head was stripped. 2022-01-10 email 2022-01-10 15:16 (Rep): No new information. 2022-02-01 email 2022-02-01 15:38 (Rep): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-21 IUD (Other): No new information. | xse screw stripped salesforce outside procedure screw stripped iud new email additional screw head stripped email new email new iud new iud new kit svc clamp link | NAV3299 | NAV2054 | NAV4090 | NAV4100 | NAV3246 | NAV3299NAV2054NAV4090NAV4100NAV3246 | True
10204 | 704820505 | XSE - Unable to Detect Stylus | 2022-01-21T00:54:54.000000+0000 | MazorX Stealth Edition | NAV4110 | NO TRACKING | 30 | KIT1284 | SVC KIT1284 X-EYE CAMERA 02R REFURB | XSE - Unable to Detect Stylus 2022-Jan-20 Salesforce C00242017 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was unable to complete the stylus calibrated. The field service engineer was able to calibrate it on 5.0 software and then it was not verifying correctly. Troubleshooting involved calibrating the stylus on 4.2 software, he downgraded and tried to recalibrate and could not get the system to get through collecting the information for the first facet. No patient was present. 2022-01-20 IUD x2 (Other): No new information. 2022-01-24 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. 2022-02-24 IUD (Other): No new information. 2022-04-04 IUD (Other): No new information. 2022-04-25 IUD (Other): No new information. SVC KIT1284 X-EYE CAMERA 02R REFURB | 2022-Jan-20 Salesforce C00242017 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was unable to complete the stylus calibrated. The field service engineer was able to calibrate it on 5.0 software and then it was not verifying correctly. Troubleshooting involved calibrating the stylus on 4.2 software, he downgraded and tried to recalibrate and could not get the system to get through collecting the information for the first facet. No patient was present. 2022-01-20 IUD x2 (Other): No new information. 2022-01-24 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. 2022-02-24 IUD (Other): No new information. 2022-04-04 IUD (Other): No new information. 2022-04-25 IUD (Other): No new information. | xse unable detect stylus salesforce outside procedure unable complete stylus calibrated field service engineer able calibrate software verifying correctly troubleshooting involved calibrating stylus software downgraded tried recalibrate collecting facet present iud new iud new iud new iud new iud new iud new svc xeye camera refurb | NAV3249 | NAV4110 | NAV4095 | NAV3299 | NAV2057 | NAV3249NAV4110NAV4095NAV3299NAV2057 | True
2275 | 703231130 | X - Shoulder issue | 2019-06-27T16:55:56.000000+0000 | MazorX | NAV4108 | SHOULDER SHIFTS | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Shoulder issue It was reported that the system has had multiple shoulder shifts over the past few cases even when the patient and system were not touched. Upon speaking with Leafar and Joel it is evident that the shoulder needs to be tightened to avoid future unnecessary shoulder shift errors. Please arrange for an engineer to come to Forbes in order to tighten the shoulder and perform shoulder calibrations. UNK MAZOR X SHOULDER | It was reported that the system has had multiple shoulder shifts over the past few cases even when the patient and system were not touched. Upon speaking with Leafar and Joel it is evident that the shoulder needs to be tightened to avoid future unnecessary shoulder shift errors. Please arrange for an engineer to come to Forbes in order to tighten the shoulder and perform shoulder calibrations. | x shoulder issue multiple shoulder shifts past cases touched speaking leafar joel evident shoulder needs tightened avoid future unnecessary shoulder shift errors arrange engineer come forbes order tighten shoulder perform shoulder calibrations unk mazor x shoulder | NAV4108 | NAV4107 | NAV3252 | NAV4106 | NAV3259 | NAV4108NAV4107NAV3252NAV4106NAV3259 | True
8085 | 704348985 | XSE - BMB Screw Stripped | 2021-04-19T20:17:06.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2188-04S | KIT SVC MAS2188-04S BM BRIDGE | XSE - BMB Screw Stripped 2021-Apr-19 Salesforce C00224901 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws on the bone mount bridge that clamp down on the Schanz bond were stripped. There was no patient involvement. 2021-04-22 IUD x3 (Other): No new information. 2021-07-28 email 2021-07-28 22:15 (Rep): Additional information received from a manufacturer representative reported the cause of the stripped screws was wear and tear from consistent use. The screws stripped a little bit every time the screw was tightened and they got to the point where they could no longer be used. KIT SVC MAS2188-04S BM BRIDGE | 2021-Apr-19 Salesforce C00224901 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws on the bone mount bridge that clamp down on the Schanz bond were stripped. There was no patient involvement. 2021-04-22 IUD x3 (Other): No new information. 2021-07-28 email 2021-07-28 22:15 (Rep): Additional information received from a manufacturer representative reported the cause of the stripped screws was wear and tear from consistent use. The screws stripped a little bit every time the screw was tightened and they got to the point where they could no longer be used. | xse bmb screw stripped salesforce outside procedure screws bone mount bridge clamp schanz bond stripped iud new email additional cause stripped screws wear tear consistent use screws stripped little bit time screw tightened got point longer kit svc bm bridge | NAV3299 | NAV4090 | NAV4100 | NAV4093 | NAV3298 | NAV3299NAV4090NAV4100NAV4093NAV3298 | True
3341 | 703319959 | XSE - Broken Screw/Loose Screwdriver | 2019-08-20T16:30:38.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1650-01 | UNIVERSAL SCREWDRIVER | XSE - Broken Screw/Loose Screwdriver 2019-08-14 SCL1902382 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the screw for the bone mount bridge was no longer working and the shaft on the screw driver had loosened from the handle. There was no patient involvement. 2019-08-29 email (Rep): Additional information received from a manufacturer representative reported the event occurred intra-op. UNIVERSAL SCREWDRIVER | 2019-08-14 SCL1902382 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the screw for the bone mount bridge was no longer working and the shaft on the screw driver had loosened from the handle. There was no patient involvement. 2019-08-29 email (Rep): Additional information received from a manufacturer representative reported the event occurred intra-op. | xse broken screwloose screwdriver prior procedure screw bone mount bridge longer working shaft screw driver loosened handle email additional occurred intraop universal screwdriver | NAV3299 | NAV3295 | NAV2054 | NAV3251 | NAV3249 | NAV3299NAV3295NAV2054NAV3251NAV3249 | True
1817 | 703218245 | X - Pressure issue | 2019-06-20T17:41:11.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Pressure issue It was reported that there is an issue with the pressure building in the system. The system would not reach the 5.0-7.3 to be fully functional and was not reach that number. We will have to repair the pneumatic panel and values. UNK MAZOR X WORKSTATION | It was reported that there is an issue with the pressure building in the system. The system would not reach the 5.0-7.3 to be fully functional and was not reach that number. We will have to repair the pneumatic panel and values. | x pressure issue issue pressure building reach fully functional reach number repair pneumatic panel values unk mazor x workstation | NAV3264 | NAV4095 | NAV2058 | NAV3261 | NAV3250 | NAV3264NAV4095NAV2058NAV3261NAV3250 | True
7038 | 704104221 | XSE - Stripped BMB Screw | 2020-12-03T23:01:34.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5196-01S | KIT SVC MEC5196-01S RNX-ACC-063 | XSE - Stripped BMB Screw 2020-12-03 00702019 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the screw used to close the alligator clamp around the schanz ball is stripped and needs to be replaced. No patient was present. 2020-12-09 email 2020-12-09 16:30, email 2020-12-10 08:01, email 2020-12-11 06:30 (Rep): No new information. 2021-01-06 email 2021-01-06 15:15 (Rep): No new information. 2021-02-04 email 2021-02-04 10:01 (Rep): No new information. KIT SVC MEC5196-01S RNX-ACC-063 | 2020-12-03 00702019 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the screw used to close the alligator clamp around the schanz ball is stripped and needs to be replaced. No patient was present. 2020-12-09 email 2020-12-09 16:30, email 2020-12-10 08:01, email 2020-12-11 06:30 (Rep): No new information. 2021-01-06 email 2021-01-06 15:15 (Rep): No new information. 2021-02-04 email 2021-02-04 10:01 (Rep): No new information. | xse stripped bmb screw accessory outside procedure screw close alligator clamp schanz ball stripped needs replaced present email email email new email new email new kit svc | NAV3299 | NAV4090 | NAV3295 | NAV2054 | NAV3252 | NAV3299NAV4090NAV3295NAV2054NAV3252 | True
9649 | 704690525 | XSE - Damaged Instrument | 2021-10-30T02:26:33.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2202 | MAS2202 CLAMP LINK | XSE - Damaged Instrument 2021-Oct-27 Salesforce C00237089 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the target extender or clamp link was damaged. There was no patient involvement. 2021-11-01 IUD (Other): The clamp link was the one that was damaged. 2021-11-08 IUD (Other): No new information. 2021-11-17 IUD (Other): No new information. 2021-11-30 IUD (Other): No new information. 2021-12-07 email 2021-12-07 14:31, image 2021-12-07 14:31 (Rep): Additional information received from a manufacturer representative reported that the cause of the damage to the clamp link was wear and tear. 2021-12-15 email 2021-12-15 09:46 (Rep): No new information. 2021-12-23 email 2021-12-23 17:30 (Rep): Additional information received from a manufacturer representative reported that the instrument could not be found. MAS2202 CLAMP LINK | 2021-Oct-27 Salesforce C00237089 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the target extender or clamp link was damaged. There was no patient involvement. 2021-11-01 IUD (Other): The clamp link was the one that was damaged. 2021-11-08 IUD (Other): No new information. 2021-11-17 IUD (Other): No new information. 2021-11-30 IUD (Other): No new information. 2021-12-07 email 2021-12-07 14:31, image 2021-12-07 14:31 (Rep): Additional information received from a manufacturer representative reported that the cause of the damage to the clamp link was wear and tear. 2021-12-15 email 2021-12-15 09:46 (Rep): No new information. 2021-12-23 email 2021-12-23 17:30 (Rep): Additional information received from a manufacturer representative reported that the instrument could not be found. | xse damaged instrument salesforce outside procedure target extender clamp link damaged iud clamp link damaged iud new iud new iud new email image additional cause damage clamp link wear tear email new email additional instrument clamp link | NAV3299 | NAV3295 | NAV2054 | NAV3258 | NAV2058 | NAV3299NAV3295NAV2054NAV3258NAV2058 | True
10417 | 704875193 | XSE - System Unintended Restart | 2022-02-21T21:08:04.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT1378 | KIT1378 USA REGIONAL KIT | XSE - System Unintended Restart 2022-Feb-16 Salesforce C00243979 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during a surgeon demo, the system displayed the message “Your PC ran into a problem and needs to restart” shown below. The system was then not functional and a hard restart was needed. 2022-02-25 IUD (Other): No new information. 2022-03-09 e1, e2, e3, IUD (Other): No new information. 2022-03-15 e4 (Rep): No new information. 2022-03-24 IUD (Other): No new information. 2022-04-04 IUD (Other): No new information. KIT1378 USA REGIONAL KIT | 2022-Feb-16 Salesforce C00243979 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during a surgeon demo, the system displayed the message “Your PC ran into a problem and needs to restart” shown below. The system was then not functional and a hard restart was needed. 2022-02-25 IUD (Other): No new information. 2022-03-09 e1, e2, e3, IUD (Other): No new information. 2022-03-15 e4 (Rep): No new information. 2022-03-24 IUD (Other): No new information. 2022-04-04 IUD (Other): No new information. | xse unintended restart salesforce outside procedure surgeon demo displayed message “your pc ran problem needs restart” shown functional hard restart needed iud new iud new new iud new iud new usa regional kit | NAV3246 | NAV3249 | NAV3299 | NAV3262 | NAV2047 | NAV3246NAV3249NAV3299NAV3262NAV2047 | False
2000 | 703223618 | XSE - screen flickering/elevator | 2019-06-24T17:18:54.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | ASM0215-04 | MAZOR-X PC STATION | XSE - screen flickering/elevator 2019-Jun-19 SCL1901807 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon screen was flicking and going blank and needed a hard restart to fix. The elevator also would no longer raise or lower. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. MAZOR-X PC STATION | 2019-Jun-19 SCL1901807 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon screen was flicking and going blank and needed a hard restart to fix. The elevator also would no longer raise or lower. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | xse screen flickeringelevator spinal procedure surgeon screen flicking going blank needed hard restart fix elevator longer raise lower additional complications anticipated mazorx pc station | NAV3262 | NAV3253 | NAV2054 | NAV3261 | NAV1006 | NAV3262NAV3253NAV2054NAV3261NAV1006 | True
8072 | 704346151 | XSE ¦ Lost Instruments | 2021-04-16T20:16:53.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 40 | MAS2271-04S | KIT SVC MAS2271-04S SCHANZ BOND | XSE ¦ Lost Instruments 2021-Apr-16 Salesforce C00224835 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that parts are lost and needed to complete trays. No patient was present. 2021-04-23 IUD x2 (Other): No new information. 2021-05-16 IUD (Other): No new information. KIT SVC MAS2271-04S SCHANZ BOND | 2021-Apr-16 Salesforce C00224835 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that parts are lost and needed to complete trays. No patient was present. 2021-04-23 IUD x2 (Other): No new information. 2021-05-16 IUD (Other): No new information. | xse lost instruments salesforce outside procedure parts lost needed complete trays present iud new iud new kit svc schanz bond | NAV2053 | NAV3256 | NAV3249 | NAV2047 | NAV2058 | NAV2053NAV3256NAV3249NAV2047NAV2058 | True
3753 | 703433474 | X - DVD won't eject | 2019-10-25T16:25:14.000000+0000 | MazorX | NAV3257 | CD DRIVE ISSUES | 20 | DSD0038-01 | DVD, PORTABLE DVD REWRITER, GP70NS50 | X - DVD won't eject 2019-Oct-23 SCL1903167 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they had to press and hold the DVD button for the disc to eject. It doesn’t eject if it is just pressed once. There was no patient harm or additional complications reported or anticipated as a result of the event. DVD, PORTABLE DVD REWRITER, GP70NS50 | 2019-Oct-23 SCL1903167 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they had to press and hold the DVD button for the disc to eject. It doesn’t eject if it is just pressed once. There was no patient harm or additional complications reported or anticipated as a result of the event. | x dvd wont eject spinal procedure press hold dvd button disc eject doesn’t eject just pressed additional complications anticipated dvd portable dvd rewriter | NAV3257 | NAV3295 | NAV2054 | NAV1006 | NAV4112 | NAV3257NAV3295NAV2054NAV1006NAV4112 | True
8966 | 704516357 | XSE - OOBF/Loss of Pressure | 2021-07-22T12:11:56.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE - OOBF/Loss of Pressure 2021-Jul-22 Salesforce C00230596 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that when pressing the manipulator or positioner buttons, the workstation hood lost pressure and started to move down. There was no patient involvement. 2021-Jul-27 IUD x2 (Other, For): No new information. 2021-08-03 e1, e2, e3, e4 (Rep, For): Additional information received from a manufacturer representative reported that the hood dropped for a few centimeters due to losing pressure when mounting the surgical system to the bed. The issue happened when releasing the manipulator button. The magnetic sensor of the hood was adjusted, but the issue was not resolved. A new pneumatic panel was ordered due to finding an air leak. The new pneumatic panel was defective, didn't provide positioner pressure, and just showed -555 on the system indications tab. The representative tried to reinstall the software, but the issue was not resolved. The old pneumatic panel was placed on the system and the zone where the air leak was coming out was taped. No further pressure drops were experienced during testing. 2021-08-10 IUD x4 (Other, For): No new information. 2021-08-24 IUD (Other, For): No new information. 2021-09-08 IUD (Other, For): No new information. 2021-10-12 e5 (Other, For): No new information. 2021-10-18 e6, e7, e8 (Rep, For): No new information. 2021-10-28 e9, e10, e11 (Rep, For): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2021-Jul-22 Salesforce C00230596 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that when pressing the manipulator or positioner buttons, the workstation hood lost pressure and started to move down. There was no patient involvement. 2021-Jul-27 IUD x2 (Other, For): No new information. 2021-08-03 e1, e2, e3, e4 (Rep, For): Additional information received from a manufacturer representative reported that the hood dropped for a few centimeters due to losing pressure when mounting the surgical system to the bed. The issue happened when releasing the manipulator button. The magnetic sensor of the hood was adjusted, but the issue was not resolved. A new pneumatic panel was ordered due to finding an air leak. The new pneumatic panel was defective, didn't provide positioner pressure, and just showed -555 on the system indications tab. The representative tried to reinstall the software, but the issue was not resolved. The old pneumatic panel was placed on the system and the zone where the air leak was coming out was taped. No further pressure drops were experienced during testing. 2021-08-10 IUD x4 (Other, For): No new information. 2021-08-24 IUD (Other, For): No new information. 2021-09-08 IUD (Other, For): No new information. 2021-10-12 e5 (Other, For): No new information. 2021-10-18 e6, e7, e8 (Rep, For): No new information. 2021-10-28 e9, e10, e11 (Rep, For): No new information. | xse oobfloss pressure salesforce outside proc edure pressing manipulator positioner buttons workstation hood lost pressure started iud new additional hood dropped centimeters losing pressure mounting surgical bed issue happened releasing manipulator button magnetic sensor hood adjusted issue resolved new pneumatic panel ordered finding air leak new pneumatic panel defective didnt provide positioner pressure just showed indications tab tried reinstall software issue resolved old pneumatic panel placed zone air leak coming taped pressure drops experienced testing iud new iud new iud new new new new kit svc pneumatic panel | NAV3264 | NAV2054 | NAV4105 | NAV3246 | NAV3251 | NAV3264NAV2054NAV4105NAV3246NAV3251 | True
2576 | 703241700 | X - Registration issue | 2019-07-03T20:01:02.000000+0000 | MazorX | NAV4112 | NO REGISTRATION | 20 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X - Registration issue Registration issues (see detailed report) UNKNOWN MAZOR X SW | Registration issues (see detailed report) | x registration issue registration issues unknown mazor x sw | NAV3242 | NAV3246 | NAV4091 | NAV4112 | NAV3247 | NAV3242NAV3246NAV4091NAV4112NAV3247 | True
8706 | 704465124 | XSE - OOBF/Black Monitor Screen | 2021-06-23T15:54:53.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 70 | ASM0275 | KIT SVC ASM0275 PNEUMATIC PANEL 02S NEW | XSE - OOBF/Black Monitor Screen 2021-Jun-23 Salesforce C00228961 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the workstation monitor was completely black and could not be turned on. There was no patient involvement. 2021-06-23 IUD x10 (Rep): Additional information received from a manufacturer representative reported that the monitor was bad at in stall. 2021-07-19 IUDx10 (other): No new information. 2021-08-09 IUD x5, e1, e2, e3, e4, e5 (Rep): Additional information received from a manufacturer representative reported the surgeon screen worked intermittently before it was replaced. A boom monitor was being used. 2021-08-25 e7, e8 (Rep): Additional information received from a manufacturer representative reported that one of the screens was used for troubleshooting and was returned to be disposed of. 2021-08-31 e9, e10 (Rep): No new information. 2021-10-20 IUD x2 (Other): No new information. 2021-10-13 00714021, IUD (Other): No new information. Information related to ENT was excluded from this PE. KIT SVC ASM0275 PNEUMATIC PANEL 02S NEW | 2021-Jun-23 Salesforce C00228961 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the workstation monitor was completely black and could not be turned on. There was no patient involvement. 2021-06-23 IUD x10 (Rep): Additional information received from a manufacturer representative reported that the monitor was bad at in stall. 2021-07-19 IUDx10 (other): No new information. 2021-08-09 IUD x5, e1, e2, e3, e4, e5 (Rep): Additional information received from a manufacturer representative reported the surgeon screen worked intermittently before it was replaced. A boom monitor was being used. 2021-08-25 e7, e8 (Rep): Additional information received from a manufacturer representative reported that one of the screens was used for troubleshooting and was returned to be disposed of. 2021-08-31 e9, e10 (Rep): No new information. 2021-10-20 IUD x2 (Other): No new information. 2021-10-13 00714021, IUD (Other): No new information. Information related to ENT was excluded from this PE. | xse oobfblack monitor screen salesforce outside procedure workstation monitor completely black turned iud additional monitor bad stall new iud additional surgeon screen worked intermittently replaced boom monitor additional screens troubleshooting returned disposed new iud new iud new related ent excluded pe kit svc pneumatic panel new | NAV3261 | NAV2054 | NAV3253 | NAV2058 | NAV3263 | NAV3261NAV2054NAV3253NAV2058NAV3263 | True
9157 | 704561871 | XSE ¦ Localizer Disconnected | 2021-08-17T21:42:46.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 20 | HRS0350-01S | KIT SVC HRS0350-01S NAV CONNECT CABLE | XSE ¦ Localizer Disconnected 2021-Aug-17 Salesforce C00233067 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the localizer was disconnected. The camera was still getting power and was "ON" in the settings. The manufacturer representative checked the internal cables and confirmed everything was securely connected. When he swapped the cable from the camera to the workstation the issue resolved. No patient was present. 2021-Sep-07 IUD (rep): It was reported that the new cable resolved the issue. KIT SVC HRS0350-01S NAV CONNECT CABLE | 2021-Aug-17 Salesforce C00233067 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the localizer was disconnected. The camera was still getting power and was "ON" in the settings. The manufacturer representative checked the internal cables and confirmed everything was securely connected. When he swapped the cable from the camera to the workstation the issue resolved. No patient was present. 2021-Sep-07 IUD (rep): It was reported that the new cable resolved the issue. | xse localizer disconnected salesforce outside procedure localizer disconnected camera getting power settings checked internal cables confirmed securely connected swapped cable camera workstation issue resolved present iud new cable resolved issue kit svc nav connect cable | NAV3256 | NAV2047 | NAV2058 | NAV3264 | NAV3257 | NAV3256NAV2047NAV2058NAV3264NAV3257 | True
173 | 703023973 | X - Incorrect Image Orientation | 2019-02-25T13:55:26.000000+0000 | MazorX | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Incorrect Image Orientation 2019-02-12 SCL1900403, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during an anterior lumbar interbody fusion (ALIF) procedure. It was reported that the oblique image was oriented incorrectly when attempting segmentation and registration during the procedure. The patient was placed in the prone position for a one level (L5 - S1) posterior fixation to back up the ALIF. The pre-operation CT to flouro protocol was followed. Prior to placing an 18 mm high by 18 degree lordotic cage, the disc space was marginally collapsed. The patient was then flipped prone and prepped for the posterior percutaneous fixation. The C-arm was calibrated with success and right sided entry was selected. The marker was loaded correctly onto the target extender and a 3Define scan was done followed by Draw Spine. The AP trajectory was selected, the arm was sent and an image was successfully captured and transferred to the workstation. The target extender was removed and the arm was sent to the oblique trajectory. The target extender was placed and an oblique image was captured and successfully transferred to the workstation. During registration and segmentation, the oblique image was found to be completely flipped right to left. The AP image was unaffected. The manufacturer representative attempted to merge the images, but they had difficulty so the plan was checked to make sure the correct alignment was chosen when planning the screws. The plan was correct. The manufacturer representative tried to match the images again, but that did not work. Another 3Define scan and Draw Spine were done followed by taking new AP and oblique images. Registration and segmentation was tried again, but the oblique image was flipped right to left again. Segmentation was able to be completed; however, the ALIF cage had altered the anatomy so the disc space had to be manually adjusted due to added disc height from the cage in order to advance to approving the levels. The manufacturing representative was able to successfully approve the L5 screws bilateral as well as the left S1 screw, but the right S1 screw was highlighted yellow. When the arm was going to be sent to trajectories, the right S1 screw had turned red. The manufacturer representative tried to manipulate the arm guide to short the distance from the field and move into the long tools category, but the screw would not turn green. The surgeon did not want to use the robot for 3 screws so he aborted the guidance system and used navigation to implant the screws. There was no patient harm, but the procedure was delayed by one hour. The case was completed using navigation to place the implants. 2019-02-24 email 2019-03-19 08:11 (Rep): Additional information received from a manufacturer representative reported the issue was resolved and the oblique image flipping when the C-arm entered from the right hand side was normal. 2019-03-28 e1 (Rep): No new information. MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-02-12 SCL1900403, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during an anterior lumbar interbody fusion (ALIF) procedure. It was reported that the oblique image was oriented incorrectly when attempting segmentation and registration during the procedure. The patient was placed in the prone position for a one level (L5 - S1) posterior fixation to back up the ALIF. The pre-operation CT to flouro protocol was followed. Prior to placing an 18 mm high by 18 degree lordotic cage, the disc space was marginally collapsed. The patient was then flipped prone and prepped for the posterior percutaneous fixation. The C-arm was calibrated with success and right sided entry was selected. The marker was loaded correctly onto the target extender and a 3Define scan was done followed by Draw Spine. The AP trajectory was selected, the arm was sent and an image was successfully captured and transferred to the workstation. The target extender was removed and the arm was sent to the oblique trajectory. The target extender was placed and an oblique image was captured and successfully transferred to the workstation. During registration and segmentation, the oblique image was found to be completely flipped right to left. The AP image was unaffected. The manufacturer representative attempted to merge the images, but they had difficulty so the plan was checked to make sure the correct alignment was chosen when planning the screws. The plan was correct. The manufacturer representative tried to match the images again, but that did not work. Another 3Define scan and Draw Spine were done followed by taking new AP and oblique images. Registration and segmentation was tried again, but the oblique image was flipped right to left again. Segmentation was able to be completed; however, the ALIF cage had altered the anatomy so the disc space had to be manually adjusted due to added disc height from the cage in order to advance to approving the levels. The manufacturing representative was able to successfully approve the L5 screws bilateral as well as the left S1 screw, but the right S1 screw was highlighted yellow. When the arm was going to be sent to trajectories, the right S1 screw had turned red. The manufacturer representative tried to manipulate the arm guide to short the distance from the field and move into the long tools category, but the screw would not turn green. The surgeon did not want to use the robot for 3 screws so he aborted the guidance system and used navigation to implant the screws. There was no patient harm, but the procedure was delayed by one hour. The case was completed using navigation to place the implants. 2019-02-24 email 2019-03-19 08:11 (Rep): Additional information received from a manufacturer representative reported the issue was resolved and the oblique image flipping when the C-arm entered from the right hand side was normal. 2019-03-28 e1 (Rep): No new information. | x incorrect image orientation anterior lumbar interbody fusion procedure oblique image oriented incorrectly attempting segmentation registration procedure placed prone position level posterior fixation alif preoperation ct flouro protocol followed prior placing mm high degree lordotic cage disc space marginally collapsed flipped prone prepped posterior percutaneous fixation carm calibrated success right sided entry selected marker loaded correctly target extender scan followed draw spine ap trajectory selected arm sent image successfully captured transferred workstation target extender removed arm sent oblique trajectory target extender placed oblique image captured successfully transferred workstation registration segmentation oblique image completely flipped right left ap image unaffected attempted merge images difficulty plan checked make sure correct alignment chosen planning screws plan correct tried match images did work scan draw spine followed taking new ap oblique images registration segmentation tried oblique image flipped right left segmentation able completed alif cage altered anatomy disc space manually adjusted added disc height cage order advance approving levels manufacturing able successfully approve screws bilateral left screw right screw highlighted yellow arm going sent trajectories right screw turned red tried manipulate arm guide short distance field long tools category screw turn green surgeon did want use robot screws aborted navigation implant screws procedure delayed hour case completed using navigation place implants email additional issue resolved oblique image flipping carm entered right hand normal new mazor x software kit | NAV3237 | NAV3242 | NAV3246 | NAV3257 | NAV3258 | NAV3237NAV3242NAV3246NAV3257NAV3258 | False
2385 | 703234016 | X- Deviation | 2019-06-28T19:39:44.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X- Deviation All right side trajectories medial. This was due to soft tissue pressure and unilateral retraction. (see report) MAZOR X SOFTWARE KIT VER-2.3 | All right side trajectories medial. This was due to soft tissue pressure and unilateral retraction. (see report) | x deviation right trajectories medial soft tissue pressure unilateral retraction mazor x software kit | NAV4091 | NAV5000 | NAV3246 | NAV3242 | NAV3247 | NAV4091NAV5000NAV3246NAV3242NAV3247 | True
4474 | 703545781 | X - Registration Issues and No Tracking | 2020-01-06T16:56:19.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Registration Issues and No Tracking 2018-07-03 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for three vertebrae on level L3 - L5. During the procedure, drawspine was skipped due to X-eye not seeing square frame. 3Define was successful. First registration failed. This was likely due to previous hardware masking beads. Pain pump was also in the images and probably complicated things as well. Trajectory was drilled accurately around existing hardware that remained implanted. No patient harm was reported. UNK MAZOR X X-EYE CAMERA | 2018-07-03 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for three vertebrae on level L3 - L5. During the procedure, drawspine was skipped due to X-eye not seeing square frame. 3Define was successful. First registration failed. This was likely due to previous hardware masking beads. Pain pump was also in the images and probably complicated things as well. Trajectory was drilled accurately around existing hardware that remained implanted. No patient harm was reported. | x registration issues tracking ccr spinal procedure case vertebrae level procedure drawspine skipped xeye seeing square frame successful registration failed likely previous hardware masking beads pain pump images probably complicated things trajectory drilled accurately existing hardware remained implanted unk mazor x xeye camera | NAV4110 | NAV4109 | NAV3254 | NAV3299 | NAV3242 | NAV4110NAV4109NAV3254NAV3299NAV3242 | True
6296 | 703906165 | XSE - SRA Fails Advanced Accuracy Test | 2020-08-17T05:36:49.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 40 | MAS2585-01S | KIT SVC MAS2585-01S RIGHT CURTAIN ASSY | XSE - SRA Fails Advanced Accuracy Test 2020-08-14 00693465, 00694660 (Rep): Medtronic received information regarding a guidance system during preventative maintenance. It was reported that the surgical arm failed the advanced accuracy due to a loose joint 6. Joint 6 could be moved left or right and had too much play in the joint to pass the advanced accuracy test. The left positioner button was broken and the handle was replaced. No patient was present. 2020-08-17 IUD x2 (Other): No new information. 2020-08-24 IUD x5 (Other): No new information. 2020-09-09 IUD x5 (Other): No new information. 2020-09-22 IUD x8 (Other): No new information. 2020-10-09 IUD x9 (Other): No new information. 2020-11-10 IUD x2 (Other): No new information. 2020-11-16 IUD x2 (Other): No new information. KIT SVC MAS2585-01S RIGHT CURTAIN ASSY | 2020-08-14 00693465, 00694660 (Rep): Medtronic received information regarding a guidance system during preventative maintenance. It was reported that the surgical arm failed the advanced accuracy due to a loose joint 6. Joint 6 could be moved left or right and had too much play in the joint to pass the advanced accuracy test. The left positioner button was broken and the handle was replaced. No patient was present. 2020-08-17 IUD x2 (Other): No new information. 2020-08-24 IUD x5 (Other): No new information. 2020-09-09 IUD x5 (Other): No new information. 2020-09-22 IUD x8 (Other): No new information. 2020-10-09 IUD x9 (Other): No new information. 2020-11-10 IUD x2 (Other): No new information. 2020-11-16 IUD x2 (Other): No new information. | xse sra fails advanced accuracy test preventative maintenance surgical arm failed advanced accuracy loose joint joint moved left right play joint pass advanced accuracy test left positioner button broken handle replaced present iud new iud new iud new iud new iud new iud new iud new kit svc right curtain assy | NAV3251 | NAV4093 | NAV2054 | NAV2058 | NAV4097 | NAV3251NAV4093NAV2054NAV2058NAV4097 | True
847 | 703178468 | X- Accessory Issues | 2019-05-29T15:24:59.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2351-03 | SCHANZ ARM | X- Accessory Issues When doctor was traying to inserting Schanz Pin through Schanz Arm, the doctor found the pin was unable to pass through the hole. The doctor swap the arm and found the pin could pass through the other arm hole. SCHANZ ARM | When doctor was traying to inserting Schanz Pin through Schanz Arm, the doctor found the pin was unable to pass through the hole. The doctor swap the arm and found the pin could pass through the other arm hole. | x accessory issues doctor traying inserting schanz pin schanz arm doctor pin unable pass hole doctor swap arm pin pass arm hole schanz arm | NAV3299 | NAV3249 | NAV4100 | NAV3295 | NAV4095 | NAV3299NAV3249NAV4100NAV3295NAV4095 | True
4898 | 703585130 | X - Deviation Left L2 | 2020-01-28T16:54:24.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation Left L2 2018-11-30 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L2-L3. It was reported that shoulder calibration was checked, advanced accuracy was completed, and the surgical arm setup was completed. Registration was successful on the basic algorithm with green values at all levels. All trajectories were sent to mark incisions. All wires were placed and all looked good in lateral x-ray. On the AP x-ray, left L2 was lateral to the plan and lateral to the pedicle. Deviation is noted to be due to skiving. Left L2 was the first trajectory placed. The surgeon placed L2 freehand. All other screws were placed safely and accurately per plan. Placements were verified in AP and lateral, all screws were stimulated to acceptable values. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-11-30 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L2-L3. It was reported that shoulder calibration was checked, advanced accuracy was completed, and the surgical arm setup was completed. Registration was successful on the basic algorithm with green values at all levels. All trajectories were sent to mark incisions. All wires were placed and all looked good in lateral x-ray. On the AP x-ray, left L2 was lateral to the plan and lateral to the pedicle. Deviation is noted to be due to skiving. Left L2 was the first trajectory placed. The surgeon placed L2 freehand. All other screws were placed safely and accurately per plan. Placements were verified in AP and lateral, all screws were stimulated to acceptable values. There was no patient harm and the procedure was not delayed over an hour. | x deviation left ccr spinal procedure treat degenerative disk disease fixation shoulder calibration checked advanced accuracy completed surgical arm setup completed registration successful basic algorithm green values levels trajectories sent mark incisions wires placed looked good lateral xray ap xray left lateral plan lateral pedicle deviation noted skiving left trajectory placed surgeon placed freehand screws placed safely accurately plan placements verified ap lateral screws stimulated acceptable values procedure delayed hour mazor x | NAV4091 | NAV5000 | NAV3242 | NAV4095 | NAV3247 | NAV4091NAV5000NAV3242NAV4095NAV3247 | True
6336 | 703917473 | XSE - CANBUS Error | 2020-08-21T21:40:28.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error 2020-08-21 00695058 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that when they were mounting the system to the patient when it was noted that the shoulder locked. The site had then done a soft restart of the system which led then to the canbus error on the system. The site had then did a hard shutdown of the system and waited 5 min. Upon boot up the system was working as intended. The manufacturer representative had recommended not to use the guidance system for the rest of the case but the doctor decided to proceed anyways. After nearly completing the case the system had done another canbus error so the system was shut down and brought back up. Site had quickly removed the system from the bed. There was a 25 min delay in the case. There was no impact on the patient outcome. 2020-08-25 IUD x3 (Other): No new information. 2020-09-10 IUD x2 (Other): No new information. 2020-09-30 IUD (Other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2020-08-21 00695058 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that when they were mounting the system to the patient when it was noted that the shoulder locked. The site had then done a soft restart of the system which led then to the canbus error on the system. The site had then did a hard shutdown of the system and waited 5 min. Upon boot up the system was working as intended. The manufacturer representative had recommended not to use the guidance system for the rest of the case but the doctor decided to proceed anyways. After nearly completing the case the system had done another canbus error so the system was shut down and brought back up. Site had quickly removed the system from the bed. There was a 25 min delay in the case. There was no impact on the patient outcome. 2020-08-25 IUD x3 (Other): No new information. 2020-09-10 IUD x2 (Other): No new information. 2020-09-30 IUD (Other): No new information. | xse canbus error intraoperatively spinal procedure mounting noted shoulder locked site soft restart led canbus error site did hard shutdown waited min boot working intended recommended use rest case doctor decided proceed anyways nearly completing case canbus error shut brought site quickly removed bed min delay case impact outcome iud new iud new iud new svc main cable | NAV2047 | NAV3261 | NAV3249 | NAV2057 | NAV3250 | NAV2047NAV3261NAV3249NAV2057NAV3250 | True
9996 | 704767004 | XSE ¦ Stripped Schanz Arm Screw | 2021-12-15T20:29:39.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 40 | MEC5750-01S | KIT SVC MEC5750-01S SCH CON SCREW | XSE ¦ Stripped Schanz Arm Screw 2021-Dec-15 Salesforce C00239785 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when he tried to get the Schanz arm off he system the screw stripped. He tried an extractor tool and was unable to remove it from the system. No patient present. 2021-12-16 IUD x2 (Other): No new information. 2022-01-05 IUD x2 (Other): No new information. 2022-01-25 IUD (Other): No new information. 2022-02-03 IUD (Other): No new information. 2022-02-10 IUD (Other): No new information. KIT SVC MEC5750-01S SCH CON SCREW | 2021-Dec-15 Salesforce C00239785 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when he tried to get the Schanz arm off he system the screw stripped. He tried an extractor tool and was unable to remove it from the system. No patient present. 2021-12-16 IUD x2 (Other): No new information. 2022-01-05 IUD x2 (Other): No new information. 2022-01-25 IUD (Other): No new information. 2022-02-03 IUD (Other): No new information. 2022-02-10 IUD (Other): No new information. | xse stripped schanz arm screw salesforce outside procedure tried schanz arm screw stripped tried extractor tool unable remove present iud new iud new iud new iud new iud new kit svc sch screw | NAV3299 | NAV4100 | NAV2054 | NAV4090 | NAV3264 | NAV3299NAV4100NAV2054NAV4090NAV3264 | True
6944 | 704074482 | XSE - Unable To Launch Software | 2020-11-17T18:51:00.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - Unable To Launch Software 2020-11-17 00701051 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that before a clinical case the site was having issues powering on the system fully. It was noted that they were unable to get passed black screens on the system. The manufacturer representative noted that the software was unable to be launched with and while speaking to the field service engineer over the phone. The case was noted to be cancelled. No patient was present. 2020-11-18 IUD x6 (Other): No new information. 2020-12-01 IUD x6 (Other): No new information. 2020-12-21 IUD x6 (Other): No new information. 2021-01-26 IUD (Other): No new information. 2021-02-04 IUD (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2020-11-17 00701051 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that before a clinical case the site was having issues powering on the system fully. It was noted that they were unable to get passed black screens on the system. The manufacturer representative noted that the software was unable to be launched with and while speaking to the field service engineer over the phone. The case was noted to be cancelled. No patient was present. 2020-11-18 IUD x6 (Other): No new information. 2020-12-01 IUD x6 (Other): No new information. 2020-12-21 IUD x6 (Other): No new information. 2021-01-26 IUD (Other): No new information. 2021-02-04 IUD (Other): No new information. | xse unable launch software outside procedure clinical case site having issues powering fully noted unable passed black screens noted software unable launched speaking field service engineer phone case noted cancelled present iud new iud new iud new iud new iud new mazor x pc station | NAV3262 | NAV4104 | NAV3247 | NAV4112 | NAV3250 | NAV3262NAV4104NAV3247NAV4112NAV3250 | False
1430 | 703200331 | X - Schanz bond screw stripped | 2019-06-11T18:07:30.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC4530-01 | SLOTTED BALL ARM TARGET SCREW | X - Schanz bond screw stripped 2019-Jun-04 SCL1901657 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Schanz Bond screw stripped out. There was no patient harm reported. There were no additional complications reported or anticipated as a result of the event. SLOTTED BALL ARM TARGET SCREW | 2019-Jun-04 SCL1901657 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Schanz Bond screw stripped out. There was no patient harm reported. There were no additional complications reported or anticipated as a result of the event. | x schanz bond screw stripped spinal procedure schanz bond screw stripped additional complications anticipated slotted ball arm target screw | NAV3295 | NAV3299 | NAV4090 | NAV4100 | NAV3246 | NAV3295NAV3299NAV4090NAV4100NAV3246 | True
4677 | 703561812 | X - Arm stuck during 3Define | 2020-01-14T20:03:10.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - Arm stuck during 3Define 2020-1-12 Intake Form, 00682027(Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm got stuck after the second image of the 3Define scan. The live view of the camera was showing on the screen, but nothing was moving forward. The 3Define scan was resent, and the scan was completed. There was no patient harm and the procedure was not delayed over an hour. 2020-01-19 Email 2020-01-02:00 (Rep): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-1-12 Intake Form, 00682027(Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm got stuck after the second image of the 3Define scan. The live view of the camera was showing on the screen, but nothing was moving forward. The 3Define scan was resent, and the scan was completed. There was no patient harm and the procedure was not delayed over an hour. 2020-01-19 Email 2020-01-02:00 (Rep): No new information. | x arm stuck spinal procedure surgical arm got stuck second image scan live view camera showing screen moving forward scan resent scan completed procedure delayed hour email new unk mazor x camera | NAV3244 | NAV4101 | NAV3299 | NAV4096 | NAV1006 | NAV3244NAV4101NAV3299NAV4096NAV1006 | True
9052 | 704540673 | XSE - SRA Joint Shift Error | 2021-08-05T00:20:25.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Joint Shift Error 2021-Aug-04 Salesforce C00231459 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while on site and doing a system checkout, the field service engineer was getting a joint shift error in joint 5. He was unable to resolve with field motor calibrations. No patient was present. 2021-08-05 IUD x5 (Other): No new information. 2021-10-20 302603222 (Other): No new information. 2021-10-21 IUD x5 (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Aug-04 Salesforce C00231459 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while on site and doing a system checkout, the field service engineer was getting a joint shift error in joint 5. He was unable to resolve with field motor calibrations. No patient was present. 2021-08-05 IUD x5 (Other): No new information. 2021-10-20 302603222 (Other): No new information. 2021-10-21 IUD x5 (Other): No new information. | xse sra joint shift error salesforce outside procedure site doing checkout field service engineer getting joint shift error joint unable resolve field motor calibrations present iud new new iud new kit svc surgical arm refurb | NAV4098 | NAV3249 | NAV2054 | NAV1006 | NAV4092 | NAV4098NAV3249NAV2054NAV1006NAV4092 | True
5325 | 703648908 | XSE - Surgical Arm Error | 2020-03-04T12:16:28.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Surgical Arm Error 2020-MAR-04 00685607 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a quote for a replacement arm was requested due to some difficult arm error received. There was no patient present when this issue was identified. 2020-03-04 IUD x3, Email 2020-03-09 15:01 (Rep): Additional information received from a manufacturer representative reported that there was an error during initial Startup where they initially had a Joint 2 error which was fixed by a hard restart. The next issue was the system was working fine and a nurse unplugged the system while working. When rebooted the system was taking in excess of 5-8 minutes to start up and the software for Mazor was not working. Multiple hard restarts were tried but did not fix the issue. The representative tried another outlet and a restart, but this too did not fix the issue. The representative finally let the system stay on for an extended period of time and they were finally able to get the software up and running. 2020-03-30 Email 2020-03-30 14:15 (Rep): Additional information received from a manufacturer representative reported that the issues occurred before the case and the error code was 4059. 2020-04-16 email 2020-04-16 11:30 x2, e1, e2 (Rep): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-MAR-04 00685607 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a quote for a replacement arm was requested due to some difficult arm error received. There was no patient present when this issue was identified. 2020-03-04 IUD x3, Email 2020-03-09 15:01 (Rep): Additional information received from a manufacturer representative reported that there was an error during initial Startup where they initially had a Joint 2 error which was fixed by a hard restart. The next issue was the system was working fine and a nurse unplugged the system while working. When rebooted the system was taking in excess of 5-8 minutes to start up and the software for Mazor was not working. Multiple hard restarts were tried but did not fix the issue. The representative tried another outlet and a restart, but this too did not fix the issue. The representative finally let the system stay on for an extended period of time and they were finally able to get the software up and running. 2020-03-30 Email 2020-03-30 14:15 (Rep): Additional information received from a manufacturer representative reported that the issues occurred before the case and the error code was 4059. 2020-04-16 email 2020-04-16 11:30 x2, e1, e2 (Rep): No new information. | xse surgical arm error outside procedure quote replacement arm requested difficult arm error present issue identified iud email additional error initial startup initially joint error fixed hard restart issue working fine nurse unplugged working rebooted taking excess minutes start software mazor working multiple hard restarts tried did fix issue tried outlet restart did fix issue finally let stay extended period time finally able software running email additional issues occurred case error code email new mazor x software installation kit | NAV3246 | NAV2047 | NAV3242 | NAV3248 | NAV3262 | NAV3246NAV2047NAV3242NAV3248NAV3262 | True
6993 | 704094787 | XSE - Lateral Deviation | 2020-11-30T14:57:55.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Lateral Deviation 2020-11-30 00701649 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system was used for a case from T10-L2 and the L1 and L2 screws on the left side were lateral by a few millimeters. New C-arm images were taken and the patient's lumbar region was re-registered. After the registration, the screws were able to be accurately placed with the guidance system. There was no patient harm and the procedure was delayed less than an hour. 2020-12-11 IUD x2 (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-11-30 00701649 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system was used for a case from T10-L2 and the L1 and L2 screws on the left side were lateral by a few millimeters. New C-arm images were taken and the patient's lumbar region was re-registered. After the registration, the screws were able to be accurately placed with the guidance system. There was no patient harm and the procedure was delayed less than an hour. 2020-12-11 IUD x2 (Other): No new information. | xse lateral deviation spinal procedure case screws left lateral millimeters new carm images taken patients lumbar region reregistered registration screws able accurately placed procedure delayed hour iud new mazor x | NAV5000 | NAV3242 | NAV2047 | NAV1012 | NAV4112 | NAV5000NAV3242NAV2047NAV1012NAV4112 | True
2893 | 703258752 | XSE - Schanz arms rusted | 2019-07-15T14:52:50.000000+0000 | MazorX Stealth Edition | NAV3296 | DISCOLORATION | 20 | MAS2351-03 | SCHANZ ARM | XSE - Schanz arms rusted 2019-Jul-10 SCL1901996 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that four Schanz arms had become rusted and they had no working Schanz arms. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. 2019-09-05 e1, e2, e3 (Rep): No new information. SCHANZ ARM | 2019-Jul-10 SCL1901996 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that four Schanz arms had become rusted and they had no working Schanz arms. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. 2019-09-05 e1, e2, e3 (Rep): No new information. | xse schanz arms rusted spinal procedure schanz arms rusted working schanz arms additional complications anticipated new schanz arm | NAV3296 | NAV3299 | NAV3295 | NAV5000 | NAV3298 | NAV3296NAV3299NAV3295NAV5000NAV3298 | True
9623 | 704684976 | XSE - Bedframe Adapter & Monitor Issues | 2021-10-27T18:24:15.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 30 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - Bedframe Adapter & Monitor Issues 2021-Oct-27 Salesforce C00236893 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw had fallen out from the bed frame adapter. It was also reported that when storing the surgeon monitor after a case, the sensors were not triggering and the software would report that it was not in the storage position. No patient was present. 2021-11-03 IUD (Other): No new information. 2021-11-03 IUD x4 (Rep): No new information. Information unrelated to this event was excluded from this PE; see PE 704691596 for the OOBF of the surgical system. 2021-11-09 IUD x5 (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-Oct-27 Salesforce C00236893 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw had fallen out from the bed frame adapter. It was also reported that when storing the surgeon monitor after a case, the sensors were not triggering and the software would report that it was not in the storage position. No patient was present. 2021-11-03 IUD (Other): No new information. 2021-11-03 IUD x4 (Rep): No new information. Information unrelated to this event was excluded from this PE; see PE 704691596 for the OOBF of the surgical system. 2021-11-09 IUD x5 (Other): No new information. | xse bedframe adapter monitor issues salesforce outside procedure screw fallen bed frame adapter storing surgeon monitor case sensors triggering software report storage position present iud new iud new unrelated excluded pe pe oobf surgical iud new kit svc surgical syst type | NAV2054 | NAV3255 | NAV3261 | NAV2057 | NAV3250 | NAV2054NAV3255NAV3261NAV2057NAV3250 | False
6368 | 703925442 | XSE - Difficulty Moving the Positioner | 2020-08-27T01:42:43.000000+0000 | MazorX Stealth Edition | NAV3239 | INACCURATE WORK VOLUMES | 30 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - Difficulty Moving the Positioner 2020-08-26 00695330, 00695329 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a T10-pelvis procedure, the manufacturer representative had difficulty moving the positioner and there seemed to be more resistance than usual. During the procedure, the 3Define scan looked normal and tracked the spine well, but the work volume did not appear to be defined correctly because the tip of the surgical arm was headed toward the patient. The surgical arm was stropped and draw spine was done again, but the same issue occurred. A generic work volume was used and the case proceeded without issue. The second segment of the procedure was registered without issue. There was no patient harm and the procedure was delayed less than an hour. 2020-09-02 IUD x5 (Other): No new information. 2020-09-11 IUD (Other): No new information. 2020-09-25 IUD (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-08-26 00695330, 00695329 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a T10-pelvis procedure, the manufacturer representative had difficulty moving the positioner and there seemed to be more resistance than usual. During the procedure, the 3Define scan looked normal and tracked the spine well, but the work volume did not appear to be defined correctly because the tip of the surgical arm was headed toward the patient. The surgical arm was stropped and draw spine was done again, but the same issue occurred. A generic work volume was used and the case proceeded without issue. The second segment of the procedure was registered without issue. There was no patient harm and the procedure was delayed less than an hour. 2020-09-02 IUD x5 (Other): No new information. 2020-09-11 IUD (Other): No new information. 2020-09-25 IUD (Other): No new information. | xse difficulty moving positioner spinal procedure procedure difficulty moving positioner resistance usual procedure scan looked normal tracked spine work volume did appear defined correctly tip surgical arm headed surgical arm stropped draw spine issue occurred generic work volume case proceeded issue second segment procedure registered issue procedure delayed hour iud new iud new iud new unk mazor x camera | NAV4104 | NAV3244 | NAV3239 | NAV3251 | NAV4110 | NAV4104NAV3244NAV3239NAV3251NAV4110 | True
5485 | 703735083 | XSE - No Connection to Controller | 2020-05-01T23:43:02.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0363-01S | KIT SVC ASM0363-01S IO MODULE | XSE - No Connection to Controller 2020-05-01 00688263 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system would not connect to the controller. Troubleshooting involved multiple attempts to hard restart with no success. The field service engineer visited site and could not get the controller to connect. No patient was present. 2020-05-08 IUD x4 (Other): No new information. 2020-05-27 IUD x2 (Other): no new information. 2020-06-21 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-07-09 IUD (Other): No new information. 2020-08-02 IUD (Other): No new information. 2022-02-27 302737442, IUD (Other): No new information. 2022-02-28 IUD (Other): No new information. KIT SVC ASM0363-01S IO MODULE | 2020-05-01 00688263 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system would not connect to the controller. Troubleshooting involved multiple attempts to hard restart with no success. The field service engineer visited site and could not get the controller to connect. No patient was present. 2020-05-08 IUD x4 (Other): No new information. 2020-05-27 IUD x2 (Other): no new information. 2020-06-21 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-07-09 IUD (Other): No new information. 2020-08-02 IUD (Other): No new information. 2022-02-27 302737442, IUD (Other): No new information. 2022-02-28 IUD (Other): No new information. | xse connection controller outside procedure connect controller troubleshooting involved multiple attempts hard restart success field service engineer visited site controller connect present iud new iud new iud new iud new iud new iud new iud new iud new kit svc io module | NAV2047 | NAV3249 | NAV1006 | NAV3244 | NAV2057 | NAV2047NAV3249NAV1006NAV3244NAV2057 | True
6506 | 703958351 | XSE - 3Define Issues | 2020-09-15T19:57:09.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - 3Define Issues 2020-09-15 00696680 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the 3define camera was not working and that the light by the camera turned red. Toggling the 3Define camera in the utilities fixed this issue. No patient was present. UNK MAZOR X 3DEFINE CAMERA | 2020-09-15 00696680 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the 3define camera was not working and that the light by the camera turned red. Toggling the 3Define camera in the utilities fixed this issue. No patient was present. | xse issues outside procedure camera working light camera turned red toggling camera utilities fixed issue present unk mazor x camera | NAV3244 | NAV4110 | NAV2054 | NAV3261 | NAV2047 | NAV3244NAV4110NAV2054NAV3261NAV2047 | True
10014 | 704771181 | XSE - Alleged Inaccuracy | 2021-12-18T00:32:15.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 20 | ASM0206-03S | SURGICAL ARM SPARE PART ASM0206-03S | XSE - Alleged Inaccuracy 2021-Dec-17 Salesforce C00239946, 00730935 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that this case was done without navigation. This open scoliosis case was planned to be T2-L4 planned but actually was from T3 to L3. The dual spine clamp was used to fixate to the patient. The surgeon performed the lower half of the procedure without issue. The surgeon registered the upper half of the anatomy and sent the arm to T3R which looked accurate but T3L was medial. The manufacturer representative noted that there was a possible platform shift. Surgeon did not troubleshoot and decided to freehand the remainder of the procedure. There was no patient impact and no delay. 2021-12-20 IUD x3 (Other): No new information. 2022-01-11 IUD (Other): No new information. 2022-01-18 IUD (Other): No new information. 2022-01-28 IUD (Other): No new information. 2022-02-08 IUD (Other): No new information. 2022-02-08 email 2022-02-08 18:45 (Rep): Trajectories were deviated over 3mm. The cause or suspected cause of the deviation was due to the patient potentially moved by sinking down further into the frame pads, or there was some type of platform shift. The manufacturer representative did not visibly notice this during the case. 2022-02-10 IUD (Other): No new information. 2022-02-22 IUD, 302733957 (Other): No new information. 2022-03-01 IUD x5 (Other): No new information. SURGICAL ARM SPARE PART ASM0206-03S | 2021-Dec-17 Salesforce C00239946, 00730935 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that this case was done without navigation. This open scoliosis case was planned to be T2-L4 planned but actually was from T3 to L3. The dual spine clamp was used to fixate to the patient. The surgeon performed the lower half of the procedure without issue. The surgeon registered the upper half of the anatomy and sent the arm to T3R which looked accurate but T3L was medial. The manufacturer representative noted that there was a possible platform shift. Surgeon did not troubleshoot and decided to freehand the remainder of the procedure. There was no patient impact and no delay. 2021-12-20 IUD x3 (Other): No new information. 2022-01-11 IUD (Other): No new information. 2022-01-18 IUD (Other): No new information. 2022-01-28 IUD (Other): No new information. 2022-02-08 IUD (Other): No new information. 2022-02-08 email 2022-02-08 18:45 (Rep): Trajectories were deviated over 3mm. The cause or suspected cause of the deviation was due to the patient potentially moved by sinking down further into the frame pads, or there was some type of platform shift. The manufacturer representative did not visibly notice this during the case. 2022-02-10 IUD (Other): No new information. 2022-02-22 IUD, 302733957 (Other): No new information. 2022-03-01 IUD x5 (Other): No new information. | xse alleged inaccuracy salesforce spinal procedure case navigation open scoliosis case planned planned actually dual spine clamp fixate surgeon performed lower half procedure issue surgeon registered upper half anatomy sent arm looked accurate medial noted possible platform shift surgeon did troubleshoot decided freehand remainder procedure impact delay iud new iud new iud new iud new iud new email trajectories deviated cause suspected cause deviation potentially moved sinking frame pads type platform shift did visibly notice case iud new iud new iud new surgical arm spare | NAV5000 | NAV4098 | NAV3244 | NAV3299 | NAV2058 | NAV5000NAV4098NAV3244NAV3299NAV2058 | False
6664 | 703999118 | XSE - SRA Not Moving/CANBUS Error | 2020-10-07T16:10:04.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - SRA Not Moving/CANBUS Error 2020-09-22 00698286, IUD (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after placing the screws, the surgical arm was sent to three trajectories for placing cages. The surgeon placed the scalpel and then the cannula down the arm guide. When pushing the instruments down, an error message was displayed that stated the surgical arm was not on trajectory. The representative tried to resend the surgical arm to the trajectory and a moving sound was mad, but the arm did not move. The surgical arm was sent to the clear up position, but the same issue occurred. A soft restart of the system was done, but the issue was not resolved. When the emergency stop button was pressed, the system did not change color to red. The representative noted that the controller status was green, but the monitor height button did not react. A hard restart of the system was done and the issue was resolved. There was no patient harm and the procedure was delayed less than an hour. 2020-10-12 IUD x4 (Other, For): No new information. 2020-11-23 IUD (Other, For): No new information. 2020-12-21 email 2020-12-21 03:15, image 2020-12-21 03:15 (Rep, For): No new information. 2021-01-13 e1, e2, e3, e4 ,e5 (Other, For): No new information. 2021-01-17 e6, e7, e8 (Other, For): No new information. 2021-01-21 e9, e10, e11, e12 (Other, For): No new information. 2021-06-07 e13 (Other, For): No new information. 2021-06-11 e14, e15 (Other, For): No new information. 2021-08-01 IUD (Other, For): No new information. 2021-08-03 302515410 (Other, For): No new information. 2021-08-08 IUD x6 (Other, For): no new information. 2021-08-23 IUD x2 (Other, For): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-09-22 00698286, IUD (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after placing the screws, the surgical arm was sent to three trajectories for placing cages. The surgeon placed the scalpel and then the cannula down the arm guide. When pushing the instruments down, an error message was displayed that stated the surgical arm was not on trajectory. The representative tried to resend the surgical arm to the trajectory and a moving sound was mad, but the arm did not move. The surgical arm was sent to the clear up position, but the same issue occurred. A soft restart of the system was done, but the issue was not resolved. When the emergency stop button was pressed, the system did not change color to red. The representative noted that the controller status was green, but the monitor height button did not react. A hard restart of the system was done and the issue was resolved. There was no patient harm and the procedure was delayed less than an hour. 2020-10-12 IUD x4 (Other, For): No new information. 2020-11-23 IUD (Other, For): No new information. 2020-12-21 email 2020-12-21 03:15, image 2020-12-21 03:15 (Rep, For): No new information. 2021-01-13 e1, e2, e3, e4 ,e5 (Other, For): No new information. 2021-01-17 e6, e7, e8 (Other, For): No new information. 2021-01-21 e9, e10, e11, e12 (Other, For): No new information. 2021-06-07 e13 (Other, For): No new information. 2021-06-11 e14, e15 (Other, For): No new information. 2021-08-01 IUD (Other, For): No new information. 2021-08-03 302515410 (Other, For): No new information. 2021-08-08 IUD x6 (Other, For): no new information. 2021-08-23 IUD x2 (Other, For): No new information. | xse sra movingcanbus error iud spinal procedure placing screws surgical arm sent trajectories placing cages surgeon placed scalpel cannula arm guide pushing instruments error message displayed stated surgical arm trajectory tried resend surgical arm trajectory moving sound mad arm did surgical arm sent clear position issue occurred soft restart issue resolved emergency stop button pressed did change color red noted controller status green monitor height button did react hard restart issue resolved procedure delayed hour iud new iud new email image new new new new new new iud new new iud new iud new kit svc surgical arm | NAV2047 | NAV4101 | NAV3249 | NAV4095 | NAV3246 | NAV2047NAV4101NAV3249NAV4095NAV3246 | True
1649 | 703208304 | X- SW Issues | 2019-06-14T19:19:47.000000+0000 | MazorX | NAV3235 | 3D SCANS | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X- SW Issues The star marker was centered more above L4, than L5‐S1. Attempted to import the scan and the star marker was not found. Upon reviewing the scan there was a significant amount of scatter that appeared to be coming from the air frame, see images. On the O‐arm were able to visually see most of the star marker beads within the scatter. Switched out the airframe for the traditional passive frame and were able to proceed with the case. Export from O‐arm is available. MAZOR X SOFTWARE KIT VER-3.0.2 | The star marker was centered more above L4, than L5‐S1. Attempted to import the scan and the star marker was not found. Upon reviewing the scan there was a significant amount of scatter that appeared to be coming from the air frame, see images. On the O‐arm were able to visually see most of the star marker beads within the scatter. Switched out the airframe for the traditional passive frame and were able to proceed with the case. Export from O‐arm is available. | x sw issues star marker centered ‐ attempted import scan star marker reviewing scan significant scatter appeared coming air frame images o‐arm able visually star marker beads scatter switched airframe traditional passive frame able proceed case export o‐arm available mazor x software kit | NAV3246 | NAV3235 | NAV3242 | NAV4091 | NAV3248 | NAV3246NAV3235NAV3242NAV4091NAV3248 | True
5002 | 703589423 | X- Registration issues & Deviation | 2020-01-30T18:24:44.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X- Registration issues & Deviation 2018-6-29 CCR(Rep, HCP):Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the plan was reviewed by the surgeon and the manufacturer representative prior to the procedure. An advanced accuracy test was completed, the shoulder calibration was checked, and all cameras were found to be functional. The schanz pin was inserted into the right PSIS and attached to a BMB. A 3Define scan and draw spine were completed. Two AP and OBL images were taken. The initial registration attempt was unsuccessful, but then the marker was moved to another body and basic registration was achieved. Left L5 and S1 had red trajectories, so the plan was adjusted until a solution was found. All trajectories were executed. Fluoro images showed that right L4 had skived lateral of the pedicle, so the screw was redirected on the plan and the surgeon placed another k-wire. An AP image was taken, but still showed that the k-wire had skived laterally. The screw trajectory was adjusted more medial, and a new k-wire was placed. All screws were placed accurately and stimulated to appropriate values. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-6-29 CCR(Rep, HCP):Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the plan was reviewed by the surgeon and the manufacturer representative prior to the procedure. An advanced accuracy test was completed, the shoulder calibration was checked, and all cameras were found to be functional. The schanz pin was inserted into the right PSIS and attached to a BMB. A 3Define scan and draw spine were completed. Two AP and OBL images were taken. The initial registration attempt was unsuccessful, but then the marker was moved to another body and basic registration was achieved. Left L5 and S1 had red trajectories, so the plan was adjusted until a solution was found. All trajectories were executed. Fluoro images showed that right L4 had skived lateral of the pedicle, so the screw was redirected on the plan and the surgeon placed another k-wire. An AP image was taken, but still showed that the k-wire had skived laterally. The screw trajectory was adjusted more medial, and a new k-wire was placed. All screws were placed accurately and stimulated to appropriate values. There was no patient harm and the procedure was not delayed over an hour. | x registration issues deviation ccrmedtronic spinal procedure plan reviewed surgeon prior procedure advanced accuracy test completed shoulder calibration checked cameras functional schanz pin inserted right psis attached bmb scan draw spine completed ap obl images taken initial registration attempt unsuccessful marker moved body basic registration achieved left red trajectories plan adjusted solution trajectories executed fluoro images showed right skived lateral pedicle screw redirected plan surgeon placed kwire ap image taken showed kwire skived laterally screw trajectory adjusted medial new kwire placed screws placed accurately stimulated appropriate values procedure delayed hour mazor x | NAV3242 | NAV4091 | NAV5000 | NAV3243 | NAV4106 | NAV3242NAV4091NAV5000NAV3243NAV4106 | True
2687 | 703246022 | X- Air leakage in Positioner | 2019-07-08T11:23:46.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | SLD0163-01 | PISTON, FESTO ADVC-25-15-A-P | X- Air leakage in Positioner 1. Air leakage in Positioner plate locking teeth piston (SLD0163) due to piston damage during standard operation. System SN 400-12-022016 pneumatic function needs to be fixed in the filed by Jeff by replacing the components below: SLD0163-01 (x1)SLD0180-01 (x1)SLD0198-01 (x3) 2. Unintended Positioner pneumatic function activation due to blockage in exhaust line PISTON, FESTO ADVC-25-15-A-P | 1. Air leakage in Positioner plate locking teeth piston (SLD0163) due to piston damage during standard operation. System SN 400-12-022016 pneumatic function needs to be fixed in the filed by Jeff by replacing the components below: SLD0163-01 (x1)SLD0180-01 (x1)SLD0198-01 (x3) 2. Unintended Positioner pneumatic function activation due to blockage in exhaust line | x air leakage positioner air leakage positioner plate locking teeth piston piston damage standard operation sn pneumatic function needs fixed filed jeff replacing components unintended positioner pneumatic function activation blockage exhaust line piston festo | NAV3251 | NAV3264 | NAV4103 | NAV4105 | NAV2058 | NAV3251NAV3264NAV4103NAV4105NAV2058 | True
10238 | 704836007 | XSE - Air Leak | 2022-01-28T21:01:52.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 20 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE - Air Leak 2022-Jan-28 Salesforce C00242412 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the ma nipulator had air leak. This caused no issue during case but was audible when moving the system. The manufacturer representative reported this caused the pneumatic panel to run excessively. There was no impact on the patient outcome. 2022-02-01 IUD (Other): No new information. 2022-02-02 email 2022-02-02 20:30 (Rep): The issue occurred during a TLIF procedure. 2022-02-08 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. UNK MAZOR X MANIPULATOR | 2022-Jan-28 Salesforce C00242412 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the ma nipulator had air leak. This caused no issue during case but was audible when moving the system. The manufacturer representative reported this caused the pneumatic panel to run excessively. There was no impact on the patient outcome. 2022-02-01 IUD (Other): No new information. 2022-02-02 email 2022-02-02 20:30 (Rep): The issue occurred during a TLIF procedure. 2022-02-08 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. | xse air leak salesforce ma nipulator air leak caused issue case audible moving caused pneumatic panel run excessively impact outcome iud new email issue occurred tlif procedure iud new iud new unk mazor x manipulator | NAV3264 | NAV3261 | NAV3244 | NAV3260 | NAV2058 | NAV3264NAV3261NAV3244NAV3260NAV2058 | True
4412 | 703541598 | X - Shoulder Shift, Deviation | 2020-01-02T20:22:10.000000+0000 | MazorX | NAV4108 | SHOULDER SHIFTS | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Shoulder Shift, Deviation 2018-05-15 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure for cement augmentation at L5. It was reported that planning for the L5 vertebroplasty was done in the OR prior to the case. This was reportedly the first case with CT/Fluoro at this facility. The engineer and representative both explained the skive potential to the surgeon but the surgeon did not want to spend any more time on the plan. The platform was mount to the patient with a schanz pin in the patient's right iliac crest. The 3Define scan was good. The manufacturer representative felt that the surgeon and surgical team need additional explanation/training about why Draw Spine is done and the importance of doing it well and to avoid manipulation of the 3D marker. Quality of the OBL image was poor but semi-automatic segmentation was achieved off L4 and registration was achieved. The surgeon quickly verified and operating began. The surgeon applied excessive force on the blunt mallet using cannula/blunt technique which caused a shoulder shift. A second registration was quickly accomplished using the basic algorithm. However, there a superior deviation at left L5 due to skiving caused by the surgeon using excessive force with instrumentation. Superior skive was confirmed with lateral and AP fluoro. The surgeon decided to perform the vertebroplasty only on the right side. No patient harm was reported and surgery was delayed less than an hour. The surgeon was happy with the overall outcome of the surgery but was annoyed about having to register twice and preferred scan and plan surgery flow. UNK MAZOR X SHOULDER | 2018-05-15 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure for cement augmentation at L5. It was reported that planning for the L5 vertebroplasty was done in the OR prior to the case. This was reportedly the first case with CT/Fluoro at this facility. The engineer and representative both explained the skive potential to the surgeon but the surgeon did not want to spend any more time on the plan. The platform was mount to the patient with a schanz pin in the patient's right iliac crest. The 3Define scan was good. The manufacturer representative felt that the surgeon and surgical team need additional explanation/training about why Draw Spine is done and the importance of doing it well and to avoid manipulation of the 3D marker. Quality of the OBL image was poor but semi-automatic segmentation was achieved off L4 and registration was achieved. The surgeon quickly verified and operating began. The surgeon applied excessive force on the blunt mallet using cannula/blunt technique which caused a shoulder shift. A second registration was quickly accomplished using the basic algorithm. However, there a superior deviation at left L5 due to skiving caused by the surgeon using excessive force with instrumentation. Superior skive was confirmed with lateral and AP fluoro. The surgeon decided to perform the vertebroplasty only on the right side. No patient harm was reported and surgery was delayed less than an hour. The surgeon was happy with the overall outcome of the surgery but was annoyed about having to register twice and preferred scan and plan surgery flow. | x shoulder shift deviation ccr spinal procedure cement augmentation planning vertebroplasty prior case reportedly case ctfluoro facility engineer explained skive potential surgeon surgeon did want spend time plan platform mount schanz pin patients right iliac crest scan good felt surgeon surgical team need additional explanationtraining draw spine importance doing avoid manipulation marker quality obl image poor semiautomatic segmentation achieved registration achieved surgeon quickly verified operating began surgeon applied excessive force blunt mallet using cannulablunt technique caused shoulder shift second registration quickly accomplished using basic algorithm superior deviation left skiving caused surgeon using excessive force instrumentation superior skive confirmed lateral ap fluoro surgeon decided perform vertebroplasty right surgery delayed hour surgeon happy overall outcome surgery annoyed having register twice preferred scan plan surgery flow unk mazor x shoulder | NAV4091 | NAV4108 | NAV3250 | NAV3244 | NAV4101 | NAV4091NAV4108NAV3250NAV3244NAV4101 | True
6959 | 704078203 | XSE - J4 Error/Cable Damaged | 2020-11-18T22:51:05.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 30 | HRS0350-01S | KIT SVC HRS0350-01S NAV CONNECT CABLE | XSE - J4 Error/Cable Damaged 2020-11-18 00701146 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that It was reported that they received a joint 4 error and the camera cart cable is damaged. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-11-18 IUD x5 (Other): No new information. 2020-12-15 IUD x3 (Other): No new information. 2021-02-08 IUD (Other): No new information. KIT SVC HRS0350-01S NAV CONNECT CABLE | 2020-11-18 00701146 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that It was reported that they received a joint 4 error and the camera cart cable is damaged. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-11-18 IUD x5 (Other): No new information. 2020-12-15 IUD x3 (Other): No new information. 2021-02-08 IUD (Other): No new information. | xse errorcable damaged outside procedure joint error camera cart cable damaged additional complications anticipated iud new iud new iud new kit svc nav connect cable | NAV3256 | NAV2054 | NAV2058 | NAV2047 | NAV3249 | NAV3256NAV2054NAV2058NAV2047NAV3249 | True
2491 | 703236589 | XSE - deviaiton | 2019-07-01T18:26:54.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - deviaiton 2019-Jun-27 SCL1901867 intake form (HCP,rep): ***Information omitted pertaining to previous deviations 703209809 703226476*** Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the post-op CT showed screws in pedicles but shifted about 3mm to the right for each trajectory. This procedure was performed Right Side up. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jun-27 SCL1901867 intake form (HCP,rep): ***Information omitted pertaining to previous deviations 703209809 703226476*** Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the post-op CT showed screws in pedicles but shifted about 3mm to the right for each trajectory. This procedure was performed Right Side up. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse deviaiton omitted pertaining previous deviations spinal procedure postop ct showed screws pedicles shifted right trajectory procedure performed right additional complications anticipated mazor x software installation kit | NAV5000 | NAV3246 | NAV2054 | NAV3243 | NAV2047 | NAV5000NAV3246NAV2054NAV3243NAV2047 | True
7984 | 704325004 | XSE ¦ Lost Screws | 2021-04-06T16:10:33.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 50 | MEC4530-02S | KIT SVC MEC4530-02S BALL ARM TRGT SCREW | XSE ¦ Lost Screws 2021-Apr-06 Salesforce C00223895 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site lost the schanz bond screws. No patient was present. 2021-04-13 IUD x3 (Other): No new information. KIT SVC MEC4530-02S BALL ARM TRGT SCREW | 2021-Apr-06 Salesforce C00223895 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site lost the schanz bond screws. No patient was present. 2021-04-13 IUD x3 (Other): No new information. | xse lost screws salesforce outside procedure site lost schanz bond screws present iud new kit svc ball arm trgt screw | NAV2053 | NAV3299 | NAV3249 | NAV2047 | NAV4090 | NAV2053NAV3299NAV3249NAV2047NAV4090 | True
1071 | 703187960 | X - 3Define scan issue | 2019-06-04T12:48:21.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define scan issue First shot of the 3define scan the system would freeze. Tried to send the scan again with the same result. Tried a soft restart and the 3Define never worked. UNK MAZOR X 3DEFINE CAMERA | First shot of the 3define scan the system would freeze. Tried to send the scan again with the same result. Tried a soft restart and the 3Define never worked. | x scan issue shot scan freeze tried send scan tried soft restart worked unk mazor x camera | NAV3244 | NAV2058 | NAV4095 | NAV3248 | NAV3250 | NAV3244NAV2058NAV4095NAV3248NAV3250 | True
2463 | 703236295 | X - ARM/COMMUNICATION ERROR ISSUE | 2019-07-01T17:20:51.000000+0000 | MazorX | NAV4096 | ARM HIT WORK VOLUME | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - ARM/COMMUNICATION ERROR ISSUE Arm Shift detected in Joint 3. Followed by "FATAL ERROR: CANBUS Connection Lost to Nodes # 1, 2, 3, 4, 5, 6, 7, 8, 10" UNK MAZOR X ARM | Arm Shift detected in Joint 3. Followed by "FATAL ERROR: CANBUS Connection Lost to Nodes # 1, 2, 3, 4, 5, 6, 7, 8, 10" | x armcommunication error issue arm shift detected joint followed fatal error canbus connection lost nodes unk mazor x arm | NAV4098 | NAV2047 | NAV4101 | NAV2054 | NAV3244 | NAV4098NAV2047NAV4101NAV2054NAV3244 | False
7160 | 704131874 | XSE - Leak In Pneumatic Tubing | 2020-12-17T21:16:01.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 20 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE - Leak In Pneumatic Tubing 2020-12-17 00702914, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a leak in the pneumatic tubing in the manipulator, the X-Eye file was missing and the camera ports needed to be fixed. There was no patient involvement. UNK MAZOR X MANIPULATOR | 2020-12-17 00702914, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a leak in the pneumatic tubing in the manipulator, the X-Eye file was missing and the camera ports needed to be fixed. There was no patient involvement. | xse leak pneumatic tubing iud outside procedure leak pneumatic tubing manipulator xeye file missing camera ports needed fixed unk mazor x manipulator | NAV3264 | NAV3260 | NAV4110 | NAV2053 | NAV2054 | NAV3264NAV3260NAV4110NAV2053NAV2054 | True
7189 | 704136214 | XSE - Deviation | 2020-12-21T17:18:25.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2020-12-21 00703126, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon was looking at the revisions of a patient for a surgery that was previously done with the guidance system. There was a platform shift and all of the screws were noted to be misplaced. There was a slight medial breach with one screw, but there was no patient harm. The procedure was delayed less than an hour. 2021-Jan-12 IUD (other): No new information. 2021-Feb-03 rep 2/3 (rep): The patient in question does not have files on the system as the computer was replaced. The revision was not due to misplaced screws, it was an extension to what was already placed. 2021-03-08 email 2021-03-08 12:16 (Rep): Additional information received from a manufacturer representative reported the screws were deviated approximately 5 mm. The suspected cause of the deviation was a patient shift at some point during the procedure. MAZOR X SYSTEM TPL0059 | 2020-12-21 00703126, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon was looking at the revisions of a patient for a surgery that was previously done with the guidance system. There was a platform shift and all of the screws were noted to be misplaced. There was a slight medial breach with one screw, but there was no patient harm. The procedure was delayed less than an hour. 2021-Jan-12 IUD (other): No new information. 2021-Feb-03 rep 2/3 (rep): The patient in question does not have files on the system as the computer was replaced. The revision was not due to misplaced screws, it was an extension to what was already placed. 2021-03-08 email 2021-03-08 12:16 (Rep): Additional information received from a manufacturer representative reported the screws were deviated approximately 5 mm. The suspected cause of the deviation was a patient shift at some point during the procedure. | xse deviation iud spinal procedure surgeon looking revisions surgery previously platform shift screws noted misplaced slight medial breach screw procedure delayed hour iud new question does files computer replaced revision misplaced screws extension placed email additional screws deviated approximately mm suspected cause deviation shift point procedure mazor x | NAV5000 | NAV4098 | NAV1012 | NAV4108 | NAV2053 | NAV5000NAV4098NAV1012NAV4108NAV2053 | True
5269 | 703634134 | XSE - Screw order plan not save in SW | 2020-02-25T08:11:32.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Screw order plan not save in SW 2020-02-25 Intake Form, 00685077 ID, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after selecting the screw order in the summary details page, the selections were not saved when advancing to the operation screen. No other issues were reported during the procedure and all trajectories were placed accurately using the guidance system. No patient harm was reported and surgery was delayed less than an hour. 2020-03-02 email 2020-03-02 19:00 (Rep): Additional information received from a manufacturer representative reported that the issue was not resolved and no troubleshooting was done since they just manually selected the trajectory. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-02-25 Intake Form, 00685077 ID, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after selecting the screw order in the summary details page, the selections were not saved when advancing to the operation screen. No other issues were reported during the procedure and all trajectories were placed accurately using the guidance system. No patient harm was reported and surgery was delayed less than an hour. 2020-03-02 email 2020-03-02 19:00 (Rep): Additional information received from a manufacturer representative reported that the issue was not resolved and no troubleshooting was done since they just manually selected the trajectory. | xse screw order plan save sw id ccr spinal procedure selecting screw order summary details page selections saved advancing operation screen issues procedure trajectories placed accurately using surgery delayed hour email additional issue resolved troubleshooting just manually selected trajectory mazor x software installation kit | NAV3246 | NAV5000 | NAV3242 | NAV3243 | NAV3248 | NAV3246NAV5000NAV3242NAV3243NAV3248 | True
3269 | 703301488 | XSE - Surgical Arm Shaking | 2019-08-08T15:02:15.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01 | SURGICAL ARM | XSE - Surgical Arm Shaking 2019-07-31 SCL1902286, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that the surgical arm would shake when the screw was being placed. The case flow started with the trajectory at left L4 down to S1 and then right S1 up to L4. The representative also noted that it was very difficult to get the screwdriver out of the cannula, possibly due to the force on the arm. No shoulder shifts were experienced. The representative thought the shaking could have been caused by the pressure of the patient's tissue or the surgeon. When the surgical arm was moved from the left side of the patient to the right side, the arm was shaking. One of the screws could not be placed using the guidance system due to the arm shaking when the screw was engaged with bone so navigation was used. The surgical arm did not shake when used in a procedure on the following day. There was no patient harm and the procedure was not delayed more than an hour. 2019-08-14 email 2019-08-14 22:15 (Rep): Additional information received from a manufacturer representative reported that during screw insertion, the shaking would occur when the screw was inserted to a certain point and ceased when removed again, making it seem that the driver/screw was putting pressure on the arm. Troubleshooting involved removing instruments from the arm guide, resending the arm to the trajectories and switching the elbow side of the arm. All of the screws were able to be placed successfully. SURGICAL ARM | 2019-07-31 SCL1902286, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported that the surgical arm would shake when the screw was being placed. The case flow started with the trajectory at left L4 down to S1 and then right S1 up to L4. The representative also noted that it was very difficult to get the screwdriver out of the cannula, possibly due to the force on the arm. No shoulder shifts were experienced. The representative thought the shaking could have been caused by the pressure of the patient's tissue or the surgeon. When the surgical arm was moved from the left side of the patient to the right side, the arm was shaking. One of the screws could not be placed using the guidance system due to the arm shaking when the screw was engaged with bone so navigation was used. The surgical arm did not shake when used in a procedure on the following day. There was no patient harm and the procedure was not delayed more than an hour. 2019-08-14 email 2019-08-14 22:15 (Rep): Additional information received from a manufacturer representative reported that during screw insertion, the shaking would occur when the screw was inserted to a certain point and ceased when removed again, making it seem that the driver/screw was putting pressure on the arm. Troubleshooting involved removing instruments from the arm guide, resending the arm to the trajectories and switching the elbow side of the arm. All of the screws were able to be placed successfully. | xse surgical arm shaking ccr spinal procedure treat degenerative disc fixation surgical arm shake screw placed case flow started trajectory left right noted difficult screwdriver cannula possibly force arm shoulder shifts experienced thought shaking caused pressure patients tissue surgeon surgical arm moved left right arm shaking screws placed using arm shaking screw engaged bone navigation surgical arm did shake procedure following day procedure delayed hour email additional screw insertion shaking occur screw inserted certain point ceased removed making driverscrew putting pressure arm troubleshooting involved removing instruments arm guide resending arm trajectories switching elbow arm screws able placed successfully surgical arm | NAV4101 | NAV4096 | NAV3249 | NAV4098 | NAV4100 | NAV4101NAV4096NAV3249NAV4098NAV4100 | True
9728 | 704707955 | XSE - System Needs Repair | 2021-11-10T16:05:08.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 50 | ASM0263-01 | RNX-WS-1040 | XSE - System Needs Repair 2021-Nov-10 Salesforce C00237642, 00725556 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system needed repair. There was no patient involvement. 2021-11-14 email 2021-11-14 03:00, email 2021-11-14 03:15, email 2021-11-14 03:30 (Rep, For): No new information. 2021-11-18 email 2021-11-18 00:15 (Rep): No new information. 2021-11-21 IUD x2 (Other): No new information. 2021-12-03 email 2021-12-02 02:46 (Rep, For): Additional information received from a manufacturer representative reported the PCs were replaced due to being faulty and a PC fan error. The electrical unit was replaced because of a communication error in the system. 2021-Dec-26 IUD, SAP 302673385 (other): No new information. RNX-WS-1040 | 2021-Nov-10 Salesforce C00237642, 00725556 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system needed repair. There was no patient involvement. 2021-11-14 email 2021-11-14 03:00, email 2021-11-14 03:15, email 2021-11-14 03:30 (Rep, For): No new information. 2021-11-18 email 2021-11-18 00:15 (Rep): No new information. 2021-11-21 IUD x2 (Other): No new information. 2021-12-03 email 2021-12-02 02:46 (Rep, For): Additional information received from a manufacturer representative reported the PCs were replaced due to being faulty and a PC fan error. The electrical unit was replaced because of a communication error in the system. 2021-Dec-26 IUD, SAP 302673385 (other): No new information. | xse needs repair salesforce outside procedure needed repair email email email new email new iud new email additional pcs replaced faulty pc fan error electrical unit replaced communication error iud sap new | NAV3249 | NAV3262 | NAV2047 | NAV2058 | NAV3299 | NAV3249NAV3262NAV2047NAV2058NAV3299 | True
1013 | 703183951 | X- WS and Manipulator Issues | 2019-05-31T21:01:38.000000+0000 | MazorX | NAV3260 | MANIPULATOR ISSUES | 30 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | X- WS and Manipulator Issues While trying to remove the manipulator from the SRS the manipulator would not move front to back. Found the pneumatic tubing was pinched. Had rep unpinch the tubing to get the system off the bed at the end of the case but the tubing will need replaced. UNK MAZOR X MANIPULATOR | While trying to remove the manipulator from the SRS the manipulator would not move front to back. Found the pneumatic tubing was pinched. Had rep unpinch the tubing to get the system off the bed at the end of the case but the tubing will need replaced. | x ws manipulator issues trying remove manipulator srs manipulator pneumatic tubing pinched unpinch tubing bed end case tubing need replaced unk mazor x manipulator | NAV3260 | NAV3264 | NAV3250 | NAV3244 | NAV4108 | NAV3260NAV3264NAV3250NAV3244NAV4108 | True
5247 | 703627407 | XSE - Pneumatic/Electrical Panel damaged | 2020-02-20T20:09:24.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 40 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Pneumatic/Electrical Panel damaged 2020-02-20 00684564 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the pneumatic and electrical panel are not working properly and need to be replaced on the system. There was no patient involvement. 2020-02-27 IUD (Other): No new information. 2020-03-11 Email 2020-03-11 13:00, IUD x7 (Rep, Other): No new information. 2020-04-20 IUD x2 (Other): No new information. 2020-04-29 IUD x3 (Other): No new information. 2020-05-15 IUD x2 (Other): No new information. 2020-05-27 IUD X2 (Other): No new information. 2021-03-23 IUD x2 (Other): No new information. 2021-05-11 IUD (Other): No new information. 2021-05-31 IUD (Other): No new information. 2021-09-29 IUD x4, 302579492 (Other): No new information. 2021-12-08 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2020-02-20 00684564 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the pneumatic and electrical panel are not working properly and need to be replaced on the system. There was no patient involvement. 2020-02-27 IUD (Other): No new information. 2020-03-11 Email 2020-03-11 13:00, IUD x7 (Rep, Other): No new information. 2020-04-20 IUD x2 (Other): No new information. 2020-04-29 IUD x3 (Other): No new information. 2020-05-15 IUD x2 (Other): No new information. 2020-05-27 IUD X2 (Other): No new information. 2021-03-23 IUD x2 (Other): No new information. 2021-05-11 IUD (Other): No new information. 2021-05-31 IUD (Other): No new information. 2021-09-29 IUD x4, 302579492 (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse pneumaticelectrical panel damaged outside procedure pneumatic electrical panel working properly need replaced iud new email iud new iud new iud new iud new iud new iud new iud new iud new iud new iud new kit svc cart elec panel assy | NAV2058 | NAV3264 | NAV2054 | NAV3244 | NAV2047 | NAV2058NAV3264NAV2054NAV3244NAV2047 | True
6277 | 703902986 | XSE - A4 Access Error | 2020-08-13T19:31:19.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - A4 Access Error 2020-08-13 00694556, IUD (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that while setting up for a case an A4 Access error occurred. They decided to use the other guidance system to continue. No patient was involved. 2020-10-19 IUD (Other): No new information. 2020-11-09 IUD (Other): No new information. 2020-11-30 IUD (Other): No new information. 2020-12-20 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-08-13 00694556, IUD (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that while setting up for a case an A4 Access error occurred. They decided to use the other guidance system to continue. No patient was involved. 2020-10-19 IUD (Other): No new information. 2020-11-09 IUD (Other): No new information. 2020-11-30 IUD (Other): No new information. 2020-12-20 IUD (Other): No new information. | xse access error iud prior procedure setting case access error occurred decided use continue involved iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV2054 | NAV4098 | NAV1006 | NAV3249NAV4095NAV2054NAV4098NAV1006 | True
8481 | 704424895 | XSE - OOBF/Missing Screw | 2021-06-01T18:59:54.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | MEC6605-01S | KIT SVC MEC6605-01S ROBOTIC FRAME SCREW | XSE - OOBF/Missing Screw 2021-May-28 Salesforce C00227575, email (Rep): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that the screw at the base of the robotic reference frame connection was missing out of the box. There was no patient involvement. 2021-06-07 IUD (Other): No new information. 2021-07-18 e1, e2, e3, e4, e5 (Rep): No new information. 2021-08-18 e6 (Other): No new information. 2021-08-26 e7 (Other): No new information. 2021-08-31 e8 (Other): No new information. KIT SVC MEC6605-01S ROBOTIC FRAME SCREW | 2021-May-28 Salesforce C00227575, email (Rep): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that the screw at the base of the robotic reference frame connection was missing out of the box. There was no patient involvement. 2021-06-07 IUD (Other): No new information. 2021-07-18 e1, e2, e3, e4, e5 (Rep): No new information. 2021-08-18 e6 (Other): No new information. 2021-08-26 e7 (Other): No new information. 2021-08-31 e8 (Other): No new information. | xse oobfmissing screw salesforce email outside pr ocedure screw base robotic reference frame connection missing box iud new new new new new kit svc robotic frame screw | NAV3299 | NAV2054 | NAV2053 | NAV4090 | NAV3251 | NAV3299NAV2054NAV2053NAV4090NAV3251 | True
2442 | 703235875 | X - Workstation Issue | 2019-07-01T13:08:32.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Workstation Issue Pressure issue deteriorating. It is becoming more challenging to attach SS to bedframe UNK MAZOR X WORKSTATION | Pressure issue deteriorating. It is becoming more challenging to attach SS to bedframe | x workstation issue pressure issue deteriorating challenging attach ss bedframe unk mazor x workstation | NAV2058 | NAV3264 | NAV2047 | NAV3255 | NAV3256 | NAV2058NAV3264NAV2047NAV3255NAV3256 | True
5778 | 703800881 | XSE - 3Define Scan Not Working | 2020-06-13T01:16:15.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - 3Define Scan Not Working 2020-06-12 00690721 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the 3Define camera was not working. There was no patient involvement. 2020-06-16 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-06-30 IUD (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-06-12 00690721 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the 3Define camera was not working. There was no patient involvement. 2020-06-16 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-06-30 IUD (Other): No new information. | xse scan working outside procedure camera working iud new iud new iud new unk mazor x camera | NAV3244 | NAV2054 | NAV4110 | NAV2058 | NAV3299 | NAV3244NAV2054NAV4110NAV2058NAV3299 | True
5443 | 703711300 | XSE - SRA Failed Advanced Accuracy Test | 2020-04-15T18:32:05.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - SRA Failed Advanced Accuracy Test 2020-04-15 00687455 (Rep): Medtronic received information regarding a guidance system outside of procedure. It was reported that the arm completed the setup test. The arm was unable to reach all 10 point of the advanced accuracy test. The arm was accurate when it reached but failed to reach all 10 points on multiple attempts. The manufacturer representative attempted a reset movement twice and finally the arm reached all 10 points of the accuracy test and all points were accurate test passed. Joint 6 was showing movement under the advanced settings but was not actually moving. No patient was present. 2020-04-22 IUD x6 (Other): No new information. 2020-07-07 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-04-15 00687455 (Rep): Medtronic received information regarding a guidance system outside of procedure. It was reported that the arm completed the setup test. The arm was unable to reach all 10 point of the advanced accuracy test. The arm was accurate when it reached but failed to reach all 10 points on multiple attempts. The manufacturer representative attempted a reset movement twice and finally the arm reached all 10 points of the accuracy test and all points were accurate test passed. Joint 6 was showing movement under the advanced settings but was not actually moving. No patient was present. 2020-04-22 IUD x6 (Other): No new information. 2020-07-07 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. | xse sra failed advanced accuracy test outside procedure arm completed setup test arm unable reach point advanced accuracy test arm accurate reached failed reach points multiple attempts attempted reset movement twice finally arm reached points accuracy test points accurate test passed joint showing movement advanced settings actually moving present iud new iud new iud new kit svc surgical arm refurb | NAV4095 | NAV4093 | NAV4092 | NAV3249 | NAV1006 | NAV4095NAV4093NAV4092NAV3249NAV1006 | True
6760 | 704026627 | XSE - Stripped Screw | 2020-10-21T22:35:13.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5196-01S | KIT SVC MEC5196-01S RNX-ACC-063 | XSE - Stripped Screw 2020-10-21 00699240 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw on the bone mount bridge was found to be stripped. There was no patient involvement. 2020-11-17 email 11/17 (Rep): Additional information received from a manufacturer representative reported that the part was discarded. KIT SVC MEC5196-01S RNX-ACC-063 | 2020-10-21 00699240 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screw on the bone mount bridge was found to be stripped. There was no patient involvement. 2020-11-17 email 11/17 (Rep): Additional information received from a manufacturer representative reported that the part was discarded. | xse stripped screw outside procedure screw bone mount bridge stripped email additional discarded kit svc | NAV3299 | NAV2054 | NAV4100 | NAV4090 | NAV3251 | NAV3299NAV2054NAV4100NAV4090NAV3251 | True
6116 | 703870829 | XSE - Registration Issues | 2020-07-25T18:50:11.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Issues 2020-05-27 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the plan was reviewed with the surgeon with close attention paid to hypoplastic anatomy. The 360 bridge in combination with a tall and two short clamps was used with the MIST bridge connected to the Thomson system at the top of the construct to mount the system to the patient. Registration for the first segment was achieved using the basic algorithm when labeling from T11. Initial attempts to register the second segment failed. The manufacturer representative was able to achieve labeling from T7. T5 had inoperable registration values, but there was no shift. The surgical arm was sent to the operable values. The representative tried to re-register T5, but it was unsuccessful from T5-T7. The representative believed it was due to the clamp washing out the anatomy. All drilled trajectories were accurate. There was no patient harm and the procedure was delayed less t han an hour. MAZOR X SYSTEM TPL0059 | 2020-05-27 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the plan was reviewed with the surgeon with close attention paid to hypoplastic anatomy. The 360 bridge in combination with a tall and two short clamps was used with the MIST bridge connected to the Thomson system at the top of the construct to mount the system to the patient. Registration for the first segment was achieved using the basic algorithm when labeling from T11. Initial attempts to register the second segment failed. The manufacturer representative was able to achieve labeling from T7. T5 had inoperable registration values, but there was no shift. The surgical arm was sent to the operable values. The representative tried to re-register T5, but it was unsuccessful from T5-T7. The representative believed it was due to the clamp washing out the anatomy. All drilled trajectories were accurate. There was no patient harm and the procedure was delayed less t han an hour. | xse registration issues ccr spinal procedure treat deformity fixation plan reviewed surgeon close attention paid hypoplastic anatomy bridge combination tall short clamps mist bridge connected thomson construct mount registration segment achieved using basic algorithm labeling initial attempts register second segment failed able achieve labeling inoperable registration values shift surgical arm sent operable values tried reregister unsuccessful believed clamp washing anatomy drilled trajectories accurate procedure delayed t han hour mazor x | NAV3242 | NAV4112 | NAV5000 | NAV3243 | NAV3295 | NAV3242NAV4112NAV5000NAV3243NAV3295 | True
8833 | 704496196 | XSE - Pipe Missing on FCA Kit | 2021-07-12T16:05:16.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | KIT1370 | KIT1370 POSITIONER VALVE REMOVAL KIT | XSE - Pipe Missing on FCA Kit 2021-Jul-08Salesforce C00229930 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proce dure. It was reported that FCA KIT came only with the valve cover, the replacement pipe was missing. No patient was present. 2021-07-08 e1, e2, e3 (Rep, For): Additional information received from a manufacturer representative reported that an old cable from a demo system was able to be used for the FCA. 2021-10-21 IUD (Other): No new information. KIT1370 POSITIONER VALVE REMOVAL KIT | 2021-Jul-08Salesforce C00229930 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proce dure. It was reported that FCA KIT came only with the valve cover, the replacement pipe was missing. No patient was present. 2021-07-08 e1, e2, e3 (Rep, For): Additional information received from a manufacturer representative reported that an old cable from a demo system was able to be used for the FCA. 2021-10-21 IUD (Other): No new information. | xse pipe missing fca kit outside proce dure fca kit came valve cover replacement pipe missing present additional old cable demo able fca iud new positioner valve removal kit | NAV2053 | NAV3246 | NAV2054 | NAV3248 | NAV3262 | NAV2053NAV3246NAV2054NAV3248NAV3262 | True
1637 | 703208246 | X- 3Define Issues | 2019-06-14T18:09:55.000000+0000 | MazorX | NAV2051 | INTRA-OP ISSUES | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X- 3Define Issues All trajectories sent accurately, to Surgeons plan At L4 LEFT the surgical arm turned all the way around where the 3 define eyes were facing the ceiling. No issues and with placement with the guide wires and actually flipped correctly on S1 RIGHT (per BJ Wood) UNK MAZOR X 3DEFINE CAMERA | All trajectories sent accurately, to Surgeons plan At L4 LEFT the surgical arm turned all the way around where the 3 define eyes were facing the ceiling. No issues and with placement with the guide wires and actually flipped correctly on S1 RIGHT (per BJ Wood) | x issues trajectories sent accurately surgeons plan left surgical arm turned way define eyes facing ceiling issues placement guide wires actually flipped correctly right unk mazor x camera | NAV3244 | NAV4110 | NAV3249 | NAV3261 | NAV3264 | NAV3244NAV4110NAV3249NAV3261NAV3264 | False
3430 | 703346876 | XSE - LED Blinking/Arm Would Not Unlock | 2019-09-05T14:57:22.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - LED Blinking/Arm Would Not Unlock 2019-09-03 SCL1902586, Intake Form (HCP, REP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while booting up the system to perform the pre-op tasks the button that lifts the hood of the system was pressed and when raised it was noticed that the LED light was continuously blinking and even when hitting the lock/unlock button it would continue to flash and the arm wouldn't unlock. A soft restart was performed and the same issue occurred, next a hard restart was performed and the same issue occurred, continuous blinking and the arm wouldn't unlock. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2019-09-17 email x2 (Rep): Additional information received from a manufacturer representative reported that the trajectory was free handed, confirming how the case was completed as well as how it changed the case. 2020-01-30 e1, e2 (Other): No new information. SURGICAL ARM | 2019-09-03 SCL1902586, Intake Form (HCP, REP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while booting up the system to perform the pre-op tasks the button that lifts the hood of the system was pressed and when raised it was noticed that the LED light was continuously blinking and even when hitting the lock/unlock button it would continue to flash and the arm wouldn't unlock. A soft restart was performed and the same issue occurred, next a hard restart was performed and the same issue occurred, continuous blinking and the arm wouldn't unlock. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2019-09-17 email x2 (Rep): Additional information received from a manufacturer representative reported that the trajectory was free handed, confirming how the case was completed as well as how it changed the case. 2020-01-30 e1, e2 (Other): No new information. | xse led blinkingarm unlock spinal procedure booting perform preop tasks button lifts hood pressed raised noticed led light continuously blinking hitting lockunlock button continue flash arm wouldnt unlock soft restart performed issue occurred hard restart performed issue occurred continuous blinking arm wouldnt unlock issue extended surgical time hour impact outcome email additional trajectory free handed confirming case completed changed case new surgical arm | NAV2047 | NAV3259 | NAV4107 | NAV4101 | NAV2058 | NAV2047NAV3259NAV4107NAV4101NAV2058 | False
2519 | 703238535 | X - Software bug | 2019-07-02T14:13:18.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X - Software bug At the beginning of our preparation for the procedure during surgical arm set up: under device availability X eye and 3 define cameras were both red. After about 30 seconds to a minute, on its own, the X eye and 3 define camera switched back over to green. There was no trouble shooting done to correct this. Plan was done by me and reviewed by Dr. Ly the morning of the surgery. L3-L5 were planned for a percutaneous surgery. There were no concerns discussed pre operatively or changes made to the plan.A single schantz pin was utilized for the procedure and placed in the right PSIS. There were no complications mounting the platform. There were no complications during the 3define process or draw spine.Registration was achieved on the second attempt after additional images were taken. the oblique image had to be adjusted a little lower towards the feet and with a little bit more wag to clear up the image. During the procedure there were no issues until after the wires were placed accurately. After the wires were placed and we were waiting for the decompression and TLIF to be completed the screen automatically went to egistration without anyone touching the work station.-as the software worked through the registration process again, we received the message "registration problems have been detected. make sure labeling is correct." Although we ere done with our part of the procedure, there were multiple attempts to re register in case a wire came out. I then backed up to the labeling screen to select a body to register off of but registration still failed even though they were he same images we originally used to register off of and place the wires. no other trouble shooting was done as the case was over. Another case was conducted after this about 2 hours later and none of these problems happened again UNKNOWN MAZOR X SW | At the beginning of our preparation for the procedure during surgical arm set up: under device availability X eye and 3 define cameras were both red. After about 30 seconds to a minute, on its own, the X eye and 3 define camera switched back over to green. There was no trouble shooting done to correct this. Plan was done by me and reviewed by Dr. Ly the morning of the surgery. L3-L5 were planned for a percutaneous surgery. There were no concerns discussed pre operatively or changes made to the plan.A single schantz pin was utilized for the procedure and placed in the right PSIS. There were no complications mounting the platform. There were no complications during the 3define process or draw spine.Registration was achieved on the second attempt after additional images were taken. the oblique image had to be adjusted a little lower towards the feet and with a little bit more wag to clear up the image. During the procedure there were no issues until after the wires were placed accurately. After the wires were placed and we were waiting for the decompression and TLIF to be completed the screen automatically went to egistration without anyone touching the work station.-as the software worked through the registration process again, we received the message "registration problems have been detected. make sure labeling is correct." Although we ere done with our part of the procedure, there were multiple attempts to re register in case a wire came out. I then backed up to the labeling screen to select a body to register off of but registration still failed even though they were he same images we originally used to register off of and place the wires. no other trouble shooting was done as the case was over. Another case was conducted after this about 2 hours later and none of these problems happened again | x software bug beginning preparation procedure surgical arm set availability x eye define cameras red seconds minute x eye define camera switched green trouble shooting correct plan reviewed dr ly morning surgery planned percutaneous surgery concerns discussed pre operatively changes plana single schantz pin utilized procedure placed right psis complications mounting platform complications process draw spineregistration achieved second attempt additional images taken oblique image adjusted little lower feet little bit wag clear image procedure issues wires placed accurately wires placed waiting decompression tlif completed screen automatically went egistration touching work stationas software worked registration process message registration problems detected make sure labeling correct ere procedure multiple attempts register case wire came backed labeling screen select body register registration failed images originally register place wires trouble shooting case case conducted hours later problems happened unknown mazor x sw | NAV3242 | NAV4112 | NAV4110 | NAV3256 | NAV1006 | NAV3242NAV4112NAV4110NAV3256NAV1006 | False
6326 | 703912456 | XSE - Lost Target Extender Screw | 2020-08-19T21:39:07.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MAS2345-03S | KIT SVC MAS2345-03S TARGET EXTEND SCREW | XSE - Lost Target Extender Screw 2020-08-19 00694910 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the target extender screw was lost. There was no patient involvement. 2020-08-20 IUD x2 (Other): No new information. 2020-09-04 IUD x2 (Other): No new information. KIT SVC MAS2345-03S TARGET EXTEND SCREW | 2020-08-19 00694910 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the target extender screw was lost. There was no patient involvement. 2020-08-20 IUD x2 (Other): No new information. 2020-09-04 IUD x2 (Other): No new information. | xse lost target extender screw outside procedure target extender screw lost iud new iud new kit svc target extend screw | NAV2053 | NAV2054 | NAV3299 | NAV4090 | NAV2047 | NAV2053NAV2054NAV3299NAV4090NAV2047 | True
4423 | 703543109 | X - Dual Clamp Breakage | 2020-01-03T15:24:27.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | MAS1355-03 | DUAL CLAMP | X - Dual Clamp Breakage 2018-08-16 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a trauma injury with fixation at T10-S1. It was reported that plan was to extend previous construction at L3-S1 to include T10-L2. The shoulder calibration was checked and surgical arm set-up completed. While placing the dual clamp, the clamp broke and was unable to be used. The patient platform attachment was changed to a schanz pin inserted in the patient's right PSIS. Fluoro segmentation was completed with green values achieved at all levels. The surgeon visually verified. All trajectories drilled and screws placed safely and accurately to plan. Accuracy of placement was verified with AP/Lateral images and neuro stimulation. No patient harm was reported and surgery was delayed less than an hour. DUAL CLAMP | 2018-08-16 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a trauma injury with fixation at T10-S1. It was reported that plan was to extend previous construction at L3-S1 to include T10-L2. The shoulder calibration was checked and surgical arm set-up completed. While placing the dual clamp, the clamp broke and was unable to be used. The patient platform attachment was changed to a schanz pin inserted in the patient's right PSIS. Fluoro segmentation was completed with green values achieved at all levels. The surgeon visually verified. All trajectories drilled and screws placed safely and accurately to plan. Accuracy of placement was verified with AP/Lateral images and neuro stimulation. No patient harm was reported and surgery was delayed less than an hour. | x dual clamp breakage ccr spinal procedure treat trauma injury fixation plan extend previous construction include shoulder calibration checked surgical arm setup completed placing dual clamp clamp broke unable platform attachment changed schanz pin inserted patients right psis fluoro segmentation completed green values achieved levels surgeon visually verified trajectories drilled screws placed safely accurately plan accuracy placement verified aplateral images neuro stimulation surgery delayed hour dual clamp | NAV3295 | NAV4091 | NAV4095 | NAV3299 | NAV3242 | NAV3295NAV4091NAV4095NAV3299NAV3242 | True
9751 | 704714406 | XSE - Electrical Issue/Blown Fuse | 2021-11-13T04:07:25.000000+0000 | MazorX Stealth Edition | NAV3263 | POWER ISSUES | 40 | MAS2017-01S | KIT SVC MAS2017-01S COMPRESSOR ASSEMBLY | XSE - Electrical Issue/Blown Fuse 2021-Nov-12 Salesforce C00237831, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system began to emit a smell like "hot brakes" during a procedure. The use of the guidance system was aborted and the system was removed from the operating room. Several hard reboots were done while the manufacturer representative waited three minutes before starting the system. The system was unable to build pressure above 4.1 PSI. The representative removed the covers and found several blown fuses, which appeared to be the source of the smell. There was no patient harm and the procedure was delayed less than an hour. 2021-11-15 IUD (Other): No new information. 2021-11-19 IUD x2 (Other): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. 2021-12-21 IUD (Other): No new information. 2022-01-05 IUD (Other): No new information. KIT SVC MAS2017-01S COMPRESSOR ASSEMBLY | 2021-Nov-12 Salesforce C00237831, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system began to emit a smell like "hot brakes" during a procedure. The use of the guidance system was aborted and the system was removed from the operating room. Several hard reboots were done while the manufacturer representative waited three minutes before starting the system. The system was unable to build pressure above 4.1 PSI. The representative removed the covers and found several blown fuses, which appeared to be the source of the smell. There was no patient harm and the procedure was delayed less than an hour. 2021-11-15 IUD (Other): No new information. 2021-11-19 IUD x2 (Other): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. 2021-12-21 IUD (Other): No new information. 2022-01-05 IUD (Other): No new information. | xse electrical issueblown fuse salesforce iud spinal procedure began emit smell like hot brakes procedure use aborted removed operating room hard reboots waited minutes starting unable build pressure psi removed covers blown fuses appeared source smell procedure delayed hour iud new iud new iud new iud new iud new iud new kit svc compressor assembly | NAV3264 | NAV3299 | NAV2058 | NAV2047 | NAV4095 | NAV3264NAV3299NAV2058NAV2047NAV4095 | False
7765 | 704269693 | XSE - Shoulder Shift/Registration Mismat | 2021-03-09T01:14:35.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | XSE - Shoulder Shift/Registration Mismat 2021-Mar-04 Salesforce C00222304 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that a registration mismatch and shoulder shift occurred during the procedure. The surgeon attempted to re-register the patient after the shoulder shift, but after 30 minutes registration was unable to be completed. The surgeon had placed four screws on the right side so they were removed, the patient was closed and the surgery was aborted. There was no patient harm and the procedure was delayed less than an hour. 2021-03-14 IUD x2 (Other): No new information. 2021-03-29 e1, e2, e3, e4, e5, e6, e7 (Rep): No new information. 2021-04-27 email 2021-04-27 09:16 (Rep): No new information. 2021-05-20 email 2021-05-20 09:02 (Rep): No new information. UNK MAZOR X SHOULDER | 2021-Mar-04 Salesforce C00222304 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that a registration mismatch and shoulder shift occurred during the procedure. The surgeon attempted to re-register the patient after the shoulder shift, but after 30 minutes registration was unable to be completed. The surgeon had placed four screws on the right side so they were removed, the patient was closed and the surgery was aborted. There was no patient harm and the procedure was delayed less than an hour. 2021-03-14 IUD x2 (Other): No new information. 2021-03-29 e1, e2, e3, e4, e5, e6, e7 (Rep): No new information. 2021-04-27 email 2021-04-27 09:16 (Rep): No new information. 2021-05-20 email 2021-05-20 09:02 (Rep): No new information. | xse shoulder shiftregistration mismat salesforce spinal p rocedure registration mismatch shoulder shift occurred procedure surgeon attempted reregister shoulder shift minutes registration unable completed surgeon placed screws right removed closed surgery aborted procedure delayed hour iud new new email new email new unk mazor x shoulder | NAV4108 | NAV2058 | NAV3252 | NAV4107 | NAV4112 | NAV4108NAV2058NAV3252NAV4107NAV4112 | True
5769 | 703798356 | XSE - X-Eye Camera Not Functioning | 2020-06-11T21:15:35.000000+0000 | MazorX Stealth Edition | NAV3254 | X-EYE ISSUES | 30 | UK |  | XSE - X-Eye Camera Not Functioning 2020-06-11 00690612, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the X-Eye camera was not functioning. There was no patient involvement. 2020-06-19 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-06-30 IUD x2 (Other): No new information. 2020-07-14 IUD x3 (Other): No new information. 2020-08-02 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. 2020-10-27 IUD x2 (Other): No new information. | 2020-06-11 00690612, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the X-Eye camera was not functioning. There was no patient involvement. 2020-06-19 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-06-30 IUD x2 (Other): No new information. 2020-07-14 IUD x3 (Other): No new information. 2020-08-02 IUD (Other): No new information. 2020-08-05 IUD (Other): No new information. 2020-10-27 IUD x2 (Other): No new information. | xse xeye camera functioning iud outside procedure xeye camera functioning iud new iud new iud new iud new iud new iud new iud new | NAV4110 | NAV3299 | NAV1006 | NAV3254 | NAV3249 | NAV4110NAV3299NAV1006NAV3254NAV3249 | True
6553 | 703976961 | XSE - OOBF/Issue with RRF Prongs | 2020-09-25T00:02:29.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | KIT0899 | FRAME KIT0899 ROBOTIC FRAME CONNECTION | XSE - OOBF/Issue with RRF Prongs 2020-09-24 00697444 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the prongs on the robotic reference frame would not fit into the holder. The issue was found during installation. There was no patient involvement. 2020-09-29 e1, e2, e3, e4, e5, e6, IUD (Rep): Additional information received from a manufacturer representative reported that the prongs could be inserted, but it was difficult. The representative thought the prongs would not fit once the drape was in place. The representative noted that one prong was bigger than the other and it looked to have a single tac weld on one of the prongs. The weld was possibly interfering with the attachment. The replacement frame was working well. 2020-10-07 email 2020-10-07 13:00, IUD x2 (Rep): Additional information received from a manufacturer representative reported that the issue was resolved. FRAME KIT0899 ROBOTIC FRAME CONNECTION | 2020-09-24 00697444 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the prongs on the robotic reference frame would not fit into the holder. The issue was found during installation. There was no patient involvement. 2020-09-29 e1, e2, e3, e4, e5, e6, IUD (Rep): Additional information received from a manufacturer representative reported that the prongs could be inserted, but it was difficult. The representative thought the prongs would not fit once the drape was in place. The representative noted that one prong was bigger than the other and it looked to have a single tac weld on one of the prongs. The weld was possibly interfering with the attachment. The replacement frame was working well. 2020-10-07 email 2020-10-07 13:00, IUD x2 (Rep): Additional information received from a manufacturer representative reported that the issue was resolved. | xse oobfissue rrf prongs outside procedure prongs robotic reference frame fit holder issue installation iud additional prongs inserted difficult thought prongs fit drape place noted prong bigger looked single tac weld prongs weld possibly interfering attachment replacement frame working email iud additional issue resolved frame robotic frame connection | NAV2054 | NAV3299 | NAV3295 | NAV1006 | NAV3249 | NAV2054NAV3299NAV3295NAV1006NAV3249 | True
6823 | 704043628 | XSE - Damaged and lost instruments | 2020-10-30T14:40:51.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 50 | MAS2241-04S | KIT SVC MAS2241-04S LONG CANNULA | XSE - Damaged and lost instruments 2020-Oct-30 00699892 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the Schanz Bond is stripped and the site has lost their target extender screw. there was no patient involved. 2020-11-06 IUD x4 (Other): No new information. 2021-01-22 email 2021-01-22 19:45 (Rep): Additional information received from a manufacturer representative reported instrument was discarded. KIT SVC MAS2241-04S LONG CANNULA | 2020-Oct-30 00699892 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the Schanz Bond is stripped and the site has lost their target extender screw. there was no patient involved. 2020-11-06 IUD x4 (Other): No new information. 2021-01-22 email 2021-01-22 19:45 (Rep): Additional information received from a manufacturer representative reported instrument was discarded. | xse damaged lost instruments outside procedure schanz bond stripped site lost target extender screw involved iud new email additional instrument discarded kit svc long cannula | NAV2053 | NAV3299 | NAV2054 | NAV3295 | NAV3256 | NAV2053NAV3299NAV2054NAV3295NAV3256 | True
8050 | 704341548 | XSE - Deviated Trajectories/SW Froze | 2021-04-15T00:14:51.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Deviated Trajectories/SW Froze 2021-Apr-14 Salesforce C00224686, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was sent to the trajectories and the patient's skin was marked. The manufacturer representative then sent the surgical arm to S1 and hit the pause button, but the software froze. The representative let the system site and then did a soft restart. After losing connection to the controller, a hard restart of the system was done. The surgical arm was then sent to the trajectories on the right side. The surgeon was using the anti-skive tool and was hitting the tool very hard. The surgeon then moved to the screws on the left side. All screws were lateral after the right side were completed. There was no patient harm and the procedure was delayed less than an hour. 2021-07-23 e1 (Rep): Additional information received from a manufacturer representative reported that the amount of deviation was unknown, but only the left side was deviated laterally. The right side was accurate. No troubleshooting was done as the surgeon brought in the O-arm and navigation to reposition the lateral screws. The cause of the deviations was suspected to be the anti-skive pin being driven too far causing the patient to shift. 2021-08-02 e2 (Rep): No new information. 2021-08-03 e3, e4 (Rep): No new information. 2021-08-19 e5 (Rep): No new information. UNK MAZOR X WORKSTATION | 2021-Apr-14 Salesforce C00224686, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was sent to the trajectories and the patient's skin was marked. The manufacturer representative then sent the surgical arm to S1 and hit the pause button, but the software froze. The representative let the system site and then did a soft restart. After losing connection to the controller, a hard restart of the system was done. The surgical arm was then sent to the trajectories on the right side. The surgeon was using the anti-skive tool and was hitting the tool very hard. The surgeon then moved to the screws on the left side. All screws were lateral after the right side were completed. There was no patient harm and the procedure was delayed less than an hour. 2021-07-23 e1 (Rep): Additional information received from a manufacturer representative reported that the amount of deviation was unknown, but only the left side was deviated laterally. The right side was accurate. No troubleshooting was done as the surgeon brought in the O-arm and navigation to reposition the lateral screws. The cause of the deviations was suspected to be the anti-skive pin being driven too far causing the patient to shift. 2021-08-02 e2 (Rep): No new information. 2021-08-03 e3, e4 (Rep): No new information. 2021-08-19 e5 (Rep): No new information. | xse deviated trajectoriessw froze salesforce iud spinal procedure surgical arm sent trajectories patients skin marked sent surgical arm hit pause button software froze let site did soft restart losing connection controller hard restart surgical arm sent trajectories right surgeon using antiskive tool hitting tool hard surgeon moved screws left screws lateral right completed procedure delayed hour additional deviation unknown left deviated laterally right accurate troubleshooting surgeon brought oarm navigation reposition lateral screws cause deviations suspected antiskive pin driven far causing shift new new new unk mazor x workstation | NAV2047 | NAV3248 | NAV5000 | NAV3264 | NAV4108 | NAV2047NAV3248NAV5000NAV3264NAV4108 | True
9308 | 704606186 | XSE - OOBF/Installation Issues | 2021-09-13T05:55:40.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 30 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE - OOBF/Installation Issues 2021-Aug-07 Salesforce C00234297 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the system failed to install software, the manipulator handle cannot firmly old the weight of the Surgical system. X eye camera software not installed. Surgical system bin damaged from Bad packing during transit. Surgeon monitor Doesn't function correctly in application. There was no patient involved. 2021-09-23 e1 (Rep, For): Additional information received from a manufacturer representative reported that the the firmware file for the X-Eye camera was not on the PC. The surgical system had rub markings from bad packaging that caused damage to the plastic covers in the storage area. The surgeon monitor was unable to be linked to the application correctly. Multiple displays had to be engaged several times for the screen to register as a display and have touch screen input into the application. 2021-10-04 IUDx2, e2, e3 (Other, For): No new information. 2021-10-18 e4 (Other, For): No new information. 2021-10-28 e5, e6 (Other, For): No new information. 2021-11-07 e7, e8 (Other, For): No new information. 2021-11-08 e9 (Rep, For): Additional information received from a manufacturer representative reported the electrical panel looked like it worked, but failed to accept the installed software. The covers in the storage area were not replaced. The issue with the manipulator was not resolved. The shoulder was unable to be moved more than 10 degrees from center when on the manipulator. 2021-11-15 e10 (Other, For): No new information. 2021-11-22 e11 (Rep, For): Additional information received from a manufacturer representative reported that the surgical arm was passible, but the shoulder could not be moved a lot while the surgical system was on the manipulator, otherwise the system would tilt from 0 to 10 degrees. The X-Eye calibration file was not on the system so they had to get the correct file and load it to restore function. After several disconnections in the Windows multi-monitor display window, function was restored to the surgeon screen. The best result was to switch the monitor to duplicate, test the touch screen input in the software, and then set the software back to multiple displays. 2021-12-07 e12 (Rep, For): No new information. 2021-12-13 e13, e14 (Rep, For): No new information. 2022-01-11 e15, e16, e17, IUD (Rep, For): No new information. 2022-02-10 302721036 (Other, For): No new information. 2022-02-11 IUD x2 (Other, For): No new information. 2022-02-13 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. UNK MAZOR X MANIPULATOR | 2021-Aug-07 Salesforce C00234297 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the system failed to install software, the manipulator handle cannot firmly old the weight of the Surgical system. X eye camera software not installed. Surgical system bin damaged from Bad packing during transit. Surgeon monitor Doesn't function correctly in application. There was no patient involved. 2021-09-23 e1 (Rep, For): Additional information received from a manufacturer representative reported that the the firmware file for the X-Eye camera was not on the PC. The surgical system had rub markings from bad packaging that caused damage to the plastic covers in the storage area. The surgeon monitor was unable to be linked to the application correctly. Multiple displays had to be engaged several times for the screen to register as a display and have touch screen input into the application. 2021-10-04 IUDx2, e2, e3 (Other, For): No new information. 2021-10-18 e4 (Other, For): No new information. 2021-10-28 e5, e6 (Other, For): No new information. 2021-11-07 e7, e8 (Other, For): No new information. 2021-11-08 e9 (Rep, For): Additional information received from a manufacturer representative reported the electrical panel looked like it worked, but failed to accept the installed software. The covers in the storage area were not replaced. The issue with the manipulator was not resolved. The shoulder was unable to be moved more than 10 degrees from center when on the manipulator. 2021-11-15 e10 (Other, For): No new information. 2021-11-22 e11 (Rep, For): Additional information received from a manufacturer representative reported that the surgical arm was passible, but the shoulder could not be moved a lot while the surgical system was on the manipulator, otherwise the system would tilt from 0 to 10 degrees. The X-Eye calibration file was not on the system so they had to get the correct file and load it to restore function. After several disconnections in the Windows multi-monitor display window, function was restored to the surgeon screen. The best result was to switch the monitor to duplicate, test the touch screen input in the software, and then set the software back to multiple displays. 2021-12-07 e12 (Rep, For): No new information. 2021-12-13 e13, e14 (Rep, For): No new information. 2022-01-11 e15, e16, e17, IUD (Rep, For): No new information. 2022-02-10 302721036 (Other, For): No new information. 2022-02-11 IUD x2 (Other, For): No new information. 2022-02-13 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. | xse oobfinstallation issues salesforce outside proc edure failed install software manipulator handle firmly old weight surgical x eye camera software installed surgical bin damaged bad packing transit surgeon monitor doesnt function correctly application involved additional firmware file xeye camera pc surgical rub markings bad packaging caused damage plastic covers storage area surgeon monitor unable linked application correctly multiple displays engaged times screen register display touch screen input application new new new new additional electrical panel looked like worked failed accept installed software covers storage area replaced issue manipulator resolved shoulder unable moved degrees center manipulator new additional surgical arm passible shoulder moved lot surgical manipulator tilt degrees xeye calibration file correct file load restore function disconnections windows multimonitor display window function restored surgeon screen best switch monitor duplicate test touch screen input software set software multiple displays new new iud new new iud new iud new iud new unk mazor x manipulator | NAV3260 | NAV2054 | NAV3261 | NAV2057 | NAV3263 | NAV3260NAV2054NAV3261NAV2057NAV3263 | True
8261 | 704381164 | XSE - Monitor/Keyboard Issues | 2021-05-06T19:47:57.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XSE - Monitor/Keyboard Issues 2021-May-06 Salesforce C00225929, 00712268 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the workstation screen intermittently went black during the case and the keyboard was unre sponsive. The manufacturer representative was not sure if the surgeon screen went black at the same time. A hard reboot of the system was done and the keyboard issue was resolved. There was no patient harm and the procedure was delayed less than an hour. 2021-05-07 IUD x2 (Other): No new information. UNK MAZOR X MONITOR | 2021-May-06 Salesforce C00225929, 00712268 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the workstation screen intermittently went black during the case and the keyboard was unre sponsive. The manufacturer representative was not sure if the surgeon screen went black at the same time. A hard reboot of the system was done and the keyboard issue was resolved. There was no patient harm and the procedure was delayed less than an hour. 2021-05-07 IUD x2 (Other): No new information. | xse monitorkeyboard issues salesforce spinal procedure workstation screen intermittently went black case keyboard unre sponsive sure surgeon screen went black time hard reboot keyboard issue resolved procedure delayed hour iud new unk mazor x monitor | NAV2058 | NAV3261 | NAV3253 | NAV2047 | NAV3247 | NAV2058NAV3261NAV3253NAV2047NAV3247 | True
1471 | 703202185 | X- 3Define Issues | 2019-06-12T12:11:34.000000+0000 | MazorX | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | X- 3Define Issues 2018-2-6 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a percutaneous spinal procedure on L3-L5 to address degenerative discs. It was reported that planning was done by the manufacturer representative and then reviewed by the surgeon the morning of the surgery. The surgical arm was set up with no complications. A schanz pin was used as the mount. During the 3Define scan, the representative received a message saying the arm did not reach the position and the arm stopped moving. The surgical arm was sent to the up position and then the 3Define scan was conducted again with success. Registration was achieved on the first attempt, and the operation had no other issues. There was no patient harm and the procedure was not delayed over an hour. MAZOR X PC STATION - REFURBISHED | 2018-2-6 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a percutaneous spinal procedure on L3-L5 to address degenerative discs. It was reported that planning was done by the manufacturer representative and then reviewed by the surgeon the morning of the surgery. The surgical arm was set up with no complications. A schanz pin was used as the mount. During the 3Define scan, the representative received a message saying the arm did not reach the position and the arm stopped moving. The surgical arm was sent to the up position and then the 3Define scan was conducted again with success. Registration was achieved on the first attempt, and the operation had no other issues. There was no patient harm and the procedure was not delayed over an hour. | x issues ccr percutaneous spinal procedure address degenerative discs planning reviewed surgeon morning surgery surgical arm set complications schanz pin mount scan message saying arm did reach position arm stopped moving surgical arm sent position scan conducted success registration achieved attempt operation issues procedure delayed hour mazor x pc station refurbished | NAV3242 | NAV4095 | NAV3244 | NAV4091 | NAV3246 | NAV3242NAV4095NAV3244NAV4091NAV3246 | False
8882 | 704502172 | X - Damaged System Parts | 2021-07-15T00:07:24.000000+0000 | MazorX | NAV3258 | IMAGE ADAPTOR ISSUES | 30 | KIT1386 | KIT1386 STANDARD IMAGE ADAPTOR SPARE PRT | X - Damaged System Parts 2021-Jul-14 Salesforce C00230210 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system 3define camera does not function. The hinge for the image adapter is damaged and stuck on the system and can not be removed. The x-eye camera is damaged and will not move without physically moving the camera and pointing it at the area of interest. No patient was present. 2021-08-03 IUD (Other): No new information. 2021-Aug-16 email 8/16 IUDx2 (rep): Additional information received indicated that the issues had not been resolved yet. All damage done to the system is normal wear and tear. This system is old and has not had an active service contract in quite a while. 2021-10-06 00723504 (Rep): No new information. 2021-10-13 IUD (Other): No new information. 2021-10-20 IUD x2 (Other): No new information. 2021-11-04 IUD (Other): No new information. 2022-03-10 IUD (Other): No new information. 2022-03-16 IUD x2 (Other): No new information. 2022-03-23 IUD (Other): No new information. 2022-03-26 IUD (Other): No new information. 2022-04-11 IUD (Other): No new information. KIT1386 STANDARD IMAGE ADAPTOR SPARE PRT | 2021-Jul-14 Salesforce C00230210 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system 3define camera does not function. The hinge for the image adapter is damaged and stuck on the system and can not be removed. The x-eye camera is damaged and will not move without physically moving the camera and pointing it at the area of interest. No patient was present. 2021-08-03 IUD (Other): No new information. 2021-Aug-16 email 8/16 IUDx2 (rep): Additional information received indicated that the issues had not been resolved yet. All damage done to the system is normal wear and tear. This system is old and has not had an active service contract in quite a while. 2021-10-06 00723504 (Rep): No new information. 2021-10-13 IUD (Other): No new information. 2021-10-20 IUD x2 (Other): No new information. 2021-11-04 IUD (Other): No new information. 2022-03-10 IUD (Other): No new information. 2022-03-16 IUD x2 (Other): No new information. 2022-03-23 IUD (Other): No new information. 2022-03-26 IUD (Other): No new information. 2022-04-11 IUD (Other): No new information. | x damaged parts salesforce outside procedure camera does function hinge image adapter damaged stuck removed xeye camera damaged physically moving camera pointing area present iud new email additional indicated issues resolved damage normal wear tear old active service contract quite new iud new iud new iud new iud new iud new iud new iud new iud new standard image adaptor spare prt | NAV3244 | NAV3258 | NAV3299 | NAV3254 | NAV2057 | NAV3244NAV3258NAV3299NAV3254NAV2057 | True
468 | 703111472 | X - Monitor Issue/Broken Positioner | 2019-04-17T19:13:01.000000+0000 | MazorX | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 40 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Monitor Issue/Broken Positioner 2019-04-17 SCL1900973 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the workstation monitor when black during the planning process, C-arm calibration and after a hard restart of the system. The manufacturer representative also reported the right positioner handle was broken due to the system being run into a door frame. Wires had been exposed. No patient harm was reported. UNK MAZOR X WORKSTATION | 2019-04-17 SCL1900973 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the workstation monitor when black during the planning process, C-arm calibration and after a hard restart of the system. The manufacturer representative also reported the right positioner handle was broken due to the system being run into a door frame. Wires had been exposed. No patient harm was reported. | x monitor issuebroken positioner spinal procedure workstation monitor black planning process carm calibration hard restart right positioner handle broken run door frame wires exposed unk mazor x workstation | NAV3261 | NAV2058 | NAV3251 | NAV2047 | NAV3264 | NAV3261NAV2058NAV3251NAV2047NAV3264 | False
7580 | 704231307 | XSE ¦ Low Pressure | 2021-02-15T21:02:03.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE ¦ Low Pressure 2021-02-15 00706617 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that they were having issues with mounting the system to the bed. The site noted that every time they attached the system or went to attach the system to the bed they were getting a low pressure error. There was no delay and impact on the patient outcome. 2021-02-16 IUD x4 (Other): No new information. 2021-03-04 IUD x2 (Other): No new information. 2021-04-12 IUD x3 (Other): No new information. 2022-03-23 IUD (Other): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2021-02-15 00706617 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that they were having issues with mounting the system to the bed. The site noted that every time they attached the system or went to attach the system to the bed they were getting a low pressure error. There was no delay and impact on the patient outcome. 2021-02-16 IUD x4 (Other): No new information. 2021-03-04 IUD x2 (Other): No new information. 2021-04-12 IUD x3 (Other): No new information. 2022-03-23 IUD (Other): No new information. | xse low pressure intraoperatively spinal procedure having issues mounting bed site noted time attached went attach bed getting low pressure error delay impact outcome iud new iud new iud new iud new kit svc pneumatic panel | NAV3264 | NAV3250 | NAV1006 | NAV2054 | NAV2047 | NAV3264NAV3250NAV1006NAV2054NAV2047 | True
6259 | 703898852 | X - Images in 3Define Scan | 2020-08-11T23:30:04.000000+0000 | MazorX | NAV3239 | INACCURATE WORK VOLUMES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Images in 3Define Scan 2020-08-11 00694422 (Rep, HPC): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the surgeon potentially put his arm in the 3Define scan adding additional artifacts to the image. The manufacturer representative was able to send to the AP but not the oblique. It could never calculate the movement. The surgeon aborted the use of the guidance system. There was no impact on the patient outcome. 2020-08-18 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-08-11 00694422 (Rep, HPC): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the surgeon potentially put his arm in the 3Define scan adding additional artifacts to the image. The manufacturer representative was able to send to the AP but not the oblique. It could never calculate the movement. The surgeon aborted the use of the guidance system. There was no impact on the patient outcome. 2020-08-18 IUD (Other): No new information. | x images scan intraoperatively spinal procedure surgeon potentially arm scan adding additional artifacts image able send ap oblique calculate movement surgeon aborted use impact outcome iud new mazor x | NAV4112 | NAV3246 | NAV3243 | NAV1006 | NAV3253 | NAV4112NAV3246NAV3243NAV1006NAV3253 | False
4732 | 703570954 | X-pressure, communication, 3Define issue | 2020-01-20T18:08:23.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | ASM0219-01 | WORKSTATION | X-pressure, communication, 3Define issue 2018-8-1 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and the surgical arm was set up. The 3Define camera had no firmware, so a soft restart was performed. Registration was achieved using the basic algorithm with green values at all surgical levels. The k-wire placement appeared accurate in the lateral x-ray, but the surgeon was not able to match the trajectory. While placing screws, left L2 kept slipping laterally so the surgeon skipped left L2. All other screws were placed accurately and verified with fluoro. The surgical arm was resent to left L2, but a CANBUS error was received and the arm became unresponsive. The surgical arm was dismounted, and troubleshooting was attempted. Node errors appeared at all joints, and the compressor became inoperable. There was no patient harm and the procedure was not delayed over an hour. WORKSTATION | 2018-8-1 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and the surgical arm was set up. The 3Define camera had no firmware, so a soft restart was performed. Registration was achieved using the basic algorithm with green values at all surgical levels. The k-wire placement appeared accurate in the lateral x-ray, but the surgeon was not able to match the trajectory. While placing screws, left L2 kept slipping laterally so the surgeon skipped left L2. All other screws were placed accurately and verified with fluoro. The surgical arm was resent to left L2, but a CANBUS error was received and the arm became unresponsive. The surgical arm was dismounted, and troubleshooting was attempted. Node errors appeared at all joints, and the compressor became inoperable. There was no patient harm and the procedure was not delayed over an hour. | xpressure communication issue ccr spinal procedure shoulder calibration checked surgical arm set camera firmware soft restart performed registration achieved using basic algorithm green values surgical levels kwire placement appeared accurate lateral xray surgeon able match trajectory placing screws left kept slipping laterally surgeon skipped left screws placed accurately verified fluoro surgical arm resent left canbus error arm unresponsive surgical arm dismounted troubleshooting attempted node errors appeared joints compressor inoperable procedure delayed hour workstation | NAV2047 | NAV4095 | NAV4091 | NAV4098 | NAV3249 | NAV2047NAV4095NAV4091NAV4098NAV3249 | False
7222 | 704149725 | XSE - Damaged BNC/Ethernet Cables | 2020-12-30T21:21:51.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 20 | KIT1229-01 | SVC KIT1229-01 CBL BNC | XSE - Damaged BNC/Ethernet Cables 2020-12-30 00703534, 00703531 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the BNC cable that ran from the guidance system to the C-arm was frayed and damaged due to over use. The cable from the guidance system to the stealth camera was frayed and damaged. There was no patient involvement. 2021-01-07 IUD x3 (Other): No new information. 2021-02-02 email 2021-02-02 07:46 (Rep): No new information. 2021-03-10 email 2021-03-10 20:30 (Rep): No new information. SVC KIT1229-01 CBL BNC | 2020-12-30 00703534, 00703531 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the BNC cable that ran from the guidance system to the C-arm was frayed and damaged due to over use. The cable from the guidance system to the stealth camera was frayed and damaged. There was no patient involvement. 2021-01-07 IUD x3 (Other): No new information. 2021-02-02 email 2021-02-02 07:46 (Rep): No new information. 2021-03-10 email 2021-03-10 20:30 (Rep): No new information. | xse damaged bncethernet cables spinal procedure bnc cable ran carm frayed damaged use cable stealth camera frayed damaged iud new email new email new svc cbl bnc | NAV3256 | NAV3295 | NAV2053 | NAV1006 | NAV3249 | NAV3256NAV3295NAV2053NAV1006NAV3249 | True
6017 | 703859623 | X - Registration Issues/Cross View Error | 2020-07-19T13:19:04.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issues/Cross View Error 2020-06-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during an L5-S1 MIS TLIF, there was a crossview error during registration due to some of the 3D marker beads missing in the first AP image. A new AP image was taken and registration was able to be completed successfully. All k-wires were accurately placed. There was no patient harm and the procedure was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2020-06-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during an L5-S1 MIS TLIF, there was a crossview error during registration due to some of the 3D marker beads missing in the first AP image. A new AP image was taken and registration was able to be completed successfully. All k-wires were accurately placed. There was no patient harm and the procedure was delayed less than an hour. | x registration issuescross view error ccr spinal procedure mis tlif crossview error registration marker beads missing ap image new ap image taken registration able completed successfully kwires accurately placed procedure delayed hour mazor x | NAV3242 | NAV4112 | NAV3237 | NAV4091 | NAV1006 | NAV3242NAV4112NAV3237NAV4091NAV1006 | True
6381 | 703932181 | XSE - Deviated Screw | 2020-08-31T21:45:46.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Screw 2020-08-31 00695574 (Rep, HCP): Medtronic received information regarding a guidance system that was used intra-operatively of a spinal procedure. It was reported that the post-op images showed that the right T10 screw was breached medial. This was the second construct they drilled the left side screws first and than T10 was the first right screw drilled and the only screw deviated. The surgeon has reported the screw breached the medial wall. There was no impact on the patient outcome. 2020-09-02 email 2020-09-02 12:30, IUD (Rep): Additional information received from a manufacturer representative reported that the t rajectories were deviated between 3.5 and 10 mm. No revision had been scheduled. The surgeon was waiting to determine what to do based on the patient's pain. MAZOR X SYSTEM TPL0059 | 2020-08-31 00695574 (Rep, HCP): Medtronic received information regarding a guidance system that was used intra-operatively of a spinal procedure. It was reported that the post-op images showed that the right T10 screw was breached medial. This was the second construct they drilled the left side screws first and than T10 was the first right screw drilled and the only screw deviated. The surgeon has reported the screw breached the medial wall. There was no impact on the patient outcome. 2020-09-02 email 2020-09-02 12:30, IUD (Rep): Additional information received from a manufacturer representative reported that the t rajectories were deviated between 3.5 and 10 mm. No revision had been scheduled. The surgeon was waiting to determine what to do based on the patient's pain. | xse deviated screw intraoperatively spinal procedure postop images showed right screw breached medial second construct drilled left screws right screw drilled screw deviated surgeon screw breached medial wall impact outcome email iud additional t rajectories deviated mm revision scheduled surgeon waiting determine based patients pain mazor x | NAV5000 | NAV2054 | NAV2050 | NAV3242 | NAV3244 | NAV5000NAV2054NAV2050NAV3242NAV3244 | True
6728 | 704015924 | XSE - Monitor Not Displaying Video | 2020-10-16T00:15:43.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | UNK_MAZORX_PC | UNK MAZOR X PC | XSE - Monitor Not Displaying Video 2020-Oct-15 00698911 (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported when they first boot up the system the system did not display a video for about 10 mins. The compressor came on but no video. They performed a hard restart and the monitor displayed a video. When they shut the system down they go the blue screen saying the system did not shut down properly. They waited until it reached 100% and rebooted. When the system came on it still did not display a video feed. There was no patient involved. 2020-10-26 IUD x2 (Other): No new information. UNK MAZOR X PC | 2020-Oct-15 00698911 (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported when they first boot up the system the system did not display a video for about 10 mins. The compressor came on but no video. They performed a hard restart and the monitor displayed a video. When they shut the system down they go the blue screen saying the system did not shut down properly. They waited until it reached 100% and rebooted. When the system came on it still did not display a video feed. There was no patient involved. 2020-10-26 IUD x2 (Other): No new information. | xse monitor displaying video outside procedure boot did display video mins compressor came video performed hard restart monitor displayed video shut blue screen saying did shut properly waited reached rebooted came did display video feed involved iud new unk mazor x pc | NAV3262 | NAV2054 | NAV3261 | NAV3247 | NAV4101 | NAV3262NAV2054NAV3261NAV3247NAV4101 | True
1373 | 703197863 | X- SW/Deviation | 2019-06-10T18:12:41.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X- SW/Deviation L2 R-Trajectory was difficult to reach due to soft tissue surgeon pushed on retractor laterally, Rep stated that this pressure may push spine lateral and tools will dock medially Upon AP L2 R was noted to be medial, surgeon free handed L2R MAZOR X SOFTWARE KIT VER-3.0.2 | L2 R-Trajectory was difficult to reach due to soft tissue surgeon pushed on retractor laterally, Rep stated that this pressure may push spine lateral and tools will dock medially Upon AP L2 R was noted to be medial, surgeon free handed L2R | x swdeviation rtrajectory difficult reach soft tissue surgeon pushed retractor laterally stated pressure push spine lateral tools dock medially ap r noted medial surgeon free handed mazor x software kit | NAV4091 | NAV3246 | NAV3242 | NAV4095 | NAV5000 | NAV4091NAV3246NAV3242NAV4095NAV5000 | True
4953 | 703587457 | X - Registration issues | 2020-01-29T19:56:42.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration issues 2018-03-05 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L2. It was reported that the surgeon and manufacturer representative reviewed the plan prior to the procedure. The system was attached to the bed, draping was completed, bone mount bridge, reference marker, and arm guide were attached. A Schanz pin was placed in the right PSIS and connected to the BMB with a Schanz bond. 3Define scan was successful. Draw spine was completed. Two AP and six OBL images were taken. On the first registration attempt L2 had a yellow matching value. A significant shift was detected on AP and OBL when approving registration. Registration was achieved on the third attempt, with the advanced algorithm. The surgeon verified matching accuracies at all levels. Left and right L2 were placed and verified in lateral. Surgeon was pleased with outcome. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-03-05 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L2. It was reported that the surgeon and manufacturer representative reviewed the plan prior to the procedure. The system was attached to the bed, draping was completed, bone mount bridge, reference marker, and arm guide were attached. A Schanz pin was placed in the right PSIS and connected to the BMB with a Schanz bond. 3Define scan was successful. Draw spine was completed. Two AP and six OBL images were taken. On the first registration attempt L2 had a yellow matching value. A significant shift was detected on AP and OBL when approving registration. Registration was achieved on the third attempt, with the advanced algorithm. The surgeon verified matching accuracies at all levels. Left and right L2 were placed and verified in lateral. Surgeon was pleased with outcome. There was no patient harm and the procedure was not delayed over an hour. | x registration issues ccr spinal procedure treat degenerative disk disease fixation surgeon reviewed plan prior procedure attached bed draping completed bone mount bridge reference marker arm guide attached schanz pin placed right psis connected bmb schanz bond scan successful draw spine completed ap obl images taken registration attempt yellow matching value significant shift detected ap obl approving registration registration achieved attempt advanced algorithm surgeon verified matching accuracies levels left right placed verified lateral surgeon pleased outcome procedure delayed hour mazor x | NAV3242 | NAV4091 | NAV4112 | NAV3247 | NAV3254 | NAV3242NAV4091NAV4112NAV3247NAV3254 | True
458 | 703111299 | X - Arm Failed 3 Point Test | 2019-04-17T17:52:54.000000+0000 | MazorX | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | X - Arm Failed 3 Point Test 2019-04-16 SCL1900968, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm failed the 3 point test. A hard restart of the system was done and then the 3 point test was able to be passed. During the procedure, the surgical arm shook so bad the patient was shaken. The manufacturer representative mentioned the surgical arm failed to reach the first trajectory and the drape around the arm had to be loosened for it to move. The shoulder would not unlock or lock, the positioner would not move, and there were no lights once the surgical system was removed from the operating table. When moving the shoulder to the storage position a "Joint 5 Shift" error message was displayed. A hard restart of the system was done three times to resolve the error message. There was no patient harm and the procedure was not delayed more than an hour. SURGICAL ARM | 2019-04-16 SCL1900968, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm failed the 3 point test. A hard restart of the system was done and then the 3 point test was able to be passed. During the procedure, the surgical arm shook so bad the patient was shaken. The manufacturer representative mentioned the surgical arm failed to reach the first trajectory and the drape around the arm had to be loosened for it to move. The shoulder would not unlock or lock, the positioner would not move, and there were no lights once the surgical system was removed from the operating table. When moving the shoulder to the storage position a "Joint 5 Shift" error message was displayed. A hard restart of the system was done three times to resolve the error message. There was no patient harm and the procedure was not delayed more than an hour. | x arm failed point test spinal procedure surgical arm failed point test hard restart point test able passed procedure surgical arm shook bad shaken mentioned surgical arm failed reach trajectory drape arm loosened shoulder unlock lock positioner lights surgical removed operating table moving shoulder storage position joint shift error message displayed hard restart times resolve error message procedure delayed hour surgical arm | NAV4095 | NAV4101 | NAV4104 | NAV3249 | NAV3259 | NAV4095NAV4101NAV4104NAV3249NAV3259 | False
8019 | 704333210 | XSE - SRA Failed 5.0 Upgrade | 2021-04-10T00:12:45.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | XSE - SRA Failed 5.0 Upgrade 2021-Apr-09 Salesforce C00224453 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during the 5.0 upgrade the system failed the accuracy check. No patient was present. 2021-04-15 IUD x5 (Other): No new information. 2021-06-09 IUD (other): No new information. 2021-08-02 302514180 (Other): No new information. 2021-08-03 IUD x2 (Other): No new information. 2021-08-05 IUD x5 (Other): No new information. 2021-09-01 IUD, Attachment (Other): No new information. SURGICAL ARM | 2021-Apr-09 Salesforce C00224453 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during the 5.0 upgrade the system failed the accuracy check. No patient was present. 2021-04-15 IUD x5 (Other): No new information. 2021-06-09 IUD (other): No new information. 2021-08-02 302514180 (Other): No new information. 2021-08-03 IUD x2 (Other): No new information. 2021-08-05 IUD x5 (Other): No new information. 2021-09-01 IUD, Attachment (Other): No new information. | xse sra failed upgrade salesforce outside procedure upgrade failed accuracy check present iud new iud new new iud new iud new iud attachment new surgical arm | NAV3249 | NAV4092 | NAV3299 | NAV4093 | NAV2047 | NAV3249NAV4092NAV3299NAV4093NAV2047 | True
8142 | 704361150 | XSE - Poor Quality O-arm Images | 2021-04-26T16:21:59.000000+0000 | MazorX Stealth Edition | NAV3235 | 3D SCANS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Poor Quality O-arm Images 2021-Apr-20 Salesforce C00225287 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the O-arm scan was acquired and transferred to the guidance system during a Scan & Plan case. The image quality was very poor when transferred vis USB to the guidance system. The quality of the scan was good on the O-arm. The surgeon was able to plan the trajectories using the poor quality images. There was no patient harm and the procedure was not delayed. 2021-05-02 e1 (Other): No new information. 2021-05-03 IUD x2 (Other): No new information. 2021-05-18 IUD x2 (Other): No new information. 2021-05-18 email 2021-05-18 14:31 (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Apr-20 Salesforce C00225287 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the O-arm scan was acquired and transferred to the guidance system during a Scan & Plan case. The image quality was very poor when transferred vis USB to the guidance system. The quality of the scan was good on the O-arm. The surgeon was able to plan the trajectories using the poor quality images. There was no patient harm and the procedure was not delayed. 2021-05-02 e1 (Other): No new information. 2021-05-03 IUD x2 (Other): No new information. 2021-05-18 IUD x2 (Other): No new information. 2021-05-18 email 2021-05-18 14:31 (Other): No new information. | xse poor quality oarm images salesforce spinal p rocedure oarm scan acquired transferred scan plan case image quality poor transferred vis usb quality scan good oarm surgeon able plan trajectories using poor quality images procedure delayed new iud new iud new email new mazorx sw installation kit | NAV3235 | NAV3246 | NAV3248 | NAV3242 | NAV2058 | NAV3235NAV3246NAV3248NAV3242NAV2058 | True
5099 | 703597644 | XSE - CANBUS Error | 2020-02-05T00:12:57.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CANBUS Error 2020-02-04 00683416, images 2020-02-24 13:32 x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system encountered a CANBUS error while the surgeon was doing a decompression. The manufacturer representative did a soft reboot of the system, but this did not clear the error. Two hard reboots were done and then the system was able to function. There was no patient harm and the procedure was not delayed over an hour. 2020-02-19 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2020-02-04 00683416, images 2020-02-24 13:32 x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system encountered a CANBUS error while the surgeon was doing a decompression. The manufacturer representative did a soft reboot of the system, but this did not clear the error. Two hard reboots were done and then the system was able to function. There was no patient harm and the procedure was not delayed over an hour. 2020-02-19 IUD (Other): No new information. | xse canbus error images spinal procedure encountered canbus error surgeon doing decompression did soft reboot did clear error hard reboots able function procedure delayed hour iud new unk mazor x workstation | NAV2047 | NAV2058 | NAV3264 | NAV3256 | NAV3244 | NAV2047NAV2058NAV3264NAV3256NAV3244 | True
3668 | 703411420 | XSE - Schanz arms not locking | 2019-10-14T15:29:44.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2351-03 | SCHANZ ARM | XSE - Schanz arms not locking 2019-Oct-10 SCL1903052 (HCP,rep); Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that three schanz arms were no longer locking properly. There was no patient harm or additional complications reported or anticipated as a result of the event. SCHANZ ARM | 2019-Oct-10 SCL1903052 (HCP,rep); Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that three schanz arms were no longer locking properly. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse schanz arms locking spinal procedure schanz arms longer locking properly additional complications anticipated schanz arm | NAV3299 | NAV3296 | NAV3295 | NAV3251 | NAV5000 | NAV3299NAV3296NAV3295NAV3251NAV5000 | True
9501 | 704658475 | XSE - Registration Issues/SW Defect | 2021-10-12T18:14:20.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Registration Issues/SW Defect 2021-Oct-12 Salesforce C00236114 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure to treat degenerative discs. It was reported that during registration there was difficulty achieving a passing registration. Multiple attempts were made, and the patient was eventually registered via scan-and-plan workflow. After the case the manufacturer representative noticed that the red dot for registration that she had placed in planning on her laptop had moved when the export was transferred to the guidance system. She had placed the dot at the base of the vertebra but when she checked after the procedure it showed next to the spinous process. There was a delay of about an hour and 20 minutes. There was no impact on the patient outcome. 2021-10-15 email 2021-10-15 10:16, image 2021-10-15 10:16 x2, email 2021-10-18 06:46, email 2021-10-19 09:32, email 2021-10-20 08:16, email 2021-10-20 08:32 x2, email 2021-10-20 08:46, email 2021-10-20 09:48, CCR, IUD (Rep): Additional information received from a manufacturer representative reported that the surgical system was mounted using a bone mount bridge connected to a Schanz screw placed in the left PSIS. Everything was normal in the case until placing the red dot during registration. Since the case was an SI fusion, L5 and S1 were segmented. There were issues with registration being denied even after getting new AP and oblique images, moving the red dot around, and trying semi-auto and manual registration a few times. When reviewing the oblique CT and fluoro shots, the red dot was found to have moved on L5. The representative pulled up the plan on their laptop and found the red dot was correctly placed and the issue occurred during the transfer from the laptop to the guidance system. The representative tested the patient's export and another patient. The representative moved the red dot tot eh correct spot again for the patient's export, but when going into a demo case the red dot moved again. The representative moved the dot to the correct spot and continued with a 3Define scan. When going into registration after acquiring images, the red dot moved again and was to the left of the spinous process. When testing the export from a different patient, there were no issues. 2021-10-25 IUD (Other): No new information. 2021-10-20, 2021-10-21, 2021-10-25, 2021-10-26, 2021-10-27 emails (Rep, Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Oct-12 Salesforce C00236114 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure to treat degenerative discs. It was reported that during registration there was difficulty achieving a passing registration. Multiple attempts were made, and the patient was eventually registered via scan-and-plan workflow. After the case the manufacturer representative noticed that the red dot for registration that she had placed in planning on her laptop had moved when the export was transferred to the guidance system. She had placed the dot at the base of the vertebra but when she checked after the procedure it showed next to the spinous process. There was a delay of about an hour and 20 minutes. There was no impact on the patient outcome. 2021-10-15 email 2021-10-15 10:16, image 2021-10-15 10:16 x2, email 2021-10-18 06:46, email 2021-10-19 09:32, email 2021-10-20 08:16, email 2021-10-20 08:32 x2, email 2021-10-20 08:46, email 2021-10-20 09:48, CCR, IUD (Rep): Additional information received from a manufacturer representative reported that the surgical system was mounted using a bone mount bridge connected to a Schanz screw placed in the left PSIS. Everything was normal in the case until placing the red dot during registration. Since the case was an SI fusion, L5 and S1 were segmented. There were issues with registration being denied even after getting new AP and oblique images, moving the red dot around, and trying semi-auto and manual registration a few times. When reviewing the oblique CT and fluoro shots, the red dot was found to have moved on L5. The representative pulled up the plan on their laptop and found the red dot was correctly placed and the issue occurred during the transfer from the laptop to the guidance system. The representative tested the patient's export and another patient. The representative moved the red dot tot eh correct spot again for the patient's export, but when going into a demo case the red dot moved again. The representative moved the dot to the correct spot and continued with a 3Define scan. When going into registration after acquiring images, the red dot moved again and was to the left of the spinous process. When testing the export from a different patient, there were no issues. 2021-10-25 IUD (Other): No new information. 2021-10-20, 2021-10-21, 2021-10-25, 2021-10-26, 2021-10-27 emails (Rep, Other): No new information. | xse registration issuessw defect salesforce spinal p rocedure treat degenerative discs registration difficulty achieving passing registration multiple attempts eventually registered scanandplan workflow case noticed red dot registration placed planning laptop moved export transferred placed dot base vertebra checked procedure showed spinous process delay hour minutes impact outcome email image email email email email email email ccr iud additional surgical mounted using bone mount bridge connected schanz screw placed left psis normal case placing red dot registration case si fusion segmented issues registration denied getting new ap oblique images moving red dot trying semiauto manual registration times reviewing oblique ct fluoro shots red dot moved pulled plan laptop red dot correctly placed issue occurred transfer laptop tested patients export moved red dot tot eh correct spot patients export going demo case red dot moved moved dot correct spot continued scan going registration acquiring images red dot moved left spinous process testing export different issues iud new emails new mazorx sw installation kit | NAV3242 | NAV3246 | NAV4112 | NAV3243 | NAV4101 | NAV3242NAV3246NAV4112NAV3243NAV4101 | False
3347 | 703320347 | XSE - Stripped Screw Hole on Arm | 2019-08-20T19:11:38.000000+0000 | MazorX Stealth Edition | NAV4100 | STRIPPED THREADS | 20 | ASM0206-01 | SURGICAL ARM | XSE - Stripped Screw Hole on Arm 2019-08-20 SCL1902419, PDF, 00672511, ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the hole on the shoulder of the surgical system where the robotic reference frame is attached was stripped and pushed into the shoulder. The representative was unable to get the robotic reference frame to connect to the shoulder so the spine reference frame was used for the procedure. There was no patien tharm and the procedure was not delayed more than an hour. 2019-08-27 IUD (Other): No new information. SURGICAL ARM | 2019-08-20 SCL1902419, PDF, 00672511, ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the hole on the shoulder of the surgical system where the robotic reference frame is attached was stripped and pushed into the shoulder. The representative was unable to get the robotic reference frame to connect to the shoulder so the spine reference frame was used for the procedure. There was no patien tharm and the procedure was not delayed more than an hour. 2019-08-27 IUD (Other): No new information. | xse stripped screw hole arm pdf id spinal procedure hole shoulder surgical robotic reference frame attached stripped pushed shoulder unable robotic reference frame connect shoulder spine reference frame procedure patien tharm procedure delayed hour iud new surgical arm | NAV4100 | NAV3299 | NAV3252 | NAV4107 | NAV3249 | NAV4100NAV3299NAV3252NAV4107NAV3249 | True
10376 | 704867370 | XSE - NAV Camera Unable to Communicate | 2022-02-16T17:03:57.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - NAV Camera Unable to Communicate 2022-Feb-16 Salesforce C00243754, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the guidance system was unable to communicate with the navigation camera during preventative maintenance. There was no patient involvement. 2022-02-23 IUD (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2022-Feb-16 Salesforce C00243754, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the guidance system was unable to communicate with the navigation camera during preventative maintenance. There was no patient involvement. 2022-02-23 IUD (Other): No new information. | xse nav camera unable communicate salesforce iud outside proc edure unable communicate navigation camera preventative maintenance iud new mazorx sw installation kit | NAV3246 | NAV2056 | NAV3244 | NAV3247 | NAV3235 | NAV3246NAV2056NAV3244NAV3247NAV3235 | True
7636 | 704241937 | XSE - Defective Instruments | 2021-02-22T17:05:37.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | 9734924 | FRAME 9734924 ROBOTIC REFERENCE | XSE - Defective Instruments 2021-02-22 00706991 (Rep, FOR): Medtronic received information regarding guidance system accessories, outside of procedure. It was reported that two instruments were defective. No patient was present. 2021-03-17 email 2021-03-17 00:15 x2 (Other, For): No new information. 2021-05-10 email 2021-05-10 04:15 (Rep, For): No new information. 2021-05-14 email 2021-05-14 04:01, IUD (Other, For): No new information. FRAME 9734924 ROBOTIC REFERENCE | 2021-02-22 00706991 (Rep, FOR): Medtronic received information regarding guidance system accessories, outside of procedure. It was reported that two instruments were defective. No patient was present. 2021-03-17 email 2021-03-17 00:15 x2 (Other, For): No new information. 2021-05-10 email 2021-05-10 04:15 (Rep, For): No new information. 2021-05-14 email 2021-05-14 04:01, IUD (Other, For): No new information. | xse defective instruments accessories outside procedure instruments defective present email new email new email iud new frame robotic reference | NAV3299 | NAV2054 | NAV2053 | NAV3249 | NAV2044 | NAV3299NAV2054NAV2053NAV3249NAV2044 | True
926 | 703180873 | XSE - arm didn't reach location/image ad | 2019-05-30T17:14:50.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | XSE - arm didn't reach location/image ad 2019-May-23 SCL1901330 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the arm would not reach predefined location. They cleared up and sent to set OB and it reached. When they sent the arm to screw trajectories, the arm would shake when moving in the Z axis. It was also noted that the C-Arm image adaptor was not recognized. The image looked fine and was calibrated properly. They changed the image adaptors with another system still saw the issue. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-May-23 SCL1901330 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the arm would not reach predefined location. They cleared up and sent to set OB and it reached. When they sent the arm to screw trajectories, the arm would shake when moving in the Z axis. It was also noted that the C-Arm image adaptor was not recognized. The image looked fine and was calibrated properly. They changed the image adaptors with another system still saw the issue. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. | xse arm didnt reach locationimage ad outside procedure arm reach predefined location cleared sent set ob reached sent arm screw trajectories arm shake moving z axis noted carm image adaptor recognized image looked fine calibrated properly changed image adaptors saw issue additional complications anticipated surgical arm | NAV4095 | NAV3258 | NAV4101 | NAV3299 | NAV2054 | NAV4095NAV3258NAV4101NAV3299NAV2054 | True
2598 | 703244325 | X - 3Define Functionality issue | 2019-07-05T21:54:15.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define Functionality issue It was reported that the 3Define camera not working, called Leafar and verified system is not recognizing device. UNK MAZOR X 3DEFINE CAMERA | It was reported that the 3Define camera not working, called Leafar and verified system is not recognizing device. | x functionality issue camera working called leafar verified recognizing unk mazor x camera | NAV3244 | NAV4110 | NAV2054 | NAV3261 | NAV3295 | NAV3244NAV4110NAV2054NAV3261NAV3295 | True
3924 | 703467814 | XSE - OOBF Arm Failed Stress Test | 2019-11-14T16:09:02.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - OOBF Arm Failed Stress Test 2019-11-14 00678356 (Rep, HCP): Medtronic received information regarding a guidance system being set up after a replacement surgical arm was received. It was reported that the arm failed the stress test. There was no patient involvement. 2019-11-26 IUD x3 (Other): No new information. SURGICAL ARM | 2019-11-14 00678356 (Rep, HCP): Medtronic received information regarding a guidance system being set up after a replacement surgical arm was received. It was reported that the arm failed the stress test. There was no patient involvement. 2019-11-26 IUD x3 (Other): No new information. | xse oobf arm failed stress test set replacement surgical arm arm failed stress test iud new surgical arm | NAV2054 | NAV3249 | NAV4095 | NAV4092 | NAV4101 | NAV2054NAV3249NAV4095NAV4092NAV4101 | True
3340 | 703319959 | XSE - Broken Screw/Loose Screwdriver | 2019-08-20T16:30:38.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | MAS2188-02 | BM BRIDGE | XSE - Broken Screw/Loose Screwdriver 2019-08-14 SCL1902382 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the screw for the bone mount bridge was no longer working and the shaft on the screw driver had loosened from the handle. There was no patient involvement. 2019-08-29 email (Rep): Additional information received from a manufacturer representative reported the event occurred intra-op. BM BRIDGE | 2019-08-14 SCL1902382 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the screw for the bone mount bridge was no longer working and the shaft on the screw driver had loosened from the handle. There was no patient involvement. 2019-08-29 email (Rep): Additional information received from a manufacturer representative reported the event occurred intra-op. | xse broken screwloose screwdriver prior procedure screw bone mount bridge longer working shaft screw driver loosened handle email additional occurred intraop bm bridge | NAV3299 | NAV3295 | NAV2054 | NAV3298 | NAV3251 | NAV3299NAV3295NAV2054NAV3298NAV3251 | True
2490 | 703236587 | XSE - deviaiton/arm froze | 2019-07-01T18:11:31.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - deviaiton/arm froze 2019-Jun-26 SCL1901865 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that after taking the oblique images, the Arm was frozen in the oblique position and refused to move. They tried sending to L2 left and the arm wouldn’t move. They did a soft restart and system said that it wasn’t within the work volume. The issue was resolved when the rep went into screen to adjust the 3d marker with manual movements and moved the arm towards the head a few times and then moved it high up and then they were able to clear arm up and proceed with the operation. They checked accuracy with both the navigation instruments and placing the dilator on bone and taking a lateral shot. Once the surgeon was happy with the confirmation shots, they proceeded with the operation. After confirmation shots at the end, they noticed L5 left was lateral and surgeon was a little heavy handed with instrumenting and they had to keep reminding him to be a little le ss aggressive. It was also noted that the patient had a BMI greater than 40 and the surgeon also said it could have been soft tissue pressure. The case was done open and it was their first case. The surgeon skipped L5 Left and pulled the screw out. L4 left and S1 Left were accurate. Time to troubleshoot took about 10 minutes. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jun-26 SCL1901865 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that after taking the oblique images, the Arm was frozen in the oblique position and refused to move. They tried sending to L2 left and the arm wouldn’t move. They did a soft restart and system said that it wasn’t within the work volume. The issue was resolved when the rep went into screen to adjust the 3d marker with manual movements and moved the arm towards the head a few times and then moved it high up and then they were able to clear arm up and proceed with the operation. They checked accuracy with both the navigation instruments and placing the dilator on bone and taking a lateral shot. Once the surgeon was happy with the confirmation shots, they proceeded with the operation. After confirmation shots at the end, they noticed L5 left was lateral and surgeon was a little heavy handed with instrumenting and they had to keep reminding him to be a little le ss aggressive. It was also noted that the patient had a BMI greater than 40 and the surgeon also said it could have been soft tissue pressure. The case was done open and it was their first case. The surgeon skipped L5 Left and pulled the screw out. L4 left and S1 Left were accurate. Time to troubleshoot took about 10 minutes. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse deviaitonarm froze spinal procedure taking oblique images arm frozen oblique position refused tried sending left arm wouldn’t did soft restart said wasn’t work volume issue resolved went screen adjust marker manual movements moved arm head times moved high able clear arm proceed operation checked accuracy navigation instruments placing dilator bone taking lateral shot surgeon happy confirmation shots proceeded operation confirmation shots end noticed left lateral surgeon little heavy handed instrumenting reminding little le ss aggressive noted bmi greater surgeon said soft tissue pressure case open case surgeon skipped left pulled screw left left accurate time troubleshoot took minutes additional complications anticipated mazor x software installation kit | NAV3246 | NAV5000 | NAV4095 | NAV3248 | NAV4101 | NAV3246NAV5000NAV4095NAV3248NAV4101 | True
8855 | 704499121 | XSE ¦ Damaged Stealth Camera Cables | 2021-07-13T20:06:28.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 30 | HRS0350-01S | KIT SVC HRS0350-01S NAV CONNECT CABLE | XSE ¦ Damaged Stealth Camera Cables 2021-Jul-13 Salesforce C00230109 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the internal yellow cable and black external cable were damaged. It was also reported that the surgeon screen was unresponsive. No patient was present. 2021-07-14 IUD x3 (Other): No new information. KIT SVC HRS0350-01S NAV CONNECT CABLE | 2021-Jul-13 Salesforce C00230109 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the internal yellow cable and black external cable were damaged. It was also reported that the surgeon screen was unresponsive. No patient was present. 2021-07-14 IUD x3 (Other): No new information. | xse damaged stealth camera cables salesforce outside procedure internal yellow cable black external cable damaged surgeon screen unresponsive present iud new kit svc nav connect cable | NAV3256 | NAV3253 | NAV2053 | NAV3248 | NAV3249 | NAV3256NAV3253NAV2053NAV3248NAV3249 | True
431 | 703108965 | X - Error Messages | 2019-04-17T00:19:16.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2310-03 | TARGET EXTENDER | X - Error Messages 2019-04-10 SCL1900925, SCL1900926, SCL1900927, SCL1900928, Intake Form, CCR, e1, e2, e3, Image 1, Image 2, Vid (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat global deformity with fixation. It was reported that while exposing the anatomy with the surgical system already docked to the bed, two error messages were displayed. A "Error 4059: Arm Shift Detected in Joint 6. Maintenance Action is Required. Stop Operation" was displayed first followed by "Error 295: Fatal Error. CANBUS Connection Lost to Nodes #7 Positioner IO Cart IO." The manufacturer representative pressed the Ok button on the screen, but the error messages reappeared. With the help of technical services, a hard restart of the system was done. When booting the system, the shoulder lock LED was blinking when the should was locked. They had to enter Utilities and press to open the shoulder. It was also difficult to tighten the 3D-marker into the target extender, and as well on the way out it was difficult to release it. They had cross view errors during Fluoro segmentation: they acquired the AP and OBL images, marked the red point and labeled the vertebra but the registration had failed. When putting the red dot on one of the clamp’s teeth on the OBL, they noticed that the red dot didn't fit to the teeth’s location on the AP. They had to acquire new AP/OBL/both images. There was no patient harm and the procedure was not delayed more than an hour. The case was completed with no changes in the course of the procedure. -###-From PE-703109759-2019-04-10 SCL1900927, Intake Form, e1, e2, e3, Vid (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that it was difficult to tighten the 3D marker into the target extender and it was difficult to remove the 3D marker. There was no patient harm and the procedure was not delayed more than an hour. The case was completed with no changes in the course of the procedure. 2019-May-06 SCL1901096 intake form, email 5/8/2019 x4 (rep): Additional information received from the manufacturer representative (rep) indicated that the complaint started after replacing the marker holder recently. They thought there was a tolerance issue with it. The defective one would be returned for analysis. TARGET EXTENDER | 2019-04-10 SCL1900925, SCL1900926, SCL1900927, SCL1900928, Intake Form, CCR, e1, e2, e3, Image 1, Image 2, Vid (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat global deformity with fixation. It was reported that while exposing the anatomy with the surgical system already docked to the bed, two error messages were displayed. A "Error 4059: Arm Shift Detected in Joint 6. Maintenance Action is Required. Stop Operation" was displayed first followed by "Error 295: Fatal Error. CANBUS Connection Lost to Nodes #7 Positioner IO Cart IO." The manufacturer representative pressed the Ok button on the screen, but the error messages reappeared. With the help of technical services, a hard restart of the system was done. When booting the system, the shoulder lock LED was blinking when the should was locked. They had to enter Utilities and press to open the shoulder. It was also difficult to tighten the 3D-marker into the target extender, and as well on the way out it was difficult to release it. They had cross view errors during Fluoro segmentation: they acquired the AP and OBL images, marked the red point and labeled the vertebra but the registration had failed. When putting the red dot on one of the clamp’s teeth on the OBL, they noticed that the red dot didn't fit to the teeth’s location on the AP. They had to acquire new AP/OBL/both images. There was no patient harm and the procedure was not delayed more than an hour. The case was completed with no changes in the course of the procedure. -###-From PE-703109759-2019-04-10 SCL1900927, Intake Form, e1, e2, e3, Vid (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that it was difficult to tighten the 3D marker into the target extender and it was difficult to remove the 3D marker. There was no patient harm and the procedure was not delayed more than an hour. The case was completed with no changes in the course of the procedure. 2019-May-06 SCL1901096 intake form, email 5/8/2019 x4 (rep): Additional information received from the manufacturer representative (rep) indicated that the complaint started after replacing the marker holder recently. They thought there was a tolerance issue with it. The defective one would be returned for analysis. | x error messages ccr image image vid spinal procedure treat global deformity fixation exposing anatomy surgical docked bed error messages displayed error arm shift detected joint maintenance action required stop operation displayed followed error fatal error canbus connection lost nodes positioner io cart io pressed ok button screen error messages reappeared help technical services hard restart booting shoulder lock led blinking locked enter utilities press open shoulder difficult tighten target extender way difficult release cross view errors fluoro segmentation acquired ap obl images marked red point labeled vertebra registration failed putting red dot clamp’s teeth obl noticed red dot didnt fit teeth’s location ap acquire new apoblboth images procedure delayed hour case completed changes course procedure vid spinal procedure difficult tighten marker target extender difficult remove marker procedure delayed hour case completed changes course procedure email additional indicated complaint started replacing marker holder recently thought tolerance issue defective returned analysis target extender | NAV2047 | NAV4103 | NAV3233 | NAV4107 | NAV3246 | NAV2047NAV4103NAV3233NAV4107NAV3246 | False
6905 | 704061632 | XSE - Stripped BMB Screws | 2020-11-10T18:47:01.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2188-04S | KIT SVC MAS2188-04S BM BRIDGE | XSE - Stripped BMB Screws 2020-11-10 00700577 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the bone mount bridge was stripped out. No patient was present. 2020-Dec-15Nov-23 IUD (other): No new information. 2020-12-10 email 2020-12-10 09:45, IUD (Rep): Additional information received from a manufacturer representative reported that the screws of the bone mount bridge were replaced. The surgeon insisted on overtightening the screw so they run into the issue with the square design. The screws were removed and discarded. KIT SVC MAS2188-04S BM BRIDGE | 2020-11-10 00700577 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the bone mount bridge was stripped out. No patient was present. 2020-Dec-15Nov-23 IUD (other): No new information. 2020-12-10 email 2020-12-10 09:45, IUD (Rep): Additional information received from a manufacturer representative reported that the screws of the bone mount bridge were replaced. The surgeon insisted on overtightening the screw so they run into the issue with the square design. The screws were removed and discarded. | xse stripped bmb screws accessory outside procedure bone mount bridge stripped present iud new email iud additional screws bone mount bridge replaced surgeon insisted overtightening screw run issue square design screws removed discarded kit svc bm bridge | NAV3299 | NAV4090 | NAV2058 | NAV3250 | NAV3295 | NAV3299NAV4090NAV2058NAV3250NAV3295 | True
15 | 702971083 | X - Comm Issue with Surgical Sys.-Case 2 | 2019-01-23T10:38:05.000000+0000 | MazorX | NAV3247 | EXITS | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Comm Issue with Surgical Sys.-Case 2 2019-01-21 SC SCL1900205, PDF 2019-01-23 05:42 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the manufacturer representative noticed the light on the shoulder was blinking; however, the shoulder was not unlocked. The manufacturer representative thought there was a miscommunication between the encoders and workstation. Sometimes the arm would lock and sometimes it would not. The light continually blinked regardless of the lock status. The system maintained performance and passed all tests so the procedure was continued. After moving the system into position, the arm was locked and the surgeon began instrumenting. L4 and L5 were marked with a pen and not cut per the surgeon's routine. While sending the arm to the L4 right trajectory, the system "booted them out" to the beginning patient folder screen indicating that the shoulder was no longer locked. The ma nufacturer representative did a soft restart, remounted the robot and registered again. After registration, the arm was sent to L5 left and the new trajectory was approximately 20 mm off of the original mark. The surgeon decided to proceed and used confirmation shots throughout the rest of the case. The surgery was completed successfully. There was no delay in the procedure over an hour. No patient harm was reported. Information unrelated to this event was included in separate PEs. See PE 702971082 for Case 1, 702971084 for X-Eye Issues, 702971085 for Arm Vibration, and 702971376 for Schanz Arm Malfunction. 2019-02-06 e1 (Rep): No new information. 2019-02-26 e2 (Rep): No new information. UNK MAZOR X SHOULDER | 2019-01-21 SC SCL1900205, PDF 2019-01-23 05:42 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the manufacturer representative noticed the light on the shoulder was blinking; however, the shoulder was not unlocked. The manufacturer representative thought there was a miscommunication between the encoders and workstation. Sometimes the arm would lock and sometimes it would not. The light continually blinked regardless of the lock status. The system maintained performance and passed all tests so the procedure was continued. After moving the system into position, the arm was locked and the surgeon began instrumenting. L4 and L5 were marked with a pen and not cut per the surgeon's routine. While sending the arm to the L4 right trajectory, the system "booted them out" to the beginning patient folder screen indicating that the shoulder was no longer locked. The ma nufacturer representative did a soft restart, remounted the robot and registered again. After registration, the arm was sent to L5 left and the new trajectory was approximately 20 mm off of the original mark. The surgeon decided to proceed and used confirmation shots throughout the rest of the case. The surgery was completed successfully. There was no delay in the procedure over an hour. No patient harm was reported. Information unrelated to this event was included in separate PEs. See PE 702971082 for Case 1, 702971084 for X-Eye Issues, 702971085 for Arm Vibration, and 702971376 for Schanz Arm Malfunction. 2019-02-06 e1 (Rep): No new information. 2019-02-26 e2 (Rep): No new information. | x comm issue surgical syscase sc pdf noticed light shoulder blinking shoulder unlocked thought miscommunication encoders workstation arm lock light continually blinked regardless lock status maintained performance passed tests procedure continued moving position arm locked surgeon began instrumenting marked pen cut surgeons routine sending arm right trajectory booted beginning folder screen indicating shoulder longer locked ma nufacturer did soft restart remounted robot registered registration arm sent left new trajectory approximately mm original mark surgeon decided proceed confirmation shots rest case surgery completed successfully delay procedure hour unrelated included separate pes pe case xeye issues arm vibration schanz arm malfunction new new unk mazor x shoulder | NAV4107 | NAV4101 | NAV2047 | NAV4108 | NAV3261 | NAV4107NAV4101NAV2047NAV4108NAV3261 | False
4029 | 703491378 | XSE - Trajectory Deviation L5 Superior | 2019-11-28T03:18:03.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 30 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XSE - Trajectory Deviation L5 Superior 2019-11-27 00679247, ID (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure for L4-S1 fusion. It was reported that when the system was started up the surgeon tablet screen was black. Also the workstation monitor was blinking approximately every 30 second and then went dim. At that point the manufacturer representative did a hard reboot. Both issues resolved and the procedure moved forward. After placing screws, one of the L5 screws was noted to be superior to plan. The representative believed the deviation was due to skiving. The surgeon revised the screw using fluoro. No patient harm was reported and surgery was delayed less than an hour. 2019-12-09 IUD x5 (Rep): No new information. 2020-01-06 IUD x4 (Other): No new information. UNK MAZOR X MONITOR | 2019-11-27 00679247, ID (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure for L4-S1 fusion. It was reported that when the system was started up the surgeon tablet screen was black. Also the workstation monitor was blinking approximately every 30 second and then went dim. At that point the manufacturer representative did a hard reboot. Both issues resolved and the procedure moved forward. After placing screws, one of the L5 screws was noted to be superior to plan. The representative believed the deviation was due to skiving. The surgeon revised the screw using fluoro. No patient harm was reported and surgery was delayed less than an hour. 2019-12-09 IUD x5 (Rep): No new information. 2020-01-06 IUD x4 (Other): No new information. | xse trajectory deviation superior id spinal procedure fusion started surgeon tablet screen black workstation monitor blinking approximately second went dim point did hard reboot issues resolved procedure moved forward placing screws screws noted superior plan believed deviation skiving surgeon revised screw using fluoro surgery delayed hour iud new iud new unk mazor x monitor | NAV3261 | NAV3253 | NAV5000 | NAV2047 | NAV4101 | NAV3261NAV3253NAV5000NAV2047NAV4101 | True
9771 | 704718640 | XSE - Software Anomalies | 2021-11-16T20:10:27.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 20 | KIT0830 | USA REGIONAL KIT KIT0830 | XSE - Software Anomalies 2021-Nov-16 Salesforce C00238234 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a case there were several issues with the software. The system software appeared to get dim early on in the case. There were issues with the system going unresponsive several times during the planning phase. The tablet was als o very dim but still functioning. There was no impact on the patient outcome. 2021-11-24 email 2021-11-24 15:16 (Rep): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-02 email 2021-12-02 12:15 (Rep): No new information. 2021-12-09 IUD (Other): No new information. 2021-12-16 IUD (Other): No new information. USA REGIONAL KIT KIT0830 | 2021-Nov-16 Salesforce C00238234 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during a case there were several issues with the software. The system software appeared to get dim early on in the case. There were issues with the system going unresponsive several times during the planning phase. The tablet was als o very dim but still functioning. There was no impact on the patient outcome. 2021-11-24 email 2021-11-24 15:16 (Rep): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-02 email 2021-12-02 12:15 (Rep): No new information. 2021-12-09 IUD (Other): No new information. 2021-12-16 IUD (Other): No new information. | xse software anomalies salesforce spinal p rocedure case issues software software appeared dim early case issues going unresponsive times planning phase tablet als o dim functioning impact outcome email new iud new email new iud new iud new usa regional kit | NAV3246 | NAV3248 | NAV3299 | NAV3249 | NAV3253 | NAV3246NAV3248NAV3299NAV3249NAV3253 | True
6318 | 703910175 | XSE - Crossview Error | 2020-08-18T21:38:37.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Crossview Error 2020-08-18 00694828 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a crossview error was displayed while registering T7. The system showed the incorrect level and was labeling the vertebrae as T6. On the AP image, the vertebrae was actually T6 and could not be registered. Another AP image was taken and registration was completed successfully. There was no patient harm and the procedure was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2020-08-18 00694828 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a crossview error was displayed while registering T7. The system showed the incorrect level and was labeling the vertebrae as T6. On the AP image, the vertebrae was actually T6 and could not be registered. Another AP image was taken and registration was completed successfully. There was no patient harm and the procedure was delayed less than an hour. | xse crossview error spinal procedure crossview error displayed registering showed incorrect level labeling vertebrae ap image vertebrae actually registered ap image taken registration completed successfully procedure delayed hour mazor x | NAV3242 | NAV4112 | NAV5000 | NAV2047 | NAV2054 | NAV3242NAV4112NAV5000NAV2047NAV2054 | True
2225 | 703230639 | X- Surgical Arm Dropped | 2019-06-27T12:11:33.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01 | SURGICAL ARM | X- Surgical Arm Dropped Coliseum Northside Hospital (HCA) Dr. Holliday L2-S1 PSF System #48, Arm #53. CSR and surgeon planned case prior to surgery. Special attention was given to mitigate soft tissue pressure. CSR explained to surgeon that patience is required during image acquisition in order to ensure all bodies are captured and to keep from needing two registrations. Surgeon confirmed that he understood. Surgeon said he would prefer 1 registration and would freehand S1 if need be. CSR informed surgeon that he felt confident that he could register L2-S1, but was sure that L2-L5 would be captured in 1 image. Prior to surgery, touch screen was not working. Soft restart did not fix. CSR performed hard restart to alleviate the issue. No further issues with touchscreen. 8:40-X mounted to bed (2 people/1min) 9:05-Draping of X (3 people/2min) 9:10-Skin Incision 9:55-Dual Clamp mounted on L3-L4, BMU at 15 degrees 9:59-First 3Define scan failed immediately. The scan was attempted again and timed out during the scan. On the third attempt, a successful 3Define scan was completed. 10:03-Successful Drawspine, L3 selected 10:04-3D marker sent to AP. AP position was not optimal. Upon using virtual keyboard to move the 3D marker right to center over the patient, the arm dropped. CSR then sent the arm to clear up and resent the AP position using the Generic trajectory. Virtual keyboard was used again and 2 images were needed. 10:10-3D marker sent to Generic position. CSR lowered the marker toward the patient and 6 images were required. CSR gained registration on the first attempt on Basic/Auto by labeling L4. Surgeon confirmed matching.10:22-Left L2 10:24-Left L3 10:26-Left L4 10:28-Left L5 10:30-Left S1. 10:32-Right L2 10:37-Right L3 10:39-Right L4 10:41-Right L5 10:42-Right S1. Verification images were taken to confirm the placement of each trajectory. Surgeon was very pleased about the results of the kwire placement. Surgeon commented that the X system was especially helpful on longer constructs. SURGICAL ARM | Coliseum Northside Hospital (HCA) Dr. Holliday L2-S1 PSF System #48, Arm #53. CSR and surgeon planned case prior to surgery. Special attention was given to mitigate soft tissue pressure. CSR explained to surgeon that patience is required during image acquisition in order to ensure all bodies are captured and to keep from needing two registrations. Surgeon confirmed that he understood. Surgeon said he would prefer 1 registration and would freehand S1 if need be. CSR informed surgeon that he felt confident that he could register L2-S1, but was sure that L2-L5 would be captured in 1 image. Prior to surgery, touch screen was not working. Soft restart did not fix. CSR performed hard restart to alleviate the issue. No further issues with touchscreen. 8:40-X mounted to bed (2 people/1min) 9:05-Draping of X (3 people/2min) 9:10-Skin Incision 9:55-Dual Clamp mounted on L3-L4, BMU at 15 degrees 9:59-First 3Define scan failed immediately. The scan was attempted again and timed out during the scan. On the third attempt, a successful 3Define scan was completed. 10:03-Successful Drawspine, L3 selected 10:04-3D marker sent to AP. AP position was not optimal. Upon using virtual keyboard to move the 3D marker right to center over the patient, the arm dropped. CSR then sent the arm to clear up and resent the AP position using the Generic trajectory. Virtual keyboard was used again and 2 images were needed. 10:10-3D marker sent to Generic position. CSR lowered the marker toward the patient and 6 images were required. CSR gained registration on the first attempt on Basic/Auto by labeling L4. Surgeon confirmed matching.10:22-Left L2 10:24-Left L3 10:26-Left L4 10:28-Left L5 10:30-Left S1. 10:32-Right L2 10:37-Right L3 10:39-Right L4 10:41-Right L5 10:42-Right S1. Verification images were taken to confirm the placement of each trajectory. Surgeon was very pleased about the results of the kwire placement. Surgeon commented that the X system was especially helpful on longer constructs. | x surgical arm dropped coliseum northside hospital dr holliday psf arm csr surgeon planned case prior surgery special attention given mitigate soft tissue pressure csr explained surgeon patience required image acquisition order ensure bodies captured needing registrations surgeon confirmed understood surgeon said prefer registration freehand need csr informed surgeon felt confident register sure captured image prior surgery touch screen working soft restart did fix csr performed hard restart alleviate issue issues touchscreen mounted bed x incision clamp mounted bmu degrees scan failed immediately scan attempted timed scan attempt successful scan completed drawspine selected marker sent ap ap position optimal using virtual keyboard marker right center arm dropped csr sent arm clear resent ap position using generic trajectory virtual keyboard images needed marker sent generic position csr lowered marker images required csr gained registration attempt basicauto labeling surgeon confirmed verification images taken confirm placement trajectory surgeon pleased results kwire placement surgeon commented x especially helpful longer constructs surgical arm | NAV4101 | NAV3244 | NAV3242 | NAV3253 | NAV3249 | NAV4101NAV3244NAV3242NAV3253NAV3249 | True
318 | 703078082 | X - Bed frame adapter screws broken | 2019-03-28T13:56:46.000000+0000 | MazorX | NAV4090 | BROKEN SCREW | 20 | MAS1574-01 | BED FRAME ADAPTOR | X - Bed frame adapter screws broken 2019-Mar-26 SCL1900739 intake form, image1, MOV (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the bed frame adaptor screws were damaged. The back screw was broken and the front two did not tighten. It was noted there was no patient involvement. There were no additional complications reported or anticipated as a result of the event. BED FRAME ADAPTOR | 2019-Mar-26 SCL1900739 intake form, image1, MOV (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the bed frame adaptor screws were damaged. The back screw was broken and the front two did not tighten. It was noted there was no patient involvement. There were no additional complications reported or anticipated as a result of the event. | x bed frame adapter screws broken mov spinal procedure bed frame adaptor screws damaged screw broken did tighten noted additional complications anticipated bed frame adaptor | NAV3255 | NAV3295 | NAV4090 | NAV3299 | NAV2054 | NAV3255NAV3295NAV4090NAV3299NAV2054 | True
5754 | 703795523 | X - Unable to Get System Off Bed | 2020-06-10T17:14:35.000000+0000 | MazorX | NAV4104 | UNABLE TO MOVE | 30 | HRS0327-01S | KIT SVC HRS0327-01S TEETH LOCKING SENSOR | X - Unable to Get System Off Bed 2020-06-10 IUD x2 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner would not move. The manufacturer representative tried to manually open the teeth in the override panel, but the issue was not resolved. A hard restart of the guidance system was done and the system utilities were checked. The right tooth was red, but it was engaged on the bed frame adapter. The representative tried to use the override panel again, but was unsuccessful. The hood of the workstation was open and the manipulator was placed on the back of the positioner. The system was removed from the bed by removing the bed frame adapter. There was no patient involvement. 2020-06-18 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-07-01 IUD (Other): No new information. 2020-07-14 IUD (Other): No new information. 2020-07-23 IUD (Other): No new information. KIT SVC HRS0327-01S TEETH LOCKING SENSOR | 2020-06-10 IUD x2 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner would not move. The manufacturer representative tried to manually open the teeth in the override panel, but the issue was not resolved. A hard restart of the guidance system was done and the system utilities were checked. The right tooth was red, but it was engaged on the bed frame adapter. The representative tried to use the override panel again, but was unsuccessful. The hood of the workstation was open and the manipulator was placed on the back of the positioner. The system was removed from the bed by removing the bed frame adapter. There was no patient involvement. 2020-06-18 IUD (Other): No new information. 2020-06-23 IUD (Other): No new information. 2020-07-01 IUD (Other): No new information. 2020-07-14 IUD (Other): No new information. 2020-07-23 IUD (Other): No new information. | x unable bed iud outside procedure positioner tried manually open teeth override panel issue resolved hard restart utilities checked right tooth red engaged bed frame adapter tried use override panel unsuccessful hood workstation open manipulator placed positioner removed bed removing bed frame adapter iud new iud new iud new iud new iud new kit svc teeth locking sensor | NAV3250 | NAV2058 | NAV4103 | NAV2047 | NAV3255 | NAV3250NAV2058NAV4103NAV2047NAV3255 | False
4215 | 703515472 | X - Software crash | 2019-12-12T15:23:21.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - Software crash 2018-01-10 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was planned prior to surgery for four vertebrae on levels T4 - T9 with open fusion. The patient had traumatic compression fracture at T7 with paralysis. CT scan was done when the patient was admitted but had to be redone because it had been done without following proper protocol. The bed mount was placed to far inferior and had to be adjusted intra-operatively. During the procedure, software crashed roughly eight times during exposure rebooting back to login menu each time. Surgeon was pleased with trajectories. No patient harm was reported. MAZOR X SOFTWARE KIT VER-2.3.1 | 2018-01-10 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was planned prior to surgery for four vertebrae on levels T4 - T9 with open fusion. The patient had traumatic compression fracture at T7 with paralysis. CT scan was done when the patient was admitted but had to be redone because it had been done without following proper protocol. The bed mount was placed to far inferior and had to be adjusted intra-operatively. During the procedure, software crashed roughly eight times during exposure rebooting back to login menu each time. Surgeon was pleased with trajectories. No patient harm was reported. | x software crash ccr spinal procedure case planned prior surgery vertebrae levels open fusion traumatic compression fracture paralysis ct scan admitted redone following proper protocol bed mount placed far inferior adjusted intraoperatively procedure software crashed roughly times exposure rebooting login menu time surgeon pleased trajectories mazor x software kit | NAV3247 | NAV3248 | NAV4091 | NAV3246 | NAV3243 | NAV3247NAV3248NAV4091NAV3246NAV3243 | True
330 | 703078615 | X - Broken Dual Clamps | 2019-03-28T15:18:30.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | MAS1355-03 | DUAL CLAMP | X - Broken Dual Clamps 2019-03-25 SCL1900730, Intake Form (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that both dual clamps cracked when tightening down the screws during the procedure. There was no patient harm and the surgery was not delayed. The surgery was completed with no changes to the course of the procedure. 2019-04-24 e2 (Rep): Additional information received from a manufacturer representative reported that no pictures or videos were taken. No additional troubleshooting was done and the case was completed using a third tray. DUAL CLAMP | 2019-03-25 SCL1900730, Intake Form (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that both dual clamps cracked when tightening down the screws during the procedure. There was no patient harm and the surgery was not delayed. The surgery was completed with no changes to the course of the procedure. 2019-04-24 e2 (Rep): Additional information received from a manufacturer representative reported that no pictures or videos were taken. No additional troubleshooting was done and the case was completed using a third tray. | x broken dual clamps spinal procedure dual clamps cracked tightening screws procedure surgery delayed surgery completed changes course procedure additional pictures videos taken additional troubleshooting case completed using tray dual clamp | NAV3295 | NAV3299 | NAV4101 | NAV3249 | NAV4090 | NAV3295NAV3299NAV4101NAV3249NAV4090 | True
7063 | 704112107 | XSE - 3Define Scan Failed | 2020-12-08T19:03:36.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | KIT0749-01 | KIT SPARE PART FOR THE 3D CAMERA | XSE - 3Define Scan Failed 2020-12-07 00702276 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera failed during a case. The surgical arm went to the first point and then stopped. A hard restart of the system was done. A second 3Define scan was done and the surgical arm got to the third point, but the camera failed and an error message was displayed. A generic work volume was used to continue with the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-09 IUD x5 (Other): No new information. KIT SPARE PART FOR THE 3D CAMERA | 2020-12-07 00702276 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera failed during a case. The surgical arm went to the first point and then stopped. A hard restart of the system was done. A second 3Define scan was done and the surgical arm got to the third point, but the camera failed and an error message was displayed. A generic work volume was used to continue with the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-09 IUD x5 (Other): No new information. | xse scan failed spinal procedure camera failed case surgical arm went point stopped hard restart second scan surgical arm got point camera failed error message displayed generic work volume continue procedure procedure delayed hour iud new kit spare camera | NAV3244 | NAV3246 | NAV4095 | NAV4104 | NAV4093 | NAV3244NAV3246NAV4095NAV4104NAV4093 | True
7388 | 704193982 | XSE - System Not Building Pressure | 2021-01-26T01:36:42.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE - System Not Building Pressure 2021-01-25 00705132 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the system would not build pressure so that the compressor turns off. The compressor was timing out and over time the system would run out of pressure and they were unable to move the shoulder. They had to restart the system to send a new signal to the compressor so they could move the system. There was no impact on the patient outcome. 2021-01-26 email 2021-01-26 16:45, IUD x4 (Rep): Additional information received from a manufacturer representative reported that the pneumatic panel was bad at install. 2021-02-15 IUD x7 (Other): No new information. 2021-03-09 e1 (Other): No new information. 2021-03-21 e2, e3 (Other): No new information. 2021-04-01 e4, e5 (Other): No new information. 2021-04-07 IUD (Other): No new information. 2021-04-12 IUD (Other): No new information. 2021-04-18 e6, e7 (Other): No new information. 2021-04-24 e8 (Other): No new information. 2021-04-25 IUD x2 (Other): No new information. 2021-05-25 e9, e10 (Other): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2021-01-25 00705132 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the system would not build pressure so that the compressor turns off. The compressor was timing out and over time the system would run out of pressure and they were unable to move the shoulder. They had to restart the system to send a new signal to the compressor so they could move the system. There was no impact on the patient outcome. 2021-01-26 email 2021-01-26 16:45, IUD x4 (Rep): Additional information received from a manufacturer representative reported that the pneumatic panel was bad at install. 2021-02-15 IUD x7 (Other): No new information. 2021-03-09 e1 (Other): No new information. 2021-03-21 e2, e3 (Other): No new information. 2021-04-01 e4, e5 (Other): No new information. 2021-04-07 IUD (Other): No new information. 2021-04-12 IUD (Other): No new information. 2021-04-18 e6, e7 (Other): No new information. 2021-04-24 e8 (Other): No new information. 2021-04-25 IUD x2 (Other): No new information. 2021-05-25 e9, e10 (Other): No new information. | xse building pressure intraoperatively spinal procedure build pressure compressor turns compressor timing time run pressure unable shoulder restart send new signal compressor impact outcome email iud additional pneumatic panel bad install iud new new new new iud new iud new new new iud new new kit svc pneumatic panel | NAV3264 | NAV2054 | NAV1006 | NAV3250 | NAV3261 | NAV3264NAV2054NAV1006NAV3250NAV3261 | True
8849 | 704498682 | XSE ¦ Monitors Not Displaying | 2021-07-13T16:06:19.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 50 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE ¦ Monitors Not Displaying 2021-Jul-13 Salesforce C00230089 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the work station monitor and surgeon screen were not displaying any video. Overall, the rest of the system was receiving power (computer powering, pneumatics enabling, work station and hood buttons illuminated). The manufacturer repr esentative tried restarting the system two times with no resolution. Additional information was received stating that reseating the cables had no effect. The manufacturer representative also clarified that initially only the surgeon monitor was out, but when a soft reboot was performed to attempt to resolve that issue both monitors went black and neither came on. The use of navigation was aborted. The issues extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-07-14 IUD x8 (Other): No new information. 2021-08-09 IUD x8 (Other): No new information. 2021-09-02 email 2021-09-02 07:01 (Rep): No new information. MAZOR X PC STATION - REFURBISHED | 2021-Jul-13 Salesforce C00230089 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the work station monitor and surgeon screen were not displaying any video. Overall, the rest of the system was receiving power (computer powering, pneumatics enabling, work station and hood buttons illuminated). The manufacturer repr esentative tried restarting the system two times with no resolution. Additional information was received stating that reseating the cables had no effect. The manufacturer representative also clarified that initially only the surgeon monitor was out, but when a soft reboot was performed to attempt to resolve that issue both monitors went black and neither came on. The use of navigation was aborted. The issues extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-07-14 IUD x8 (Other): No new information. 2021-08-09 IUD x8 (Other): No new information. 2021-09-02 email 2021-09-02 07:01 (Rep): No new information. | xse monitors displaying salesforce spinal p rocedure work station monitor surgeon screen displaying video overall rest receiving power repr esentative tried restarting times resolution additional stating reseating cables effect clarified initially surgeon monitor soft reboot performed attempt resolve issue monitors went black came use navigation aborted issues extended surgical time hour impact outcome iud new iud new email new mazor x pc station refurbished | NAV3262 | NAV3261 | NAV3253 | NAV4112 | NAV5000 | NAV3262NAV3261NAV3253NAV4112NAV5000 | True
1674 | 703208480 | X - Software issue | 2019-06-14T21:19:59.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Software issue Dr. Wang/Dr. Olivero Open PLIF T11-L3 (planned for T10 and L4 but they were not clinically needed). Started up system, completed 3 step setup and attached to bed with no issues. Reviewed plan with Dr.Oliverio and Dr. Wang-approved. Dr. wanted to take AP and Lat shots and requested the images to bezoomed in. Had to recalibrate c-arm adapter plateonce he finished with taking images. Dr. Wang exposed then used the dual clamp to clamp onto T12 and L1. Attached with bmb and then mounted system. Looked at arm and saw thatangle was not ideal, unmounted and had PA bend the arm to be in line with the Spinous process. Re-mounted. Completed 3 define scan successfully, did not use draw spine. Sent for AP-imaging was poor and ribs were not clearlyvisible to anyone in theroom to determine what levels we were at. Took a lateral-determined levelby placing pin at T12-moved back to AP and was able to get all 5 levels in with 3d marker. Moved to obli but had difficultywith hitting the armboards. Was able to get in position and get all 5 levels of T11-L3 in shot with3d marker. Completed registration successfully. Came to send arm screen and saw L2 on the left had a red screw. Clicked on that trajectory and got error message 295 -no solution found for screw. Wasunable to send to that particular trajectory. Proceeded with starting at top and sent arm to T11 first. While surgeon was working on T11, I went intoplanning for L2 on left and double clicked on theword description at bottom left. The system founda trajectory at a 13 degree angle (instead of 11)that it could send to and the screw was no longerred. I played through this new screw trajectory and surgeon approved. Continued sending to all trajectories and this issue caused no delay. Brought in x-ray and all guidewires looked good except L2 on the left. The guidewire was in the correct position but just hadn't been pushed in deep enough and was falling out. Resent the arm to L2 onthe left and guidewire was placed deeper into thevertebral body. Cannulated Solara Medtronic screws were placed over guidewires, confirmed all were correct by x-ray. Removed Mazor from bed with no issues, case completed successfully. MAZOR X SOFTWARE KIT VER-3.0.2 | Dr. Wang/Dr. Olivero Open PLIF T11-L3 (planned for T10 and L4 but they were not clinically needed). Started up system, completed 3 step setup and attached to bed with no issues. Reviewed plan with Dr.Oliverio and Dr. Wang-approved. Dr. wanted to take AP and Lat shots and requested the images to bezoomed in. Had to recalibrate c-arm adapter plateonce he finished with taking images. Dr. Wang exposed then used the dual clamp to clamp onto T12 and L1. Attached with bmb and then mounted system. Looked at arm and saw thatangle was not ideal, unmounted and had PA bend the arm to be in line with the Spinous process. Re-mounted. Completed 3 define scan successfully, did not use draw spine. Sent for AP-imaging was poor and ribs were not clearlyvisible to anyone in theroom to determine what levels we were at. Took a lateral-determined levelby placing pin at T12-moved back to AP and was able to get all 5 levels in with 3d marker. Moved to obli but had difficultywith hitting the armboards. Was able to get in position and get all 5 levels of T11-L3 in shot with3d marker. Completed registration successfully. Came to send arm screen and saw L2 on the left had a red screw. Clicked on that trajectory and got error message 295 -no solution found for screw. Wasunable to send to that particular trajectory. Proceeded with starting at top and sent arm to T11 first. While surgeon was working on T11, I went intoplanning for L2 on left and double clicked on theword description at bottom left. The system founda trajectory at a 13 degree angle (instead of 11)that it could send to and the screw was no longerred. I played through this new screw trajectory and surgeon approved. Continued sending to all trajectories and this issue caused no delay. Brought in x-ray and all guidewires looked good except L2 on the left. The guidewire was in the correct position but just hadn't been pushed in deep enough and was falling out. Resent the arm to L2 onthe left and guidewire was placed deeper into thevertebral body. Cannulated Solara Medtronic screws were placed over guidewires, confirmed all were correct by x-ray. Removed Mazor from bed with no issues, case completed successfully. | x software issue dr wangdr olivero open plif started completed step setup attached bed issues reviewed plan droliverio dr wangapproved dr wanted ap lat shots requested images bezoomed recalibrate carm adapter plateonce finished taking images dr wang exposed dual clamp clamp attached bmb mounted looked arm saw thatangle ideal unmounted pa bend arm line spinous process remounted completed define scan successfully did use draw spine sent apimaging poor ribs clearlyvisible theroom determine levels took lateraldetermined levelby placing pin ap able levels marker moved obli difficultywith hitting armboards able position levels shot marker completed registration successfully came send arm screen saw left red screw clicked trajectory got error message solution screw wasunable send particular trajectory proceeded starting sent arm surgeon working went intoplanning left double clicked theword description left founda trajectory degree angle send screw longerred played new screw trajectory surgeon approved continued sending trajectories issue caused delay brought xray guidewires looked good left guidewire correct position just hadnt pushed deep falling resent arm onthe left guidewire placed deeper thevertebral body cannulated solara screws placed guidewires confirmed correct xray removed mazor bed issues case completed successfully mazor x software kit | NAV3246 | NAV3242 | NAV1012 | NAV3243 | NAV3247 | NAV3246NAV3242NAV1012NAV3243NAV3247 | True
10169 | 704815194 | XSE - Alleged Inaccuracy | 2022-01-18T16:51:37.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Alleged Inaccuracy 2022-Jan-17 Salesforce C00241808, IUD x2, email 2022-01-19 10:32 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during an L5-S1 they had difficulties registering the patient. The surgeon planned alone and then went over it before the case. They were eventually able to register. They went to left L5 and looked accurate. Then when going to S1 on the left the axial looked accurate but sagittal looked inferior 3-5mm to the plan. They took another snapshot and still looked inaccurate. They then re-registered and went back to S1 and were still inferior. The surgeon aborted the use of the guidance system and proceeded with normal navigation and imaging. The cause of the inaccuracy was unknown. There was about a 25 minute delay. There was no impact on the patient outcome. 2022-02-08 IUD (Other): No new information. 2022-02-14 IUD (Other): No new information. 2022-02-25 IUD (Rep): Additional information received from a manufacturer representative reported that the surgical arm passed an ac curacy test and the system had been used successfully in later cases. 2022-03-29 e1, e2 (Other): No new information. 2022-04-04 e3 (Other): No new information. 2022-04-5 e4 (Rep): Additional information received from a manufacturer representative reported that the trajectory was found to be inaccurate after registration. L5 was accurate, S1 was inaccurate. No screws were inserted with the guidance system as the use of the guidance system was aborted after three different registrations. An interbody was placed prior to registration since the procedure was an ALIF. Navigation accuracy was checked and was found to be accurate, so the patient was re-registered. MAZOR X SYSTEM TPL0059 | 2022-Jan-17 Salesforce C00241808, IUD x2, email 2022-01-19 10:32 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during an L5-S1 they had difficulties registering the patient. The surgeon planned alone and then went over it before the case. They were eventually able to register. They went to left L5 and looked accurate. Then when going to S1 on the left the axial looked accurate but sagittal looked inferior 3-5mm to the plan. They took another snapshot and still looked inaccurate. They then re-registered and went back to S1 and were still inferior. The surgeon aborted the use of the guidance system and proceeded with normal navigation and imaging. The cause of the inaccuracy was unknown. There was about a 25 minute delay. There was no impact on the patient outcome. 2022-02-08 IUD (Other): No new information. 2022-02-14 IUD (Other): No new information. 2022-02-25 IUD (Rep): Additional information received from a manufacturer representative reported that the surgical arm passed an ac curacy test and the system had been used successfully in later cases. 2022-03-29 e1, e2 (Other): No new information. 2022-04-04 e3 (Other): No new information. 2022-04-5 e4 (Rep): Additional information received from a manufacturer representative reported that the trajectory was found to be inaccurate after registration. L5 was accurate, S1 was inaccurate. No screws were inserted with the guidance system as the use of the guidance system was aborted after three different registrations. An interbody was placed prior to registration since the procedure was an ALIF. Navigation accuracy was checked and was found to be accurate, so the patient was re-registered. | xse alleged inaccuracy salesforce iud email spinal procedure difficulties registering surgeon planned went case eventually able register went left looked accurate going left axial looked accurate sagittal looked inferior plan took snapshot looked inaccurate reregistered went inferior surgeon aborted use proceeded normal navigation imaging cause inaccuracy unknown minute delay impact outcome iud new iud new iud additional surgical arm passed ac curacy test successfully later cases new new additional trajectory inaccurate registration accurate inaccurate screws inserted use aborted different registrations interbody placed prior registration procedure alif navigation accuracy checked accurate reregistered mazor x | NAV5000 | NAV1012 | NAV5001 | NAV4112 | NAV3243 | NAV5000NAV1012NAV5001NAV4112NAV3243 | False
2911 | 703258886 | XSE - arm froze/error messages | 2019-07-15T16:28:30.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 40 | ASM0206-01 | SURGICAL ARM | XSE - arm froze/error messages 2019-Jul-11 SCL1902018 intake formx2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they did two part registration and finished the first segment with no issues. During the second segment, the rep sent the arm from L1 right to L2 right and received the error 8798 Arm Failed to reach final destination (Driver 4 reports error: overspeed) also displayed (Failed to settle under MaxClosing tolerance). They tried to send the arm to the predefined locations but it wouldn’t reach. The rep performed a soft restart and received error 295 Fatal Error (CANBUS Connection Lost to Nodes #Positioner IO, 8 Cart IO) also (Gains lost on driver. 1 Maintance required). They tried to send to predefined but did not work and received (Failed to settle under MaxClosing tolerance). They could not unlock the shoulder through the software or the shoulder button. Did the hard restart and received the Fatal CANBUS error to nodes 1,2,3,7 and Positioner IO. After about 10 mins the shoulder unlocked by itself due to the positioner reestablishing connection and arm flung across the surgical field and almost hit the mayo stand. They were able to get it off the bed and stored but when they went into superuser all the nodes were red. The last trajectory was not drilled with the robot but there was no patient harm. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Jul-11 SCL1902018 intake formx2 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they did two part registration and finished the first segment with no issues. During the second segment, the rep sent the arm from L1 right to L2 right and received the error 8798 Arm Failed to reach final destination (Driver 4 reports error: overspeed) also displayed (Failed to settle under MaxClosing tolerance). They tried to send the arm to the predefined locations but it wouldn’t reach. The rep performed a soft restart and received error 295 Fatal Error (CANBUS Connection Lost to Nodes #Positioner IO, 8 Cart IO) also (Gains lost on driver. 1 Maintance required). They tried to send to predefined but did not work and received (Failed to settle under MaxClosing tolerance). They could not unlock the shoulder through the software or the shoulder button. Did the hard restart and received the Fatal CANBUS error to nodes 1,2,3,7 and Positioner IO. After about 10 mins the shoulder unlocked by itself due to the positioner reestablishing connection and arm flung across the surgical field and almost hit the mayo stand. They were able to get it off the bed and stored but when they went into superuser all the nodes were red. The last trajectory was not drilled with the robot but there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse arm frozeerror messages spinal procedure did registration finished segment issues second segment sent arm right right error arm failed reach final destination displayed tried send arm predefined locations wouldn’t reach performed soft restart error fatal error tried send predefined did work unlock shoulder software shoulder button did hard restart fatal canbus error nodes positioner io mins shoulder unlocked positioner reestablishing connection arm flung surgical field hit mayo stand able bed stored went superuser nodes red trajectory drilled robot additional complications anticipated surgical arm | NAV4095 | NAV2047 | NAV4107 | NAV3249 | NAV1006 | NAV4095NAV2047NAV4107NAV3249NAV1006 | True
2376 | 703234005 | X - SW, x-eye, WS, & 3Define issue | 2019-06-28T19:55:54.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X - SW, x-eye, WS, & 3Define issue 2017-6-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on T5-T8. It was reported that the bed frame was mounted too low and the angle of the shoulder was too open. A 3Define scan was done, but it had poor quality so a second scan had to be done. The x-eye could not see the reference marker, so predefined AP and OBL trajectories were sent for the acquisition. The predefined trajectories were manually adjusted to get the optimal images. Registration was completed on the first attempt off of T8. Fluoro verification was accurate. All trajectories were sent and drilled, and the surgeon was pleased with the outcome. The manufacturer representative attempted to test the 3Define from utilities. A timeout pop-up appeared after 20 seconds, and the 3Define image froze. They refreshed the 3Define image and the timeout pop-up appeared again. While dismounting the surgical system from the bed the pneumatic system stopped working. The arm almost fell on the patient but the clinical engineer held it up. The representative performed a hard restart to try and fix the pressure issue, but the software crashed en dlessly. The representative had to return the surgical system to the workstation manually. During this process, the hood unexpectedly fell, so the clinical engineer inserted a drape near the hydraulic cylinders to keep the hood open. There was no patient harm a nd the procedure was not delayed over an hour. MAZOR X SOFTWARE KIT VER-3.0 | 2017-6-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on T5-T8. It was reported that the bed frame was mounted too low and the angle of the shoulder was too open. A 3Define scan was done, but it had poor quality so a second scan had to be done. The x-eye could not see the reference marker, so predefined AP and OBL trajectories were sent for the acquisition. The predefined trajectories were manually adjusted to get the optimal images. Registration was completed on the first attempt off of T8. Fluoro verification was accurate. All trajectories were sent and drilled, and the surgeon was pleased with the outcome. The manufacturer representative attempted to test the 3Define from utilities. A timeout pop-up appeared after 20 seconds, and the 3Define image froze. They refreshed the 3Define image and the timeout pop-up appeared again. While dismounting the surgical system from the bed the pneumatic system stopped working. The arm almost fell on the patient but the clinical engineer held it up. The representative performed a hard restart to try and fix the pressure issue, but the software crashed en dlessly. The representative had to return the surgical system to the workstation manually. During this process, the hood unexpectedly fell, so the clinical engineer inserted a drape near the hydraulic cylinders to keep the hood open. There was no patient harm a nd the procedure was not delayed over an hour. | x sw xeye ws issue ccr spinal procedure bed frame mounted low angle shoulder open scan poor quality second scan xeye reference marker predefined ap obl trajectories sent acquisition predefined trajectories manually adjusted optimal images registration completed attempt fluoro verification accurate trajectories sent drilled surgeon pleased outcome attempted test utilities timeout popup appeared seconds image froze refreshed image timeout popup appeared dismounting surgical bed pneumatic stopped working arm fell clinical engineer held performed hard restart try fix pressure issue software crashed en dlessly return surgical workstation manually process hood unexpectedly fell clinical engineer inserted drape near hydraulic cylinders hood open nd procedure delayed hour mazor x software kit | NAV3248 | NAV1006 | NAV3264 | NAV3242 | NAV3239 | NAV3248NAV1006NAV3264NAV3242NAV3239 | False
10091 | 704798106 | XSE - Positioner Stuck | 2022-01-07T16:42:45.000000+0000 | MazorX Stealth Edition | NAV4104 | UNABLE TO MOVE | 20 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE - Positioner Stuck 2022-Jan-07 Salesforce C00240596 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner was stuck all the way down. The manufacturer representative believes the spring was stuck fully compressed. No patient was present. 2022-01-11 IUD (Other): No new information. 2022-02-01 IUD (Other): No new information. 2022-02-08 IUD x2 (Other): No new information. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2022-Jan-07 Salesforce C00240596 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner was stuck all the way down. The manufacturer representative believes the spring was stuck fully compressed. No patient was present. 2022-01-11 IUD (Other): No new information. 2022-02-01 IUD (Other): No new information. 2022-02-08 IUD x2 (Other): No new information. | xse positioner stuck salesforce outside procedure positioner stuck way believes spring stuck fully compressed present iud new iud new iud new kit svc surgical new | NAV3251 | NAV4104 | NAV3299 | NAV2047 | NAV4095 | NAV3251NAV4104NAV3299NAV2047NAV4095 | True
10011 | 704770813 | XSE - Stripped Spinous Process Clamp | 2021-12-17T20:31:56.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1355 | MAS1355 DUAL CLAMP | XSE - Stripped Spinous Process Clamp 2021-Dec-17 Salesforce C00239931 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the double spinous process clamp in the their tray is stripped and the clamping mechanism no longer works. This was discovered outside of a case. There was no patient involved. 2022-01-06 IUD x2 (Other): No new information. MAS1355 DUAL CLAMP | 2021-Dec-17 Salesforce C00239931 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the double spinous process clamp in the their tray is stripped and the clamping mechanism no longer works. This was discovered outside of a case. There was no patient involved. 2022-01-06 IUD x2 (Other): No new information. | xse stripped spinous process clamp salesforce outside procedure double spinous process clamp tray stripped clamping mechanism longer works discovered outside case involved iud new dual clamp | NAV3299 | NAV3261 | NAV2057 | NAV4100 | NAV2054 | NAV3299NAV3261NAV2057NAV4100NAV2054 | True
7567 | 704228215 | XSE - Damaged Ethernet Cord | 2021-02-12T17:47:11.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 20 | HRS0350-01S | KIT SVC HRS0350-01S NAV CONNECT CABLE | XSE - Damaged Ethernet Cord 2021-02-11 00706460 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the ethernet between the main cart and the navigation system camera was damaged and a replacement was required. No patient was present. 2021-02-22 IUD x2 (Other): No new information. 2021-03-16 IUD (Other): No new information. 2021-05-10 IUD (Other): No new information. 2021-08-04 IUD (Other): No new information. KIT SVC HRS0350-01S NAV CONNECT CABLE | 2021-02-11 00706460 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the ethernet between the main cart and the navigation system camera was damaged and a replacement was required. No patient was present. 2021-02-22 IUD x2 (Other): No new information. 2021-03-16 IUD (Other): No new information. 2021-05-10 IUD (Other): No new information. 2021-08-04 IUD (Other): No new information. | xse damaged ethernet cord outside procedure ethernet main cart navigation camera damaged replacement required present iud new iud new iud new iud new kit svc nav connect cable | NAV3256 | NAV2058 | NAV2047 | NAV3295 | NAV2054 | NAV3256NAV2058NAV2047NAV3295NAV2054 | True
7318 | 704174483 | XSE - 3Define Camera Failure/Blank Scree | 2021-01-14T21:29:40.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - 3Define Camera Failure/Blank Scree 2021-Jan-14 Salesforce C00219306, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that surgical arm went to the first position for the 3Define scan, but then became unresponsive. A hard reboot of the system was done, but the 3Define camera became unresponsive again. A second hard reboot of the system was done. Upon start up of the guidance system, both the monitor and surgeon screen were black and only the lave out of the external monitor functioned. A third hard reboot of the system was done and the issue was resolved. The system was working as intended. There was no patient harm and the procedure was delayed less than an hour. 2021-Jan-28 IUD (other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2021-Jan-14 Salesforce C00219306, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that surgical arm went to the first position for the 3Define scan, but then became unresponsive. A hard reboot of the system was done, but the 3Define camera became unresponsive again. A second hard reboot of the system was done. Upon start up of the guidance system, both the monitor and surgeon screen were black and only the lave out of the external monitor functioned. A third hard reboot of the system was done and the issue was resolved. The system was working as intended. There was no patient harm and the procedure was delayed less than an hour. 2021-Jan-28 IUD (other): No new information. | xse camera failureblank scree salesforce iud spinal procedure surgical arm went position scan unresponsive hard reboot camera unresponsive second hard reboot start monitor surgeon screen black lave external monitor functioned hard reboot issue resolved working intended procedure delayed hour iud new unk mazor x camera | NAV3244 | NAV3253 | NAV3261 | NAV3248 | NAV3262 | NAV3244NAV3253NAV3261NAV3248NAV3262 | True
5981 | 703851855 | XSE - Deviation | 2020-07-15T02:57:19.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2020-07-01 CCR, Attachment (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the system mounted using a single clamp from the dual clamp kit. The surgeon wanted a single clamp due to the large size of dual clamp. Registration was achieved with no issues. L2 left was sent first with no visual issues. L1 left was sent and appeared lateral. The manufacturer representative told the surgeon they were leaning on the patient causing the lateral screw. The surgeon backed off the patient and the screw was placed with no issue. A slight retraction was needed for thoracic screws but no visible issue on navigation detected. Final images were taken and L2 left was very superior. The surgeon pulled the screw and informed the manufacturer representative that 5.5 tap hole was perfect and imaging showed the same. It was unknown as to why the L2 screw was superior. T10 left screw was slightly medial, and manufacturer representative and the surgeon agreed that it w as a medial skive. The surgeon also noted that patient had very unstable joints. The surgeon noted not all screws were perfect but looked and felt acceptable. There was no known impact to the patient outcome. 2020-07-24 e1 (Rep): Additional information was received stating that surgery was delayed roughly 20 minutes due to this incident. L2 trajectory was deviated between 3.5 and 10mm. There was no impact on the patient outcome. 2020-08-03 email 2020-08-03 14:00 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2020-07-01 CCR, Attachment (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the system mounted using a single clamp from the dual clamp kit. The surgeon wanted a single clamp due to the large size of dual clamp. Registration was achieved with no issues. L2 left was sent first with no visual issues. L1 left was sent and appeared lateral. The manufacturer representative told the surgeon they were leaning on the patient causing the lateral screw. The surgeon backed off the patient and the screw was placed with no issue. A slight retraction was needed for thoracic screws but no visible issue on navigation detected. Final images were taken and L2 left was very superior. The surgeon pulled the screw and informed the manufacturer representative that 5.5 tap hole was perfect and imaging showed the same. It was unknown as to why the L2 screw was superior. T10 left screw was slightly medial, and manufacturer representative and the surgeon agreed that it w as a medial skive. The surgeon also noted that patient had very unstable joints. The surgeon noted not all screws were perfect but looked and felt acceptable. There was no known impact to the patient outcome. 2020-07-24 e1 (Rep): Additional information was received stating that surgery was delayed roughly 20 minutes due to this incident. L2 trajectory was deviated between 3.5 and 10mm. There was no impact on the patient outcome. 2020-08-03 email 2020-08-03 14:00 (Rep): No new information. | xse deviation ccr attachment intraoperatively spinal procedure mounted using single clamp dual clamp kit surgeon wanted single clamp large size dual clamp registration achieved issues left sent visual issues left sent appeared lateral told surgeon leaning causing lateral screw surgeon backed screw placed issue slight retraction needed thoracic screws visible issue navigation detected final images taken left superior surgeon pulled screw informed tap hole perfect imaging showed unknown screw superior left screw slightly medial surgeon agreed w medial skive surgeon noted unstable joints surgeon noted screws perfect looked felt acceptable known impact outcome additional stating surgery delayed roughly minutes incident trajectory deviated impact outcome email new mazor x | NAV5000 | NAV4091 | NAV3242 | NAV3299 | NAV5001 | NAV5000NAV4091NAV3242NAV3299NAV5001 | True
1380 | 703197983 | X- Workstation Issues | 2019-06-10T20:16:24.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- Workstation Issues Planning: The planning was done by the surgeon prior to the day of surgery on a personal computer. The workstation would not recognize the surgeon's usb drive. Many attempts had to be made to get the usb drive to mount even though the drive mounted on to every other computer with no issue. Platform: The surgeon chose to the use the tall clamp from tray 762. Registration: Registration was fast and accurate.Operation: The procedure went forward with some issues. An arm shift occurred prior to sending the first trajectory. Following this, a new registration was done.On one trajectory, the system alarmed that the arm was not on trajectory. The registration was not lost, so the arm was resent and the surgeon carried on. The surgeon decided to instrument the L4 trajectories by hand. After registration before sending the arm to any trajectories, there was an immediate shoulder shift. The surgeon decided to register the spine again. Later, the arm came off trajectory without being touched, but the registration remained intact. The arm was resent to the trajectory and instrumented. Following this, the surgeon did not trust the arm and instrumented the rest of the scews by hand. Additionally, there was great difficulty attempting to import and export cases from the workstation. The workstation did not want to mount the usb drive. UNK MAZOR X WORKSTATION | Planning: The planning was done by the surgeon prior to the day of surgery on a personal computer. The workstation would not recognize the surgeon's usb drive. Many attempts had to be made to get the usb drive to mount even though the drive mounted on to every other computer with no issue. Platform: The surgeon chose to the use the tall clamp from tray 762. Registration: Registration was fast and accurate.Operation: The procedure went forward with some issues. An arm shift occurred prior to sending the first trajectory. Following this, a new registration was done.On one trajectory, the system alarmed that the arm was not on trajectory. The registration was not lost, so the arm was resent and the surgeon carried on. The surgeon decided to instrument the L4 trajectories by hand. After registration before sending the arm to any trajectories, there was an immediate shoulder shift. The surgeon decided to register the spine again. Later, the arm came off trajectory without being touched, but the registration remained intact. The arm was resent to the trajectory and instrumented. Following this, the surgeon did not trust the arm and instrumented the rest of the scews by hand. Additionally, there was great difficulty attempting to import and export cases from the workstation. The workstation did not want to mount the usb drive. | x workstation issues planning planning surgeon prior day surgery personal computer workstation recognize surgeons usb drive attempts usb drive mount drive mounted computer issue platform surgeon chose use tall clamp tray registration registration fast accurateoperation procedure went forward issues arm shift occurred prior sending trajectory following new registration doneon trajectory alarmed arm trajectory registration lost arm resent surgeon carried surgeon decided instrument trajectories hand registration sending arm trajectories immediate shoulder shift surgeon decided register spine later arm came trajectory touched registration remained intact arm resent trajectory instrumented following surgeon did trust arm instrumented rest scews hand additionally great difficulty attempting import export cases workstation workstation did want mount usb drive unk mazor x workstation | NAV2058 | NAV3257 | NAV4098 | NAV4108 | NAV4095 | NAV2058NAV3257NAV4098NAV4108NAV4095 | True
10248 | 704840329 | XSE - Deviation | 2022-02-01T17:03:49.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2022-Jan-31 Salesforce C00242505, 00731345 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were deviations during a L3-S1 MIS TLIF procedure. Deviations occurred at left L3, right L4 and left L5 and the screws were deviated by about 5 mm. A bone mount bridge was used to mount the surgical system to the patient. Registration was completed without issue and a drill used during the procedure. The surgeon liked to plan screws on top of facets. The manufacturer representative noted that proper technique and using the drill at full speed was discussed with the surgeon. Navigation looked accurate when the screws were placed. A post-op spin was done and the deviations were found. The deviations occurred in the middle of the construct and all other screws were accurate so a patient shift was ruled out. The surgeon removed the deviated screws and re-registered the patient. Registration was completed off of L3 and there were green values for L4 and yellow values for L5. The surgeon confirmed that there were no shifts during registration and the screws were replaced. Navigation accuracy was che cked using the dilator and it was in the divot. During the second attempt to place L4 and left L5, the dilator was in the divot, but slightly lower and to the left. A new snapshot was done and navigation looked better. Screws were placed again and both L4 screws a nd left L5 was medial again. The surgeon decided to abort the use of the guidance system and complete the procedure freehand. There was no patient harm and the procedure was delayed less than an hour. 2022-02-03 email 2022-02-03 15:01 (Rep): The manufacturer representative suspects the cause of the deviation to be due to using ATS screws and attempting to place them only after using the navigated midas with a 3.0mm tip. These screws appear to be overly aggressive and have a hard time allowing a 5.5mm screw or larger in the pilot hole while actively guiding it in. 2022-02-15 IUD (Other): No new information. 2022-03-01 email 2022-03-03 10:10 (Rep): The surgeon placed the screws before the cage insertion due to the last case issue experienced when we placed the cages first. The surgeon places interbody prior to any bony work. The surgeon preferred to take a more open trajectory even for MIS cases and plans each screw on the lateral facets ranging from 8-12 degrees lateral rather than aiming for the “gutters” as in most MIS cases. When asked how ATS screws that are 5.5mm, and the pilot holes that were drilled at 3mm contributed to the deviation, the manufacturer representative stated that he has seen and heard that the aggressive nature of the ATS screws can lead to misplaced screws. If the surgeon doesn’t follow proper technique or if the docking site isn’t the best and has possible skive potential it can cause the screw to go wherever it wants . With such an aggressive screw it can be difficult to have a 5.5mm screws find and easily follow a 3.0mm pilot hole. If it does not find and follow it exactly it will create its own trajectory when in c omparison to the more dulled voyager and solera screws it is forced to follow the tract due to its less aggressive tip and the fact that the pilot hole was further tapped and expanded to accommodate that screw size. 2022-03-01 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2022-Jan-31 Salesforce C00242505, 00731345 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were deviations during a L3-S1 MIS TLIF procedure. Deviations occurred at left L3, right L4 and left L5 and the screws were deviated by about 5 mm. A bone mount bridge was used to mount the surgical system to the patient. Registration was completed without issue and a drill used during the procedure. The surgeon liked to plan screws on top of facets. The manufacturer representative noted that proper technique and using the drill at full speed was discussed with the surgeon. Navigation looked accurate when the screws were placed. A post-op spin was done and the deviations were found. The deviations occurred in the middle of the construct and all other screws were accurate so a patient shift was ruled out. The surgeon removed the deviated screws and re-registered the patient. Registration was completed off of L3 and there were green values for L4 and yellow values for L5. The surgeon confirmed that there were no shifts during registration and the screws were replaced. Navigation accuracy was che cked using the dilator and it was in the divot. During the second attempt to place L4 and left L5, the dilator was in the divot, but slightly lower and to the left. A new snapshot was done and navigation looked better. Screws were placed again and both L4 screws a nd left L5 was medial again. The surgeon decided to abort the use of the guidance system and complete the procedure freehand. There was no patient harm and the procedure was delayed less than an hour. 2022-02-03 email 2022-02-03 15:01 (Rep): The manufacturer representative suspects the cause of the deviation to be due to using ATS screws and attempting to place them only after using the navigated midas with a 3.0mm tip. These screws appear to be overly aggressive and have a hard time allowing a 5.5mm screw or larger in the pilot hole while actively guiding it in. 2022-02-15 IUD (Other): No new information. 2022-03-01 email 2022-03-03 10:10 (Rep): The surgeon placed the screws before the cage insertion due to the last case issue experienced when we placed the cages first. The surgeon places interbody prior to any bony work. The surgeon preferred to take a more open trajectory even for MIS cases and plans each screw on the lateral facets ranging from 8-12 degrees lateral rather than aiming for the “gutters” as in most MIS cases. When asked how ATS screws that are 5.5mm, and the pilot holes that were drilled at 3mm contributed to the deviation, the manufacturer representative stated that he has seen and heard that the aggressive nature of the ATS screws can lead to misplaced screws. If the surgeon doesn’t follow proper technique or if the docking site isn’t the best and has possible skive potential it can cause the screw to go wherever it wants . With such an aggressive screw it can be difficult to have a 5.5mm screws find and easily follow a 3.0mm pilot hole. If it does not find and follow it exactly it will create its own trajectory when in c omparison to the more dulled voyager and solera screws it is forced to follow the tract due to its less aggressive tip and the fact that the pilot hole was further tapped and expanded to accommodate that screw size. 2022-03-01 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. | xse deviation salesforce spinal procedure deviations mis tlif procedure deviations occurred left right left screws deviated mm bone mount bridge mount surgical registration completed issue drill procedure surgeon liked plan screws facets noted proper technique using drill speed discussed surgeon navigation looked accurate screws placed postop spin deviations deviations occurred middle construct screws accurate shift ruled surgeon removed deviated screws reregistered registration completed green values yellow values surgeon confirmed shifts registration screws replaced navigation accuracy che cked using dilator divot second attempt place left dilator divot slightly lower left new snapshot navigation looked better screws placed screws nd left medial surgeon decided abort use complete procedure freehand procedure delayed hour email suspects cause deviation using ats screws attempting place using navigated midas tip screws appear overly aggressive hard time allowing screw larger pilot hole actively guiding iud new email surgeon placed screws cage insertion case issue experienced placed cages surgeon places interbody prior bony work surgeon preferred open trajectory mis cases plans screw lateral facets ranging degrees lateral aiming “gutters” mis cases asked ats screws pilot holes drilled contributed deviation stated seen heard aggressive nature ats screws lead misplaced screws surgeon doesn’t follow proper technique docking site isn’t best possible skive potential cause screw wants aggressive screw difficult screws easily follow pilot hole does follow exactly create trajectory c omparison dulled voyager solera screws forced follow tract aggressive tip fact pilot hole tapped expanded accommodate screw size iud new iud new mazor x | NAV5000 | NAV5001 | NAV4091 | NAV1012 | NAV3241 | NAV5000NAV5001NAV4091NAV1012NAV3241 | True
9374 | 704620182 | XSE - 3Define Volume Off | 2021-09-20T22:00:54.000000+0000 | MazorX Stealth Edition | NAV3239 | INACCURATE WORK VOLUMES | 20 | KIT1318 | SEN0088 SPARE PART KIT KIT1318 | XSE - 3Define Volume Off 2021-Sep-17 Salesforce C00234709 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the 3Define scan was completed during the procedure and the work volume was translated toward the floor by about an inch. The surgical arm was about to contact the patient when being sent and the emergency stop button had to be pressed. A generic work volume was used to move forward and complete the procedure. There were no accuracy issues during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2021-09-21 email 2021-09-21 19:30 (Other): No new information. 2021-09-23 IUD x2 (Other): No new information. SEN0088 SPARE PART KIT KIT1318 | 2021-Sep-17 Salesforce C00234709 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the 3Define scan was completed during the procedure and the work volume was translated toward the floor by about an inch. The surgical arm was about to contact the patient when being sent and the emergency stop button had to be pressed. A generic work volume was used to move forward and complete the procedure. There were no accuracy issues during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2021-09-21 email 2021-09-21 19:30 (Other): No new information. 2021-09-23 IUD x2 (Other): No new information. | xse volume salesforce spinal p rocedure scan completed procedure work volume translated floor inch surgical arm contact sent emergency stop button pressed generic work volume forward complete procedure accuracy issues procedure procedure delayed hour email new iud new spare kit | NAV3244 | NAV4096 | NAV3246 | NAV5000 | NAV2047 | NAV3244NAV4096NAV3246NAV5000NAV2047 | False
5624 | 703767408 | XSE - System Did Not Accept Images | 2020-05-22T22:16:09.000000+0000 | MazorX Stealth Edition | NAV2050 | IMAGES NOT AVAILABLE | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - System Did Not Accept Images 2020-05-21 00689448 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a scan and plan workflow the system did not accept images. The surgeon wanted to leave the retractors in for the case. The surgeon used the guidance system for trajectory and placing Stryker pedicle markers. During the first spin there was too much metal and system did not except images. The marker was moved and system accepted second spin. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-05-28 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-05-21 00689448 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a scan and plan workflow the system did not accept images. The surgeon wanted to leave the retractors in for the case. The surgeon used the guidance system for trajectory and placing Stryker pedicle markers. During the first spin there was too much metal and system did not except images. The marker was moved and system accepted second spin. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-05-28 IUD (Other): No new information. | xse did accept images intraoperatively spinal procedure scan plan workflow did accept images surgeon wanted leave retractors case surgeon trajectory placing stryker pedicle markers spin metal did images marker moved accepted second spin issue extended surgical time hour impact outcome iud new mazor x | NAV3242 | NAV5000 | NAV3235 | NAV2050 | NAV4112 | NAV3242NAV5000NAV3235NAV2050NAV4112 | True
4945 | 703587281 | X - Deviation, Platform Shift | 2020-01-29T18:45:54.000000+0000 | MazorX | NAV3241 | PLATFORM STABILITY ISSUES | 30 | MAS2351-01 | SCHANZ ARM | X - Deviation, Platform Shift 2018-06-26 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed for a posterior spinal fusion of the patient's L5-S1 in lateral position. Pre-surgical arm accuracy tests and camera reviews were completed successfully. The surgeon and the manufacturer representative reviewed intra-operative plan prior to the procedure; S1 trajectories were discussed specifically noting medial skive potential due to the facet joint tilt. Skive potentials, screw sizing, and placements were all discussed. The system was mounted to the bed and draped. The system was connected to the patient by schanz arm platform and attached to the patient via schanz pin inserted in the left iliac crest. The 3Define scan was completed. The right L5 trajectory was sent; to better access the drill guid trajectories, the surgeon requested at that time that anesthesia airplane the table away from the operative side. In tilting the table, the rotation was incorrectly adjusted which lead to a major shift of the patient on the the table. The surgeon was questions but made the decision to continue with right S1. Right S1 trajectory was drilled and screw placed. Right S1 was checked with fluoroscopy in the lateral position and the implant was noted to be placed inferior to plan. The surgeon opted to proceed to left L5. The left L5 trajectory was sent as well and and placed. Fluoroscopy showed left L5 was also inferior to the pedicle. The platform was check and observed to be loose. The assisting surgeon was able to tighten the platform joint known one full turn. Once the arm was tightened and secure, new registration images were obtained, new 3Define completed, and registration was verified. Left L5 was sent. Fluoroscopy used to check left L5 placement which was deemed accurate. Right S1 was sent to fix the previous inferior screw. While adjusting the height of the surgical arm relative to the patient, the arm collided with the MIS screw tower and caused a shift of the surgical arm. Registration had to be redone again and new 3Define was done; registration was verified. Right S1 was sent and placed. Left S1 was sent and placed. The robot was unmounted from the patient, then removed from the bed. Final screw placements were deemed accurate with fluoro imaging and confirmed by the surgeon and the representative, and engineer. No patient harm was reported and surgery was delayed less than an hour. It was noted that the surgeon felt the system should be aware of the height of the towers in the software to prevent accidental collision with the arm guide. SCHANZ ARM | 2018-06-26 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed for a posterior spinal fusion of the patient's L5-S1 in lateral position. Pre-surgical arm accuracy tests and camera reviews were completed successfully. The surgeon and the manufacturer representative reviewed intra-operative plan prior to the procedure; S1 trajectories were discussed specifically noting medial skive potential due to the facet joint tilt. Skive potentials, screw sizing, and placements were all discussed. The system was mounted to the bed and draped. The system was connected to the patient by schanz arm platform and attached to the patient via schanz pin inserted in the left iliac crest. The 3Define scan was completed. The right L5 trajectory was sent; to better access the drill guid trajectories, the surgeon requested at that time that anesthesia airplane the table away from the operative side. In tilting the table, the rotation was incorrectly adjusted which lead to a major shift of the patient on the the table. The surgeon was questions but made the decision to continue with right S1. Right S1 trajectory was drilled and screw placed. Right S1 was checked with fluoroscopy in the lateral position and the implant was noted to be placed inferior to plan. The surgeon opted to proceed to left L5. The left L5 trajectory was sent as well and and placed. Fluoroscopy showed left L5 was also inferior to the pedicle. The platform was check and observed to be loose. The assisting surgeon was able to tighten the platform joint known one full turn. Once the arm was tightened and secure, new registration images were obtained, new 3Define completed, and registration was verified. Left L5 was sent. Fluoroscopy used to check left L5 placement which was deemed accurate. Right S1 was sent to fix the previous inferior screw. While adjusting the height of the surgical arm relative to the patient, the arm collided with the MIS screw tower and caused a shift of the surgical arm. Registration had to be redone again and new 3Define was done; registration was verified. Right S1 was sent and placed. Left S1 was sent and placed. The robot was unmounted from the patient, then removed from the bed. Final screw placements were deemed accurate with fluoro imaging and confirmed by the surgeon and the representative, and engineer. No patient harm was reported and surgery was delayed less than an hour. It was noted that the surgeon felt the system should be aware of the height of the towers in the software to prevent accidental collision with the arm guide. | x deviation platform shift ccr spinal procedure planning completed posterior spinal fusion patients lateral position presurgical arm accuracy tests camera reviews completed successfully surgeon reviewed intraoperative plan prior procedure trajectories discussed specifically noting medial skive potential facet joint tilt skive potentials screw sizing placements discussed mounted bed draped connected schanz arm platform attached schanz pin inserted left iliac crest scan completed right trajectory sent better access drill guid trajectories surgeon requested time anesthesia airplane table away operative tilting table rotation incorrectly adjusted lead major shift table surgeon questions decision continue right right trajectory drilled screw placed right checked fluoroscopy lateral position implant noted placed inferior plan surgeon opted proceed left left trajectory sent placed fluoroscopy showed left inferior pedicle platform check observed loose assisting surgeon able tighten platform joint known turn arm tightened secure new registration images obtained new completed registration verified left sent fluoroscopy check left placement deemed accurate right sent fix previous inferior screw adjusting height surgical arm relative arm collided mis screw tower caused shift surgical arm registration redone new registration verified right sent placed left sent placed robot unmounted removed bed final screw placements deemed accurate fluoro imaging confirmed surgeon engineer surgery delayed hour noted surgeon felt aware height towers software prevent accidental collision arm guide schanz arm | NAV4091 | NAV4096 | NAV4098 | NAV3299 | NAV5000 | NAV4091NAV4096NAV4098NAV3299NAV5000 | False
9839 | 704730816 | XSE - Stripped Schanz Arm Screw | 2021-11-24T04:17:11.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5750-01S | KIT SVC MEC5750-01S SCH CON SCREW | XSE - Stripped Schanz Arm Screw 2021-Nov-23 Salesforce C00238644 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the the Schanz arm screw was found to be stripped in sterile processing. There was no patient involvement. 2021-11-23 mPXR 899838 (Rep, HCP): Additional information received from a manufacturer representative reported that the Schanz arm was unable to be removed due to the screw being stripped. The issue was discovered after the case was completed and the system was being put away. 2021-12-03 IUD x2 (Other): No new information. 2021-12-17 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. 2021-12-28 email 2021-12-28 12:45 (Rep): No new information. 2022-01-07 email 2022-01-07 14:31 (Rep): No new information. KIT SVC MEC5750-01S SCH CON SCREW | 2021-Nov-23 Salesforce C00238644 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the the Schanz arm screw was found to be stripped in sterile processing. There was no patient involvement. 2021-11-23 mPXR 899838 (Rep, HCP): Additional information received from a manufacturer representative reported that the Schanz arm was unable to be removed due to the screw being stripped. The issue was discovered after the case was completed and the system was being put away. 2021-12-03 IUD x2 (Other): No new information. 2021-12-17 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. 2021-12-28 email 2021-12-28 12:45 (Rep): No new information. 2022-01-07 email 2022-01-07 14:31 (Rep): No new information. | xse stripped schanz arm screw salesforce outside procedure schanz arm screw stripped sterile processing mpxr additional schanz arm unable removed screw stripped issue discovered case completed away iud new iud new iud new email new email new kit svc sch screw | NAV3299 | NAV4100 | NAV4090 | NAV2054 | NAV4095 | NAV3299NAV4100NAV4090NAV2054NAV4095 | True
1260 | 703194774 | X - DEVIATION | 2019-06-07T12:59:25.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 40 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - DEVIATION This complaint is separated to two complaints:Medial deviation - this complaint. 3Define camera not working- SCL1801777. 3define camera was not working, soft restart done 2-3 times in order to get the camera functioning. L5 right was medial – surgeon requests an engineer come look at the robot./5.7.18 Dr. Francis L5-S1 TLIF Preop: Cameras checked, AA completed, shoulder checked. Plan: Done by Dr. Francis, axial angles were narrow and uniform from level to level. Schanz arm used and schanz pin placed in right PSIS. 3Define not completed on the firs try, RTC started 3define, but the 3define cameras were not working-RTC did a soft restart, and tried to start the 3define again, but the 3define camera was not responding again. '-RTC did another soft restart, this time it worked and 3Define was completed. -Drawspine completed on 2nd try -first try, the surgeon interfered with the the x-eye. Registration: Completed on the first try with green values and visually verified by surgeon.Operation: Operation seemed quick and accurate, and surgeon did not seem like anything seemed inaccurate. When final AP was taken surgeon said that the right L5 was medial. Surgeon removed the robot and said he would free hand the right L5. Surgeon requests that the engineers come and check the robots accuracy. UNK MAZOR X 3DEFINE CAMERA | This complaint is separated to two complaints:Medial deviation - this complaint. 3Define camera not working- SCL1801777. 3define camera was not working, soft restart done 2-3 times in order to get the camera functioning. L5 right was medial – surgeon requests an engineer come look at the robot./5.7.18 Dr. Francis L5-S1 TLIF Preop: Cameras checked, AA completed, shoulder checked. Plan: Done by Dr. Francis, axial angles were narrow and uniform from level to level. Schanz arm used and schanz pin placed in right PSIS. 3Define not completed on the firs try, RTC started 3define, but the 3define cameras were not working-RTC did a soft restart, and tried to start the 3define again, but the 3define camera was not responding again. '-RTC did another soft restart, this time it worked and 3Define was completed. -Drawspine completed on 2nd try -first try, the surgeon interfered with the the x-eye. Registration: Completed on the first try with green values and visually verified by surgeon.Operation: Operation seemed quick and accurate, and surgeon did not seem like anything seemed inaccurate. When final AP was taken surgeon said that the right L5 was medial. Surgeon removed the robot and said he would free hand the right L5. Surgeon requests that the engineers come and check the robots accuracy. | x deviation complaint separated complaintsmedial deviation complaint camera working camera working soft restart times order camera functioning right medial – surgeon requests engineer come look dr francis tlif preop cameras checked completed shoulder checked plan dr francis axial angles narrow uniform level level schanz arm schanz pin placed right psis completed firs try rtc started cameras workingrtc did soft restart tried start camera responding rtc did soft restart time worked completed drawspine completed try try surgeon interfered xeye registration completed try green values visually verified surgeonoperation operation quick accurate surgeon did like inaccurate final ap taken surgeon said right medial surgeon removed robot said free hand right surgeon requests engineers come check robots accuracy unk mazor x camera | NAV3244 | NAV4091 | NAV4110 | NAV3233 | NAV3241 | NAV3244NAV4091NAV4110NAV3233NAV3241 | True
4496 | 703546442 | X- Monitor Issue, CANBUS error, L4 Skive | 2020-01-06T21:11:36.000000+0000 | MazorX | NAV3261 | MONITOR ISSUES | 30 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | X- Monitor Issue, CANBUS error, L4 Skive 2018-04-19 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a global deformity with fixation at L4-L5. It was reported that the surgeon approved the preoperative plan; the plan attempted to address potential for lateral skiving at right L4 with left angling. There were issues with booting up; the screen would not come on. The representative was able to power it on through the PC console. This resulted in an 2 minutes delay. Just before completing Draw Spine, an error displayed: "FATAL ERROR 295 CANBUS CONNECTION LOST NODE #4." The representative reported it did not affect the case; a generic Draw Spine was used instead. There were no issues with registration. The surgeon did skive once at right L4 but it was adjusted slightly and placed accurately. All final placements were accurate. No patient harm was reported and surgery was delayed less than an hour. UNK MAZOR X MONITOR | 2018-04-19 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a global deformity with fixation at L4-L5. It was reported that the surgeon approved the preoperative plan; the plan attempted to address potential for lateral skiving at right L4 with left angling. There were issues with booting up; the screen would not come on. The representative was able to power it on through the PC console. This resulted in an 2 minutes delay. Just before completing Draw Spine, an error displayed: "FATAL ERROR 295 CANBUS CONNECTION LOST NODE #4." The representative reported it did not affect the case; a generic Draw Spine was used instead. There were no issues with registration. The surgeon did skive once at right L4 but it was adjusted slightly and placed accurately. All final placements were accurate. No patient harm was reported and surgery was delayed less than an hour. | x monitor issue canbus error skive ccr spinal procedure treat global deformity fixation surgeon approved preoperative plan plan attempted address potential lateral skiving right left angling issues booting screen come able power pc console resulted minutes delay just completing draw spine error displayed fatal error canbus connection lost node did affect case generic draw spine instead issues registration surgeon did skive right adjusted slightly placed accurately final placements accurate surgery delayed hour unk mazor x monitor | NAV2047 | NAV4091 | NAV3261 | NAV3244 | NAV4107 | NAV2047NAV4091NAV3261NAV3244NAV4107 | True
10411 | 704875172 | XSE - SRA Cross Threaded Screws | 2022-02-21T21:07:35.000000+0000 | MazorX Stealth Edition | NAV4100 | STRIPPED THREADS | 20 | ASM0206-03S | SURGICAL ARM SPARE PART ASM0206-03S | XSE - SRA Cross Threaded Screws 2022-Feb-21 Salesforce C00243962 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the threading on the surgical arm where the reference frame attached was "countersunk" and the frame could no longer be connected. This prevented navigation from being used. There was no patient involvement. 2022-02-24 IUD (Other): Every joint on the replacement arm failed. 2022-03-06 email, attachment 2022-03-07 10:17 (Rep): Approximately 15 minutes Into the stress test the arm stopped and an error code popped up on the monitor indicating that there were joint failures in joints 1-6. 2022-03-03 IUD x3 (Other): No new information. 2022-03-29 IUD (Other): No new information. 2022-04-03 e3, e4 (Rep): Additional information received from a manufacturer representative reported that the surgical arm exceeded all max static allowance on every joint. 2022-04-06 e6 (Rep): Additional information was received stating that there wasn’t an error code, the system showed that the stress test failed, and once the report popped up it shows a pass or fail on each joint, all joints showed failed. 2022-04-05 IUD (Other): No new information. 2022-04-12 IUD (Other): No new information. 2022-04-14 302786411 (Other): No new information. 2022-04-20 e7, e8, e9 (Other): No new information. 2022-05-02 IUD x3 (Other): No new information. SURGICAL ARM SPARE PART ASM0206-03S | 2022-Feb-21 Salesforce C00243962 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the threading on the surgical arm where the reference frame attached was "countersunk" and the frame could no longer be connected. This prevented navigation from being used. There was no patient involvement. 2022-02-24 IUD (Other): Every joint on the replacement arm failed. 2022-03-06 email, attachment 2022-03-07 10:17 (Rep): Approximately 15 minutes Into the stress test the arm stopped and an error code popped up on the monitor indicating that there were joint failures in joints 1-6. 2022-03-03 IUD x3 (Other): No new information. 2022-03-29 IUD (Other): No new information. 2022-04-03 e3, e4 (Rep): Additional information received from a manufacturer representative reported that the surgical arm exceeded all max static allowance on every joint. 2022-04-06 e6 (Rep): Additional information was received stating that there wasn’t an error code, the system showed that the stress test failed, and once the report popped up it shows a pass or fail on each joint, all joints showed failed. 2022-04-05 IUD (Other): No new information. 2022-04-12 IUD (Other): No new information. 2022-04-14 302786411 (Other): No new information. 2022-04-20 e7, e8, e9 (Other): No new information. 2022-05-02 IUD x3 (Other): No new information. | xse sra cross threaded screws salesforce outside procedure threading surgical arm reference frame attached countersunk frame longer connected prevented navigation iud joint replacement arm failed email attachment approximately minutes stress test arm stopped error code popped monitor indicating joint failures joints iud new iud new additional surgical arm exceeded max static allowance joint additional stating wasn’t error code showed stress test failed report popped shows pass fail joint joints showed failed iud new iud new new new iud new surgical arm spare | NAV3249 | NAV2054 | NAV4092 | NAV4098 | NAV3261 | NAV3249NAV2054NAV4092NAV4098NAV3261 | False
7350 | 704185034 | XSE - Failed Pneumatic Test | 2021-01-20T21:32:49.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE - Failed Pneumatic Test 2021-Jan-20 Salesforce C00219630 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while performing the Positioner Type II FCA on the system, the pneumatic test failed. There was no patient present. 2021-Jan-25 IUDx4 (other): No new information. 2021-02-22 IUD x2 (Other): No new information. 2021-03-23 IUD (Other): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2021-Jan-20 Salesforce C00219630 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while performing the Positioner Type II FCA on the system, the pneumatic test failed. There was no patient present. 2021-Jan-25 IUDx4 (other): No new information. 2021-02-22 IUD x2 (Other): No new information. 2021-03-23 IUD (Other): No new information. | xse failed pneumatic test salesforce outside procedure performing positioner type ii fca pneumatic test failed present new iud new iud new kit svc pneumatic panel | NAV3264 | NAV2054 | NAV3251 | NAV4093 | NAV4092 | NAV3264NAV2054NAV3251NAV4093NAV4092 | True
7204 | 704141345 | XSE ¦ Clamp Link stripped | 2020-12-23T17:19:34.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5206-01S | KIT SVC MEC5206-01S RNX-ACC-021 | XSE ¦ Clamp Link stripped 2020-Dec-23 00703275 (rep): Medtronic received information from a manufacturer representative (rep) regarding a device being used outside a procedure. It was reported that the screw on the clamp link is stripped. The rep tried to get it removed, but was unable to. There was no patient harm or reported as a result of the event. 2020-Dec-23 IUDx3 (other): They were able to get the screw out, but the clamp link is damaged. 2021-Feb-01 IUD (other): No new information. 2021-03-18 IUD (Other): No new information. 2021-04-16 IUD x4 (Other): No new information. 2021-05-21 IUD x3 (Other): No new information. 2021-07-08 IUD (Other): No new information. KIT SVC MEC5206-01S RNX-ACC-021 | 2020-Dec-23 00703275 (rep): Medtronic received information from a manufacturer representative (rep) regarding a device being used outside a procedure. It was reported that the screw on the clamp link is stripped. The rep tried to get it removed, but was unable to. There was no patient harm or reported as a result of the event. 2020-Dec-23 IUDx3 (other): They were able to get the screw out, but the clamp link is damaged. 2021-Feb-01 IUD (other): No new information. 2021-03-18 IUD (Other): No new information. 2021-04-16 IUD x4 (Other): No new information. 2021-05-21 IUD x3 (Other): No new information. 2021-07-08 IUD (Other): No new information. | xse clamp link stripped outside procedure screw clamp link stripped tried removed unable able screw clamp link damaged iud new iud new iud new iud new iud new kit svc | NAV3299 | NAV2054 | NAV4100 | NAV3295 | NAV4090 | NAV3299NAV2054NAV4100NAV3295NAV4090 | True
10368 | 704863050 | XSE - SRA Shift Error | 2022-02-14T21:01:58.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Shift Error 2022-Feb-14 Salesforce C00243652 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that after the screws were placed the surgeon got a joint shift error. They were able to get the robot off the bed and then it would no longer move. They were able to get it out of the operating field and have the system powered off. There was no impact on the patient outcome. 2022-02-21 IUD (Other): No new information. 2022-03-02 IUD (Other): No new information. 2022-03-10 IUD x2 (Other): No new information. 2022-04-01 IUD (Other): No new information. 2022-04-15 IUD (Other): No new information. 2022-04-18 email 2022-04-18 12:46 (Rep): No new information. 2022-04-22 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2022-Feb-14 Salesforce C00243652 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that after the screws were placed the surgeon got a joint shift error. They were able to get the robot off the bed and then it would no longer move. They were able to get it out of the operating field and have the system powered off. There was no impact on the patient outcome. 2022-02-21 IUD (Other): No new information. 2022-03-02 IUD (Other): No new information. 2022-03-10 IUD x2 (Other): No new information. 2022-04-01 IUD (Other): No new information. 2022-04-15 IUD (Other): No new information. 2022-04-18 email 2022-04-18 12:46 (Rep): No new information. 2022-04-22 IUD (Other): No new information. | xse sra shift error salesforce spinal p rocedure screws placed surgeon got joint shift error able robot bed longer able operating field powered impact outcome iud new iud new iud new iud new iud new email new iud new kit svc surgical arm refurb | NAV4098 | NAV1006 | NAV3249 | NAV4095 | NAV4092 | NAV4098NAV1006NAV3249NAV4095NAV4092 | True
7053 | 704109248 | XSE - No pressure | 2020-12-07T19:02:09.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - No pressure 2020-Dec-07 00702159 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that when on site the rep had noted that the system was not the way the system was left. It was noted that the arm was not back in the storage unit, the hood was propped and the side panel to the pneumatic portion of the system was not getting any pressure. Pneumatic panel was noted to be showing readings in the software but no on the panel. The rep was unable to check any status of the arm or attach the system to the bed with out pressure to the system. No patient was present. No delay noted. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-14 IUD x3 (Other): No new information. 2021-01-12 IUD (Other): No new information. 2021-01-06 email 2021-01-06 23:45 (Rep): No new information. 2021-02-02 email 2021-02-02 13:00 (Rep): The pneumatics of the system were recovered and they had a successful case. UNK MAZOR X WORKSTATION | 2020-Dec-07 00702159 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that when on site the rep had noted that the system was not the way the system was left. It was noted that the arm was not back in the storage unit, the hood was propped and the side panel to the pneumatic portion of the system was not getting any pressure. Pneumatic panel was noted to be showing readings in the software but no on the panel. The rep was unable to check any status of the arm or attach the system to the bed with out pressure to the system. No patient was present. No delay noted. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-14 IUD x3 (Other): No new information. 2021-01-12 IUD (Other): No new information. 2021-01-06 email 2021-01-06 23:45 (Rep): No new information. 2021-02-02 email 2021-02-02 13:00 (Rep): The pneumatics of the system were recovered and they had a successful case. | xse pressure outside procedure site noted way left noted arm storage unit hood propped panel pneumatic portion getting pressure pneumatic panel noted showing readings software panel unable check status arm attach bed pressure present delay noted additional complications anticipated iud new iud new email new email pneumatics recovered successful case unk mazor x workstation | NAV3264 | NAV3250 | NAV2054 | NAV2058 | NAV3263 | NAV3264NAV3250NAV2054NAV2058NAV3263 | True
7500 | 704215205 | XSE - Image Adaptor Screws | 2021-02-05T13:42:32.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 20 | KIT0153-03 | ACCESSORIES, KIT,ATTACHMENT KIT FOR IMAG | XSE - Image Adaptor Screws 2021-02-05 00705944 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the image adaptor screws were damaged. No patient was present. 2021-02-15 IUD x3 (Other): No new information. 2021-03-10 IUD (Other): No new information. 2021-03-23 email 2021-03-23 14:16 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined. The screws were discarded at the site after being replaced. ACCESSORIES, KIT,ATTACHMENT KIT FOR IMAG | 2021-02-05 00705944 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the image adaptor screws were damaged. No patient was present. 2021-02-15 IUD x3 (Other): No new information. 2021-03-10 IUD (Other): No new information. 2021-03-23 email 2021-03-23 14:16 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined. The screws were discarded at the site after being replaced. | xse image adaptor screws outside procedure image adaptor screws damaged present iud new iud new email additional cause determined screws discarded site replaced accessories kitattachment kit imag | NAV3299 | NAV3258 | NAV2054 | NAV2053 | NAV4090 | NAV3299NAV3258NAV2054NAV2053NAV4090 | True
6725 | 704014719 | XSE - 3Define Camera Issue/WV Inaccurate | 2020-10-15T16:15:07.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - 3Define Camera Issue/WV Inaccurate 2020-10-15 00698837, 00698856, PDF, IUD x4 (Rep, HPC): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera was not functioning during the procedure. Setup of the guidance system was completed without issue. Instruments were setup and verified without issues. The surgeon attached the surgical system to the patient and the first 3Define scan was attempted. The surgical arm moved to the first position, but the froze with no image visible on the monitor. The surgeon waited for a minute, but there was no change. A new 3Define scan was attempted, but the same issue occurred. The manufacturer representative checked the USB connection and the base of the cart and the surgeon screen was noted to be working. The surgeon decided to proceed with the case using a generic work volume. The surgical arm was sent to the snapshot position and the snapshot was acquired with no issue. The representative noted that the visual representation of the surgical arm on the monitor was accurate and correctly showed the physical position of the arm. The position for star marker was marked by the surgeon and the surgical arm was sent. O-arm scan and planning was completed without issue. The surgeon decided to complete the facetectomy before placing screws due to the challenging nature of the anatomy. The surgeon knew they may need a new spin if the anatomy shifted. The surgeon completed the facetectomy on both sides and wanted to complete an accuracy test before proceeding. The passive planar was placed on the divot of the surgical arm, while the arm was at right L4. The accuracy test window opened, but only the passive planar and shoulder were rendered on the monitor. The surgical arm was not showing up. A new snapshot was acquired and the surgeon placed the passive planar into the arm guide divot again when the arm was at the snapshot position. The visual representation on the monitor was different than the actual position of the surgical arm. The representative backed out of the current operation and started a new one in order to prepare for a new O-arm scan. A generic work volume was used and the surgical arm was sent to the snapshot position, but the same issue occurred. The surgeon decided to abort the use of the guidance system and place the screws freehand. After removing the system from the bed while keeping the surgical arm in the same position, the setup screen was opened to find the visual representation of the surgical arm to be correct. In a new location, the system was restarted. In the surgical arm setup window, t he shoulder was inverted compared to the actual position. The shoulder was recalibrated and the visual representation of the surgical arm was accurate. The procedure was delayed less than an hour and there was no patient harm. 2020-10-20 IUD x6 (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-10-15 00698837, 00698856, PDF, IUD x4 (Rep, HPC): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera was not functioning during the procedure. Setup of the guidance system was completed without issue. Instruments were setup and verified without issues. The surgeon attached the surgical system to the patient and the first 3Define scan was attempted. The surgical arm moved to the first position, but the froze with no image visible on the monitor. The surgeon waited for a minute, but there was no change. A new 3Define scan was attempted, but the same issue occurred. The manufacturer representative checked the USB connection and the base of the cart and the surgeon screen was noted to be working. The surgeon decided to proceed with the case using a generic work volume. The surgical arm was sent to the snapshot position and the snapshot was acquired with no issue. The representative noted that the visual representation of the surgical arm on the monitor was accurate and correctly showed the physical position of the arm. The position for star marker was marked by the surgeon and the surgical arm was sent. O-arm scan and planning was completed without issue. The surgeon decided to complete the facetectomy before placing screws due to the challenging nature of the anatomy. The surgeon knew they may need a new spin if the anatomy shifted. The surgeon completed the facetectomy on both sides and wanted to complete an accuracy test before proceeding. The passive planar was placed on the divot of the surgical arm, while the arm was at right L4. The accuracy test window opened, but only the passive planar and shoulder were rendered on the monitor. The surgical arm was not showing up. A new snapshot was acquired and the surgeon placed the passive planar into the arm guide divot again when the arm was at the snapshot position. The visual representation on the monitor was different than the actual position of the surgical arm. The representative backed out of the current operation and started a new one in order to prepare for a new O-arm scan. A generic work volume was used and the surgical arm was sent to the snapshot position, but the same issue occurred. The surgeon decided to abort the use of the guidance system and place the screws freehand. After removing the system from the bed while keeping the surgical arm in the same position, the setup screen was opened to find the visual representation of the surgical arm to be correct. In a new location, the system was restarted. In the surgical arm setup window, t he shoulder was inverted compared to the actual position. The shoulder was recalibrated and the visual representation of the surgical arm was accurate. The procedure was delayed less than an hour and there was no patient harm. 2020-10-20 IUD x6 (Other): No new information. | xse camera issuewv inaccurate pdf iud spinal procedure camera functioning procedure setup completed issue instruments setup verified issues surgeon attached surgical scan attempted surgical arm moved position froze image visible monitor surgeon waited minute change new scan attempted issue occurred checked usb connection base cart surgeon screen noted working surgeon decided proceed case using generic work volume surgical arm sent snapshot position snapshot acquired issue noted visual representation surgical arm monitor accurate correctly showed physical position arm position star marker marked surgeon surgical arm sent oarm scan planning completed issue surgeon decided complete facetectomy placing screws challenging nature anatomy surgeon knew need new spin anatomy shifted surgeon completed facetectomy sides wanted complete accuracy test proceeding passive planar placed divot surgical arm arm right accuracy test window opened passive planar shoulder rendered monitor surgical arm showing new snapshot acquired surgeon placed passive planar arm guide divot arm snapshot position visual representation monitor different actual position surgical arm backed current operation started new order prepare new oarm scan generic work volume surgical arm sent snapshot position issue occurred surgeon decided abort use place screws freehand removing bed keeping surgical arm position setup screen opened visual representation surgical arm correct new location restarted surgical arm setup window t shoulder inverted compared actual position shoulder recalibrated visual representation surgical arm accurate procedure delayed hour iud new unk mazor x camera | NAV3244 | NAV3253 | NAV4101 | NAV3239 | NAV2054 | NAV3244NAV3253NAV4101NAV3239NAV2054 | True
2506 | 703236949 | X - surgical arm issue | 2019-07-01T20:48:13.000000+0000 | MazorX | NAV4106 | CALIBRATION ISSUES | 20 | ASM0206-01 | SURGICAL ARM | X - surgical arm issue J5 comm error, J5 mechanical issue SURGICAL ARM | J5 comm error, J5 mechanical issue | x surgical arm issue comm error mechanical issue surgical arm | NAV3249 | NAV4101 | NAV4098 | NAV4095 | NAV2054 | NAV3249NAV4101NAV4098NAV4095NAV2054 | False
10466 | 704885310 | XSE - Back Handle and Door Collision | 2022-02-26T01:14:05.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 40 | MEC5416-01 | CART MAIN HANDLE | XSE - Back Handle and Door Collision 2022-Feb-25 Salesforce C00244283 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the main cart handle on the unit broke during use and the there is a a loud popping noise coming from the surgical system storage door when closing. No patient was present. 2022-02-28 IUD x2 (Other): No new information. 2022-03-04 IUD x3 (Other): No new information. 2022-03-31 IUD x2 (Other): No new information. 2022-04-28 IUD (Other): No new information. CART MAIN HANDLE | 2022-Feb-25 Salesforce C00244283 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the main cart handle on the unit broke during use and the there is a a loud popping noise coming from the surgical system storage door when closing. No patient was present. 2022-02-28 IUD x2 (Other): No new information. 2022-03-04 IUD x3 (Other): No new information. 2022-03-31 IUD x2 (Other): No new information. 2022-04-28 IUD (Other): No new information. | xse handle door collision salesforce outside procedure main cart handle unit broke use loud popping noise coming surgical storage door closing present iud new iud new iud new iud new cart main handle | NAV2058 | NAV3252 | NAV3253 | NAV2057 | NAV2047 | NAV2058NAV3252NAV3253NAV2057NAV2047 | True
6214 | 703892903 | XSE ¦ Surgeon Screen Cracked | 2020-08-07T15:27:54.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 30 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE ¦ Surgeon Screen Cracked 2020-08-06 00694191 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the wired surgeon screen is cracked and not responding correctly. No patient was present. 2020-08-10 IUD x5 (Other): No new information. 2020-08-21 IUD x2 (Other): No new information. 2020-09-07 IUD x3 (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2020-08-06 00694191 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the wired surgeon screen is cracked and not responding correctly. No patient was present. 2020-08-10 IUD x5 (Other): No new information. 2020-08-21 IUD x2 (Other): No new information. 2020-09-07 IUD x3 (Other): No new information. | xse surgeon screen cracked outside procedure wired surgeon screen cracked responding correctly present iud new iud new iud new kit svc srgn screen arm | NAV3253 | NAV2054 | NAV2047 | NAV3249 | NAV1006 | NAV3253NAV2054NAV2047NAV3249NAV1006 | True
10448 | 704882949 | XSE - Unresponsive & Worn Instruments | 2022-02-25T01:12:36.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Unresponsive & Worn Instruments 2022-Feb-24 Salesforce C00244200 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system was used for a case and when they went to unmount the system from the bed they hit the unmount button and the system stopped responding and spinning circle appeared next to the cursor. Then the system soft restarted itself. After that they were able to unmount the system. It was also reported that the target extender and C-arm image adapter screw are worn out and could not be used anymore. There was no impact on the patient outcome. 2022-02-25 IUD x2 (Other): No new information. 2022-03-09 IUD (Other): No new information. 2022-04-04 IUD (Other): No new information. 2022-04-08 IUD (Other): No new information. 2022-04-18 IUD (Other): No new information. 2022-04-20 IUD (Other): No new information. 2022-05-03 IUD (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2022-Feb-24 Salesforce C00244200 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system was used for a case and when they went to unmount the system from the bed they hit the unmount button and the system stopped responding and spinning circle appeared next to the cursor. Then the system soft restarted itself. After that they were able to unmount the system. It was also reported that the target extender and C-arm image adapter screw are worn out and could not be used anymore. There was no impact on the patient outcome. 2022-02-25 IUD x2 (Other): No new information. 2022-03-09 IUD (Other): No new information. 2022-04-04 IUD (Other): No new information. 2022-04-08 IUD (Other): No new information. 2022-04-18 IUD (Other): No new information. 2022-04-20 IUD (Other): No new information. 2022-05-03 IUD (Other): No new information. | xse unresponsive worn instruments salesforce spinal p rocedure case went unmount bed hit unmount button stopped responding spinning circle appeared cursor soft restarted able unmount target extender carm image adapter screw worn anymore impact outcome iud new iud new iud new iud new iud new iud new iud new mazorx sw installation kit | NAV3248 | NAV3299 | NAV3250 | NAV3246 | NAV3247 | NAV3248NAV3299NAV3250NAV3246NAV3247 | True
7121 | 704123074 | XSE - Damaged Instrumentation Trays | 2020-12-14T17:13:08.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 190 | KIT0899 | FRAME KIT0899 ROBOTIC FRAME CONNECTION | XSE - Damaged Instrumentation Trays 2020-12-11 00702590, e1, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instrumentation trays were damaged and the instruments failed tracking during preventative maintenance. The trajectory error ranged between 0.554 and 0.976 which was above the maximum amount allowed. Troubleshooting involved trying different spheres, trying different Navlock trackers and snapshot tracker, using long and short tools, adjusting the mounted position of the surgical system, hard restarts of the guidance system and trying different surgical arms. The stealth camera was viewing all of the spheres well during testing. The most probable cause was noted to be the washing and cleaning by the sterile processing department. The instruments were in rough condition and the edges. The edges of the trackers were chipped and the paint was faded. The trackers were difficult to detect during testing. There was no patient involvement. 2021-01-10 IUD (Other): No new information. 2021-01-10 IUD x2 (Other): No new information. 2021-02-02 email 2021-02-02 13:45, IUD (Other): No new information. 2021-02-24 IUD x2, email 2021-02-23 11:31 (Rep): No new information. 2021-03-09 IUD (Rep): Additional information received from a manufacturer representative reported that the issue was determined to be the camera. The camera was replaced and the issue was resolved. 2021-03-16 email 2021-03-16 20:45 (Rep): Additional information received from a manufacturer representative reported that all of the instruments were tested and found to be functioning as designed. FRAME KIT0899 ROBOTIC FRAME CONNECTION | 2020-12-11 00702590, e1, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the instrumentation trays were damaged and the instruments failed tracking during preventative maintenance. The trajectory error ranged between 0.554 and 0.976 which was above the maximum amount allowed. Troubleshooting involved trying different spheres, trying different Navlock trackers and snapshot tracker, using long and short tools, adjusting the mounted position of the surgical system, hard restarts of the guidance system and trying different surgical arms. The stealth camera was viewing all of the spheres well during testing. The most probable cause was noted to be the washing and cleaning by the sterile processing department. The instruments were in rough condition and the edges. The edges of the trackers were chipped and the paint was faded. The trackers were difficult to detect during testing. There was no patient involvement. 2021-01-10 IUD (Other): No new information. 2021-01-10 IUD x2 (Other): No new information. 2021-02-02 email 2021-02-02 13:45, IUD (Other): No new information. 2021-02-24 IUD x2, email 2021-02-23 11:31 (Rep): No new information. 2021-03-09 IUD (Rep): Additional information received from a manufacturer representative reported that the issue was determined to be the camera. The camera was replaced and the issue was resolved. 2021-03-16 email 2021-03-16 20:45 (Rep): Additional information received from a manufacturer representative reported that all of the instruments were tested and found to be functioning as designed. | xse damaged instrumentation trays iud outside procedure instrumentation trays damaged instruments failed tracking preventative maintenance trajectory error ranged maximum allowed troubleshooting involved trying different spheres trying different navlock trackers snapshot tracker using long short tools adjusting mounted position surgical hard restarts trying different surgical arms stealth camera viewing spheres testing probable cause noted washing cleaning sterile processing department instruments rough condition edges edges trackers chipped paint faded trackers difficult detect testing iud new iud new email iud new iud email new iud additional issue determined camera camera replaced issue resolved email additional instruments tested functioning designed frame robotic frame connection | NAV3299 | NAV3244 | NAV3295 | NAV3249 | NAV3262 | NAV3299NAV3244NAV3295NAV3249NAV3262 | True
5095 | 703597189 | XSE - CANBUS Error | 2020-02-04T20:12:53.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CANBUS Error 2020-02-04 00683379 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that following mounting the shoulder, a CANBUS error was seen immediately. The case was able to continue but troubleshooting was needed to remove the system from the bed. First technical services had the manufacturer representative log into superuser and go into service tools. All status bars were green except for positioner IO 10 and 11. However, this is a type I system so red for 11 is normal, but red for 10 does not. The representative was instructed to reset the movement and send the arm to a predefined position. Then they were instructed to go to the general tab and try to open the shoulder, positioner y and z. Positioner y opened, but the shoulder and z did not. When attempting to open the shoulder through the overrides panel, it again did not open. The right side of the panel was removed so the electric panel at the breaker could be seen. One of the breakers popped, so the system was powered off and the breaker was reset. After waiting 5 mins, the system was powered on again and it was able to be used. It is likely that there is a short in the positioner and a FSE will be sent to place a replacement surgical system. There was no patient harm and the procedure was not delayed over an hour. 2020-02-06 IUD x2 (Rep): No new information. UNK MAZOR X WORKSTATION | 2020-02-04 00683379 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that following mounting the shoulder, a CANBUS error was seen immediately. The case was able to continue but troubleshooting was needed to remove the system from the bed. First technical services had the manufacturer representative log into superuser and go into service tools. All status bars were green except for positioner IO 10 and 11. However, this is a type I system so red for 11 is normal, but red for 10 does not. The representative was instructed to reset the movement and send the arm to a predefined position. Then they were instructed to go to the general tab and try to open the shoulder, positioner y and z. Positioner y opened, but the shoulder and z did not. When attempting to open the shoulder through the overrides panel, it again did not open. The right side of the panel was removed so the electric panel at the breaker could be seen. One of the breakers popped, so the system was powered off and the breaker was reset. After waiting 5 mins, the system was powered on again and it was able to be used. It is likely that there is a short in the positioner and a FSE will be sent to place a replacement surgical system. There was no patient harm and the procedure was not delayed over an hour. 2020-02-06 IUD x2 (Rep): No new information. | xse canbus error spinal procedure following mounting shoulder canbus error seen immediately case able continue troubleshooting needed remove bed technical services log superuser service tools status bars green positioner io type red normal red does instructed reset movement send arm predefined position instructed general tab try open shoulder positioner y z positioner y opened shoulder z did attempting open shoulder overrides panel did open right panel removed electric panel breaker seen breakers popped powered breaker reset waiting mins powered able likely short positioner fse sent place replacement surgical procedure delayed hour iud new unk mazor x workstation | NAV4107 | NAV3263 | NAV2047 | NAV3251 | NAV4106 | NAV4107NAV3263NAV2047NAV3251NAV4106 | True
5162 | 703608924 | XSE - SRA Failed Acc./Stress Test | 2020-02-12T00:03:30.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - SRA Failed Acc./Stress Test 2020-02-11 00683883 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the surgical arm failed the advanced accuracy test and stress test during the stealth upgrade. There was no patient involvement. 2020-02-12 IUD x4 (Other): No new information. 2020-03-19 IUD x4 (Other): No new information. 2020-04-19 IUD (Other): No new information. 2020-05-18 IUD (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-02-11 00683883 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the surgical arm failed the advanced accuracy test and stress test during the stealth upgrade. There was no patient involvement. 2020-02-12 IUD x4 (Other): No new information. 2020-03-19 IUD x4 (Other): No new information. 2020-04-19 IUD (Other): No new information. 2020-05-18 IUD (Other): No new information. | xse sra failed accstress test outside procedure surgical arm failed advanced accuracy test stress test stealth upgrade iud new iud new iud new iud new kit svc surgical arm | NAV3249 | NAV4093 | NAV4092 | NAV2054 | NAV1006 | NAV3249NAV4093NAV4092NAV2054NAV1006 | True
654 | 703169638 | XSE - error message | 2019-05-23T15:28:37.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | XSE - error message 2019-May-17 SCL1901264 intake form, msg, jpg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they mounted the device to the bed and sent the arm to clear up then to drape. On the way to drape, twice it gave the error 3836. A soft restart was performed and sent arm to clear up and then to drape and received the same error twice on the way to drape. In inter-op, they sent arm to drape after 3define and did not get the error. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-May-17 SCL1901264 intake form, msg, jpg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they mounted the device to the bed and sent the arm to clear up then to drape. On the way to drape, twice it gave the error 3836. A soft restart was performed and sent arm to clear up and then to drape and received the same error twice on the way to drape. In inter-op, they sent arm to drape after 3define and did not get the error. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse error message msg jpg spinal procedure mounted bed sent arm clear drape way drape twice gave error soft restart performed sent arm clear drape error twice way drape interop sent arm drape did error additional complications anticipated surgical arm | NAV3232 | NAV4098 | NAV3249 | NAV2054 | NAV4095 | NAV3232NAV4098NAV3249NAV2054NAV4095 | True
6235 | 703893960 | XSE - Positioner Dropped | 2020-08-07T23:28:21.000000+0000 | MazorX Stealth Edition | NAV4105 | UNEXPECTED POSITIONER MOVEMENT | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | XSE - Positioner Dropped 2020-08-06 00694243 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while the manufacturer representative was putting the system away in the hood, the positioner dropped slightly and became un-homed in the z-axis. He had to get it back onto the bed to get it into the home positioner again. He was then able to store the system in the hood. No patient was present. 2020-08-14 IUD (Rep): Additional information received from a manufacturer representative reported that the following case was comple ted without issue. UNK MAZOR X POSITIONER | 2020-08-06 00694243 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while the manufacturer representative was putting the system away in the hood, the positioner dropped slightly and became un-homed in the z-axis. He had to get it back onto the bed to get it into the home positioner again. He was then able to store the system in the hood. No patient was present. 2020-08-14 IUD (Rep): Additional information received from a manufacturer representative reported that the following case was comple ted without issue. | xse positioner dropped outside procedure putting away hood positioner dropped slightly unhomed zaxis bed home positioner able store hood present iud additional following case comple ted issue unk mazor x positioner | NAV4105 | NAV3251 | NAV4104 | NAV2058 | NAV4101 | NAV4105NAV3251NAV4104NAV2058NAV4101 | True
1528 | 703205006 | X - SURGICAL ARM ISSUE | 2019-06-13T11:23:11.000000+0000 | MazorX | NAV4106 | CALIBRATION ISSUES | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - SURGICAL ARM ISSUE It was reported that they received a fault on shoulder rear encoder. UNK MAZOR X SHOULDER | It was reported that they received a fault on shoulder rear encoder. | x surgical arm issue fault shoulder rear encoder unk mazor x shoulder | NAV4101 | NAV3244 | NAV2054 | NAV4106 | NAV4107 | NAV4101NAV3244NAV2054NAV4106NAV4107 | True
8277 | 704384892 | X - Multiple System Issues | 2021-05-09T11:48:15.000000+0000 | MazorX | NAV4099 | PLASTIC COVER ISSUES | 20 | MEC4654-01 | J6 - INTEL COVER | X - Multiple System Issues 2021-May-09 Salesforce C00226027, IUD, image 2021-05-11 10:42 x11 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there were multiple issues with the guidance system. There was an issue with the paint, the pinch sticker needed to be replaced, the 3Define cover needed to be replaced as it was attached with tape, the mouse was scratched and the main screen was loose. The manufacturer representative tried to tighten the screen, but the screws were loose on the lower main screen assembly. There was no patient involvement. 2021-05-12 IUD (Other, For): No new information. 2021-06-01 IUD (Other, For): no new information. 2021-07-05 email 2021-07-05 02:46 (Rep, For): No new information. 2021-07-11 IUD x2 (Other, For): No new information. 2021-07-28 email 2021-07-28 02:16 (Other, For): No new information. 2021-08-26 email 2021-08-26 09:31 (Other, For): No new information. 2021-09-12 IUD (Other, For): No new information. 2022-05-12 IUD (Other): No new information. J6 - INTEL COVER | 2021-May-09 Salesforce C00226027, IUD, image 2021-05-11 10:42 x11 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there were multiple issues with the guidance system. There was an issue with the paint, the pinch sticker needed to be replaced, the 3Define cover needed to be replaced as it was attached with tape, the mouse was scratched and the main screen was loose. The manufacturer representative tried to tighten the screen, but the screws were loose on the lower main screen assembly. There was no patient involvement. 2021-05-12 IUD (Other, For): No new information. 2021-06-01 IUD (Other, For): no new information. 2021-07-05 email 2021-07-05 02:46 (Rep, For): No new information. 2021-07-11 IUD x2 (Other, For): No new information. 2021-07-28 email 2021-07-28 02:16 (Other, For): No new information. 2021-08-26 email 2021-08-26 09:31 (Other, For): No new information. 2021-09-12 IUD (Other, For): No new information. 2022-05-12 IUD (Other): No new information. | x multiple issues salesforce iud image outside procedure multiple issues issue paint pinch sticker needed replaced cover needed replaced attached tape mouse scratched main screen loose tried tighten screen screws loose lower main screen assembly iud new iud new email new iud new email new email new iud new iud new intel cover | NAV4099 | NAV2058 | NAV2054 | NAV3299 | NAV2057 | NAV4099NAV2058NAV2054NAV3299NAV2057 | True
8959 | 704515141 | XSE - Navigation Inaccuracy | 2021-07-22T00:11:50.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 20 | M450001B020 | NAVIGATION SOFTWARE MODULE 1.1 | XSE - Navigation Inaccuracy 2021-Jul-21 Salesforce C00230587, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the instruments appeared inaccurate during a L5-S1 TLIF case. The screws on the right side were placed with no issues. When the surgeon moved to the left side, navigation was noticed to be off. Another snapshot was done and the issue was resolved. There was no impact on the screw accuracy and the case was completed successfully. There was no patient harm and the procedure was delayed less than an hour. 2021-08-20 IUD x2 (Other): No new information. 2021-09-12 email 2021-09-12 21:45 (Rep): No new information. 2021-09-14 email 2021-09-14 14:01 (Rep): Additional information was received stating that the deviation was less than 3.5mm and it may have been due to bad spheres or navlocks. NAVIGATION SOFTWARE MODULE 1.1 | 2021-Jul-21 Salesforce C00230587, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the instruments appeared inaccurate during a L5-S1 TLIF case. The screws on the right side were placed with no issues. When the surgeon moved to the left side, navigation was noticed to be off. Another snapshot was done and the issue was resolved. There was no impact on the screw accuracy and the case was completed successfully. There was no patient harm and the procedure was delayed less than an hour. 2021-08-20 IUD x2 (Other): No new information. 2021-09-12 email 2021-09-12 21:45 (Rep): No new information. 2021-09-14 email 2021-09-14 14:01 (Rep): Additional information was received stating that the deviation was less than 3.5mm and it may have been due to bad spheres or navlocks. | xse navigation inaccuracy salesforce iud spinal procedure instruments appeared inaccurate tlif case screws right placed issues surgeon moved left navigation noticed snapshot issue resolved impact screw accuracy case completed successfully procedure delayed hour iud new email new email additional stating deviation bad spheres navlocks navigation software module | NAV5001 | NAV5000 | NAV3299 | NAV2056 | NAV1012 | NAV5001NAV5000NAV3299NAV2056NAV1012 | True
5457 | 703714994 | XSE - Clamp Link Screw Stripped | 2020-04-17T18:33:16.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5206-01S | KIT SVC MEC5206-01S RNX-ACC-021 | XSE - Clamp Link Screw Stripped 2020-APR-17 00687553 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the clamp link screw was stripped during the in-service. It was noted that it had never been used for a case. There was no patient present when this issue was identified. 2020-04-20 email 2020-04-20 10:00, IUD (Other): No new information. 2020-05-13 Email (Rep): No new information. 2020-05-20 email 2020-05-20 01:30 x2, email 2020-05-20 17:00 (Rep): No new information. 2020-06-03 email (Rep): No new information 2020-06-09 IUD x2 (Other): No new information. 2020-06-24 IUD (Other): No new information. 20200-7-06 IUD (Other): No new information. KIT SVC MEC5206-01S RNX-ACC-021 | 2020-APR-17 00687553 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the clamp link screw was stripped during the in-service. It was noted that it had never been used for a case. There was no patient present when this issue was identified. 2020-04-20 email 2020-04-20 10:00, IUD (Other): No new information. 2020-05-13 Email (Rep): No new information. 2020-05-20 email 2020-05-20 01:30 x2, email 2020-05-20 17:00 (Rep): No new information. 2020-06-03 email (Rep): No new information 2020-06-09 IUD x2 (Other): No new information. 2020-06-24 IUD (Other): No new information. 20200-7-06 IUD (Other): No new information. | xse clamp link screw stripped outside procedure clamp link screw stripped inservice noted case present issue identified email iud new email new email email new email new iud new iud new iud new kit svc | NAV3299 | NAV4090 | NAV2054 | NAV3295 | NAV3246 | NAV3299NAV4090NAV2054NAV3295NAV3246 | True
309 | 703076180 | XSE - Software Crash | 2019-03-27T12:36:59.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Software Crash 2019-03-20 SCL1900721, Intake Form, Image 1, Image 2, e1, e2, e3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system crashed when an AP image was attempting to be sent to the system. Prior to the procedure, the manufacturer representative thoroughly checked the system and everything was functioning correctly. During the procedure, the arm and software froze on the first trajectory during the surface scan. The manufacturer representative attempted to stop the scan and restart it, but the same issue occurred. The arm was then sent to the clear up position, but the same issue occurred when starting the scan. A soft restart was done and the camera started to work. The scan was able to be completed. Later in the procedure, the software completely blacked out and crashed when accepting an AP image. The system did not attempt to restart like it normally would so a manual hard restart was done. The system prompted the manufacturer representative for all of the codes and setup. The surgeon agreed to start the case from the beginning. Upon the second attempt for the case, the same issue occurred and a second hard restart was done. There was no patient harm, but the surgery was delayed more than an hour. The case was not completed. After the procedure, the field engineer was able to use the system to successfully complete four demo cases. Information unrelated to this event was included in three separate PEs; See PE 703076217 for XSE - Slow Performance/SW Bug, 703066185 for XSE - No Pressure, and 703076216 for XSE - Camera Issues. 2019-03-31 e5 (Rep): Additional information received from a manufacturer representative reported that troubleshooting was done and the issue was able to be recreated. The representative was confident the issues were related to the C-arm. 2019-04-22 e7 (Rep): Additional information received from a manufacturer representative reported the issue had been resolved and no further information was available. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-03-20 SCL1900721, Intake Form, Image 1, Image 2, e1, e2, e3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system crashed when an AP image was attempting to be sent to the system. Prior to the procedure, the manufacturer representative thoroughly checked the system and everything was functioning correctly. During the procedure, the arm and software froze on the first trajectory during the surface scan. The manufacturer representative attempted to stop the scan and restart it, but the same issue occurred. The arm was then sent to the clear up position, but the same issue occurred when starting the scan. A soft restart was done and the camera started to work. The scan was able to be completed. Later in the procedure, the software completely blacked out and crashed when accepting an AP image. The system did not attempt to restart like it normally would so a manual hard restart was done. The system prompted the manufacturer representative for all of the codes and setup. The surgeon agreed to start the case from the beginning. Upon the second attempt for the case, the same issue occurred and a second hard restart was done. There was no patient harm, but the surgery was delayed more than an hour. The case was not completed. After the procedure, the field engineer was able to use the system to successfully complete four demo cases. Information unrelated to this event was included in three separate PEs; See PE 703076217 for XSE - Slow Performance/SW Bug, 703066185 for XSE - No Pressure, and 703076216 for XSE - Camera Issues. 2019-03-31 e5 (Rep): Additional information received from a manufacturer representative reported that troubleshooting was done and the issue was able to be recreated. The representative was confident the issues were related to the C-arm. 2019-04-22 e7 (Rep): Additional information received from a manufacturer representative reported the issue had been resolved and no further information was available. | xse software crash image image spinal procedure crashed ap image attempting sent prior procedure thoroughly checked functioning correctly procedure arm software froze trajectory surface scan attempted stop scan restart issue occurred arm sent clear position issue occurred starting scan soft restart camera started work scan able completed later procedure software completely blacked crashed accepting ap image did attempt restart like normally manual hard restart prompted codes setup surgeon agreed start case beginning second attempt case issue occurred second hard restart surgery delayed hour case completed procedure field engineer able use successfully complete demo cases unrelated included separate pes pe xse slow performancesw bug xse pressure xse camera issues additional troubleshooting issue able recreated confident issues related carm additional issue resolved available mazor x software installation kit | NAV3247 | NAV4095 | NAV3244 | NAV5000 | NAV3246 | NAV3247NAV4095NAV3244NAV5000NAV3246 | False
3573 | 703378361 | XSE - Arm Failed Advanced Accuracy Test | 2019-09-24T16:56:06.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Arm Failed Advanced Accuracy Test 2019-09-24 SCL1902883, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported hat planning for the procedure was completed by the surgeon. The surgeon first completed a L3-S1 OLIF and then flipped the patient to a prone position. The surgical system was mounted to the patient using the Schanz bond and BMB connected to a Schanz pin inserted into the right PSIS. 3Define scan and Draw Spine were successfully completed. Registration was completed on the first attempt using auto registration and the basic algorithm. The surgeon decided to drill all of the trajectories first prior to placing screws. All of the left side trajectories were drilled starting at the top followed by the right side trajectories. Following drilling of the last trajectory, the surgeon began placing screws working backwards in drilling order. When inserting the screw at right L3, the screw appeared to be inferior to the prior drilled hole. Navigation showed the trajectory to be inferior to plan. A platform shift was suspected and a new registration was done with the towers and screws already at right L4-S1. Registration was unsuccessful after multiple attempts so the screws were removed in order to achieve registration. A third registration was performed quickly using the auto registration and basic algorithm. The surgical arm was sent to all trajectories to place k-wires and then screws. After the procedure, the surgical arm failed an advanced accuracy test. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. 2019-09-24 email 2019-09-24 12:30 (Rep): Additional information received from a manufacturer representative reported that during the surgical procedure the surgeon noted the surgical arm wrist was slightly loose. By wrist, the representative was referring to the point where the arm guide attached to the arm. The representative was waiting for post-op CT to confirm accuracy of the screws compared to the plan. Final AP and lateral fluoro images appeared to show L4-S1 screws were accurately placed, but L3 trajectories were slightly inferior to plan. During the procedure, the surgeon's glove got stuck between the tracker and the drill connection on two separate occasions. The representative also mentioned the long cannula was stuck in the arm guide. The surgeon had to apply excessive force to the cannula in order to remove it from the arm guide. These events and the wrist being loose, led the representative to perform a 10 point accuracy test after the procedure. MAZOR X SYSTEM TPL0059 | 2019-09-24 SCL1902883, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with fixation. It was reported hat planning for the procedure was completed by the surgeon. The surgeon first completed a L3-S1 OLIF and then flipped the patient to a prone position. The surgical system was mounted to the patient using the Schanz bond and BMB connected to a Schanz pin inserted into the right PSIS. 3Define scan and Draw Spine were successfully completed. Registration was completed on the first attempt using auto registration and the basic algorithm. The surgeon decided to drill all of the trajectories first prior to placing screws. All of the left side trajectories were drilled starting at the top followed by the right side trajectories. Following drilling of the last trajectory, the surgeon began placing screws working backwards in drilling order. When inserting the screw at right L3, the screw appeared to be inferior to the prior drilled hole. Navigation showed the trajectory to be inferior to plan. A platform shift was suspected and a new registration was done with the towers and screws already at right L4-S1. Registration was unsuccessful after multiple attempts so the screws were removed in order to achieve registration. A third registration was performed quickly using the auto registration and basic algorithm. The surgical arm was sent to all trajectories to place k-wires and then screws. After the procedure, the surgical arm failed an advanced accuracy test. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. 2019-09-24 email 2019-09-24 12:30 (Rep): Additional information received from a manufacturer representative reported that during the surgical procedure the surgeon noted the surgical arm wrist was slightly loose. By wrist, the representative was referring to the point where the arm guide attached to the arm. The representative was waiting for post-op CT to confirm accuracy of the screws compared to the plan. Final AP and lateral fluoro images appeared to show L4-S1 screws were accurately placed, but L3 trajectories were slightly inferior to plan. During the procedure, the surgeon's glove got stuck between the tracker and the drill connection on two separate occasions. The representative also mentioned the long cannula was stuck in the arm guide. The surgeon had to apply excessive force to the cannula in order to remove it from the arm guide. These events and the wrist being loose, led the representative to perform a 10 point accuracy test after the procedure. | xse arm failed advanced accuracy test ccr spinal procedure treat degenerative disc fixation hat planning procedure completed surgeon surgeon completed olif flipped prone position surgical mounted using schanz bond bmb connected schanz pin inserted right psis scan draw spine successfully completed registration completed attempt using auto registration basic algorithm surgeon decided drill trajectories prior placing screws left trajectories drilled starting followed right trajectories following drilling trajectory surgeon began placing screws working backwards drilling order inserting screw right screw appeared inferior prior drilled hole navigation showed trajectory inferior plan platform shift suspected new registration towers screws right registration unsuccessful multiple attempts screws removed order achieve registration registration performed quickly using auto registration basic algorithm surgical arm sent trajectories place kwires screws procedure surgical arm failed advanced accuracy test procedure delayed hour case completed email additional surgical procedure surgeon noted surgical arm wrist slightly loose wrist referring point arm guide attached arm waiting postop ct confirm accuracy screws compared plan final ap lateral fluoro images appeared screws accurately placed trajectories slightly inferior plan procedure surgeons glove got stuck tracker drill connection separate occasions mentioned long cannula stuck arm guide surgeon apply excessive force cannula order remove arm guide events wrist loose led perform point accuracy test procedure mazor x | NAV5000 | NAV4093 | NAV3299 | NAV3243 | NAV4097 | NAV5000NAV4093NAV3299NAV3243NAV4097 | True
8469 | 704420437 | X - SRA Movement Error | 2021-05-28T06:58:12.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 30 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | X - SRA Movement Error 2021-05-28 00713990, IUD (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after moving the surgical arm to right S2 and placing the cannula down the arm guide, the system displayed a message that the arm was not on trajectory. The surgeon was holding the cannula and no force was applied to the surgical arm. The surgical arm was resent to the right S2 trajectory, and the arm could not reach the trajectory. When sending the surgical arm to the clear up position, CANBUS error 295 and 8798 were displayed. A soft reboot of the system did not resolve the issue so a hard reboot was done. After the hard reboot, the surgical arm was able to be moved, but Error message 295 was displayed. There was no patient harm and the procedure was delayed less than an hour. 2021-08-04 IUD x3 (Other, For): No new information. 2021-09-02 email 2021-09-02 08:31 (Other, For): No new information. 2021-09-10 email 2021-09-10 05:45 (Other, For): No new information. 2021-09-14 email 2021-09-14 13:45, email 2021-09-15 01:46 (Rep, For): No new information. 2021-09-17 email 2021-09-17 04:01, IUD (Rep, For): Additional information received from a manufacturer representative reported that the main cable was replaced to resolve the communication issues. The faulty cable affected the surgical arm so both the arm and cable were replaced. 2021-01-13 IUD (Other, For): No new information. 2022-05-09 email 2022-05-09 05:16 (Rep): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2021-05-28 00713990, IUD (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after moving the surgical arm to right S2 and placing the cannula down the arm guide, the system displayed a message that the arm was not on trajectory. The surgeon was holding the cannula and no force was applied to the surgical arm. The surgical arm was resent to the right S2 trajectory, and the arm could not reach the trajectory. When sending the surgical arm to the clear up position, CANBUS error 295 and 8798 were displayed. A soft reboot of the system did not resolve the issue so a hard reboot was done. After the hard reboot, the surgical arm was able to be moved, but Error message 295 was displayed. There was no patient harm and the procedure was delayed less than an hour. 2021-08-04 IUD x3 (Other, For): No new information. 2021-09-02 email 2021-09-02 08:31 (Other, For): No new information. 2021-09-10 email 2021-09-10 05:45 (Other, For): No new information. 2021-09-14 email 2021-09-14 13:45, email 2021-09-15 01:46 (Rep, For): No new information. 2021-09-17 email 2021-09-17 04:01, IUD (Rep, For): Additional information received from a manufacturer representative reported that the main cable was replaced to resolve the communication issues. The faulty cable affected the surgical arm so both the arm and cable were replaced. 2021-01-13 IUD (Other, For): No new information. 2022-05-09 email 2022-05-09 05:16 (Rep): No new information. | x sra movement error iud spinal procedure moving surgical arm right placing cannula arm guide displayed message arm trajectory surgeon holding cannula force applied surgical arm surgical arm resent right trajectory arm reach trajectory sending surgical arm clear position canbus error displayed soft reboot did resolve issue hard reboot hard reboot surgical arm able moved error message displayed procedure delayed hour iud new email new email new email email new email iud additional main cable replaced resolve communication issues faulty cable affected surgical arm arm cable replaced iud new email new svc main cable | NAV2047 | NAV4095 | NAV3256 | NAV4098 | NAV1006 | NAV2047NAV4095NAV3256NAV4098NAV1006 | True
2327 | 703232925 | X - Accessory Issue | 2019-06-28T13:38:51.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | MAS1355-03 | DUAL CLAMP | X - Accessory Issue While setting up Mazor dual clamp scrub tech said the dual clamp snapped like a pretzel - Picture attached. DUAL CLAMP | While setting up Mazor dual clamp scrub tech said the dual clamp snapped like a pretzel - Picture attached. | x accessory issue setting mazor dual clamp scrub tech said dual clamp snapped like pretzel picture attached dual clamp | NAV3295 | NAV4091 | NAV4090 | NAV3246 | NAV3299 | NAV3295NAV4091NAV4090NAV3246NAV3299 | True
2251 | 703230939 | X - Accessory issue | 2019-06-27T15:24:26.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | X - Accessory issue It was reported that the knife blade dislodged in patient after surgeon made a twisting motion with the handle below the patient's skin level. Able to retrieve the blade after a few minutes. UNK MAZOR X ACCESSORIES | It was reported that the knife blade dislodged in patient after surgeon made a twisting motion with the handle below the patient's skin level. Able to retrieve the blade after a few minutes. | x accessory issue knife blade dislodged surgeon twisting motion handle patients skin level able retrieve blade minutes unk mazor x accessories | NAV3299 | NAV3295 | NAV2044 | NAV3260 | NAV2058 | NAV3299NAV3295NAV2044NAV3260NAV2058 | True
540 | 703140661 | X - low trajectories | 2019-05-07T13:32:47.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - low trajectories 2019-May-02 SCL1901059 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the HCP performed an L3-S1 fusion using the Mazor X system. Due to other cases lately at the account with inaccuracies, the HCP decided to perform an O-Arm intraoperative spin after placing guidewires and placing screws. After each pedicle was drilled and guidewires were placed it was determined via an O-Arm spin that every single trajectory was low in the pedicle and approximately 3mm off from their planned trajectories. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE KIT VER-4.0.0 | 2019-May-02 SCL1901059 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the HCP performed an L3-S1 fusion using the Mazor X system. Due to other cases lately at the account with inaccuracies, the HCP decided to perform an O-Arm intraoperative spin after placing guidewires and placing screws. After each pedicle was drilled and guidewires were placed it was determined via an O-Arm spin that every single trajectory was low in the pedicle and approximately 3mm off from their planned trajectories. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | x low trajectories spinal procedure hcp performed fusion using mazor x cases lately account inaccuracies hcp decided perform oarm intraoperative spin placing guidewires placing screws pedicle drilled guidewires placed determined oarm spin single trajectory low pedicle approximately planned trajectories additional complications anticipated mazor x software kit | NAV5000 | NAV1006 | NAV3242 | NAV3235 | NAV3247 | NAV5000NAV1006NAV3242NAV3235NAV3247 | True
7435 | 704203982 | XSE - Error 4059 Arm Shift | 2021-01-30T01:38:55.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 30 | ASM0206-01 | SURGICAL ARM | XSE - Error 4059 Arm Shift 2021-Jan-29 Salesforce C00220197, intake form, e1 (rep, hcp): Medtronic received information regarding a guidance system. It was reported that during system setup accuracy test prior to the procedure, the arm reached accuracy point and moved into clear up to finish accuracy check movement, and it failed before reaching final point with Error 4059 – “Arm Shit detected in joint 2. Maintenance action is required. Stop operation.” They performed soft restart and upon restart, the same message Error 4059 popped up. Hard restart performed. Trouble shooting was performed using Service Tools – Reset Movement to move arm to predefined location (Clear up). The arm reached this destination. They performed Advanced Accuracy test which it passed. They then performed stress test which it failed 19 second into the stress test with Error 4059 – Joint 2 failure. They reset movement and sent arm to Clear Up. The arm failed several times with same error message about Joint 2. Joint 2 behaved abnormally: moving varying distances before failing, moving backwards then forwards and then failing, and moving *past* the Clear Up position. Final attempt to trouble shoot was with the Field Motor Calibration of Joint 2 which it failed several times before completion. Robotic Fusion case was aborted. Surgeon and patient decided to reschedule for when the robotic arm is fixed. Next surgery is scheduled for Wednesday, February 3rd. There was no patient harm reported. 2021-02-02 IUD x5 (Other): No new information. 2021-02-22 IUD (Other): No new information. 2021-04-19 IUD, 302404134 (Other): No new information. 2021-04-22 IUD x6 (Other): No new information. 2021-04-27 IUD x2 (Other): No new information. 2021-06-24 IUD x3 (Other): No new information. SURGICAL ARM | 2021-Jan-29 Salesforce C00220197, intake form, e1 (rep, hcp): Medtronic received information regarding a guidance system. It was reported that during system setup accuracy test prior to the procedure, the arm reached accuracy point and moved into clear up to finish accuracy check movement, and it failed before reaching final point with Error 4059 – “Arm Shit detected in joint 2. Maintenance action is required. Stop operation.” They performed soft restart and upon restart, the same message Error 4059 popped up. Hard restart performed. Trouble shooting was performed using Service Tools – Reset Movement to move arm to predefined location (Clear up). The arm reached this destination. They performed Advanced Accuracy test which it passed. They then performed stress test which it failed 19 second into the stress test with Error 4059 – Joint 2 failure. They reset movement and sent arm to Clear Up. The arm failed several times with same error message about Joint 2. Joint 2 behaved abnormally: moving varying distances before failing, moving backwards then forwards and then failing, and moving *past* the Clear Up position. Final attempt to trouble shoot was with the Field Motor Calibration of Joint 2 which it failed several times before completion. Robotic Fusion case was aborted. Surgeon and patient decided to reschedule for when the robotic arm is fixed. Next surgery is scheduled for Wednesday, February 3rd. There was no patient harm reported. 2021-02-02 IUD x5 (Other): No new information. 2021-02-22 IUD (Other): No new information. 2021-04-19 IUD, 302404134 (Other): No new information. 2021-04-22 IUD x6 (Other): No new information. 2021-04-27 IUD x2 (Other): No new information. 2021-06-24 IUD x3 (Other): No new information. | xse error arm shift salesforce setup accuracy test prior procedure arm reached accuracy point moved clear finish accuracy check movement failed reaching final point error – “arm shit detected joint maintenance action required stop operation” performed soft restart restart message error popped hard restart performed trouble shooting performed using service tools – reset movement arm predefined location arm reached destination performed advanced accuracy test passed performed stress test failed second stress test error – joint failure reset movement sent arm clear arm failed times error message joint joint behaved abnormally moving varying distances failing moving backwards forwards failing moving past clear position final attempt trouble shoot field motor calibration joint failed times completion robotic fusion case aborted surgeon decided reschedule robotic arm fixed surgery scheduled wednesday february iud new iud new iud new iud new iud new iud new surgical arm | NAV4098 | NAV4092 | NAV4101 | NAV4095 | NAV2058 | NAV4098NAV4092NAV4101NAV4095NAV2058 | True
3084 | 703264237 | X - Short in Arm / OOBF | 2019-07-17T20:04:30.000000+0000 | MazorX | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0206-01 | SURGICAL ARM | X - Short in Arm / OOBF Derrick put arm SRA0072 onto the system. This caused a loss to node 7 and breaker 5A 36V was tripped. Once the arm was removed connection to node 7 was restored. Believed to have a short in the arm./ The first arm I received had a short in it tripping the 36V breaker and did not have connection to the xeye drive (node 7) when the arm was attached. Upon removing the arm the node was connected again. I put the old arm back on and it was functioning fine again. The green LED that is constantly on when the system is on was also not working. I tried the old LED button but the result was the same. No wires were seen loose inside of the botton assembly or in the LEMO. This was an out of the box failure in the field. It also had multiple draping plastic pieces still inside the screw holes. The customer was very upset that the arm did not work upon arrival and wanted to know why so many arms have been replaced for their site. SURGICAL ARM | Derrick put arm SRA0072 onto the system. This caused a loss to node 7 and breaker 5A 36V was tripped. Once the arm was removed connection to node 7 was restored. Believed to have a short in the arm./ The first arm I received had a short in it tripping the 36V breaker and did not have connection to the xeye drive (node 7) when the arm was attached. Upon removing the arm the node was connected again. I put the old arm back on and it was functioning fine again. The green LED that is constantly on when the system is on was also not working. I tried the old LED button but the result was the same. No wires were seen loose inside of the botton assembly or in the LEMO. This was an out of the box failure in the field. It also had multiple draping plastic pieces still inside the screw holes. The customer was very upset that the arm did not work upon arrival and wanted to know why so many arms have been replaced for their site. | x short arm oobf derrick arm caused loss node breaker tripped arm removed connection node restored believed short arm arm short tripping breaker did connection xeye drive arm attached removing arm node connected old arm functioning fine green led constantly working tried old led button wires seen loose inside botton assembly lemo box failure field multiple draping plastic pieces inside screw holes customer upset arm did work arrival wanted know arms replaced site surgical arm | NAV3259 | NAV2054 | NAV3249 | NAV4100 | NAV3263 | NAV3259NAV2054NAV3249NAV4100NAV3263 | True
1866 | 703220276 | X - Software Issue | 2019-06-21T11:24:16.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X - Software Issue two software crashed backing the system out of operate mode and asking of a kit code to be entered MAZOR X SOFTWARE KIT VER-2.3 | two software crashed backing the system out of operate mode and asking of a kit code to be entered | x software issue software crashed backing operate mode asking kit code entered mazor x software kit | NAV3246 | NAV3247 | NAV4091 | NAV3248 | NAV3242 | NAV3246NAV3247NAV4091NAV3248NAV3242 | True
10408 | 704874771 | XSE - Touch Screen Not Functioning | 2022-02-21T17:07:38.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - Touch Screen Not Functioning 2022-Feb-21 Salesforce C00243955 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the touch functionality of the surgeon screen was not working. There was no patient involvement. 2022-02-25 IUD (Other): No new information. 2022-03-10 IUD (Other); No new information. 2022-03-24 IUD (Other): No new information. 2022-04-05 IUD x2 (Other): No new information. SURGEON SCREEN | 2022-Feb-21 Salesforce C00243955 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the touch functionality of the surgeon screen was not working. There was no patient involvement. 2022-02-25 IUD (Other): No new information. 2022-03-10 IUD (Other); No new information. 2022-03-24 IUD (Other): No new information. 2022-04-05 IUD x2 (Other): No new information. | xse touch screen functioning salesforce outside procedure touch functionality surgeon screen working iud new iud new iud new iud new surgeon screen | NAV3253 | NAV3244 | NAV3249 | NAV2058 | NAV3261 | NAV3253NAV3244NAV3249NAV2058NAV3261 | True
9301 | 704605821 | XSE - Registration Issue/Shift | 2021-09-12T09:55:34.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Issue/Shift 2021-Sep-02 Salesforce C00234291 (rep, hcp, for): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after one failed registration with the O-Arm due to not recognizing the star marker, they were able to successfully register the anatomy with the O-Arm using a higher dose. But when they placed the dilator on the anatomy, there was a clear shift of at least 2 centimeters. The surgeon decided to perform the surgery freehand. There was no patient harm and a delay of less than one hour. MAZOR X SYSTEM TPL0059 | 2021-Sep-02 Salesforce C00234291 (rep, hcp, for): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after one failed registration with the O-Arm due to not recognizing the star marker, they were able to successfully register the anatomy with the O-Arm using a higher dose. But when they placed the dilator on the anatomy, there was a clear shift of at least 2 centimeters. The surgeon decided to perform the surgery freehand. There was no patient harm and a delay of less than one hour. | xse registration issueshift salesforce spinal procedure failed registration oarm recognizing star marker able successfully register anatomy oarm using higher dose placed dilator anatomy clear shift centimeters surgeon decided perform surgery freehand delay hour mazor x | NAV4112 | NAV5000 | NAV3242 | NAV1012 | NAV3264 | NAV4112NAV5000NAV3242NAV1012NAV3264 | True
4176 | 703510997 | XSE - Target Extender Stripped | 2019-12-10T19:23:11.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | MAS2310-03 | TARGET EXTENDER | XSE - Target Extender Stripped 2019-12-10 00679340 (REP): Medtronic received information regarding a guidance system being used outside procedure. It was reported that the target extender was stripped where the knob attaches. No patient was present. 2019-12-12 email 2019-12-12 15:46 x3, image 2019-12-12 15:46, IUD x2 (Rep): No new information. TARGET EXTENDER | 2019-12-10 00679340 (REP): Medtronic received information regarding a guidance system being used outside procedure. It was reported that the target extender was stripped where the knob attaches. No patient was present. 2019-12-12 email 2019-12-12 15:46 x3, image 2019-12-12 15:46, IUD x2 (Rep): No new information. | xse target extender stripped outside procedure target extender stripped knob attaches present email image iud new target extender | NAV3299 | NAV3295 | NAV2054 | NAV4090 | NAV3252 | NAV3299NAV3295NAV2054NAV4090NAV3252 | True
1053 | 703186448 | XSE - Deviations | 2019-06-03T19:34:19.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Deviations 2019-May-30 SCL1901503 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were inaccurate trajectories. The plan was L4 and 5, and used the CT to plan. The surgeon did an OLIF and placed cages on the left side. They did registration with manual segmentation. They sent to L5 on the left and it was lateral to plan. The navigated dilator was lateral to plan. They sent to L4 on the left and it was also lateral with the dilator. They took the chicken foot in the divot and verified the navigation was accurate but could not do it to the bony anatomy. They did a new snapshot and all the navigation was inaccurate. They sent to the right side for both L4 and L5 and everything was accurate. They re-registered with the same AP and lateral images and it was still lateral. They then took new AP and Oblique images and re-registered and the left was still lateral in the same way. Then they brought in a O-arm for scan and plan and everything was accurate and placed with the robot. It took over an hour to do the different registration troubleshooting but there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-May-30 SCL1901503 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were inaccurate trajectories. The plan was L4 and 5, and used the CT to plan. The surgeon did an OLIF and placed cages on the left side. They did registration with manual segmentation. They sent to L5 on the left and it was lateral to plan. The navigated dilator was lateral to plan. They sent to L4 on the left and it was also lateral with the dilator. They took the chicken foot in the divot and verified the navigation was accurate but could not do it to the bony anatomy. They did a new snapshot and all the navigation was inaccurate. They sent to the right side for both L4 and L5 and everything was accurate. They re-registered with the same AP and lateral images and it was still lateral. They then took new AP and Oblique images and re-registered and the left was still lateral in the same way. Then they brought in a O-arm for scan and plan and everything was accurate and placed with the robot. It took over an hour to do the different registration troubleshooting but there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse deviations spinal procedure inaccurate trajectories plan ct plan surgeon did olif placed cages left did registration manual segmentation sent left lateral plan navigated dilator lateral plan sent left lateral dilator took chicken foot divot verified navigation accurate bony anatomy did new snapshot navigation inaccurate sent right accurate reregistered ap lateral images lateral took new ap oblique images reregistered left lateral way brought oarm scan plan accurate placed robot took hour different registration troubleshooting additional complications anticipated mazor x software installation kit | NAV4091 | NAV5000 | NAV3242 | NAV1012 | NAV3246 | NAV4091NAV5000NAV3242NAV1012NAV3246 | True
10425 | 704877642 | XSE - Issues Connecting Manipulator | 2022-02-22T21:09:06.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 30 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE - Issues Connecting Manipulator 2022-Feb-22 Salesforce C00244023 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical system was placed on the bed and then was unable to be removed. The manufacturer representative had to remove the surgical system with the bed frame adapter attached. The light on the handle was not illuminating. There was no patient involvement. 2022-02-24 IUD x5 (Other): No new information. 2022-03-03 IUD x2 (Other): No new information. 2022-03-11 IUD (Other): No new information. 2022-03-21 IUD (Other): No new information. 2022-04-01 IUD x2 (Other): No new information. 2022-04-12 IUD (Other): No new information. 2022-04-28 IUD (Other): No new information UNK MAZOR X MANIPULATOR | 2022-Feb-22 Salesforce C00244023 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical system was placed on the bed and then was unable to be removed. The manufacturer representative had to remove the surgical system with the bed frame adapter attached. The light on the handle was not illuminating. There was no patient involvement. 2022-02-24 IUD x5 (Other): No new information. 2022-03-03 IUD x2 (Other): No new information. 2022-03-11 IUD (Other): No new information. 2022-03-21 IUD (Other): No new information. 2022-04-01 IUD x2 (Other): No new information. 2022-04-12 IUD (Other): No new information. 2022-04-28 IUD (Other): No new information | xse issues connecting manipulator salesforce outside procedure surgical placed bed unable removed remove surgical bed frame adapter attached light handle illuminating iud new iud new iud new iud new iud new iud new iud new unk mazor x manipulator | NAV3260 | NAV3250 | NAV4110 | NAV3244 | NAV3264 | NAV3260NAV3250NAV4110NAV3244NAV3264 | True
3137 | 703266424 | XSE - OOB failures | 2019-07-18T18:30:38.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | FAN0023-01 | FAN-DC, 12V | XSE - OOB failures 2019-Jul-16 SCL1902073 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the fan was found to have loud bearings and a slower RPM. The arm popped up stating arm shift in joint 1. Engineer tried to send arm to other default trajectories with the same response. All joints were functional with the exception being joint 1. Engineer tried to reset movement and hard restart. That was also unsuccessful. Engineer performed field motor calibration. Motor in joint one felt very stiff and grinding. Arm shift error was still popping up along with new error "power_stage_fault" and failed to close under max tolerance. The replacement arm failed verification on the right side during stealth integration tests. All 4 snapshot trackers were used with similar results. The right side of the arm had significantly higher angles of deviation than the left. New spheres were used after the first failed test. Multiple setups were attempted with different snapshots all with similar results. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. FAN-DC, 12V | 2019-Jul-16 SCL1902073 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the fan was found to have loud bearings and a slower RPM. The arm popped up stating arm shift in joint 1. Engineer tried to send arm to other default trajectories with the same response. All joints were functional with the exception being joint 1. Engineer tried to reset movement and hard restart. That was also unsuccessful. Engineer performed field motor calibration. Motor in joint one felt very stiff and grinding. Arm shift error was still popping up along with new error "power_stage_fault" and failed to close under max tolerance. The replacement arm failed verification on the right side during stealth integration tests. All 4 snapshot trackers were used with similar results. The right side of the arm had significantly higher angles of deviation than the left. New spheres were used after the first failed test. Multiple setups were attempted with different snapshots all with similar results. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. | xse oob failures outside procedure fan loud bearings slower rpm arm popped stating arm shift joint engineer tried send arm default trajectories response joints functional exception joint engineer tried reset movement hard restart unsuccessful engineer performed field motor calibration motor joint felt stiff grinding arm shift error popping new error powerstagefault failed close max tolerance replacement arm failed verification right stealth integration tests snapshot trackers similar results right arm significantly higher angles deviation left new spheres failed test multiple setups attempted different snapshots similar results additional complications anticipated fandc | NAV2054 | NAV2058 | NAV4098 | NAV4095 | NAV3299 | NAV2054NAV2058NAV4098NAV4095NAV3299 | True
5772 | 703800494 | XSE - Low Pressure Error | 2020-06-12T21:16:04.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Low Pressure Error 2020-06-12 00690685, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the guidance system had low system pressure and a low pressure error was displayed, but the system worked fine. During the procedure, there was a crossview error during registration. A second registration attempt was made and that was successful. The case was completed. There was no patient harm and the procedure was delayed less than an hour. 2020-06-19 IUD (Other): No new information. 2020-06-26 IUD (Other): No new information. 2020-7-06 IUD (Other): No new information. 2020-07-13 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-06-12 00690685, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the guidance system had low system pressure and a low pressure error was displayed, but the system worked fine. During the procedure, there was a crossview error during registration. A second registration attempt was made and that was successful. The case was completed. There was no patient harm and the procedure was delayed less than an hour. 2020-06-19 IUD (Other): No new information. 2020-06-26 IUD (Other): No new information. 2020-7-06 IUD (Other): No new information. 2020-07-13 IUD (Other): No new information. | xse low pressure error ccr spinal procedure treat degenerative discs fixation low pressure low pressure error displayed worked fine procedure crossview error registration second registration attempt successful case completed procedure delayed hour iud new iud new iud new iud new mazor x | NAV3264 | NAV3242 | NAV2047 | NAV4112 | NAV5000 | NAV3264NAV3242NAV2047NAV4112NAV5000 | True
5563 | 703752450 | XSE - System Not Building Pressure | 2020-05-13T22:47:27.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE - System Not Building Pressure 2020-05-13 00688843 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the system will not build pressure. No patient was present. 2020-05-19 IUD x4 (Other): No new information. 2020-05-27 IUD x2 (Other): No new information. 2020-06-17 IUD (Other): No new information. 2020-07-05 IUD x2 (Other): No new information. 2020-07-09 IUD (Other): No new information. 2020-07-20 IUD (Other): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2020-05-13 00688843 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the system will not build pressure. No patient was present. 2020-05-19 IUD x4 (Other): No new information. 2020-05-27 IUD x2 (Other): No new information. 2020-06-17 IUD (Other): No new information. 2020-07-05 IUD x2 (Other): No new information. 2020-07-09 IUD (Other): No new information. 2020-07-20 IUD (Other): No new information. | xse building pressure outside procedure build pressure present iud new iud new iud new iud new iud new iud new kit svc pneumatic panel | NAV3264 | NAV2054 | NAV2047 | NAV1006 | NAV3251 | NAV3264NAV2054NAV2047NAV1006NAV3251 | True
2425 | 703235145 | X - SHOULDER ISSUE | 2019-06-29T18:40:49.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - SHOULDER ISSUE Shoulder will not lock after pressing the control on the arm. To complete case, CSR had to lock shoulder from general screen in SW. See video attachment." UNK MAZOR X SHOULDER | Shoulder will not lock after pressing the control on the arm. To complete case, CSR had to lock shoulder from general screen in SW. See video attachment." | x shoulder issue shoulder lock pressing control arm complete case csr lock shoulder general screen sw video attachment unk mazor x shoulder | NAV4107 | NAV4101 | NAV4106 | NAV2054 | NAV2058 | NAV4107NAV4101NAV4106NAV2054NAV2058 | True
5022 | 703591547 | X - Mounting/Arm issues | 2020-01-31T16:22:54.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | X - Mounting/Arm issues 2018-04-19 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the advanced accuracy test completed. Shoulder calibration and camera function checked. Surgical plan was reviewed with surgeon prior to surgery and no issues were noted. The system was attached to the bed by one person with no issues. Representative had ancillary staff monitor system indication for left plate sensor to turn green. Plate sensor indicating light still not fixed. 120 mm schanz pin placed in patient's right PSIS. BMB was attached and the guidance system mounted. 3Define scan and draw spine were completed. Middle body labeled L5. Two x-rays were taken in AP. Guidance system bumped with C-arm when obtaining the oblique image. The system got an arm shift error. The representative checked the stability of the schanz pin. The guidance system was remou nted to the patient. 3Define scan and draw spine were repeated. One x-ray was taken to obtain AP. Two x-rays were taken to obtain oblique image. Registration was successful for bodies L3 - S1. Arm was sent to mark skin in all trajectories. Surgeon exposed on left side and arm was re-sent and k wires placed. Surgeon repeated procedure on the right side. Lateral x-rays were taken throughout to verify. AP image was taken once all k wires were inserted. Surgeon was pleased with placement but was unsure whether k wire was in pedicle at L5 on the right. All trajectories tapped and stimmed at or greater than 20. Surgeon proceeded with decompression and interbody placement. The guidance system was removed from the bed with no issues prior to decompression per the surgeon’s request. There was no patient harm. SURGICAL ARM | 2018-04-19 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the advanced accuracy test completed. Shoulder calibration and camera function checked. Surgical plan was reviewed with surgeon prior to surgery and no issues were noted. The system was attached to the bed by one person with no issues. Representative had ancillary staff monitor system indication for left plate sensor to turn green. Plate sensor indicating light still not fixed. 120 mm schanz pin placed in patient's right PSIS. BMB was attached and the guidance system mounted. 3Define scan and draw spine were completed. Middle body labeled L5. Two x-rays were taken in AP. Guidance system bumped with C-arm when obtaining the oblique image. The system got an arm shift error. The representative checked the stability of the schanz pin. The guidance system was remou nted to the patient. 3Define scan and draw spine were repeated. One x-ray was taken to obtain AP. Two x-rays were taken to obtain oblique image. Registration was successful for bodies L3 - S1. Arm was sent to mark skin in all trajectories. Surgeon exposed on left side and arm was re-sent and k wires placed. Surgeon repeated procedure on the right side. Lateral x-rays were taken throughout to verify. AP image was taken once all k wires were inserted. Surgeon was pleased with placement but was unsure whether k wire was in pedicle at L5 on the right. All trajectories tapped and stimmed at or greater than 20. Surgeon proceeded with decompression and interbody placement. The guidance system was removed from the bed with no issues prior to decompression per the surgeon’s request. There was no patient harm. | x mountingarm issues ccr spinal procedure advanced accuracy test completed shoulder calibration camera function checked surgical plan reviewed surgeon prior surgery issues noted attached bed person issues ancillary staff monitor indication left plate sensor turn green plate sensor indicating light fixed mm schanz pin placed patients right psis bmb attached mounted scan draw spine completed middle body labeled xrays taken ap bumped carm obtaining oblique image got arm shift error checked stability schanz pin remou nted scan draw spine repeated xray taken obtain ap xrays taken obtain oblique image registration successful bodies arm sent mark skin trajectories surgeon exposed left arm resent k wires placed surgeon repeated procedure right lateral xrays taken verify ap image taken k wires inserted surgeon pleased placement unsure k wire pedicle right trajectories tapped stimmed greater surgeon proceeded decompression interbody placement removed bed issues prior decompression surgeon’s request surgical arm | NAV3250 | NAV4091 | NAV3242 | NAV2051 | NAV3249 | NAV3250NAV4091NAV3242NAV2051NAV3249 | False
8292 | 704389811 | XSE ¦ Manipulator Leaking | 2021-05-12T03:50:05.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 20 | KIT1259 | SVC KIT1259 PNM PUSH-IN CON QSM3 02 NEW | XSE ¦ Manipulator Leaking 2021-May-11 Salesforce C00226267 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manipulator was leaking air. The manufacturer representative can see a hole in the tubing and can hear the air hissing from the hole. No patient was present. 2021-05-12 IUD x3 (Other): No new information. 2021-05-28 14:31 email 2021-05-28 14:31 (Rep): The cause of the leaking air was due to the pneumatic line being pinched at the D5 where the manipulator sensor is located. SVC KIT1259 PNM PUSH-IN CON QSM3 02 NEW | 2021-May-11 Salesforce C00226267 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manipulator was leaking air. The manufacturer representative can see a hole in the tubing and can hear the air hissing from the hole. No patient was present. 2021-05-12 IUD x3 (Other): No new information. 2021-05-28 14:31 email 2021-05-28 14:31 (Rep): The cause of the leaking air was due to the pneumatic line being pinched at the D5 where the manipulator sensor is located. | xse manipulator leaking salesforce outside procedure manipulator leaking air hole tubing hear air hissing hole present iud new email cause leaking air pneumatic line pinched manipulator sensor located svc pnm pushin new | NAV3264 | NAV3260 | NAV4100 | NAV3253 | NAV3249 | NAV3264NAV3260NAV4100NAV3253NAV3249 | True
9259 | 704591011 | XSE - Intermittent Black Monitors | 2021-09-02T14:10:20.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 80 | PCH0078-01 | USB 3.0 REPEATER | XSE - Intermittent Black Monitors 2021-Sep-01 Salesforce C00233868 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the workstation and surgeon screens would intermittently go black and then return after between 10 and 15 seconds. There was no patient harm and the procedure was not delayed. 2021-Sep-03 IUDx5 (other): No new information. 2021-09-21 IUD x2 (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-18 IUD x2 (Other): No new information. 2021-11-21 IUD x4 (Other): No new information. 2021-11-29 IUD x2 (Other): No new information. 2021-12-05 IUD x2 (Other): No new information. 2021-12-09 IUD (Other): No new information. USB 3.0 REPEATER | 2021-Sep-01 Salesforce C00233868 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the workstation and surgeon screens would intermittently go black and then return after between 10 and 15 seconds. There was no patient harm and the procedure was not delayed. 2021-Sep-03 IUDx5 (other): No new information. 2021-09-21 IUD x2 (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-18 IUD x2 (Other): No new information. 2021-11-21 IUD x4 (Other): No new information. 2021-11-29 IUD x2 (Other): No new information. 2021-12-05 IUD x2 (Other): No new information. 2021-12-09 IUD (Other): No new information. | xse intermittent black monitors salesforce spinal p rocedure workstation surgeon screens intermittently black return seconds procedure delayed new iud new iud new iud new iud new iud new iud new iud new usb repeater | NAV3253 | NAV3261 | NAV2058 | NAV2047 | NAV3237 | NAV3253NAV3261NAV2058NAV2047NAV3237 | True
10003 | 704768966 | XSE ¦ Damaged Positioner Handle | 2021-12-16T20:30:32.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 30 | ASM0364 | KIT SVC ASM0364 RIGHT HANDLE 02S NEW | XSE ¦ Damaged Positioner Handle 2021-Dec-16 Salesforce C00239862 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the side handle on the positioner was broken off when on site today. It looks like it was ran into something but he's not positive what happened. It is not causing any functional issues. No patient present. 2021-12-21 IUD x2 (Other): No new information. 2022-01-05 IUD x2 (Other): No new information. 2022-01-14 IUD x3 (Other): No new information. KIT SVC ASM0364 RIGHT HANDLE 02S NEW | 2021-Dec-16 Salesforce C00239862 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the side handle on the positioner was broken off when on site today. It looks like it was ran into something but he's not positive what happened. It is not causing any functional issues. No patient present. 2021-12-21 IUD x2 (Other): No new information. 2022-01-05 IUD x2 (Other): No new information. 2022-01-14 IUD x3 (Other): No new information. | xse damaged positioner handle salesforce outside procedure handle positioner broken site today looks like ran hes positive happened causing functional issues present iud new iud new iud new kit svc right handle new | NAV3251 | NAV2058 | NAV4092 | NAV2055 | NAV3252 | NAV3251NAV2058NAV4092NAV2055NAV3252 | True
2102 | 703226068 | X - SOFTWARE ISSUE | 2019-06-25T16:09:52.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - SOFTWARE ISSUE Service Issue Request 3rd August, 2017. New York Presbyterian. New York, NY. Workstation #34 Arm #38. 1) Software Crashed x2. The first crash was when bovies were being used the screen went black for about 30 seconds and then brought us out of operate to patient folder. 25 Minutes later the screen went black again this time for about 5 seconds and kept us on the same page. The surgeon’s icon in the top right hand of the screen was missing after these two occurrences. 2) X Eye communication was lost after both crashes. 3) Software continued to run very slowly when transitioning form page to page./L2-S1 OLIF w/ PSF. Surgeon reviewed the plan with CSR prior to the procedure. Potential skive issues were discussed along with making sure Bilateral retraction would be done when using the single schanz. Platform chosen was the single schanz due to Doctors preference. Prior to the procedure @ 2:39 the screen went completely black and after about 30 seconds it came back on and redirected us to the patient folder. CSR then noticed that even though the 3 check marks remained when going back into the operate process the surgeons image on the top right of the screen was not there anymore. After about 25 more minutes of the surgeon exposing and using multiple bovies the screen went black again this time for only a few seconds and broughtus right back to the current screen we were on.1241 3pt check 1243 10 pt check 1245 shoulder calibration checked 1356 Mounted 2 people 45 seconds 1359 draped 1520 Attached to the patient 1522 3Define scan 1524 Drawspine X eyes wasn't communicating went right to generic AP and Oblique 1527 Ap and OB images taken 1529 Registered 1530 S1 left followed by clear up 1533 L5 left followed by clear up 1537 L4 Left 1540 L3 left 1542 L2 left 1550 S1 right 1553 L5 right 1556 L4 right 1600 L3 right sent multiple times the arm had soft tissue pressure forcing the arm to shake to keep is trajectory sent it back 2 additional times before execution 1606 L2 right 1614 Confirmation shot 1620 off the bed. The software overall ran very slow after the two software crashes. Especially moving between pages like image acquisition to fluoro segmentation to registration. Overall the surgeons was satisfied with the outcome off all drilled trajectories were accurate to their planned position but would like to see the system run smoother with less delays. MAZOR X SOFTWARE KIT VER-2.3.1 | Service Issue Request 3rd August, 2017. New York Presbyterian. New York, NY. Workstation #34 Arm #38. 1) Software Crashed x2. The first crash was when bovies were being used the screen went black for about 30 seconds and then brought us out of operate to patient folder. 25 Minutes later the screen went black again this time for about 5 seconds and kept us on the same page. The surgeon’s icon in the top right hand of the screen was missing after these two occurrences. 2) X Eye communication was lost after both crashes. 3) Software continued to run very slowly when transitioning form page to page./L2-S1 OLIF w/ PSF. Surgeon reviewed the plan with CSR prior to the procedure. Potential skive issues were discussed along with making sure Bilateral retraction would be done when using the single schanz. Platform chosen was the single schanz due to Doctors preference. Prior to the procedure @ 2:39 the screen went completely black and after about 30 seconds it came back on and redirected us to the patient folder. CSR then noticed that even though the 3 check marks remained when going back into the operate process the surgeons image on the top right of the screen was not there anymore. After about 25 more minutes of the surgeon exposing and using multiple bovies the screen went black again this time for only a few seconds and broughtus right back to the current screen we were on.1241 3pt check 1243 10 pt check 1245 shoulder calibration checked 1356 Mounted 2 people 45 seconds 1359 draped 1520 Attached to the patient 1522 3Define scan 1524 Drawspine X eyes wasn't communicating went right to generic AP and Oblique 1527 Ap and OB images taken 1529 Registered 1530 S1 left followed by clear up 1533 L5 left followed by clear up 1537 L4 Left 1540 L3 left 1542 L2 left 1550 S1 right 1553 L5 right 1556 L4 right 1600 L3 right sent multiple times the arm had soft tissue pressure forcing the arm to shake to keep is trajectory sent it back 2 additional times before execution 1606 L2 right 1614 Confirmation shot 1620 off the bed. The software overall ran very slow after the two software crashes. Especially moving between pages like image acquisition to fluoro segmentation to registration. Overall the surgeons was satisfied with the outcome off all drilled trajectories were accurate to their planned position but would like to see the system run smoother with less delays. | x software issue service issue request august new york presbyterian new york ny workstation arm software crashed crash bovies screen went black seconds brought operate folder minutes later screen went black time seconds kept page surgeon’s icon right hand screen missing occurrences x eye communication lost crashes software continued run slowly transitioning page olif w psf surgeon reviewed plan csr prior procedure potential skive issues discussed making sure bilateral retraction using single schanz platform chosen single schanz doctors preference prior procedure screen went completely black seconds came redirected folder csr noticed check marks remained going operate process surgeons image right screen anymore minutes surgeon exposing using multiple bovies screen went black time seconds broughtus right current screen check pt check shoulder calibration checked mounted people seconds draped attached scan drawspine x eyes wasnt communicating went right generic ap oblique ap ob images taken registered left followed clear left followed clear left left left right right right right sent multiple times arm soft tissue pressure forcing arm shake trajectory sent additional times execution right confirmation shot bed software overall ran slow software crashes especially moving pages like image acquisition fluoro segmentation registration overall surgeons satisfied outcome drilled trajectories accurate planned position like run smoother delays mazor x software kit | NAV3248 | NAV3247 | NAV4101 | NAV3253 | NAV3263 | NAV3248NAV3247NAV4101NAV3253NAV3263 | True
8151 | 704362308 | XSE - 5.0 Upgrade Failure | 2021-04-27T00:21:57.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 80 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - 5.0 Upgrade Failure 2021-Apr-26 Salesforce C00225326 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the software failed the security patch update when updating the system software. There was no patient involvement. 2021-04-27 IUD x11 (Other): No new information. 2021-05-21 IUD (Other): No new information. 2021-06-16 IUD (Other): No new information. 2021-07-13 IUD x3 (Other): No new information. 2021-09-22 IUD (Other): No new information. 2021-11-03 302618444, IUD x2 (Other): No new information. 2021-11-03 IUD x2 (Other): No new information. 2021-12-08 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-Apr-26 Salesforce C00225326 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the software failed the security patch update when updating the system software. There was no patient involvement. 2021-04-27 IUD x11 (Other): No new information. 2021-05-21 IUD (Other): No new information. 2021-06-16 IUD (Other): No new information. 2021-07-13 IUD x3 (Other): No new information. 2021-09-22 IUD (Other): No new information. 2021-11-03 302618444, IUD x2 (Other): No new information. 2021-11-03 IUD x2 (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse upgrade failure salesforce outside procedure software failed security patch update updating software iud new iud new iud new iud new iud new iud new iud new iud new kit svc cart elec panel assy | NAV3264 | NAV2047 | NAV2058 | NAV3249 | NAV3299 | NAV3264NAV2047NAV2058NAV3249NAV3299 | True
3826 | 703451600 | XSE - BICE/CANBUS Error Lost Connection | 2019-11-06T00:04:54.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - BICE/CANBUS Error Lost Connection 2019-Nov-06 00677669 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there was a Canbus error lost connection to all nodes. Performed a hard restart and the software would not appear. System had to be removed with the manual pump. It was noted that navigation was aborted. Several times before attaching to patient, the rep noted that the message would pop up on the workstation monitor requesting if they want to change to the tablet control. This occurred five times when no one was touching the tablet. Everything else was running smooth quick registration sent to first trajectory S2 on the right surgeon place trajectory accurate through the arm. The rep sent to S2 on the left and Fellow began instrumentation, he placed instruments down and surgeon used a bovie which hit the drill guide and outer cannula, then the system detected the arm moved out of trajectory. The rep then resent the Arm and the arm did not move. Fellow began to place instruments back down the arm when he noticed the arm slightly shifted and the SW did not detect this. The rep attempted to Clear up to reset everything after clicking clear up the arm did not move and the CANBUS error message appeared and all nodes were lost. The rep attempted a few other movements with no success. The rep communicated to surgeon about this and explained a hard restart would be the only option. The rep shut system down turned back on after 3:15. The PC turned on but the SW never came back up. After 6 minutes of waiting the rep attempted this again. Same issue occurred. Associate was directed to go into task manager and end task and restart it, after multiple attempts and a delay of the surgery the rep decided to manually remove the system and brought in the older Regular X system to complete the case. After the case rep powered on the system and still found no solutions on getting the SW up and running. CRR’s and EXPORT were unable to be obtained due to the issues above. It was also noted that when the rep removed the surgical system from the bed they placed it on the floor and the surgical system fell on the surgical arm. The surgical arm failed accuracy tests. Discovered pneumatic lead at the D5 sensor of the manipulator.There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-06 Email, IUD x2: Additional information was received stating that the system was shut down for three minutes before powering back on. The dell logo appeared on power up however the screen remained black after and the software never loaded. The system was turned on after the case but the same results upon start up were received. The dell logo appeared but then the screen was black. Both times they could see the mouse on the screen as well as access the task manager so it appeared that the operating system was working the software would not load. 2019-11-13 IUD x 8 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2019-Nov-06 00677669 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there was a Canbus error lost connection to all nodes. Performed a hard restart and the software would not appear. System had to be removed with the manual pump. It was noted that navigation was aborted. Several times before attaching to patient, the rep noted that the message would pop up on the workstation monitor requesting if they want to change to the tablet control. This occurred five times when no one was touching the tablet. Everything else was running smooth quick registration sent to first trajectory S2 on the right surgeon place trajectory accurate through the arm. The rep sent to S2 on the left and Fellow began instrumentation, he placed instruments down and surgeon used a bovie which hit the drill guide and outer cannula, then the system detected the arm moved out of trajectory. The rep then resent the Arm and the arm did not move. Fellow began to place instruments back down the arm when he noticed the arm slightly shifted and the SW did not detect this. The rep attempted to Clear up to reset everything after clicking clear up the arm did not move and the CANBUS error message appeared and all nodes were lost. The rep attempted a few other movements with no success. The rep communicated to surgeon about this and explained a hard restart would be the only option. The rep shut system down turned back on after 3:15. The PC turned on but the SW never came back up. After 6 minutes of waiting the rep attempted this again. Same issue occurred. Associate was directed to go into task manager and end task and restart it, after multiple attempts and a delay of the surgery the rep decided to manually remove the system and brought in the older Regular X system to complete the case. After the case rep powered on the system and still found no solutions on getting the SW up and running. CRR’s and EXPORT were unable to be obtained due to the issues above. It was also noted that when the rep removed the surgical system from the bed they placed it on the floor and the surgical system fell on the surgical arm. The surgical arm failed accuracy tests. Discovered pneumatic lead at the D5 sensor of the manipulator.There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-06 Email, IUD x2: Additional information was received stating that the system was shut down for three minutes before powering back on. The dell logo appeared on power up however the screen remained black after and the software never loaded. The system was turned on after the case but the same results upon start up were received. The dell logo appeared but then the screen was black. Both times they could see the mouse on the screen as well as access the task manager so it appeared that the operating system was working the software would not load. 2019-11-13 IUD x 8 (Rep): No new information. | xse bicecanbus error lost connection spinal procedure canbus error lost connection nodes performed hard restart software appear removed manual pump noted navigation aborted times attaching noted message pop workstation monitor requesting want change tablet control occurred times touching tablet running smooth quick registration sent trajectory right surgeon place trajectory accurate arm sent left fellow began instrumentation placed instruments surgeon bovie hit drill guide outer cannula detected arm moved trajectory resent arm arm did fellow began place instruments arm noticed arm slightly shifted sw did detect attempted clear reset clicking clear arm did canbus error message appeared nodes lost attempted movements success communicated surgeon explained hard restart option shut turned pc turned sw came minutes waiting attempted issue occurred associate directed task manager end task restart multiple attempts delay surgery decided manually remove brought older regular x complete case case powered solutions getting sw running crr’s export unable obtained issues noted removed surgical bed placed floor surgical fell surgical arm surgical arm failed accuracy tests discovered pneumatic lead sensor manipulatorthere additional complications anticipated email iud additional stating shut minutes powering dell logo appeared power screen remained black software loaded turned case results start dell logo appeared screen black times mouse screen access task manager appeared operating working software load iud x new mazor x | NAV2054 | NAV3249 | NAV3246 | NAV2047 | NAV3236 | NAV2054NAV3249NAV3246NAV2047NAV3236 | True
9003 | 704524754 | XSE ¦ Will Not Release From Bed Frame | 2021-07-27T20:14:55.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_COMP | UNKNOWN MAZOR X COMPONENT | XSE ¦ Will Not Release From Bed Frame 2021-Jul-27 Salesforce C00230929 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system is stating that it is in storage state, but will not release from the bed frame. Performed a soft reboot and a hard reboot without resolution. When it booted back up it then had a error message "lost connection with controller". No patient present. 2021-Jul-27 IUDx2 (other): No new information. UNKNOWN MAZOR X COMPONENT | 2021-Jul-27 Salesforce C00230929 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system is stating that it is in storage state, but will not release from the bed frame. Performed a soft reboot and a hard reboot without resolution. When it booted back up it then had a error message "lost connection with controller". No patient present. 2021-Jul-27 IUDx2 (other): No new information. | xse release bed frame salesforce outside procedure stating storage state release bed frame performed soft reboot hard reboot resolution booted error message lost connection controller present new unknown mazor x component | NAV2047 | NAV3255 | NAV3250 | NAV3262 | NAV3242 | NAV2047NAV3255NAV3250NAV3262NAV3242 | True
1294 | 703195399 | X - Deviations | 2019-06-07T17:02:15.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviations Dr. Francis 5.4.18 L5-S1 TLIF Preop: Cameras checked, AA completed, Shoulder checked Plan: Dr. Francis completed the plan prior to surgery. Axial angles were narrow and uniform from level to level. Screw angles were under 12 degrees for open procedure. Arm attached to bed -1 person Arm draped -1 person RTC had scrub tech check the Schanz arms prior to draping. Schanz arm was solid and fixated after tightening Schanz arm was used on right side PSIS 3define scan completed -clear Drawspine -OBL trajectory could never be found on first try, surgeon redid the drawspine, this time we were able to acquire an OBL trajectory-Labeled L5 -Arm sent to AP And OBL. RTC had to manually manipulate the arm for the OBL to bring it closer to the patient Registration: Completed on the first try with green values and visually verified by surgeon Operation: Open operation -Operation was quick and everything appeared fine until Dr. Francis got to the the last trajectory at S1 and he said that this was way medial, and that he felt the rest were medial as well. Surgeon told RTC to translate the screw out laterally and resend to S1. All K-wires were placed with no fluoro. AP was taken after all four k-wires were in, all k-wires were medial and Dr. Francis pulled all four K-wires and opted to freehand. -RTC felt that there was likely a system shift even though nothing was detected. Surgeon was rough with the rongeur while removing bone after registration. Surgeon was also hammering in the drill guide very hard, and when removing the drill guide he exerted a lot of force to remove the tools, causing the bed, patient, and robot to move. No alarms went off, RTC asked surgeon to ease up on the tools, but surgeon did not listen to RTC, and continued to hammer the tools down and forcibly remove the tools after placing k-wires. MAZOR X SOFTWARE KIT VER-3.0.2 | Dr. Francis 5.4.18 L5-S1 TLIF Preop: Cameras checked, AA completed, Shoulder checked Plan: Dr. Francis completed the plan prior to surgery. Axial angles were narrow and uniform from level to level. Screw angles were under 12 degrees for open procedure. Arm attached to bed -1 person Arm draped -1 person RTC had scrub tech check the Schanz arms prior to draping. Schanz arm was solid and fixated after tightening Schanz arm was used on right side PSIS 3define scan completed -clear Drawspine -OBL trajectory could never be found on first try, surgeon redid the drawspine, this time we were able to acquire an OBL trajectory-Labeled L5 -Arm sent to AP And OBL. RTC had to manually manipulate the arm for the OBL to bring it closer to the patient Registration: Completed on the first try with green values and visually verified by surgeon Operation: Open operation -Operation was quick and everything appeared fine until Dr. Francis got to the the last trajectory at S1 and he said that this was way medial, and that he felt the rest were medial as well. Surgeon told RTC to translate the screw out laterally and resend to S1. All K-wires were placed with no fluoro. AP was taken after all four k-wires were in, all k-wires were medial and Dr. Francis pulled all four K-wires and opted to freehand. -RTC felt that there was likely a system shift even though nothing was detected. Surgeon was rough with the rongeur while removing bone after registration. Surgeon was also hammering in the drill guide very hard, and when removing the drill guide he exerted a lot of force to remove the tools, causing the bed, patient, and robot to move. No alarms went off, RTC asked surgeon to ease up on the tools, but surgeon did not listen to RTC, and continued to hammer the tools down and forcibly remove the tools after placing k-wires. | x deviations dr francis tlif preop cameras checked completed shoulder checked plan dr francis completed plan prior surgery axial angles narrow uniform level level screw angles degrees open procedure arm attached bed person arm draped person rtc scrub tech check schanz arms prior draping schanz arm solid fixated tightening schanz arm right psis scan completed clear drawspine obl trajectory try surgeon redid drawspine time able acquire obl trajectorylabeled arm sent ap obl rtc manually manipulate arm obl bring closer registration completed try green values visually verified surgeon operation open operation operation quick appeared fine dr francis got trajectory said way medial felt rest medial surgeon told rtc translate screw laterally resend kwires placed fluoro ap taken kwires kwires medial dr francis pulled kwires opted freehand rtc felt likely shift detected surgeon rough rongeur removing bone registration surgeon hammering drill guide hard removing drill guide exerted lot force remove tools causing bed robot alarms went rtc asked surgeon ease tools surgeon did listen rtc continued hammer tools forcibly remove tools placing kwires mazor x software kit | NAV3246 | NAV3299 | NAV3242 | NAV3247 | NAV3233 | NAV3246NAV3299NAV3242NAV3247NAV3233 | False
9579 | 704675631 | XSE - Monitor Stand-off Issues | 2021-10-21T22:20:38.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 30 | ASM0275-01S | KIT SVC ASM0275-01S MONITOR | XSE - Monitor Stand-off Issues 2021-Oct-21 Salesforce C00236623 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the main system monitor stand was was broken, which caused the monitor to not be secure. The manufacturer representative noted the damage was not due to negligence by the site. There was no patient involvement. 2021-10-25 IUD x6 (Other): The replacement monitor was bad at install. 2021-11-04 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 2021-11-28 IUD x3 (Other): No new information. 2021-12-22 IUD (Other): No new information. KIT SVC ASM0275-01S MONITOR | 2021-Oct-21 Salesforce C00236623 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the main system monitor stand was was broken, which caused the monitor to not be secure. The manufacturer representative noted the damage was not due to negligence by the site. There was no patient involvement. 2021-10-25 IUD x6 (Other): The replacement monitor was bad at install. 2021-11-04 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 2021-11-28 IUD x3 (Other): No new information. 2021-12-22 IUD (Other): No new information. | xse monitor standoff issues salesforce outside procedure main monitor stand broken caused monitor secure noted damage negligence site iud replacement monitor bad install iud new iud new iud new iud new kit svc monitor | NAV3261 | NAV2054 | NAV2055 | NAV2058 | NAV3299 | NAV3261NAV2054NAV2055NAV2058NAV3299 | True
4662 | 703559329 | XSE - Teeth Sensor Damaged | 2020-01-13T20:02:56.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 30 | ASM0214-01 | SURGICAL SYSTEM | XSE - Teeth Sensor Damaged 2020-01-13 00681967 (Rep, HCP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the teeth sensor is damaged on the positioner type I. There was no patient involvement. 2020-01-13 IUD x2 (Other): No new information. 2020-02-11 IUS x2 (Other): Additional information received from a field service engineer reporting that the system was serviced on-site and the complaint was confirmed. The right positioner cleat/tooth sensor was broken. 2020-05-24 IUD (Other): No new information. SURGICAL SYSTEM | 2020-01-13 00681967 (Rep, HCP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the teeth sensor is damaged on the positioner type I. There was no patient involvement. 2020-01-13 IUD x2 (Other): No new information. 2020-02-11 IUS x2 (Other): Additional information received from a field service engineer reporting that the system was serviced on-site and the complaint was confirmed. The right positioner cleat/tooth sensor was broken. 2020-05-24 IUD (Other): No new information. | xse teeth sensor damaged outside procedure teeth sensor damaged positioner type iud new ius additional field service engineer reporting serviced onsite complaint confirmed right positioner cleattooth sensor broken iud new surgical | NAV3251 | NAV3249 | NAV4103 | NAV2058 | NAV4104 | NAV3251NAV3249NAV4103NAV2058NAV4104 | True
8939 | 704514253 | XSE - Positioner Sensor Tab Bent | 2021-07-21T16:12:08.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 60 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE - Positioner Sensor Tab Bent 2021-Jul-20 Salesforce C00230534, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that that when going to store the surgical system, the tab on the center sensor tooth broke off. The manufacturer representative was able to place it back on the positioner, but it was bent. There was no patient involvement. 2021-Jul-22 IUDx9 (other): No new information. 2021-Aug-09 IUDx3 (other): No new information. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2021-Jul-20 Salesforce C00230534, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that that when going to store the surgical system, the tab on the center sensor tooth broke off. The manufacturer representative was able to place it back on the positioner, but it was bent. There was no patient involvement. 2021-Jul-22 IUDx9 (other): No new information. 2021-Aug-09 IUDx3 (other): No new information. | xse positioner sensor tab bent salesforce iud outside proc edure going store surgical tab center sensor tooth broke able place positioner bent new new kit svc surgical new | NAV3251 | NAV2047 | NAV4103 | NAV4104 | NAV2054 | NAV3251NAV2047NAV4103NAV4104NAV2054 | False
882 | 703178952 | X - Deviation/Accessory/Shoulder issues | 2019-05-29T20:05:35.000000+0000 | MazorX | NAV4108 | SHOULDER SHIFTS | 30 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Deviation/Accessory/Shoulder issues It was reported that the flouro segmentation wtih basic green values at all surgical levels. Surgeon visually verified matching accuracy at all surgical levels. Sent to L5 R, L5 L, S1 L when sending to S1 Right arm guide collided with retractor and caused a shoulder shift Reg #2- Flouro seg w/basic. Green values at all surgical levels. Surgeon visually verified matching accuracy at all surgical levels. Sent to S1 reappeared good on lateral and AP. Stimulated at 9, surgeon removed screw and felt a slight medial breach. Used starting point and placed s crew slightly more lateral. UNK MAZOR X SHOULDER | It was reported that the flouro segmentation wtih basic green values at all surgical levels. Surgeon visually verified matching accuracy at all surgical levels. Sent to L5 R, L5 L, S1 L when sending to S1 Right arm guide collided with retractor and caused a shoulder shift Reg #2- Flouro seg w/basic. Green values at all surgical levels. Surgeon visually verified matching accuracy at all surgical levels. Sent to S1 reappeared good on lateral and AP. Stimulated at 9, surgeon removed screw and felt a slight medial breach. Used starting point and placed s crew slightly more lateral. | x deviationaccessoryshoulder issues flouro segmentation wtih basic green values surgical levels surgeon visually verified matching accuracy surgical levels sent r l l sending right arm guide collided retractor caused shoulder shift reg flouro seg wbasic green values surgical levels surgeon visually verified matching accuracy surgical levels sent reappeared good lateral ap stimulated surgeon removed screw felt slight medial breach starting point placed s crew slightly lateral unk mazor x shoulder | NAV4108 | NAV3244 | NAV3294 | NAV4101 | NAV4093 | NAV4108NAV3244NAV3294NAV4101NAV4093 | True
378 | 703094742 | XSE - arm didn't reach position | 2019-04-08T13:21:15.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - arm didn't reach position 2019-Apr-03 SCL1900808 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during system setup advanced accuracy test, the arm shuddered as it reached 1st position and gave a Joint 2 error. The rep performed a superuser reset movement, but it did not resolve. During the call, the arm moved to known location and superuser reset movement performed after a soft restart. They were able to move each joint individually. They performed clear up movement and there was no error. They then performed the advanced accuracy test successfully. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Apr-03 SCL1900808 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during system setup advanced accuracy test, the arm shuddered as it reached 1st position and gave a Joint 2 error. The rep performed a superuser reset movement, but it did not resolve. During the call, the arm moved to known location and superuser reset movement performed after a soft restart. They were able to move each joint individually. They performed clear up movement and there was no error. They then performed the advanced accuracy test successfully. There was no patient involvement. There were no additional complications reported or anticipated as a result of the event. | xse arm didnt reach position spinal procedure setup advanced accuracy test arm shuddered reached position gave joint error performed superuser reset movement did resolve arm moved known location superuser reset movement performed soft restart able joint individually performed clear movement error performed advanced accuracy test successfully additional complications anticipated mazor x software installation kit | NAV3246 | NAV4098 | NAV4095 | NAV3247 | NAV5000 | NAV3246NAV4098NAV4095NAV3247NAV5000 | True
2759 | 703249219 | X - Accessories Issue | 2019-07-09T16:13:07.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 60 | MAS2249-03 | SHORT CANNULA | X - Accessories Issue SC opened following Elad report in attached mail regarding tools related issues: "Marker extender (carbon)- 3D Marker screw was much too tight. Short Drill bit got stuck in the sort drill guide. Long drill bit got stuck in the long drill guide. Both were released using pliers and lots of force. Long drill guide got damaged (bent teeth), as well as both drill bits. RNS arm tube (14-8) was used with RNS tools. Cannulas had too much freedom in the arm tube. New Arm Tube needs to be tested and provided. Accuracy pointer is missing from the surgical set. Needs to be provided. Tablet adaptor (tablet side) too tight on the tablet side while on the drape - hard to tighten the screws. SHORT CANNULA | SC opened following Elad report in attached mail regarding tools related issues: "Marker extender (carbon)- 3D Marker screw was much too tight. Short Drill bit got stuck in the sort drill guide. Long drill bit got stuck in the long drill guide. Both were released using pliers and lots of force. Long drill guide got damaged (bent teeth), as well as both drill bits. RNS arm tube (14-8) was used with RNS tools. Cannulas had too much freedom in the arm tube. New Arm Tube needs to be tested and provided. Accuracy pointer is missing from the surgical set. Needs to be provided. Tablet adaptor (tablet side) too tight on the tablet side while on the drape - hard to tighten the screws. | x accessories issue sc opened following elad report attached mail tools related issues marker extender marker screw tight short drill bit got stuck sort drill guide long drill bit got stuck long drill guide released using pliers lots force long drill guide got damaged drill bits rns arm tube rns tools cannulas freedom arm tube new arm tube needs tested provided accuracy pointer missing surgical set needs provided tablet adaptor tight tablet drape hard tighten screws short cannula | NAV3299 | NAV3295 | NAV2053 | NAV3253 | NAV3297 | NAV3299NAV3295NAV2053NAV3253NAV3297 | True
4358 | 703537937 | X - Registration Issues | 2019-12-30T14:02:33.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issues 2018-08-10 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a trauma injury including fracture with fixation at T1-T5. It was reported that planning for a T1-T5 laminectomy with PSF was completed by the surgeon prior to the procedure. The manufacturer representative and the surgeon reviewed the plan with special attention paid to skin incision, axial angles, screw sizes, and rod alignment. The dual clamp mount was chosen for the platform based on clinical indication and the surgeon's preference. The system was mounted to the bed and attached to the patient without issue. A 12in C-Arm was used to acquire 3 AP and 5 OBL shots for registration. Automatic segmentation was successful but values were red for all trajectories except T5 and T6. Labeling failed at when tried off a different vertebral body. Manual adjustment was done but all values were red. Another OBL shot was taken and automatic segmentation was again successful but values were again red and manual adjustment was unsuccessful. Another OBL shot was taken but automatic segmentation was not successful. Manual adjustments made were successful with good values at all levels except the fractured vertebra which was planned to not have any screws placed. All trajectories were placed successfully. After the patient was stimulated and motor testing, the surgeon used a probe to check each site and, at the time, felt that T4 and T5 felt slightly medial. The surgeon inserted the screws anyway and took confirmation fluoro images. The surgeon was happy with the position of the screws and the outcome of the case. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2018-08-10 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a trauma injury including fracture with fixation at T1-T5. It was reported that planning for a T1-T5 laminectomy with PSF was completed by the surgeon prior to the procedure. The manufacturer representative and the surgeon reviewed the plan with special attention paid to skin incision, axial angles, screw sizes, and rod alignment. The dual clamp mount was chosen for the platform based on clinical indication and the surgeon's preference. The system was mounted to the bed and attached to the patient without issue. A 12in C-Arm was used to acquire 3 AP and 5 OBL shots for registration. Automatic segmentation was successful but values were red for all trajectories except T5 and T6. Labeling failed at when tried off a different vertebral body. Manual adjustment was done but all values were red. Another OBL shot was taken and automatic segmentation was again successful but values were again red and manual adjustment was unsuccessful. Another OBL shot was taken but automatic segmentation was not successful. Manual adjustments made were successful with good values at all levels except the fractured vertebra which was planned to not have any screws placed. All trajectories were placed successfully. After the patient was stimulated and motor testing, the surgeon used a probe to check each site and, at the time, felt that T4 and T5 felt slightly medial. The surgeon inserted the screws anyway and took confirmation fluoro images. The surgeon was happy with the position of the screws and the outcome of the case. No patient harm was reported and surgery was delayed less than an hour. | x registration issues ccr spinal procedure treat trauma injury including fracture fixation planning laminectomy psf completed surgeon prior procedure surgeon reviewed plan special attention paid skin incision axial angles screw sizes rod alignment dual clamp mount chosen platform based clinical indication surgeons preference mounted bed attached issue carm acquire ap obl shots registration automatic segmentation successful values red trajectories labeling failed tried different vertebral body manual adjustment values red obl shot taken automatic segmentation successful values red manual adjustment unsuccessful obl shot taken automatic segmentation successful manual adjustments successful good values levels fractured vertebra planned screws placed trajectories placed successfully stimulated motor testing surgeon probe check site time felt felt slightly medial surgeon inserted screws took confirmation fluoro images surgeon happy position screws outcome case surgery delayed hour mazor x | NAV3242 | NAV5000 | NAV4091 | NAV3295 | NAV3241 | NAV3242NAV5000NAV4091NAV3295NAV3241 | True
3651 | 703397297 | XSE - Deviation/Failed Accuracy Test | 2019-10-04T14:39:46.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01 | SURGICAL ARM | XSE - Deviation/Failed Accuracy Test 2019-10-02 SCL1902984, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs and scoliosis. It was reported that the surgeon decided to do Scan &Plan for a L2-L5 MIS procedure. Prior to inserting the screws, the surgeon planned to perform 3 PEEK XLIFs and then flipping the patient in the prone position. During the procedure, a Schanz arm was used on the right side of the surgical system with a Schanz pin placed in the right PSIS. Blue, black and orange trackers were then registered. One O-arm spin was required for registration. After the star marker was found, planning was completed. Due to the case being minimally invasive, skiving was an issue regarding the docking surfaces. Careful attention was paid to tulip placement to make passing the rods easier. The surgeon started at left L2, then went down the left side to L5 and then repeated the same order on the right side. The skin at L2 as marked with the dilator and a skin knife was used for the incisions. The outer trocar was introduced with the dilator and the skin incision was lengthened to accommodate the tools. The dilator was then removed and the drill guide was inserted followed by the drill guide anchor. The drill guide was then tapped into place. The anchor was removed and the trajectory was then drilled, tapped and the screw was inserted under power. This process continued for the rest of the screws. Skiving was detected at L5 right, but was corrected by the surgeon. The surgical arm moved off trajectory while removing the driver at left L5. After all of the screws were placed, the C-arm was brought in for confirmation shots. The images showed that L2, L3 and L4 right were not as planned. The surgeon decided to remove the screws and use a gear shirt with fluoro to correct the deviated screws. The representative noted the system later failed an advanced accuracy test. There was no patient harm and the procedure was not delayed more than an hour. 2019-10-05 email 2019-10-05 13:15 (Rep): Additional information received from a manufacturer representative reported the trajectories were deviated between 3-5 mm laterally and greater than 5 mm superior. All screws were placed during the procedure and then a confirmation shot was taken. After the confirmation shot, the surgeon decided to freehand the screws. Navigation showed the screws were accurate as they were being placed. The representative thought that some soft tissue pressure and skiving were the cause of the lateral deviation, but the cause of the superior deviation was unknown. 2019-10-10 e3, e4, e5 (Rep): Additional information received from a manufacturer representative reported that this procedure was the first robotic procedure for the surgeon. The case was minimally invasive using a new screw system. The surgeon was trying to line up the tulip heads as best as possible to make rod passing easier. The representative discussed risks of skiving and soft tissue pressure with the surgeon during the planning stage after the O-arm spin. During the procedure, the surgeon was reminded about making the skin incisions larger to accomodate the tools and screw towers. The surgeon used a drill guide anchor to minimize the skiving risk. A couple times during the procedure, the surgical arm shifted due to pressure and the arm had to be resent to the trajectory. After all of the screws were inserted, confirmation shots were taken and some screws were deviated. The representative believed the inaccuracy was due to soft tissue pressure and force that was used on the tools due to small incisions. SURGICAL ARM | 2019-10-02 SCL1902984, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs and scoliosis. It was reported that the surgeon decided to do Scan &Plan for a L2-L5 MIS procedure. Prior to inserting the screws, the surgeon planned to perform 3 PEEK XLIFs and then flipping the patient in the prone position. During the procedure, a Schanz arm was used on the right side of the surgical system with a Schanz pin placed in the right PSIS. Blue, black and orange trackers were then registered. One O-arm spin was required for registration. After the star marker was found, planning was completed. Due to the case being minimally invasive, skiving was an issue regarding the docking surfaces. Careful attention was paid to tulip placement to make passing the rods easier. The surgeon started at left L2, then went down the left side to L5 and then repeated the same order on the right side. The skin at L2 as marked with the dilator and a skin knife was used for the incisions. The outer trocar was introduced with the dilator and the skin incision was lengthened to accommodate the tools. The dilator was then removed and the drill guide was inserted followed by the drill guide anchor. The drill guide was then tapped into place. The anchor was removed and the trajectory was then drilled, tapped and the screw was inserted under power. This process continued for the rest of the screws. Skiving was detected at L5 right, but was corrected by the surgeon. The surgical arm moved off trajectory while removing the driver at left L5. After all of the screws were placed, the C-arm was brought in for confirmation shots. The images showed that L2, L3 and L4 right were not as planned. The surgeon decided to remove the screws and use a gear shirt with fluoro to correct the deviated screws. The representative noted the system later failed an advanced accuracy test. There was no patient harm and the procedure was not delayed more than an hour. 2019-10-05 email 2019-10-05 13:15 (Rep): Additional information received from a manufacturer representative reported the trajectories were deviated between 3-5 mm laterally and greater than 5 mm superior. All screws were placed during the procedure and then a confirmation shot was taken. After the confirmation shot, the surgeon decided to freehand the screws. Navigation showed the screws were accurate as they were being placed. The representative thought that some soft tissue pressure and skiving were the cause of the lateral deviation, but the cause of the superior deviation was unknown. 2019-10-10 e3, e4, e5 (Rep): Additional information received from a manufacturer representative reported that this procedure was the first robotic procedure for the surgeon. The case was minimally invasive using a new screw system. The surgeon was trying to line up the tulip heads as best as possible to make rod passing easier. The representative discussed risks of skiving and soft tissue pressure with the surgeon during the planning stage after the O-arm spin. During the procedure, the surgeon was reminded about making the skin incisions larger to accomodate the tools and screw towers. The surgeon used a drill guide anchor to minimize the skiving risk. A couple times during the procedure, the surgical arm shifted due to pressure and the arm had to be resent to the trajectory. After all of the screws were inserted, confirmation shots were taken and some screws were deviated. The representative believed the inaccuracy was due to soft tissue pressure and force that was used on the tools due to small incisions. | xse deviationfailed accuracy test ccr spinal procedure treat degenerative discs scoliosis surgeon decided scan plan mis procedure prior inserting screws surgeon planned perform peek xlifs flipping prone position procedure schanz arm right surgical schanz pin placed right psis blue black orange trackers registered oarm spin required registration star marker planning completed case minimally invasive skiving issue docking surfaces careful attention paid tulip placement make passing rods easier surgeon started left went left repeated order right skin marked dilator skin knife incisions outer trocar introduced dilator skin incision lengthened accommodate tools dilator removed drill guide inserted followed drill guide anchor drill guide tapped place anchor removed trajectory drilled tapped screw inserted power process continued rest screws skiving detected right corrected surgeon surgical arm moved trajectory removing driver left screws placed carm brought confirmation shots images showed right planned surgeon decided remove screws use gear shirt fluoro correct deviated screws noted later failed advanced accuracy test procedure delayed hour email additional trajectories deviated mm laterally greater mm superior screws placed procedure confirmation shot taken confirmation shot surgeon decided freehand screws navigation showed screws accurate placed thought soft tissue pressure skiving cause lateral deviation cause superior deviation unknown additional procedure robotic procedure surgeon case minimally invasive using new screw surgeon trying line tulip heads best possible make rod passing easier discussed risks skiving soft tissue pressure surgeon planning stage oarm spin procedure surgeon reminded making skin incisions larger accomodate tools screw towers surgeon drill guide anchor minimize skiving risk couple times procedure surgical arm shifted pressure arm resent trajectory screws inserted confirmation shots taken screws deviated believed inaccuracy soft tissue pressure force tools small incisions surgical arm | NAV4093 | NAV5000 | NAV3244 | NAV3249 | NAV3236 | NAV4093NAV5000NAV3244NAV3249NAV3236 | False
9763 | 704718189 | XSE - Shipping Damage | 2021-11-16T16:10:32.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 30 | MAS2052-01S | KIT SVC MAS2052-01S POSIT PADDING BRCKT | XSE - Shipping Damage 2021-Nov-16 Salesforce C00238210 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the bracket near where the positioner is stored got damaged during shipping. No patient was present. 2021-11-18 IUD (Other): No new information. 2021-12-01 IUD (Other): No new information. 2021-12-21 IUD x3 (Other): No new information. KIT SVC MAS2052-01S POSIT PADDING BRCKT | 2021-Nov-16 Salesforce C00238210 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the bracket near where the positioner is stored got damaged during shipping. No patient was present. 2021-11-18 IUD (Other): No new information. 2021-12-01 IUD (Other): No new information. 2021-12-21 IUD x3 (Other): No new information. | xse shipping damage salesforce outside procedure bracket near positioner stored got damaged shipping present iud new iud new iud new kit svc posit padding brckt | NAV2055 | NAV3264 | NAV3249 | NAV2058 | NAV3251 | NAV2055NAV3264NAV3249NAV2058NAV3251 | True
5588 | 703758580 | XSE - Damaged Schanz Arms | 2020-05-18T22:14:16.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Damaged Schanz Arms 2020-05-18 00689106 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arms were damaged. There was no patient involvement. 2020-05-20 IUD (Other): No new information. 2020-05-28 IUD (Other): No new information. 2020-06-15 email 2020-06-15 08:15, email 2020-06-22 11:45 (Rep): No new information. 2022-05-16 IUD (Other): No new information. KIT SVC MAS2351-04S SCHANZ ARM | 2020-05-18 00689106 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arms were damaged. There was no patient involvement. 2020-05-20 IUD (Other): No new information. 2020-05-28 IUD (Other): No new information. 2020-06-15 email 2020-06-15 08:15, email 2020-06-22 11:45 (Rep): No new information. 2022-05-16 IUD (Other): No new information. | xse damaged schanz arms outside procedure schanz arms damaged iud new iud new email email new iud new kit svc schanz arm | NAV3299 | NAV3296 | NAV3295 | NAV2053 | NAV3256 | NAV3299NAV3296NAV3295NAV2053NAV3256 | True
9303 | 704605841 | XSE - Noise/No Connection to Controller | 2021-09-12T13:55:49.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Noise/No Connection to Controller 2021-Sep-12 Salesforce C00234294 (rep, hcp, for): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while booting the system they heard a noise coming from the internal compartment sounds like a minor "explosion ". The screen was uploaded but they received an error "no connection to controller." They performed hard shutdown and waited a few minutes. but after rebooting same error message was displayed. They opened the cover of the controller and verified that the short cable is connected properly, but they saw that there is no indication of power supply to any of the parts of the electrical panel. They verified that all switches are pressed and all surgical system cables are connected properly. There was no patient harm and a delay of less than one hour. The surgery was aborted. 2021-09-29 IUD x5 (Other, For): No new information. 2021-10-05 IUD x3 (Other): No new information. 2021-10-21 IUD x6, 302611538 (Other): No new information. 2021-12-07 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-Sep-12 Salesforce C00234294 (rep, hcp, for): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while booting the system they heard a noise coming from the internal compartment sounds like a minor "explosion ". The screen was uploaded but they received an error "no connection to controller." They performed hard shutdown and waited a few minutes. but after rebooting same error message was displayed. They opened the cover of the controller and verified that the short cable is connected properly, but they saw that there is no indication of power supply to any of the parts of the electrical panel. They verified that all switches are pressed and all surgical system cables are connected properly. There was no patient harm and a delay of less than one hour. The surgery was aborted. 2021-09-29 IUD x5 (Other, For): No new information. 2021-10-05 IUD x3 (Other): No new information. 2021-10-21 IUD x6, 302611538 (Other): No new information. 2021-12-07 IUD (Other): No new information. | xse noiseno connection controller salesforce spinal procedure booting heard noise coming internal compartment sounds like minor explosion screen uploaded error connection controller performed hard shutdown waited minutes rebooting error message displayed opened cover controller verified short cable connected properly saw indication power supply parts electrical panel verified switches pressed surgical cables connected properly delay hour surgery aborted iud new iud new iud new iud new kit svc cart elec panel assy | NAV2047 | NAV3264 | NAV2057 | NAV3263 | NAV2055 | NAV2047NAV3264NAV2057NAV3263NAV2055 | True
5334 | 703649948 | XSE - Arm Cant Move out of Work Volume | 2020-03-04T20:17:13.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Arm Cant Move out of Work Volume 2020-03-04 00685656 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the arm was sent very close to the patient when taking the oblique image. When the manufacturer representative tried to move the arm away, the arm could not move due to the system saying the arm had entered the working volume. Technical services explained this was normal system operation and the system was functioning correctly. They instructed the representative to use the arrows to move the arm away from the patient before continuing on to operation mode. No patient harm was reported and surgery was delayed less than an hour. 2020-03-13 IUD x4 (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-03-04 00685656 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the arm was sent very close to the patient when taking the oblique image. When the manufacturer representative tried to move the arm away, the arm could not move due to the system saying the arm had entered the working volume. Technical services explained this was normal system operation and the system was functioning correctly. They instructed the representative to use the arrows to move the arm away from the patient before continuing on to operation mode. No patient harm was reported and surgery was delayed less than an hour. 2020-03-13 IUD x4 (Other): No new information. | xse arm work volume id spinal procedure arm sent close taking oblique image tried arm away arm saying arm entered working volume technical services explained normal operation functioning correctly instructed use arrows arm away continuing operation mode surgery delayed hour iud new kit svc surgical arm | NAV4101 | NAV4098 | NAV3249 | NAV4096 | NAV4095 | NAV4101NAV4098NAV3249NAV4096NAV4095 | True
374 | 703089964 | XSE - arm rotated 90 degrees | 2019-04-04T14:04:11.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 40 | ASM0206-01 | SURGICAL ARM | XSE - arm rotated 90 degrees 2019-Apr-02 SCL1900804 intake form, jpgx3 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that Joint 4 was rotated 90 degrees of where it should be. They ran the accuracy test and the pointer was not even close to the accuracy point. They realigned joint 4 and attempted a hard restart. The joint was still rotated. It was noted that the use of the robot was aborted and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Additional information indicated that the replacement arms had Error 3017, "At axis A4. Axis following error.", Task Stess.prg.line: NONe. Modlule: Motion. SURGICAL ARM | 2019-Apr-02 SCL1900804 intake form, jpgx3 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that Joint 4 was rotated 90 degrees of where it should be. They ran the accuracy test and the pointer was not even close to the accuracy point. They realigned joint 4 and attempted a hard restart. The joint was still rotated. It was noted that the use of the robot was aborted and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Additional information indicated that the replacement arms had Error 3017, "At axis A4. Axis following error.", Task Stess.prg.line: NONe. Modlule: Motion. | xse arm rotated degrees spinal procedure joint rotated degrees ran accuracy test pointer close accuracy point realigned joint attempted hard restart joint rotated noted use robot aborted additional complications anticipated additional indicated replacement arms error axis axis following error task stessprgline modlule motion surgical arm | NAV3249 | NAV4092 | NAV4095 | NAV4098 | NAV2054 | NAV3249NAV4092NAV4095NAV4098NAV2054 | True
6503 | 703957497 | X - SRA Vibrating | 2020-09-15T15:56:43.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 30 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | X - SRA Vibrating 2020-09-15 00696638 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was vibrating intensely when the arm moves from the clear up to storage position. The issue was found during testing. There was no patient involvement. 2020-09-18 IUD x5 (Other): No new information. 2020-10-22 IUD x3 (Other): No new information. 2021-03-16 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-09-15 00696638 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was vibrating intensely when the arm moves from the clear up to storage position. The issue was found during testing. There was no patient involvement. 2020-09-18 IUD x5 (Other): No new information. 2020-10-22 IUD x3 (Other): No new information. 2021-03-16 IUD (Other): No new information. | x sra vibrating outside procedure surgical arm vibrating intensely arm moves clear storage position issue testing iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4101 | NAV4095 | NAV2054 | NAV1006 | NAV3249NAV4101NAV4095NAV2054NAV1006 | True
2211 | 703230473 | X - SOFTWARE ISSUE | 2019-06-27T11:56:29.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - SOFTWARE ISSUE When uploading the CT to the WS, error appeared that no images were found. Restarted the WS, was able to upload CT images immediately./ Diagnosis: 65 year old female with lumbar stenosis Planning: Surgeon planned 4 bilateral pedicle screws prior to procedure. Special attention was given to bilateral inferior skive potential at L3 and hypoplastic pedicles at L2. Surgeon wanted to plan screws over prior hardware L4-S1 to determine size of screws placed 14 years ago, and rod length before deleting additional screws. Platform Mounting: Single clamp was chosen per clinical indication and surgeon preference. Clamp was placed on L2 SP. Registration: 3 AP's were taken to remove pain pump from the image, and 2 Oblique images were acquired, using the virtual keyboard to give additional space between the 3D marker and drape. Automatic registration was achieved on first attempt using the basic algorithm, registering off L3. During 3Define the Surgeons Assistant moved the towel causing the software to freeze midscan. XCSR pushed Cntrl+Alt+R but when prompted to restart, WS unfroze, so XCSR chose not to do re-start and simply restarted the 3Define. Operation: Surgeon commenced with exposure prior to using the X. Surgeon wanted to remove prior hardware from L4 prior to registering to avoid shoulder shift intraoperatively. Surgeon used Maxor X to place 4 accurate wires then proceeded with laminectomy and screw placement. All screws stimmed >20. Times: Bed frame adaptor on: 9:20 Arm on bed: 9:47 (1 person) Draping: 9:50 Attachments: 9:54 Incision: 10:04 BMB to clamp: 10:48 3 Define: 10:52, software freeze, went to do a soft restart but it unfroze, so resent. Draw Spine: 10:55, labeled L2. AP: 10:58 OBL: 10:59 (used virtual keyboard to move away from the drape) Registration: 11:05 labeled off L3 Right L2: 11:06 Right L3: 11:07 Left L2: 11:09 Left L3: 11:10 Left L2: 11:15 Remove BMB from clamp: 11:20 Arm removed from bed: 11:24 MAZOR X SOFTWARE KIT VER-2.3.1 | When uploading the CT to the WS, error appeared that no images were found. Restarted the WS, was able to upload CT images immediately./ Diagnosis: 65 year old female with lumbar stenosis Planning: Surgeon planned 4 bilateral pedicle screws prior to procedure. Special attention was given to bilateral inferior skive potential at L3 and hypoplastic pedicles at L2. Surgeon wanted to plan screws over prior hardware L4-S1 to determine size of screws placed 14 years ago, and rod length before deleting additional screws. Platform Mounting: Single clamp was chosen per clinical indication and surgeon preference. Clamp was placed on L2 SP. Registration: 3 AP's were taken to remove pain pump from the image, and 2 Oblique images were acquired, using the virtual keyboard to give additional space between the 3D marker and drape. Automatic registration was achieved on first attempt using the basic algorithm, registering off L3. During 3Define the Surgeons Assistant moved the towel causing the software to freeze midscan. XCSR pushed Cntrl+Alt+R but when prompted to restart, WS unfroze, so XCSR chose not to do re-start and simply restarted the 3Define. Operation: Surgeon commenced with exposure prior to using the X. Surgeon wanted to remove prior hardware from L4 prior to registering to avoid shoulder shift intraoperatively. Surgeon used Maxor X to place 4 accurate wires then proceeded with laminectomy and screw placement. All screws stimmed >20. Times: Bed frame adaptor on: 9:20 Arm on bed: 9:47 (1 person) Draping: 9:50 Attachments: 9:54 Incision: 10:04 BMB to clamp: 10:48 3 Define: 10:52, software freeze, went to do a soft restart but it unfroze, so resent. Draw Spine: 10:55, labeled L2. AP: 10:58 OBL: 10:59 (used virtual keyboard to move away from the drape) Registration: 11:05 labeled off L3 Right L2: 11:06 Right L3: 11:07 Left L2: 11:09 Left L3: 11:10 Left L2: 11:15 Remove BMB from clamp: 11:20 Arm removed from bed: 11:24 | x software issue uploading ct ws error appeared images restarted ws able upload ct images immediately diagnosis year old female lumbar stenosis planning surgeon planned bilateral pedicle screws prior procedure special attention given bilateral inferior skive potential hypoplastic pedicles surgeon wanted plan screws prior hardware determine size screws placed years ago rod length deleting additional screws platform mounting single clamp chosen clinical indication surgeon preference clamp placed sp registration aps taken remove pain pump image oblique images acquired using virtual keyboard additional space marker drape automatic registration achieved attempt using basic algorithm registering surgeons assistant moved towel causing software freeze midscan xcsr pushed cntrlaltr prompted restart ws unfroze xcsr chose restart simply restarted operation surgeon commenced exposure prior using x surgeon wanted remove prior hardware prior registering avoid shoulder shift intraoperatively surgeon maxor x place accurate wires proceeded laminectomy screw placement screws stimmed times bed frame adaptor arm bed draping attachments incision bmb clamp define software freeze went soft restart unfroze resent draw spine labeled ap obl registration labeled right right left left left remove bmb clamp arm removed bed mazor x software kit | NAV3236 | NAV3248 | NAV3242 | NAV3255 | NAV2050 | NAV3236NAV3248NAV3242NAV3255NAV2050 | True
7228 | 704151986 | XSE - Failed Registration | 2021-01-04T13:22:56.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Failed Registration 2020-12-31 00703585, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that planning for a T2-L1 procedure with an open incision was completed by the surgeon and manufacturing representative. For segment one, a clamp was attached to the spinous process of L1 and attached to the bone mount bridge. The clamp was placed at the bottom due to the patient's severe curvature and rotation of the spine. The surgeon believed that clamping higher would cause the tools to hit the bone mount bridge or clamp when the surgical arm was sent to the trajectories. A few AP and oblique images were acquired to attempt to get as many vertebrae as possible in one registration. The first segment included everything from T8-L1. Segmentation and registration was successful with green values at all levels. All trajectories were sent, drilled and screws placed into the pilot holes. All screws stimulated with good values and were shown to be accurate with fluoro imaging. For the second segment, the clamp was repositioned to the T8 spinous process to eliminate the tendency of the tools to hit the bone mount bridge or clamp. A 3Define scan was done and draw spine was completed. Registration was attempted several times after taking several AP and oblique images at different angles. Sometimes, registration would be successful, but then have red and yellow values. Sometime registration would fail and the system would state to check the placement of the red marker. The manufacturer representative noted that the red marker was clearly labeled correctly in every attempt when counting from the T8 screws. After several attempts, the surgeon decided to freehand the rest of the screws. There was no patient harm and the procedure was delayed less than an hour. 2021-Jan-07 e1 IUD (rep): Additional information indicated there was no patient shift. The cause of registration failure was unknown, however after the case was finished, the rep uploaded the scan into the demo software to see if they ever could get it to register, and after much trial and error they were able to label off of L6 with green values at every level. An identical attempt was tried during the operation (before the surgeon decided to abort mazor usage) and was unsuccessful. MAZOR X SYSTEM TPL0059 | 2020-12-31 00703585, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that planning for a T2-L1 procedure with an open incision was completed by the surgeon and manufacturing representative. For segment one, a clamp was attached to the spinous process of L1 and attached to the bone mount bridge. The clamp was placed at the bottom due to the patient's severe curvature and rotation of the spine. The surgeon believed that clamping higher would cause the tools to hit the bone mount bridge or clamp when the surgical arm was sent to the trajectories. A few AP and oblique images were acquired to attempt to get as many vertebrae as possible in one registration. The first segment included everything from T8-L1. Segmentation and registration was successful with green values at all levels. All trajectories were sent, drilled and screws placed into the pilot holes. All screws stimulated with good values and were shown to be accurate with fluoro imaging. For the second segment, the clamp was repositioned to the T8 spinous process to eliminate the tendency of the tools to hit the bone mount bridge or clamp. A 3Define scan was done and draw spine was completed. Registration was attempted several times after taking several AP and oblique images at different angles. Sometimes, registration would be successful, but then have red and yellow values. Sometime registration would fail and the system would state to check the placement of the red marker. The manufacturer representative noted that the red marker was clearly labeled correctly in every attempt when counting from the T8 screws. After several attempts, the surgeon decided to freehand the rest of the screws. There was no patient harm and the procedure was delayed less than an hour. 2021-Jan-07 e1 IUD (rep): Additional information indicated there was no patient shift. The cause of registration failure was unknown, however after the case was finished, the rep uploaded the scan into the demo software to see if they ever could get it to register, and after much trial and error they were able to label off of L6 with green values at every level. An identical attempt was tried during the operation (before the surgeon decided to abort mazor usage) and was unsuccessful. | xse failed registration spinal p rocedure planning procedure open incision completed surgeon manufacturing segment clamp attached spinous process attached bone mount bridge clamp placed patients severe curvature rotation spine surgeon believed clamping higher cause tools hit bone mount bridge clamp surgical arm sent trajectories ap oblique images acquired attempt vertebrae possible registration segment included segmentation registration successful green values levels trajectories sent drilled screws placed pilot holes screws stimulated good values shown accurate fluoro imaging second segment clamp repositioned spinous process eliminate tendency tools hit bone mount bridge clamp scan draw spine completed registration attempted times taking ap oblique images different angles registration successful red yellow values registration fail state check placement red marker noted red marker clearly labeled correctly attempt counting screws attempts surgeon decided freehand rest screws procedure delayed hour iud additional indicated shift cause registration failure unknown case finished uploaded scan demo software register trial error able label green values level identical attempt tried operation unsuccessful mazor x | NAV3242 | NAV4112 | NAV3246 | NAV3243 | NAV2051 | NAV3242NAV4112NAV3246NAV3243NAV2051 | True
2263 | 703231026 | XSE - deviation/work volume | 2019-06-27T16:03:53.000000+0000 | MazorX Stealth Edition | NAV4096 | ARM HIT WORK VOLUME | 30 | ASM0206-01 | SURGICAL ARM | XSE - deviation/work volume 2019-Jun-26 SCL1901851 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that upon first turning on the workstation, the Mazor X tablet was working fine and then they attempted the ten point test and after the first 2 attempts, the tablet screen went black. The rep completed the test, shut down the system and changed outlet. This seemed to have fixed the issue. The monitor touchscreen was also not working throughout the case. After attaching to the patient, they did a 3define scan. During this time they told the surgeon to verify instruments. The rep set up instruments prior to case to case specific and all instruments quickly verified. However, after registration and sending to the first trajectory, none of the instruments were coming up and they needed to re-verify all the instruments. The surgeon had arm sent to L3 left and drilled, placed implant. Sent to L2 and surgeon felt the instruments were lateral. The arm was removed and surgeon used a pedagard to place implant. Arm when moving to next trajectory collided with the BMB. The rep recommended a new 3Define and surgeon agreed. Surgeon then continued and placed all trajectory. There was still a lot of difficulty placing the screw into the arm for snapshot and for the 3dmarker and as being tightened down the surgical system was moving slightly. There was no patient harm. T here were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Jun-26 SCL1901851 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that upon first turning on the workstation, the Mazor X tablet was working fine and then they attempted the ten point test and after the first 2 attempts, the tablet screen went black. The rep completed the test, shut down the system and changed outlet. This seemed to have fixed the issue. The monitor touchscreen was also not working throughout the case. After attaching to the patient, they did a 3define scan. During this time they told the surgeon to verify instruments. The rep set up instruments prior to case to case specific and all instruments quickly verified. However, after registration and sending to the first trajectory, none of the instruments were coming up and they needed to re-verify all the instruments. The surgeon had arm sent to L3 left and drilled, placed implant. Sent to L2 and surgeon felt the instruments were lateral. The arm was removed and surgeon used a pedagard to place implant. Arm when moving to next trajectory collided with the BMB. The rep recommended a new 3Define and surgeon agreed. Surgeon then continued and placed all trajectory. There was still a lot of difficulty placing the screw into the arm for snapshot and for the 3dmarker and as being tightened down the surgical system was moving slightly. There was no patient harm. T here were no additional complications reported or anticipated as a result of the event. | xse deviationwork volume spinal procedure turning workstation mazor x tablet working fine attempted point test attempts tablet screen went black completed test shut changed outlet fixed issue monitor touchscreen working case attaching did scan time told surgeon verify instruments set instruments prior case case specific instruments quickly verified registration sending trajectory instruments coming needed reverify instruments surgeon arm sent left drilled placed implant sent surgeon felt instruments lateral arm removed surgeon pedagard place implant arm moving trajectory collided bmb recommended new surgeon agreed surgeon continued placed trajectory lot difficulty placing screw arm snapshot tightened surgical moving slightly t additional complications anticipated surgical arm | NAV3253 | NAV5000 | NAV1012 | NAV4096 | NAV4101 | NAV3253NAV5000NAV1012NAV4096NAV4101 | True
678 | 703170262 | X - Workstation/Surgical System Issues | 2019-05-23T21:33:39.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 60 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Workstation/Surgical System Issues - 2.2.1.14 build of the SW installed successfully and Site ATP initiated. - Shoulder issues - experienced many issues with the shoulder locking and unlocking which eventually lead to the decision not to initiate surgery with the Mazor-X and revert to the RNS. Detailed description of the issue: o Many variations of discrepancy between the actual state of the shoulder (locked/unlocked), the LED light indication for the state of the shoulder and the software indication for the state of the shoulder. o The shoulder would spontaneously 'unlock' on the LED light indication and the software indication even though it was actually locked. This generated issues that manifested in many variations and didn’t allow advancing in the operation flow. The issue was explained in detail to the R&D team and troubleshooting attempts ensued. It is believed to be a hardware issue with the pneumatic piston of the shoulder – piston sensors and air-flow are believed to be causing the problems. It was decided by Eli that later this week an R&D team m ember will travel to THP to resolve the issue and if successful Asaf will probably travel there to support cases. - Surgical Arm setup and advanced setup – Setup and advance setup were performed many times: o Accuracy was very good both with the setup arm guide and the intra-op arm guide – at times the pointer would brush against the side of the hole but was accurately in it. o Setup was smooth every time but the following issue was experienced with the advanced setup: out of 10-15 times advanced setup was performed, there were 1-2 trajectories on average that the surgical arm 'couldn’t reach' and after the error message was 'approved' it continued to the subsequent trajectory (recurring known SW bug). - Tablet: o Touch is far from working smoothly – sometimes unresponsive even if clicked 5 times with different intensity. This was experienced without a drape and with the naked finger. It is also frequently slow to react. o The charging is very slow and works abruptly from the storage position and practically inexistent from the intraoperative position – even before draping. - Surgical-Arm issues – a couple of times we experienced the following: the Arm would not move at all after getting stuck in the middle of motion. The first time a 'hard-restart' enabled motion and the second time it had to be m oved to storage position from the ICO controller interface directly after three 'hard-restarts' didn’t work. - Accessories: o 'Older' version of the 5mm-shaft drill-bits were removed from the site and replaced with 'newer' version – long 80mm, short 80mm and short 2.4mm. Their interface with the RN drill-guides was manually checked and found to be satisfactory. o Only one arm-tube for the SA tools was left on site. Its' interface with the cannulas of one of the KIT0150 on site was tested and found to be satisfactory. o The navigation stylus (the older version, not PEEK) is in the Mazor-X accessories tray. - WS issues: o The front handle is disassembled for some reason – it can be found in the black cart next to the WS o The manipulator – surgical system interface is not in the best shape – there is a rotational movement/freedom of a few degrees in the connection (where the 'mushrooms' attach to the surgical system). o Occasionally, and more often when the compressor stops working, there is a drop in the system pressure and the manipulat0or with the surgical system 'collapses' a little bit. Increasing the pneumatic pressure didn’t resolve this. o The wheels, especially the front ones, don’t rotate smoothly and moving the WS around requires much effort. UNK MAZOR X WORKSTATION | - 2.2.1.14 build of the SW installed successfully and Site ATP initiated. - Shoulder issues - experienced many issues with the shoulder locking and unlocking which eventually lead to the decision not to initiate surgery with the Mazor-X and revert to the RNS. Detailed description of the issue: o Many variations of discrepancy between the actual state of the shoulder (locked/unlocked), the LED light indication for the state of the shoulder and the software indication for the state of the shoulder. o The shoulder would spontaneously 'unlock' on the LED light indication and the software indication even though it was actually locked. This generated issues that manifested in many variations and didn’t allow advancing in the operation flow. The issue was explained in detail to the R&D team and troubleshooting attempts ensued. It is believed to be a hardware issue with the pneumatic piston of the shoulder – piston sensors and air-flow are believed to be causing the problems. It was decided by Eli that later this week an R&D team m ember will travel to THP to resolve the issue and if successful Asaf will probably travel there to support cases. - Surgical Arm setup and advanced setup – Setup and advance setup were performed many times: o Accuracy was very good both with the setup arm guide and the intra-op arm guide – at times the pointer would brush against the side of the hole but was accurately in it. o Setup was smooth every time but the following issue was experienced with the advanced setup: out of 10-15 times advanced setup was performed, there were 1-2 trajectories on average that the surgical arm 'couldn’t reach' and after the error message was 'approved' it continued to the subsequent trajectory (recurring known SW bug). - Tablet: o Touch is far from working smoothly – sometimes unresponsive even if clicked 5 times with different intensity. This was experienced without a drape and with the naked finger. It is also frequently slow to react. o The charging is very slow and works abruptly from the storage position and practically inexistent from the intraoperative position – even before draping. - Surgical-Arm issues – a couple of times we experienced the following: the Arm would not move at all after getting stuck in the middle of motion. The first time a 'hard-restart' enabled motion and the second time it had to be m oved to storage position from the ICO controller interface directly after three 'hard-restarts' didn’t work. - Accessories: o 'Older' version of the 5mm-shaft drill-bits were removed from the site and replaced with 'newer' version – long 80mm, short 80mm and short 2.4mm. Their interface with the RN drill-guides was manually checked and found to be satisfactory. o Only one arm-tube for the SA tools was left on site. Its' interface with the cannulas of one of the KIT0150 on site was tested and found to be satisfactory. o The navigation stylus (the older version, not PEEK) is in the Mazor-X accessories tray. - WS issues: o The front handle is disassembled for some reason – it can be found in the black cart next to the WS o The manipulator – surgical system interface is not in the best shape – there is a rotational movement/freedom of a few degrees in the connection (where the 'mushrooms' attach to the surgical system). o Occasionally, and more often when the compressor stops working, there is a drop in the system pressure and the manipulat0or with the surgical system 'collapses' a little bit. Increasing the pneumatic pressure didn’t resolve this. o The wheels, especially the front ones, don’t rotate smoothly and moving the WS around requires much effort. | x workstationsurgical issues build sw installed successfully site atp initiated shoulder issues experienced issues shoulder locking unlocking eventually lead decision initiate surgery mazorx revert rns detailed description issue o variations discrepancy actual state shoulder led light indication state shoulder software indication state shoulder o shoulder spontaneously unlock led light indication software indication actually locked generated issues manifested variations didn’t allow advancing operation flow issue explained rd team troubleshooting attempts ensued believed hardware issue pneumatic piston shoulder – piston sensors airflow believed causing problems decided eli later week rd team m ember travel thp resolve issue successful asaf probably travel support cases surgical arm setup advanced setup – setup advance setup performed times o accuracy good setup arm guide intraop arm guide – times pointer brush hole accurately o setup smooth time following issue experienced advanced setup times advanced setup performed trajectories average surgical arm couldn’t reach error message approved continued subsequent trajectory tablet o touch far working smoothly – unresponsive clicked times different intensity experienced drape naked finger frequently slow react o charging slow works abruptly storage position practically inexistent intraoperative position – draping surgicalarm issues – couple times experienced following arm getting stuck middle motion time hardrestart enabled motion second time m oved storage position ico controller interface directly hardrestarts didn’t work accessories o older version drillbits removed site replaced newer version – long short short interface rn drillguides manually checked satisfactory o armtube sa tools left site interface cannulas site tested satisfactory o navigation stylus mazorx accessories tray ws issues o handle disassembled reason – black cart ws o manipulator – surgical interface best shape – rotational movementfreedom degrees connection o occasionally compressor stops working drop pressure surgical collapses little bit increasing pneumatic pressure didn’t resolve o wheels especially ones don’t rotate smoothly moving ws requires effort unk mazor x workstation | NAV3264 | NAV4095 | NAV4107 | NAV2058 | NAV3253 | NAV3264NAV4095NAV4107NAV2058NAV3253 | True
9519 | 704662617 | XSE - SW Froze/SRA Hit BMB | 2021-10-14T10:15:09.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 30 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - SW Froze/SRA Hit BMB 2021-Oct-12 Salesforce C00236230, C00236233 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported three segments were planned for a L3-%5 procedure. The first segment from L3-T12 and second segment from T11-T6 were completed successfully. The third segment from T5-T3 was challenging when acquiring a good oblique image. Once a good oblique image was taken, the surgical arm was lowered to one of the trajectories, the elbow collided with the bone mount bridge. The surgeon decided to redo registration. The second registration attempt was very challenging so the surgeon decided to switch to Scan & Plan. A successful O-arm scan was completed, but when trying to import the scan, the software froze and did not respond. A soft restart was done and the operation was entered again. The previous 3Define scan was not saved so a new generic work volume was created. The snapshot was done and the previous scan was imported. The star marker was located and planning was completed. When sending the surgical arm to the first trajectory, the arm was out of range and not on the trajectory. The surgeon realized this was due to the fact that the second time after rebooting, the surgical arm was not in the same place as when the scan was previously taken. The surgeon did not want to do a second scan so the use of the guidance system was aborted. The remaining three vertebral bodies were completed freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-10-18 e1 (Rep, For): No new information. 2021-11-02 email 2021-11-02 05:15 (Rep, For): Additional information received from a manufacturer representative reported that it was unknown how much the surgical arm was off. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Oct-12 Salesforce C00236230, C00236233 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported three segments were planned for a L3-%5 procedure. The first segment from L3-T12 and second segment from T11-T6 were completed successfully. The third segment from T5-T3 was challenging when acquiring a good oblique image. Once a good oblique image was taken, the surgical arm was lowered to one of the trajectories, the elbow collided with the bone mount bridge. The surgeon decided to redo registration. The second registration attempt was very challenging so the surgeon decided to switch to Scan & Plan. A successful O-arm scan was completed, but when trying to import the scan, the software froze and did not respond. A soft restart was done and the operation was entered again. The previous 3Define scan was not saved so a new generic work volume was created. The snapshot was done and the previous scan was imported. The star marker was located and planning was completed. When sending the surgical arm to the first trajectory, the arm was out of range and not on the trajectory. The surgeon realized this was due to the fact that the second time after rebooting, the surgical arm was not in the same place as when the scan was previously taken. The surgeon did not want to do a second scan so the use of the guidance system was aborted. The remaining three vertebral bodies were completed freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-10-18 e1 (Rep, For): No new information. 2021-11-02 email 2021-11-02 05:15 (Rep, For): Additional information received from a manufacturer representative reported that it was unknown how much the surgical arm was off. | xse sw frozesra hit bmb salesforce spinal procedure segments planned procedure segment second segment completed successfully segment challenging acquiring good oblique image good oblique image taken surgical arm lowered trajectories elbow collided bone mount bridge surgeon decided redo registration second registration attempt challenging surgeon decided switch scan plan successful oarm scan completed trying import scan software froze did respond soft restart operation entered previous scan saved new generic work volume created snapshot previous scan imported star marker located planning completed sending surgical arm trajectory arm range trajectory surgeon realized fact second time rebooting surgical arm place scan previously taken surgeon did want second scan use aborted remaining vertebral bodies completed freehand procedure delayed hour new email additional unknown surgical arm mazorx sw installation kit | NAV3246 | NAV3248 | NAV4096 | NAV5000 | NAV3239 | NAV3246NAV3248NAV4096NAV5000NAV3239 | True
4517 | 703548568 | X - 3Define Failed | 2020-01-07T17:21:17.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define Failed 2018-10-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure T11-L3. It was reported that the 3 Define scan failed. All trajectories were reached accurately. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X 3DEFINE CAMERA | 2018-10-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure T11-L3. It was reported that the 3 Define scan failed. All trajectories were reached accurately. There was no patient harm and the procedure was not delayed over an hour. | x failed ccr spinal procedure define scan failed trajectories reached accurately procedure delayed hour unk mazor x camera | NAV3244 | NAV3264 | NAV3235 | NAV3239 | NAV4095 | NAV3244NAV3264NAV3235NAV3239NAV4095 | True
7463 | 704210887 | XSE - 3define Camera Failure | 2021-02-03T17:41:39.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 30 | KIT0749-01 | KIT SPARE PART FOR THE 3D CAMERA | XSE - 3define Camera Failure 2021-Feb-02 Salesforce C00220387 (rep): Medtronic received information regarding a guidance system being used in a procedure. It was reported that the 3define camera failed. They did not perform a reboot, they used a generic work volume to complete the case. There was no patient involved. 2021-Feb-16 IUDx2 (other): No new information. 2021-Mar-16 IUD (other): No new information. 2021-04-13 IUD x2 (Other): No new information. KIT SPARE PART FOR THE 3D CAMERA | 2021-Feb-02 Salesforce C00220387 (rep): Medtronic received information regarding a guidance system being used in a procedure. It was reported that the 3define camera failed. They did not perform a reboot, they used a generic work volume to complete the case. There was no patient involved. 2021-Feb-16 IUDx2 (other): No new information. 2021-Mar-16 IUD (other): No new information. 2021-04-13 IUD x2 (Other): No new information. | xse camera failure salesforce procedure camera failed did perform reboot generic work volume complete case involved new iud new iud new kit spare camera | NAV3244 | NAV3246 | NAV4112 | NAV4095 | NAV2054 | NAV3244NAV3246NAV4112NAV4095NAV2054 | True
723 | 703174828 | x - X-eye issue | 2019-05-27T13:51:12.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | SEN0030-01 | SENSOR | x - X-eye issue This complaint will be separated to two complaint. This complaint will handle the X-eye issue. The ar m did not reached location will be handled in: Comp-0000515 The x-eye was not working. There was no video feed under the camera tabs. No green power light on the front of the X-eye. CS checked the connections in the cart but it did not resolve the issue. When moving the arm to clear up received error cant reach predefined location. Was able to move the arm to clear right and left without error. SENSOR | This complaint will be separated to two complaint. This complaint will handle the X-eye issue. The ar m did not reached location will be handled in: Comp-0000515 The x-eye was not working. There was no video feed under the camera tabs. No green power light on the front of the X-eye. CS checked the connections in the cart but it did not resolve the issue. When moving the arm to clear up received error cant reach predefined location. Was able to move the arm to clear right and left without error. | x xeye issue complaint separated complaint complaint handle xeye issue ar m did reached location handled xeye working video feed camera tabs green power light xeye cs checked connections cart did resolve issue moving arm clear error reach predefined location able arm clear right left error sensor | NAV4110 | NAV4095 | NAV4109 | NAV3254 | NAV2058 | NAV4110NAV4095NAV4109NAV3254NAV2058 | True
1506 | 703202822 | X - Software Issue | 2019-06-12T19:06:48.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - Software Issue Separated from SCL1801145 - this complaint handel the SW freeze Generic AP sent - Surgeon tilted forward headlight close to rbt arm - RTC pressed emergency stop to not interfere with surgeon - screen froze on processing for over 1min RTC soft restart - Screen backed out and 3-Define needed to be completed again - Reference frame now seen Screen then blacked out after 3 of 4 of the trajectories and returned to home screen - There was NO bovie equipment use during this time. MAZOR X SOFTWARE KIT VER-2.3.1 | Separated from SCL1801145 - this complaint handel the SW freeze Generic AP sent - Surgeon tilted forward headlight close to rbt arm - RTC pressed emergency stop to not interfere with surgeon - screen froze on processing for over 1min RTC soft restart - Screen backed out and 3-Define needed to be completed again - Reference frame now seen Screen then blacked out after 3 of 4 of the trajectories and returned to home screen - There was NO bovie equipment use during this time. | x software issue separated complaint handel sw freeze generic ap sent surgeon tilted forward headlight close rbt arm rtc pressed emergency stop interfere surgeon screen froze processing rtc soft restart screen backed needed completed reference frame seen screen blacked trajectories returned home screen bovie equipment use time mazor x software kit | NAV3248 | NAV3247 | NAV1006 | NAV2047 | NAV3246 | NAV3248NAV3247NAV1006NAV2047NAV3246 | True
4868 | 703583068 | X - Deviations (Left & T11R), Setup Fail | 2020-01-27T19:48:47.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviations (Left & T11R), Setup Fail 2018-02-16 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure for posterior spinal fusion at T11-L3 and L2 Kyphoplasty. It was reported that planning was completed preoperatively by the surgeon. The manufacturer representative and surgeon reviewed the plan with special attention paid to skin incisions, screw alignment, and rod alignment. The representative and the surgeon spoke before the case and decided to use dual Schanz pins with the dynamic Schanz arms for the system mount. This would reach the top of the construct and provide the most stability. The representative attempted a 10 point test and the arm was shaking as it was moving to position. The representative went to superuser and reset movements and gains. The system was then shut down, waited for 5 min, and turned back on. The representative then performed a 10 point test with no issues. A 3 point test was accurate as well as the shoulder calibration was checked. System was put on the bed easily. CSR assisted the scrub tech with draping instructions. Dynamic Schanz was put on either side of the arm. Two Schanz pins were inserted by the surgeon into the PSIS. The system was brought into place and the dynamic Schanz arms were inserted into both Schanz pins. The system was tightened. 3Define was performed. Drawspine was performed at L3. The representative asked the surgeon to create a larger DrawSpine and they did so. Sent to the AP. Sent to the Oblique and got images. For the first registration, 1 AP and 2 Oblique shots were necessary to achieve registration. Labeling was performed of of L3 using the Basic Algorithm. Green values were seen at each level. The representative and surgeon confirmed no shift. Sent to T11 left. As the surgeon was inserting the knife handle through the arm guide, dry blood came out of the arm guide. The surgeon asked for the nurse and told her what was going on. The representative told the surgeon that they would have to undrape everything and get a new instrument set for the system. The surgeon decided to clean the arm guide and proceed with the case. K-wires were inserted into all trajectories on the left side. At T11 right, the surgeon noted that the reduction tube did not feel right. Lateral and AP shots were taken. It was noted that T11 right was lateral, and every trajectory on the left side had deviated inferior. The surgeon removed all k-wires. The surgeon wished to re-register the system. New images were acquired. For the second registration, 2 AP and 2 Oblique shots were necessary to achieve registration. Labeling was performed off of T12 using the Basic Algorithm. Green values were seen at each level. Representative and surgeon confirmed no shift. The representative sent to T11 left. Lateral and AP shots were taken. The surgeon was pleased with the placement. The surgeon placed k-wires at T12 left, L2 left, and L3 left. Another lateral and AP shot were taken. All 3 of these wires had deviated inferiorly. The surgeon removed the k-wires and freehanded the rest of the trajectories. They also asked the representative to send the case to Israel and get a response on what occurred. MAZOR X SYSTEM TPL0059 | 2018-02-16 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure for posterior spinal fusion at T11-L3 and L2 Kyphoplasty. It was reported that planning was completed preoperatively by the surgeon. The manufacturer representative and surgeon reviewed the plan with special attention paid to skin incisions, screw alignment, and rod alignment. The representative and the surgeon spoke before the case and decided to use dual Schanz pins with the dynamic Schanz arms for the system mount. This would reach the top of the construct and provide the most stability. The representative attempted a 10 point test and the arm was shaking as it was moving to position. The representative went to superuser and reset movements and gains. The system was then shut down, waited for 5 min, and turned back on. The representative then performed a 10 point test with no issues. A 3 point test was accurate as well as the shoulder calibration was checked. System was put on the bed easily. CSR assisted the scrub tech with draping instructions. Dynamic Schanz was put on either side of the arm. Two Schanz pins were inserted by the surgeon into the PSIS. The system was brought into place and the dynamic Schanz arms were inserted into both Schanz pins. The system was tightened. 3Define was performed. Drawspine was performed at L3. The representative asked the surgeon to create a larger DrawSpine and they did so. Sent to the AP. Sent to the Oblique and got images. For the first registration, 1 AP and 2 Oblique shots were necessary to achieve registration. Labeling was performed of of L3 using the Basic Algorithm. Green values were seen at each level. The representative and surgeon confirmed no shift. Sent to T11 left. As the surgeon was inserting the knife handle through the arm guide, dry blood came out of the arm guide. The surgeon asked for the nurse and told her what was going on. The representative told the surgeon that they would have to undrape everything and get a new instrument set for the system. The surgeon decided to clean the arm guide and proceed with the case. K-wires were inserted into all trajectories on the left side. At T11 right, the surgeon noted that the reduction tube did not feel right. Lateral and AP shots were taken. It was noted that T11 right was lateral, and every trajectory on the left side had deviated inferior. The surgeon removed all k-wires. The surgeon wished to re-register the system. New images were acquired. For the second registration, 2 AP and 2 Oblique shots were necessary to achieve registration. Labeling was performed off of T12 using the Basic Algorithm. Green values were seen at each level. Representative and surgeon confirmed no shift. The representative sent to T11 left. Lateral and AP shots were taken. The surgeon was pleased with the placement. The surgeon placed k-wires at T12 left, L2 left, and L3 left. Another lateral and AP shot were taken. All 3 of these wires had deviated inferiorly. The surgeon removed the k-wires and freehanded the rest of the trajectories. They also asked the representative to send the case to Israel and get a response on what occurred. | x deviations setup fail ccr spinal procedure posterior spinal fusion kyphoplasty planning completed preoperatively surgeon surgeon reviewed plan special attention paid skin incisions screw alignment rod alignment surgeon spoke case decided use dual schanz pins dynamic schanz arms mount reach construct provide stability attempted point test arm shaking moving position went superuser reset movements gains shut waited min turned performed point test issues point test accurate shoulder calibration checked bed easily csr assisted scrub tech draping instructions dynamic schanz arm schanz pins inserted surgeon psis brought place dynamic schanz arms inserted schanz pins tightened performed drawspine performed asked surgeon create larger drawspine did sent ap sent oblique got images registration ap oblique shots necessary achieve registration labeling performed using basic algorithm green values seen level surgeon confirmed shift sent left surgeon inserting knife handle arm guide dry blood came arm guide surgeon asked nurse told going told surgeon undrape new instrument set surgeon decided clean arm guide proceed case kwires inserted trajectories left right surgeon noted reduction tube did feel right lateral ap shots taken noted right lateral trajectory left deviated inferior surgeon removed kwires surgeon wished reregister new images acquired second registration ap oblique shots necessary achieve registration labeling performed using basic algorithm green values seen level surgeon confirmed shift sent left lateral ap shots taken surgeon pleased placement surgeon placed kwires left left left lateral ap shot taken wires deviated inferiorly surgeon removed kwires freehanded rest trajectories asked send case israel response occurred mazor x | NAV4091 | NAV2046 | NAV4093 | NAV5000 | NAV4101 | NAV4091NAV2046NAV4093NAV5000NAV4101 | True
5214 | 703619117 | XSE - Damage During Transport/WS Issue | 2020-02-17T16:06:16.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 70 | MEC5057-01S | KIT SVC MEC5057-01S RIGHT CURTAIN COVER | XSE - Damage During Transport/WS Issue 2020-02-12 00684214, Intake Form (REP, For): Medtronic received information regarding a guidance system outside of procedure. It was reported that the although the front side wheel was broken, after opening the wooden box they found that the remaining one front side wheel was bent. The monitor screen also got damaged. There was no patient involvement. 2020-04-17 email 2020-04-17 00:45 (Rep, For): Additional information received from a manufacturer representative reported that the issue was not resolved yet. 2021-03-10 IUD (Other, For): No new information. 2022-05-17 IUD (Other, For): No new information. KIT SVC MEC5057-01S RIGHT CURTAIN COVER | 2020-02-12 00684214, Intake Form (REP, For): Medtronic received information regarding a guidance system outside of procedure. It was reported that the although the front side wheel was broken, after opening the wooden box they found that the remaining one front side wheel was bent. The monitor screen also got damaged. There was no patient involvement. 2020-04-17 email 2020-04-17 00:45 (Rep, For): Additional information received from a manufacturer representative reported that the issue was not resolved yet. 2021-03-10 IUD (Other, For): No new information. 2022-05-17 IUD (Other, For): No new information. | xse damage transportws issue outside procedure wheel broken opening wooden box remaining wheel bent monitor screen got damaged email additional issue resolved iud new iud new kit svc right curtain cover | NAV2055 | NAV2058 | NAV2054 | NAV3261 | NAV3251 | NAV2055NAV2058NAV2054NAV3261NAV3251 | True
6589 | 703982205 | XSE - X-Eye Not Functioning | 2020-09-29T00:03:53.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 30 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - X-Eye Not Functioning 2020-09-28 00697603 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the X-eye camera was not connecting or functioning. The issue was discovered prior to the a procedure. There was no patient involvement. 2020-10-02 IUD x3 (Other): No new information. 2020-11-04 IUD x6 (Other): No new information. 2021-06-20 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2020-09-28 00697603 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the X-eye camera was not connecting or functioning. The issue was discovered prior to the a procedure. There was no patient involvement. 2020-10-02 IUD x3 (Other): No new information. 2020-11-04 IUD x6 (Other): No new information. 2021-06-20 IUD (Other): No new information. | xse xeye functioning outside procedure xeye camera connecting functioning issue discovered prior procedure iud new iud new iud new mazor x pc station refurbished | NAV4110 | NAV3262 | NAV2054 | NAV3242 | NAV1006 | NAV4110NAV3262NAV2054NAV3242NAV1006 | True
9205 | 704574794 | XSE - Schanz Bond Screw Stripped | 2021-08-25T01:47:07.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MEC4530-02S | KIT SVC MEC4530-02S BALL ARM TRGT SCREW | XSE - Schanz Bond Screw Stripped 2021-Aug-24 Salesforce C00233440 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz bond screw was stripped. There was no patient involvement. 2021-09-03 IUD x5 (Other): No new information. 2021-09-20 email 2021-09-20 15:45 (Rep): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-08 email 2021-10-08 16:45 (Rep): No new information. 2021-10-26 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 202-12-05 IUD x3 (Other): No new information. KIT SVC MEC4530-02S BALL ARM TRGT SCREW | 2021-Aug-24 Salesforce C00233440 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz bond screw was stripped. There was no patient involvement. 2021-09-03 IUD x5 (Other): No new information. 2021-09-20 email 2021-09-20 15:45 (Rep): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-08 email 2021-10-08 16:45 (Rep): No new information. 2021-10-26 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 202-12-05 IUD x3 (Other): No new information. | xse schanz bond screw stripped salesforce outside procedure schanz bond screw stripped iud new email new iud new email new iud new iud new iud new kit svc ball arm trgt screw | NAV3299 | NAV4090 | NAV4100 | NAV3295 | NAV2047 | NAV3299NAV4090NAV4100NAV3295NAV2047 | True
5096 | 703597189 | XSE - CANBUS Error | 2020-02-04T20:12:53.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 30 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | XSE - CANBUS Error 2020-02-04 00683379 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that following mounting the shoulder, a CANBUS error was seen immediately. The case was able to continue but troubleshooting was needed to remove the system from the bed. First technical services had the manufacturer representative log into superuser and go into service tools. All status bars were green except for positioner IO 10 and 11. However, this is a type I system so red for 11 is normal, but red for 10 does not. The representative was instructed to reset the movement and send the arm to a predefined position. Then they were instructed to go to the general tab and try to open the shoulder, positioner y and z. Positioner y opened, but the shoulder and z did not. When attempting to open the shoulder through the overrides panel, it again did not open. The right side of the panel was removed so the electric panel at the breaker could be seen. One of the breakers popped, so the system was powered off and the breaker was reset. After waiting 5 mins, the system was powered on again and it was able to be used. It is likely that there is a short in the positioner and a FSE will be sent to place a replacement surgical system. There was no patient harm and the procedure was not delayed over an hour. 2020-02-06 IUD x2 (Rep): No new information. UNK MAZOR X POSITIONER | 2020-02-04 00683379 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that following mounting the shoulder, a CANBUS error was seen immediately. The case was able to continue but troubleshooting was needed to remove the system from the bed. First technical services had the manufacturer representative log into superuser and go into service tools. All status bars were green except for positioner IO 10 and 11. However, this is a type I system so red for 11 is normal, but red for 10 does not. The representative was instructed to reset the movement and send the arm to a predefined position. Then they were instructed to go to the general tab and try to open the shoulder, positioner y and z. Positioner y opened, but the shoulder and z did not. When attempting to open the shoulder through the overrides panel, it again did not open. The right side of the panel was removed so the electric panel at the breaker could be seen. One of the breakers popped, so the system was powered off and the breaker was reset. After waiting 5 mins, the system was powered on again and it was able to be used. It is likely that there is a short in the positioner and a FSE will be sent to place a replacement surgical system. There was no patient harm and the procedure was not delayed over an hour. 2020-02-06 IUD x2 (Rep): No new information. | xse canbus error spinal procedure following mounting shoulder canbus error seen immediately case able continue troubleshooting needed remove bed technical services log superuser service tools status bars green positioner io type red normal red does instructed reset movement send arm predefined position instructed general tab try open shoulder positioner y z positioner y opened shoulder z did attempting open shoulder overrides panel did open right panel removed electric panel breaker seen breakers popped powered breaker reset waiting mins powered able likely short positioner fse sent place replacement surgical procedure delayed hour iud new unk mazor x positioner | NAV4107 | NAV3251 | NAV4103 | NAV4104 | NAV4105 | NAV4107NAV3251NAV4103NAV4104NAV4105 | True
9004 | 704524756 | XSE ¦ Damaged Stylus | 2021-07-27T20:14:52.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2176-04S | KIT SVC MAS2176-04S STYLUS MARKER | XSE ¦ Damaged Stylus 2021-Jul-27 Salesforce C00230926 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the stylus was damaged. No patient was present. 2021-08-03 IUD x2 (Other): No new information. 2021-Sep-02 email 9/2 (other): It was damaged in the washer. The site discarded the stylus. KIT SVC MAS2176-04S STYLUS MARKER | 2021-Jul-27 Salesforce C00230926 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the stylus was damaged. No patient was present. 2021-08-03 IUD x2 (Other): No new information. 2021-Sep-02 email 9/2 (other): It was damaged in the washer. The site discarded the stylus. | xse damaged stylus salesforce outside procedure stylus damaged present iud new email damaged washer site discarded stylus kit svc stylus marker | NAV3299 | NAV2053 | NAV3295 | NAV2058 | NAV3264 | NAV3299NAV2053NAV3295NAV2058NAV3264 | True
6269 | 703902254 | XSE - CANBUS Error | 2020-08-13T15:31:00.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error 2020-08-11 00694530, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error was displayed and the surgeon decided to abort the use of the guidance system. The surgeon screen was black and not working properly during the procedure. During the S1 screw on the left side, the surgeon was using bovie with the cannula just lifted up a little. The manufacturer representative noted that the bovie appeared to arch to the tool and and cause the CANBUS error. There was no patient harm and the procedure was delayed less than an hour. 2020-08-18 IUD x3 (Other): No new information. 2020-08-25 IUD (Other): No new information. 2020-08-28 IUD x3 (Other): No new information. 2020-09-09 email 2020-09-09 00:00, IUD x4 (Other): No new information. 2020-09-30 IUD x2 (Other): No new information. 2020-11-06 email 2020-11-06 13:15 (Rep): No new information. 2020-12-03 email 2020-12-03 11:00, email 2020-12-14 20:45 (Rep): Additional information received from a manufacturer representative reported that the use of the guidance system was aborted due to the CANBUS error and surgical arm shift with two screws left. The procedure was successfully completed freehand. All screws placed with the guidance system were accurate. SVC KIT1317 SYSTEM MAIN CABLE | 2020-08-11 00694530, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error was displayed and the surgeon decided to abort the use of the guidance system. The surgeon screen was black and not working properly during the procedure. During the S1 screw on the left side, the surgeon was using bovie with the cannula just lifted up a little. The manufacturer representative noted that the bovie appeared to arch to the tool and and cause the CANBUS error. There was no patient harm and the procedure was delayed less than an hour. 2020-08-18 IUD x3 (Other): No new information. 2020-08-25 IUD (Other): No new information. 2020-08-28 IUD x3 (Other): No new information. 2020-09-09 email 2020-09-09 00:00, IUD x4 (Other): No new information. 2020-09-30 IUD x2 (Other): No new information. 2020-11-06 email 2020-11-06 13:15 (Rep): No new information. 2020-12-03 email 2020-12-03 11:00, email 2020-12-14 20:45 (Rep): Additional information received from a manufacturer representative reported that the use of the guidance system was aborted due to the CANBUS error and surgical arm shift with two screws left. The procedure was successfully completed freehand. All screws placed with the guidance system were accurate. | xse canbus error iud spinal procedure canbus error displayed surgeon decided abort use surgeon screen black working properly procedure screw left surgeon using bovie cannula just lifted little noted bovie appeared arch tool cause canbus error procedure delayed hour iud new iud new iud new email iud new iud new email new email email additional use aborted canbus error surgical arm shift screws left procedure successfully completed freehand screws placed accurate svc main cable | NAV2047 | NAV4098 | NAV2054 | NAV3253 | NAV1006 | NAV2047NAV4098NAV2054NAV3253NAV1006 | True
8850 | 704498691 | XSE - Joint Shift Error | 2021-07-13T16:06:36.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Joint Shift Error 2021-Jul-13 Salesforce C00230099 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative ran a pre-surgery checklist, and it said a joint shift was detected on joint 3. He was unable to clear the joint shift error and stated that the are needs to be replaced. No patient was present. 2021-07-16 IUD x2 (Other): No new information. 2021-09-21 email 2021-09-21 11:46 (Rep): No new information. 2021-09-24 IUD (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-10-21 302605194, IUD (Other): No new information. 2021-10-25 IUD x6 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2021-Jul-13 Salesforce C00230099 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative ran a pre-surgery checklist, and it said a joint shift was detected on joint 3. He was unable to clear the joint shift error and stated that the are needs to be replaced. No patient was present. 2021-07-16 IUD x2 (Other): No new information. 2021-09-21 email 2021-09-21 11:46 (Rep): No new information. 2021-09-24 IUD (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-10-21 302605194, IUD (Other): No new information. 2021-10-25 IUD x6 (Other): No new information. | xse joint shift error salesforce outside procedure ran presurgery checklist said joint shift detected joint unable clear joint shift error stated needs replaced present iud new email new iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV4098 | NAV2054 | NAV3251 | NAV4101 | NAV4095 | NAV4098NAV2054NAV3251NAV4101NAV4095 | True
9302 | 704605831 | XSE - Black Surgeon Monitor | 2021-09-12T09:55:33.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE - Black Surgeon Monitor 2021-Sep-11 Salesforce C00234292 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon monitor was black. A reboot did not resolve the issue. This occurred after a power surge last night in the mobile lab. No patient was present. 2021-09-15 IUD x2 (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-12 IDU x2 (Other): No new information. 2021-12-09 IUD x2 (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2021-Sep-11 Salesforce C00234292 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon monitor was black. A reboot did not resolve the issue. This occurred after a power surge last night in the mobile lab. No patient was present. 2021-09-15 IUD x2 (Other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-12 IDU x2 (Other): No new information. 2021-12-09 IUD x2 (Other): No new information. | xse black surgeon monitor salesforce outside procedure surgeon monitor black reboot did resolve issue occurred power surge night mobile lab present iud new iud new idu new iud new kit svc srgn screen arm | NAV3253 | NAV3261 | NAV3249 | NAV3247 | NAV2047 | NAV3253NAV3261NAV3249NAV3247NAV2047 | True
9705 | 704701638 | XSE - Skiving/Unable to Register | 2021-11-05T22:31:55.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Skiving/Unable to Register 2021-Oct-26 Salesforce C00237474 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there were registration issues during a L3/L4 and S1-S3 procedure. Registration was successfully completed for the first segment from L3-L4. Screws were placed at L3, but there was a lot of skiving at L4 due to holes from a previous procedure. The surgeon moved forward with the case, but the second segment for the pelvis region was unable to be registered. There was no patient harm and the procedure was delayed less than an hour. 2021-11-10 IUD, e1 (Rep): Additional information received from a manufacturer representative reported the L3 screw was lower in the pedicle than planned. 2021-11-15 email 2021-11-15 19:45 (Rep): No new information. 2021-11-17 IUD x2 (Other): No new information. 2021-12-02 e1, e2 (Rep): Additional information received from a manufacturer representative reported the screw were placed percutane ously and a bone mount bridge connected to a PSIS pin placed in the right PSIS was used as the mount. 2021-12-06 e3, e4 (Rep): Additional information received from a manufacturer representative reported that the lower levels were open. The surgeon did the exposure on the levels prior to putting in the Schanz screw or mounting the surgical system. The surgeon did this to ensure they did not cause a patient or platform shift after registration. 2021-12-10 email 2021-12-10 10:46 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2021-Oct-26 Salesforce C00237474 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that there were registration issues during a L3/L4 and S1-S3 procedure. Registration was successfully completed for the first segment from L3-L4. Screws were placed at L3, but there was a lot of skiving at L4 due to holes from a previous procedure. The surgeon moved forward with the case, but the second segment for the pelvis region was unable to be registered. There was no patient harm and the procedure was delayed less than an hour. 2021-11-10 IUD, e1 (Rep): Additional information received from a manufacturer representative reported the L3 screw was lower in the pedicle than planned. 2021-11-15 email 2021-11-15 19:45 (Rep): No new information. 2021-11-17 IUD x2 (Other): No new information. 2021-12-02 e1, e2 (Rep): Additional information received from a manufacturer representative reported the screw were placed percutane ously and a bone mount bridge connected to a PSIS pin placed in the right PSIS was used as the mount. 2021-12-06 e3, e4 (Rep): Additional information received from a manufacturer representative reported that the lower levels were open. The surgeon did the exposure on the levels prior to putting in the Schanz screw or mounting the surgical system. The surgeon did this to ensure they did not cause a patient or platform shift after registration. 2021-12-10 email 2021-12-10 10:46 (Rep): No new information. | xse skivingunable register salesforce spinal p rocedure registration issues procedure registration successfully completed segment screws placed lot skiving holes previous procedure surgeon moved forward case second segment pelvis region unable registered procedure delayed hour iud additional screw lower pedicle planned email new iud new additional screw placed percutane ously bone mount bridge connected psis pin placed right psis mount additional lower levels open surgeon did exposure levels prior putting schanz screw mounting surgical surgeon did ensure did cause platform shift registration email new mazor x | NAV5000 | NAV3242 | NAV4112 | NAV3246 | NAV2054 | NAV5000NAV3242NAV4112NAV3246NAV2054 | True
1292 | 703195394 | X - Xeye cam issue | 2019-06-07T17:35:52.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | ASM0224-01 | RNX-PO-030 | X - Xeye cam issue It was reported that eddie tried to upgrade the SW from 2.3 to 3.0.2 and noticed the X-Eye camera would not work. We tried copying over the camera calibration files manually but this still did not fix the issue. We were able to verify that the green light on the X-Eye was blinking and not solid green. Solid green indicates camera is functioning while blinking light indicates camera failure. System will need a new camera. RNX-PO-030 | It was reported that eddie tried to upgrade the SW from 2.3 to 3.0.2 and noticed the X-Eye camera would not work. We tried copying over the camera calibration files manually but this still did not fix the issue. We were able to verify that the green light on the X-Eye was blinking and not solid green. Solid green indicates camera is functioning while blinking light indicates camera failure. System will need a new camera. | x xeye cam issue eddie tried upgrade sw noticed xeye camera work tried copying camera calibration files manually did fix issue able verify green light xeye blinking solid green solid green indicates camera functioning blinking light indicates camera failure need new camera | NAV4109 | NAV3254 | NAV4110 | NAV3246 | NAV2047 | NAV4109NAV3254NAV4110NAV3246NAV2047 | True
2947 | 703259283 | X - Surgical Arm Joint Shift | 2019-07-15T19:49:57.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical Arm Joint Shift "Arm Shift Detected in Joint 1" error received after a case at Asante Rogue Regional. Restart attempted by RTC with no success. Reset movement performed and Arm began functioning again as normal. Passed Advanced accuracy test. SURGICAL ARM | "Arm Shift Detected in Joint 1" error received after a case at Asante Rogue Regional. Restart attempted by RTC with no success. Reset movement performed and Arm began functioning again as normal. Passed Advanced accuracy test. | x surgical arm joint shift arm shift detected joint error case asante rogue regional restart attempted rtc success reset movement performed arm began functioning normal passed advanced accuracy test surgical arm | NAV4098 | NAV4101 | NAV4097 | NAV4095 | NAV2054 | NAV4098NAV4101NAV4097NAV4095NAV2054 | True
7377 | 704192588 | XSE - SRA Failed BIST Test | 2021-01-25T17:34:52.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-03R | SURGICAL ARM - REFURBISHED ASM0206-03R | XSE - SRA Failed BIST Test 2021-Jan-25 Salesforce C00219834 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed a built in self test during the procedure. After the second failed test, the manufacturer representative mounted the surgical system to the bed frame adapter, but not the bed. The surgical arm then move past the test and completed setup. The case was able to be completed. There was no patient involvement. 2021-01-27 IUD x2 (Other): No new information. SURGICAL ARM - REFURBISHED ASM0206-03R | 2021-Jan-25 Salesforce C00219834 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed a built in self test during the procedure. After the second failed test, the manufacturer representative mounted the surgical system to the bed frame adapter, but not the bed. The surgical arm then move past the test and completed setup. The case was able to be completed. There was no patient involvement. 2021-01-27 IUD x2 (Other): No new information. | xse sra failed bist test salesforce outside procedure surgical arm failed built self test procedure second failed test mounted surgical bed frame adapter bed surgical arm past test completed setup case able completed iud new surgical arm refurbished | NAV3249 | NAV2054 | NAV4092 | NAV3264 | NAV3246 | NAV3249NAV2054NAV4092NAV3264NAV3246 | True
6660 | 703999104 | XSE - Difficult to Mount System on Bed | 2020-10-07T16:10:04.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | ASM0214-01 | SURGICAL SYSTEM | XSE - Difficult to Mount System on Bed 2020-Aug-26 Salesforce C00214247, IUD (rep, for): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was difficult to mount on the operating table. More than usual force needed to be applied to bring the surgical system down to the bed frame adapter. There was no patient involvement. 2020-11-23 IUD (Other): No new information. SURGICAL SYSTEM | 2020-Aug-26 Salesforce C00214247, IUD (rep, for): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was difficult to mount on the operating table. More than usual force needed to be applied to bring the surgical system down to the bed frame adapter. There was no patient involvement. 2020-11-23 IUD (Other): No new information. | xse difficult mount bed salesforce iud outside procedure difficult mount operating table usual force needed applied bring surgical bed frame adapter iud new surgical | NAV3250 | NAV3251 | NAV3299 | NAV3262 | NAV2054 | NAV3250NAV3251NAV3299NAV3262NAV2054 | True
3026 | 703263385 | X - Arm Issues / OOBF | 2019-07-17T14:44:05.000000+0000 | MazorX | NAV4099 | PLASTIC COVER ISSUES | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - Arm Issues / OOBF When Andrew was on site to replace the SRA at WVU Ruby Memorial he found multiple issues with the SRA that was sent to him. SRA0025. Bottom plate was in bad shape and the screws on the plate were too small. In the hole where you attached the reference marker it was stuffed full of plastic which if not cleaned out causes the threads to strip easily. The paint is in bad shape as well. UNK MAZOR X ARM | When Andrew was on site to replace the SRA at WVU Ruby Memorial he found multiple issues with the SRA that was sent to him. SRA0025. Bottom plate was in bad shape and the screws on the plate were too small. In the hole where you attached the reference marker it was stuffed full of plastic which if not cleaned out causes the threads to strip easily. The paint is in bad shape as well. | x arm issues oobf andrew site replace sra wvu ruby memorial multiple issues sra sent plate bad shape screws plate small hole attached reference marker stuffed plastic cleaned causes threads strip easily paint bad shape unk mazor x arm | NAV4100 | NAV2054 | NAV3264 | NAV4099 | NAV3249 | NAV4100NAV2054NAV3264NAV4099NAV3249 | True
2105 | 703226068 | X - SOFTWARE ISSUE | 2019-06-25T16:09:52.000000+0000 | MazorX | NAV4110 | NO TRACKING | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - SOFTWARE ISSUE Service Issue Request 3rd August, 2017. New York Presbyterian. New York, NY. Workstation #34 Arm #38. 1) Software Crashed x2. The first crash was when bovies were being used the screen went black for about 30 seconds and then brought us out of operate to patient folder. 25 Minutes later the screen went black again this time for about 5 seconds and kept us on the same page. The surgeon’s icon in the top right hand of the screen was missing after these two occurrences. 2) X Eye communication was lost after both crashes. 3) Software continued to run very slowly when transitioning form page to page./L2-S1 OLIF w/ PSF. Surgeon reviewed the plan with CSR prior to the procedure. Potential skive issues were discussed along with making sure Bilateral retraction would be done when using the single schanz. Platform chosen was the single schanz due to Doctors preference. Prior to the procedure @ 2:39 the screen went completely black and after about 30 seconds it came back on and redirected us to the patient folder. CSR then noticed that even though the 3 check marks remained when going back into the operate process the surgeons image on the top right of the screen was not there anymore. After about 25 more minutes of the surgeon exposing and using multiple bovies the screen went black again this time for only a few seconds and broughtus right back to the current screen we were on.1241 3pt check 1243 10 pt check 1245 shoulder calibration checked 1356 Mounted 2 people 45 seconds 1359 draped 1520 Attached to the patient 1522 3Define scan 1524 Drawspine X eyes wasn't communicating went right to generic AP and Oblique 1527 Ap and OB images taken 1529 Registered 1530 S1 left followed by clear up 1533 L5 left followed by clear up 1537 L4 Left 1540 L3 left 1542 L2 left 1550 S1 right 1553 L5 right 1556 L4 right 1600 L3 right sent multiple times the arm had soft tissue pressure forcing the arm to shake to keep is trajectory sent it back 2 additional times before execution 1606 L2 right 1614 Confirmation shot 1620 off the bed. The software overall ran very slow after the two software crashes. Especially moving between pages like image acquisition to fluoro segmentation to registration. Overall the surgeons was satisfied with the outcome off all drilled trajectories were accurate to their planned position but would like to see the system run smoother with less delays. UNK MAZOR X X-EYE CAMERA | Service Issue Request 3rd August, 2017. New York Presbyterian. New York, NY. Workstation #34 Arm #38. 1) Software Crashed x2. The first crash was when bovies were being used the screen went black for about 30 seconds and then brought us out of operate to patient folder. 25 Minutes later the screen went black again this time for about 5 seconds and kept us on the same page. The surgeon’s icon in the top right hand of the screen was missing after these two occurrences. 2) X Eye communication was lost after both crashes. 3) Software continued to run very slowly when transitioning form page to page./L2-S1 OLIF w/ PSF. Surgeon reviewed the plan with CSR prior to the procedure. Potential skive issues were discussed along with making sure Bilateral retraction would be done when using the single schanz. Platform chosen was the single schanz due to Doctors preference. Prior to the procedure @ 2:39 the screen went completely black and after about 30 seconds it came back on and redirected us to the patient folder. CSR then noticed that even though the 3 check marks remained when going back into the operate process the surgeons image on the top right of the screen was not there anymore. After about 25 more minutes of the surgeon exposing and using multiple bovies the screen went black again this time for only a few seconds and broughtus right back to the current screen we were on.1241 3pt check 1243 10 pt check 1245 shoulder calibration checked 1356 Mounted 2 people 45 seconds 1359 draped 1520 Attached to the patient 1522 3Define scan 1524 Drawspine X eyes wasn't communicating went right to generic AP and Oblique 1527 Ap and OB images taken 1529 Registered 1530 S1 left followed by clear up 1533 L5 left followed by clear up 1537 L4 Left 1540 L3 left 1542 L2 left 1550 S1 right 1553 L5 right 1556 L4 right 1600 L3 right sent multiple times the arm had soft tissue pressure forcing the arm to shake to keep is trajectory sent it back 2 additional times before execution 1606 L2 right 1614 Confirmation shot 1620 off the bed. The software overall ran very slow after the two software crashes. Especially moving between pages like image acquisition to fluoro segmentation to registration. Overall the surgeons was satisfied with the outcome off all drilled trajectories were accurate to their planned position but would like to see the system run smoother with less delays. | x software issue service issue request august new york presbyterian new york ny workstation arm software crashed crash bovies screen went black seconds brought operate folder minutes later screen went black time seconds kept page surgeon’s icon right hand screen missing occurrences x eye communication lost crashes software continued run slowly transitioning page olif w psf surgeon reviewed plan csr prior procedure potential skive issues discussed making sure bilateral retraction using single schanz platform chosen single schanz doctors preference prior procedure screen went completely black seconds came redirected folder csr noticed check marks remained going operate process surgeons image right screen anymore minutes surgeon exposing using multiple bovies screen went black time seconds broughtus right current screen check pt check shoulder calibration checked mounted people seconds draped attached scan drawspine x eyes wasnt communicating went right generic ap oblique ap ob images taken registered left followed clear left followed clear left left left right right right right sent multiple times arm soft tissue pressure forcing arm shake trajectory sent additional times execution right confirmation shot bed software overall ran slow software crashes especially moving pages like image acquisition fluoro segmentation registration overall surgeons satisfied outcome drilled trajectories accurate planned position like run smoother delays unk mazor x xeye camera | NAV4101 | NAV3248 | NAV3261 | NAV3253 | NAV3247 | NAV4101NAV3248NAV3261NAV3253NAV3247 | False
2388 | 703234131 | X - Registration issue | 2019-06-28T20:44:21.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0639-10 | MAZOR X SOFTWARE KIT VER-3.0.1 | X - Registration issue in the first registration, there was a significant issue with trajectories. Out of the 10 trajectories registered, 5 were red. Dr. Lieberman and I went through the planning to find an acceptable trajectory, but 20 minutes of adjustment were required. More importantly, we were able to find reachable trajectories in the planning screen, but the software would say that the trajectories were unreachable on the send surgical arm screen. When this occurred, we were having to fly blind with not knowing what to be done to correctly adjust the trajectory. It ended up being easier to readjust the platform than to find an acceptable solution. MAZOR X SOFTWARE KIT VER-3.0.1 | in the first registration, there was a significant issue with trajectories. Out of the 10 trajectories registered, 5 were red. Dr. Lieberman and I went through the planning to find an acceptable trajectory, but 20 minutes of adjustment were required. More importantly, we were able to find reachable trajectories in the planning screen, but the software would say that the trajectories were unreachable on the send surgical arm screen. When this occurred, we were having to fly blind with not knowing what to be done to correctly adjust the trajectory. It ended up being easier to readjust the platform than to find an acceptable solution. | x registration issue registration significant issue trajectories trajectories registered red dr lieberman went planning acceptable trajectory minutes adjustment required importantly able reachable trajectories planning screen software say trajectories unreachable send surgical arm screen occurred having fly blind knowing correctly adjust trajectory ended easier readjust platform acceptable solution mazor x software kit | NAV3246 | NAV3243 | NAV3242 | NAV4091 | NAV3247 | NAV3246NAV3243NAV3242NAV4091NAV3247 | True
8599 | 704447836 | XSE - Low Pressure Errors | 2021-06-14T15:48:14.000000+0000 | MazorX Stealth Edition | NAV4104 | UNABLE TO MOVE | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Low Pressure Errors 2021-Jun-14 Salesforce C00228307, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that low pressure errors were displayed when using the guidance system and the positioner was difficult to move up and down. There was no patient involvement. 2021-06-28 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Jun-14 Salesforce C00228307, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that low pressure errors were displayed when using the guidance system and the positioner was difficult to move up and down. There was no patient involvement. 2021-06-28 IUD (Other): No new information. | xse low pressure errors salesforce iud outside proc edure low pressure errors displayed using positioner difficult iud new unk mazor x workstation | NAV3264 | NAV4104 | NAV3262 | NAV2047 | NAV3256 | NAV3264NAV4104NAV3262NAV2047NAV3256 | True
9077 | 704547659 | XSE - Software Unresponsive | 2021-08-09T21:37:33.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 40 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - Software Unresponsive 2021-Aug-09 Salesforce C00231912, IUD x4 (rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the software went unresponsive and was not able to track navigation instruments. No patient was present. 2021-08-11 IUD x4, email 2021-08-13 08:46 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined, but replacing the PC resolved all of the issues. 2021-08-26 IUD (Other): No new information. 2021-09-03 IUD x3 (Other): No new information. 2021-09-22 IUD x3 (Other): No new information. 2021-12-06 IUD (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2021-Aug-09 Salesforce C00231912, IUD x4 (rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the software went unresponsive and was not able to track navigation instruments. No patient was present. 2021-08-11 IUD x4, email 2021-08-13 08:46 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined, but replacing the PC resolved all of the issues. 2021-08-26 IUD (Other): No new information. 2021-09-03 IUD x3 (Other): No new information. 2021-09-22 IUD x3 (Other): No new information. 2021-12-06 IUD (Other): No new information. | xse software unresponsive salesforce iud outside p rocedure software went unresponsive able track navigation instruments present iud email additional cause determined replacing pc resolved issues iud new iud new iud new iud new mazor x pc station | NAV3262 | NAV3248 | NAV2047 | NAV3246 | NAV3256 | NAV3262NAV3248NAV2047NAV3246NAV3256 | True
5014 | 703590076 | X - Registration issues | 2020-01-30T21:26:44.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration issues 2018-05-17 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3D marker was loose for initial registration attempts. The surgeon noted lose marker and tightened down screw. New images were taken and registration achieved. There was no patient harm. MAZOR X SYSTEM TPL0059 | 2018-05-17 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3D marker was loose for initial registration attempts. The surgeon noted lose marker and tightened down screw. New images were taken and registration achieved. There was no patient harm. | x registration issues ccr spinal procedure marker loose initial registration attempts surgeon noted lose marker tightened screw new images taken registration achieved mazor x | NAV3242 | NAV4112 | NAV4091 | NAV3246 | NAV3244 | NAV3242NAV4112NAV4091NAV3246NAV3244 | True
4409 | 703541520 | XSE - Back Shoulder Shift / Screen | 2020-01-02T19:57:33.000000+0000 | MazorX Stealth Edition | NAV4108 | SHOULDER SHIFTS | 30 | ASM0207-02 | SHOULDER | XSE - Back Shoulder Shift / Screen 2019-12-26 00681251 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was sent to L5 and received the error back shoulder shift was detected. Had to re-register, took a new AP image and when sent to oblique image the screen and tablet went black. A hard restart was preformed and brought back to the oblique step. A new image was taken, registration was achieved but the tablet went black again. Sent the arm to clear up and detected another back shoulder shift. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-01-07 IUD x2 (Other): No new information. SHOULDER | 2019-12-26 00681251 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was sent to L5 and received the error back shoulder shift was detected. Had to re-register, took a new AP image and when sent to oblique image the screen and tablet went black. A hard restart was preformed and brought back to the oblique step. A new image was taken, registration was achieved but the tablet went black again. Sent the arm to clear up and detected another back shoulder shift. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-01-07 IUD x2 (Other): No new information. | xse shoulder shift screen spinal procedure surgical arm sent error shoulder shift detected reregister took new ap image sent oblique image screen tablet went black hard restart preformed brought oblique step new image taken registration achieved tablet went black sent arm clear detected shoulder shift issue extended surgical time hour impact outcome iud new shoulder | NAV3253 | NAV4108 | NAV3252 | NAV4107 | NAV3261 | NAV3253NAV4108NAV3252NAV4107NAV3261 | True
3844 | 703453996 | XSE - Arm, SW, Wkstn, CANBUS Errors | 2019-11-07T00:05:35.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Arm, SW, Wkstn, CANBUS Errors 2019-Nov-06 00677816 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a case the arm displayed a message that could not reach destination, they tried to resend the trajectories without resolution and got the same message. They then performed a hard restart and got a 295 error CANBUS. They cleared that error with a soft restart and then multiple errors were displayed, emergency stop error, driver not found error, pneumatic pressure not building error. They had to use the foot pump to remove the system from the bed. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-11 IUD x3 (Other): No new information. MAZOR X SYSTEM TPL0059 | 2019-Nov-06 00677816 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a case the arm displayed a message that could not reach destination, they tried to resend the trajectories without resolution and got the same message. They then performed a hard restart and got a 295 error CANBUS. They cleared that error with a soft restart and then multiple errors were displayed, emergency stop error, driver not found error, pneumatic pressure not building error. They had to use the foot pump to remove the system from the bed. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-11 IUD x3 (Other): No new information. | xse arm sw wkstn canbus errors spinal procedure case arm displayed message reach destination tried resend trajectories resolution got message performed hard restart got error canbus cleared error soft restart multiple errors displayed emergency stop error driver error pneumatic pressure building error use foot pump remove bed additional complications anticipated iud new mazor x | NAV2047 | NAV3246 | NAV4095 | NAV3264 | NAV1006 | NAV2047NAV3246NAV4095NAV3264NAV1006 | True
5332 | 703649946 | XSE - Schanz Arm & Bond Missing | 2020-03-04T20:17:09.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Schanz Arm & Bond Missing 2020-MAR-04 00685647 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arm and Schanz bond were missing from the trays. There was no patient present when this issue was identified. 2020-03-13 IUD x2 (Other): No new information. KIT SVC MAS2351-04S SCHANZ ARM | 2020-MAR-04 00685647 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz arm and Schanz bond were missing from the trays. There was no patient present when this issue was identified. 2020-03-13 IUD x2 (Other): No new information. | xse schanz arm bond missing outside procedure schanz arm schanz bond missing trays present issue identified iud new kit svc schanz arm | NAV2053 | NAV3299 | NAV3298 | NAV3295 | NAV4093 | NAV2053NAV3299NAV3298NAV3295NAV4093 | True
7736 | 704261078 | X ¦ Lock axis and unable to reach | 2021-03-03T21:11:38.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | X ¦ Lock axis and unable to reach 2021-Mar-03 Salesforce C00222019 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while setting up for a case. He stated that he was doing the initial arm setup and got a message stating to lock Y axis. He said that the y positioner also was not green. He attached it to the bed and was able to do the arm setup, but then got a message stating the 'D5 position is not reached'. He was able to dismiss this message and continue. Doing a second case with the system, booted up fine and completed the arm set up without a warning of y axis. Site was able to complete case plus another without issue. Will continue to monitor. There was no patient involved. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2021-Mar-03 Salesforce C00222019 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while setting up for a case. He stated that he was doing the initial arm setup and got a message stating to lock Y axis. He said that the y positioner also was not green. He attached it to the bed and was able to do the arm setup, but then got a message stating the 'D5 position is not reached'. He was able to dismiss this message and continue. Doing a second case with the system, booted up fine and completed the arm set up without a warning of y axis. Site was able to complete case plus another without issue. Will continue to monitor. There was no patient involved. | x lock axis unable reach salesforce outside procedure setting case stated doing initial arm setup got message stating lock y axis said y positioner green attached bed able arm setup got message stating position reached able dismiss message continue doing second case booted fine completed arm set warning y axis site able complete case plus issue continue monitor involved kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV4104 | NAV3244 | NAV3259 | NAV3249NAV4095NAV4104NAV3244NAV3259 | True
1940 | 703221386 | X - 3Define camera udpate | 2019-06-21T21:51:54.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define camera udpate 3Define camera failed, went into settings, No firmware was present/ soft restart firmware present , 2nd attempt succesful UNK MAZOR X 3DEFINE CAMERA | 3Define camera failed, went into settings, No firmware was present/ soft restart firmware present , 2nd attempt succesful | x camera udpate camera failed went settings firmware present soft restart firmware present attempt succesful unk mazor x camera | NAV3244 | NAV2047 | NAV3261 | NAV2053 | NAV3249 | NAV3244NAV2047NAV3261NAV2053NAV3249 | True
4511 | 703548267 | X - 3Define, X-Eye & Arm Issues | 2020-01-07T15:58:01.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 40 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - 3Define, X-Eye & Arm Issues 2018-07-20 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a global deformity with fixation at T3-L1. It was reported that the case was planned by the surgeon. A clamp was used to secure the surgical system to the patient. Draw spine did not work. The 3Define camera also failed. It went through the right side of the surface scan perfectly, but when it looked like the left side of the scan was almost complete, the robot would beep (signaling that a trajectory could not be reached) and there was a blank image on the workstation. All trajectories were reached accurately. It is noted that the arm shook when it was finished sending a trajectory. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X ARM | 2018-07-20 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a global deformity with fixation at T3-L1. It was reported that the case was planned by the surgeon. A clamp was used to secure the surgical system to the patient. Draw spine did not work. The 3Define camera also failed. It went through the right side of the surface scan perfectly, but when it looked like the left side of the scan was almost complete, the robot would beep (signaling that a trajectory could not be reached) and there was a blank image on the workstation. All trajectories were reached accurately. It is noted that the arm shook when it was finished sending a trajectory. There was no patient harm and the procedure was not delayed over an hour. | x xeye arm issues ccr spinal procedure treat global deformity fixation case planned surgeon clamp secure surgical draw spine did work camera failed went right surface scan perfectly looked like left scan complete robot beep blank image workstation trajectories reached accurately noted arm shook finished sending trajectory procedure delayed hour unk mazor x arm | NAV4095 | NAV3244 | NAV4110 | NAV3239 | NAV3264 | NAV4095NAV3244NAV4110NAV3239NAV3264 | False
132 | 703005936 | X - X-Eye Issues | 2019-02-13T20:58:40.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 20 | ELB0057-01 | MAZOR X FIRE WIRE CARD | X - X-Eye Issues 2019-02-07 SCL1900346, Intake Form 2019-02-13 15:05 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that X-eye showed green connection during setup, but was not recognizing the stylus during draw spine and there was no live feed from the utilities menu. There was no patient harm and the surgery was not delayed more than an hour. The case was completed. 2019-02-11 e1 (Rep, HCP): Additional information received from a manufacturer representative reported that a L4 to S1 transforaminal lumbar interbody fusion (TLIF) and posterior lateral spinal fusion (PSF) procedure was being done. During setup, the shoulder calibration was checked and then the X-Eye camera was checked, but there was no output. An advanced accuracy check was completed and the X-Eye was checked again and it worked. The rest of the system was setup successfully. After mounting the surgical system to the bed, the X-Eye camera was checked again and the screen remained black with no input. The cables were checked including unplugging and replugging in the lemo and X-Eye cables. At this point, the software started crashing and was not responding. When the software did respond a CANBUS and 4059 (Arm Shift Detected in Joint 2) error messages were displayed. A soft restart and hard restart were done. The manufacturer representative attempted to set up the surgical arm again with red lights for the X-Eye, but the 4059 error was displayed again. Technical services had the manufacturer representative manually enable arm movement in the software and complete an advanced accuracy check again. The surgical arm setup was able to be completed with there being a red light for the X-Eye. The surgeon was able to successfully place all of the screws in the correct position. Accuracy was verified visually and with lateral and AP images. MAZOR X FIRE WIRE CARD | 2019-02-07 SCL1900346, Intake Form 2019-02-13 15:05 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that X-eye showed green connection during setup, but was not recognizing the stylus during draw spine and there was no live feed from the utilities menu. There was no patient harm and the surgery was not delayed more than an hour. The case was completed. 2019-02-11 e1 (Rep, HCP): Additional information received from a manufacturer representative reported that a L4 to S1 transforaminal lumbar interbody fusion (TLIF) and posterior lateral spinal fusion (PSF) procedure was being done. During setup, the shoulder calibration was checked and then the X-Eye camera was checked, but there was no output. An advanced accuracy check was completed and the X-Eye was checked again and it worked. The rest of the system was setup successfully. After mounting the surgical system to the bed, the X-Eye camera was checked again and the screen remained black with no input. The cables were checked including unplugging and replugging in the lemo and X-Eye cables. At this point, the software started crashing and was not responding. When the software did respond a CANBUS and 4059 (Arm Shift Detected in Joint 2) error messages were displayed. A soft restart and hard restart were done. The manufacturer representative attempted to set up the surgical arm again with red lights for the X-Eye, but the 4059 error was displayed again. Technical services had the manufacturer representative manually enable arm movement in the software and complete an advanced accuracy check again. The surgical arm setup was able to be completed with there being a red light for the X-Eye. The surgeon was able to successfully place all of the screws in the correct position. Accuracy was verified visually and with lateral and AP images. | x xeye issues spinal procedure xeye showed green connection setup recognizing stylus draw spine live feed utilities menu surgery delayed hour case completed additional transforaminal lumbar interbody fusion posterior lateral spinal fusion procedure setup shoulder calibration checked xeye camera checked output advanced accuracy check completed xeye checked worked rest setup successfully mounting surgical bed xeye camera checked screen remained black input cables checked including unplugging replugging lemo xeye cables point software started crashing responding software did respond canbus error messages displayed soft restart hard restart attempted set surgical arm red lights xeye error displayed technical services manually enable arm movement software complete advanced accuracy check surgical arm setup able completed red light xeye surgeon able successfully place screws correct position accuracy verified visually lateral ap images mazor x wire card | NAV4110 | NAV3254 | NAV3248 | NAV4109 | NAV1012 | NAV4110NAV3254NAV3248NAV4109NAV1012 | False
2195 | 703229048 | X - DEVIATION/REGISTRATION ISSUE | 2019-06-26T20:55:15.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X - DEVIATION/REGISTRATION ISSUE had received a message from the surgeon stating the patient is not feeling well, and showing lower signals to her feet. after post op CT performed this is his assessment. Case at SZMC System #86 arm #94 Red -X-eye crash due to bowie usage. Platform : Dual Clamp Preparation : 7:35 Grounding the system to the Bed frame. 7:45 Planning performed before the surgery by the surgeon and reviewed by the FCE prior to the case 9:10 10 point check passed. 9:15 Passed 3 point check. 9:30 C-arm set up done. 9:35 Draping done -BMB sticker appear on top of the arm instead of under the arm. Please see image attached. 9:45 Surgeon screen attached -while screwing the Arm adaptor to the surgeon screen arm the head ofone of the screws broke. please see image attached. we continued with screwing only two screws. While using bowie ,the X-eye crashed. We decided to restart the system (post exposure) to re connectthe X-eye. First segment T3-T8. 11:42 Re-doing 3point check by the surgeon. 11:44 Mount locked over dual clamp (T4-T7). 11:48 3Define scan done 11:50 Draw spine , Mid-segment at T5. 11:51 Arm sent to AP : 2 images. 11:57 Arm sent to OBL : 3 images. Registration 12:00 1st-3rd Registrations Failed due to bad images and miss-labeling. 12:03 Arm sent to new AP : 1 image. 12:06 4th Registration achieved on T4 with basic algorithm.12:13 T3 L -surgeon felt trajectory was too medial and did not drill. 12:18 T3 R 12:27 T3 L -Resent and drilled. 12:27 T4 R 12:30 T5 R 12:33 T4 L 12:36 T5 R -Collision with mount -Axial angle opened 12:39 T6 L 12:43 T6 R 12:49 T7 R 12:52 T7 L 12:55 T7 R 12:58 T8 L -Surgeon commented that new Mazor-X screw design looks good but slippery compered to pervious one. 13:00 T8 R -Not drilled by surgeon intra operative decision. As X-eye crashed again due to use of bowie System was restarted. Second segment T9-L3 13:05 Mount locked over dual clamp T12-L2 13:07 3Define done 13:09 Draw spine , Mid-segment at T11. 13:11 Arm sent to AP : 4 images. 13:20 Arm sent to OBL : 4 images. 13:23 Registration achieved on T4 with basic algorithm on T10. 13:27 T9 L 13:30 T9 R 13:34 T10 L 13:37 T10 R 13:40 T11L 13:44 T11R 13:46 T12 L 13:50 T12 R 13:52 L1 L 13:55 L1 R 13:57 L3 L 13:59 L3 R Third segment L4-S2. 14:05 Mount locked over dual clamp T12-L2 14:07 3Define -1st time aborted as 3Define plastic cover popped out. 2nd scan was successful. 14:09 Draw spine , Mid-segment at L5. 14:11 Arm sent to AP : 2 images. 14:20 Arm sent to generic OBL : 2 images. 14:18 Registration achieved on L4 with basic algorithm on T10. 14:24 L4 R 14:26 L4 L 14:28 L5 R 14:30 L5 L 14:32 S1 R 14:34 S2 R 14:37 S2 L. After screws placement, surgeon removed T6R and also removed T3 R and re-insert T3 R in the same trajectory. The surgeon didn't feel comfortable with the system accuracy through out the the operation. MAZOR X SOFTWARE KIT VER-3.0 | had received a message from the surgeon stating the patient is not feeling well, and showing lower signals to her feet. after post op CT performed this is his assessment. Case at SZMC System #86 arm #94 Red -X-eye crash due to bowie usage. Platform : Dual Clamp Preparation : 7:35 Grounding the system to the Bed frame. 7:45 Planning performed before the surgery by the surgeon and reviewed by the FCE prior to the case 9:10 10 point check passed. 9:15 Passed 3 point check. 9:30 C-arm set up done. 9:35 Draping done -BMB sticker appear on top of the arm instead of under the arm. Please see image attached. 9:45 Surgeon screen attached -while screwing the Arm adaptor to the surgeon screen arm the head ofone of the screws broke. please see image attached. we continued with screwing only two screws. While using bowie ,the X-eye crashed. We decided to restart the system (post exposure) to re connectthe X-eye. First segment T3-T8. 11:42 Re-doing 3point check by the surgeon. 11:44 Mount locked over dual clamp (T4-T7). 11:48 3Define scan done 11:50 Draw spine , Mid-segment at T5. 11:51 Arm sent to AP : 2 images. 11:57 Arm sent to OBL : 3 images. Registration 12:00 1st-3rd Registrations Failed due to bad images and miss-labeling. 12:03 Arm sent to new AP : 1 image. 12:06 4th Registration achieved on T4 with basic algorithm.12:13 T3 L -surgeon felt trajectory was too medial and did not drill. 12:18 T3 R 12:27 T3 L -Resent and drilled. 12:27 T4 R 12:30 T5 R 12:33 T4 L 12:36 T5 R -Collision with mount -Axial angle opened 12:39 T6 L 12:43 T6 R 12:49 T7 R 12:52 T7 L 12:55 T7 R 12:58 T8 L -Surgeon commented that new Mazor-X screw design looks good but slippery compered to pervious one. 13:00 T8 R -Not drilled by surgeon intra operative decision. As X-eye crashed again due to use of bowie System was restarted. Second segment T9-L3 13:05 Mount locked over dual clamp T12-L2 13:07 3Define done 13:09 Draw spine , Mid-segment at T11. 13:11 Arm sent to AP : 4 images. 13:20 Arm sent to OBL : 4 images. 13:23 Registration achieved on T4 with basic algorithm on T10. 13:27 T9 L 13:30 T9 R 13:34 T10 L 13:37 T10 R 13:40 T11L 13:44 T11R 13:46 T12 L 13:50 T12 R 13:52 L1 L 13:55 L1 R 13:57 L3 L 13:59 L3 R Third segment L4-S2. 14:05 Mount locked over dual clamp T12-L2 14:07 3Define -1st time aborted as 3Define plastic cover popped out. 2nd scan was successful. 14:09 Draw spine , Mid-segment at L5. 14:11 Arm sent to AP : 2 images. 14:20 Arm sent to generic OBL : 2 images. 14:18 Registration achieved on L4 with basic algorithm on T10. 14:24 L4 R 14:26 L4 L 14:28 L5 R 14:30 L5 L 14:32 S1 R 14:34 S2 R 14:37 S2 L. After screws placement, surgeon removed T6R and also removed T3 R and re-insert T3 R in the same trajectory. The surgeon didn't feel comfortable with the system accuracy through out the the operation. | x deviationregistration issue message surgeon stating feeling showing lower signals feet post op ct performed assessment case szmc arm red xeye crash bowie usage platform dual clamp preparation grounding bed frame planning performed surgery surgeon reviewed fce prior case point check passed passed point check carm set draping bmb sticker appear arm instead arm image attached surgeon screen attached screwing arm adaptor surgeon screen arm head ofone screws broke image attached continued screwing screws using bowie xeye crashed decided restart connectthe xeye segment redoing check surgeon mount locked dual clamp scan draw spine midsegment arm sent ap images arm sent obl images registration registrations failed bad images misslabeling arm sent new ap image registration achieved basic l surgeon felt trajectory medial did drill r l resent drilled r r l r collision mount axial angle opened l r r l r l surgeon commented new mazorx screw design looks good slippery compered pervious r drilled surgeon intra operative decision xeye crashed use bowie restarted second segment mount locked dual clamp draw spine midsegment arm sent ap images arm sent obl images registration achieved basic algorithm l r l r l r l r l r segment mount locked dual clamp time aborted plastic cover popped scan successful draw spine midsegment arm sent ap images arm sent generic obl images registration achieved basic algorithm r l r l r r l screws placement surgeon removed removed r reinsert r trajectory surgeon didnt feel comfortable accuracy operation mazor x software kit | NAV4091 | NAV3237 | NAV4096 | NAV4099 | NAV4090 | NAV4091NAV3237NAV4096NAV4099NAV4090 | False
6346 | 703920097 | XSE - Missing Part/OOBF | 2020-08-24T21:41:09.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | MAS2438-01S | KIT SVC MAS2438-01S MIST BRIDGE | XSE - Missing Part/OOBF 2020-08-24 00695134 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that it has been requested that a MAS2438-01S be shipped to the site due to the site receiving the KIT0762 without the MAS2438-01S included. No patient was present. 2020-09-01 IUD (Other): No new information. KIT SVC MAS2438-01S MIST BRIDGE | 2020-08-24 00695134 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that it has been requested that a MAS2438-01S be shipped to the site due to the site receiving the KIT0762 without the MAS2438-01S included. No patient was present. 2020-09-01 IUD (Other): No new information. | xse missing partoobf accessory outside procedure requested shipped site site receiving included present iud new kit svc mist bridge | NAV2053 | NAV3299 | NAV2054 | NAV1006 | NAV3249 | NAV2053NAV3299NAV2054NAV1006NAV3249 | True
4262 | 703520557 | X - 3Define failure | 2019-12-16T17:15:48.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define failure 2018-08-13 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a degenerative focal deformity wit fixation at C2-T2. It was reported that the surgeon completed planning prior to the procedure. The clamp was chosen for mounting the system and attached to the patient at the C6 vertebra. The manufacturer representative reported that the surface scan was not working due to suspected software problems. However, registration was fast and accurate. The case was completed with all trajectories placed accurately. No patient harm was reported and surgery was delayed less than an hour. UNK MAZOR X 3DEFINE CAMERA | 2018-08-13 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a degenerative focal deformity wit fixation at C2-T2. It was reported that the surgeon completed planning prior to the procedure. The clamp was chosen for mounting the system and attached to the patient at the C6 vertebra. The manufacturer representative reported that the surface scan was not working due to suspected software problems. However, registration was fast and accurate. The case was completed with all trajectories placed accurately. No patient harm was reported and surgery was delayed less than an hour. | x failure ccr spinal procedure treat degenerative focal deformity wit fixation surgeon completed planning prior procedure clamp chosen mounting attached vertebra surface scan working suspected software problems registration fast accurate case completed trajectories placed accurately surgery delayed hour unk mazor x camera | NAV3244 | NAV4095 | NAV4110 | NAV3260 | NAV3250 | NAV3244NAV4095NAV4110NAV3260NAV3250 | True
3334 | 703317559 | XSE - Broken Positioner Handle | 2019-08-19T15:10:25.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 30 | ASM0293-01 | RNX-PO-119 | XSE - Broken Positioner Handle 2019-08-19 SCL1902409, Intake Form, C00193524, ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the left handle of the surgical system broke off during placement of the surgical system on the bed. The system functioned as expected during the procedure and was able to be removed from the bed after the procedure without incident. There was no patient harm and the procedure was not delayed more than an hour. 2019-09-13 IUD (Rep): No new information. RNX-PO-119 | 2019-08-19 SCL1902409, Intake Form, C00193524, ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the left handle of the surgical system broke off during placement of the surgical system on the bed. The system functioned as expected during the procedure and was able to be removed from the bed after the procedure without incident. There was no patient harm and the procedure was not delayed more than an hour. 2019-09-13 IUD (Rep): No new information. | xse broken positioner handle id spinal procedure left handle surgical broke placement surgical bed functioned expected procedure able removed bed procedure incident procedure delayed hour iud new | NAV3251 | NAV2058 | NAV5000 | NAV3249 | NAV1006 | NAV3251NAV2058NAV5000NAV3249NAV1006 | True
9909 | 704745264 | XSE - Black Workstation Monitor | 2021-12-02T23:10:38.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 30 | MAS1574-01S | KIT SVC MAS1574-01S BED FRAME ADAPTOR | XSE - Black Workstation Monitor 2021-Dec-02 Salesforce C00239076 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the workstation monitor went black when the surgeon used Bovie during the procedure. The picture returned when Bovie was powered off. There was no patient harm and the procedure was not delayed. 2021-12-08 IUD x2 (Other): No new information. 2021-12-16 IUD (Other): No new information. KIT SVC MAS1574-01S BED FRAME ADAPTOR | 2021-Dec-02 Salesforce C00239076 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the workstation monitor went black when the surgeon used Bovie during the procedure. The picture returned when Bovie was powered off. There was no patient harm and the procedure was not delayed. 2021-12-08 IUD x2 (Other): No new information. 2021-12-16 IUD (Other): No new information. | xse black workstation monitor salesforce spinal p rocedure workstation monitor went black surgeon bovie procedure picture returned bovie powered procedure delayed iud new iud new kit svc bed frame adaptor | NAV3247 | NAV3261 | NAV2058 | NAV2047 | NAV3255 | NAV3247NAV3261NAV2058NAV2047NAV3255 | True
6646 | 703995990 | XSE - CANBUS Error | 2020-10-06T12:09:18.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE - CANBUS Error 2020-10-06 00698169 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when setting up for a Scan & Plan case, a CANBUS error was displayed when booting up the system. The manufacturer representative did three hard restarts of the system, but received the error at the same time during boot process. The error did occur one time before the splash screen and one time after. There was no patient involvement. 2020-10-09 IUD x8 (Other): No new information. 2020-10-26 IUD x9 (Other): No new information. 2021-02-24 IUD (Other): No new information. 2021-06-01 IUD x2 (Other): No new information. 2021-06-21 IUD (other): No new information. 2021-07-19 IUD (Other): No new information. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2020-10-06 00698169 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when setting up for a Scan & Plan case, a CANBUS error was displayed when booting up the system. The manufacturer representative did three hard restarts of the system, but received the error at the same time during boot process. The error did occur one time before the splash screen and one time after. There was no patient involvement. 2020-10-09 IUD x8 (Other): No new information. 2020-10-26 IUD x9 (Other): No new information. 2021-02-24 IUD (Other): No new information. 2021-06-01 IUD x2 (Other): No new information. 2021-06-21 IUD (other): No new information. 2021-07-19 IUD (Other): No new information. | xse canbus error outside procedure setting scan plan case canbus error displayed booting did hard restarts error time boot process error did occur time splash screen time iud new iud new iud new iud new iud new iud new kit svc surgical new | NAV2047 | NAV2054 | NAV2057 | NAV3248 | NAV3261 | NAV2047NAV2054NAV2057NAV3248NAV3261 | True
905 | 703180742 | X - 3Define Issue, Registration Issue | 2019-05-30T13:53:25.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | ASM0206-01 | SURGICAL ARM | X - 3Define Issue, Registration Issue 2018-10-03 CCR, SCL1804489 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon planned the case on the day of surgery, prior to the case for ALIF/PSF at L4-S1. The schanz arm platform was used for the procedure and attached to the patient with a schanz pin inserted in the patient's right iliac crest. The 3Define scan failed and, therefore, Draw Spine could not be completed. Generic work volume was used. 3 AP and 3 OBL images were acquired. First images and segmentation attempts were not successful but registration was still achieved within only 3 minutes. All trajectories were placed accurately and confirmed with fluoro imaging. No patient harm was reported and surgery was delayed less than an hour. After the case, field service determined the 3Define camera was not working for the scan because the firmware version was incorrect. SURGICAL ARM | 2018-10-03 CCR, SCL1804489 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon planned the case on the day of surgery, prior to the case for ALIF/PSF at L4-S1. The schanz arm platform was used for the procedure and attached to the patient with a schanz pin inserted in the patient's right iliac crest. The 3Define scan failed and, therefore, Draw Spine could not be completed. Generic work volume was used. 3 AP and 3 OBL images were acquired. First images and segmentation attempts were not successful but registration was still achieved within only 3 minutes. All trajectories were placed accurately and confirmed with fluoro imaging. No patient harm was reported and surgery was delayed less than an hour. After the case, field service determined the 3Define camera was not working for the scan because the firmware version was incorrect. | x issue registration issue ccr spinal procedure surgeon planned case day surgery prior case alifpsf schanz arm platform procedure attached schanz pin inserted patients right iliac crest scan failed draw spine completed generic work volume ap obl images acquired images segmentation attempts successful registration achieved minutes trajectories placed accurately confirmed fluoro imaging surgery delayed hour case field service determined camera working scan firmware version incorrect surgical arm | NAV3244 | NAV3242 | NAV3249 | NAV4091 | NAV3248 | NAV3244NAV3242NAV3249NAV4091NAV3248 | True
10064 | 704792592 | XSE - CANBUS Error | 2022-01-05T20:34:39.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 20 | ASM0270 | KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | XSE - CANBUS Error 2021-Dec-30 Salesforce C00240316, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that an error stating "Fatal CANBUS error. Lost communication to node 7" was displayed. There was no patient involvement. 2022-01-05 IUD x3 (Rep): Additional information received from a manufacturer representative reported that there was a recurring breaker tripping at the 24V 4-amp circuit and the electrical panel needed to be replaced. 2022-01-16 IUD x2 (Other): No new information. 2022-01-26 IUD (Other): No new information. 2022-02-08 IUD x2 (Other): No new information. 2022-02-23 IUD x3 (Other): No new information. 2022-03-07 IUD x5 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-05-10 IUD (Other): No new information. KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | 2021-Dec-30 Salesforce C00240316, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that an error stating "Fatal CANBUS error. Lost communication to node 7" was displayed. There was no patient involvement. 2022-01-05 IUD x3 (Rep): Additional information received from a manufacturer representative reported that there was a recurring breaker tripping at the 24V 4-amp circuit and the electrical panel needed to be replaced. 2022-01-16 IUD x2 (Other): No new information. 2022-01-26 IUD (Other): No new information. 2022-02-08 IUD x2 (Other): No new information. 2022-02-23 IUD x3 (Other): No new information. 2022-03-07 IUD x5 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-05-10 IUD (Other): No new information. | xse canbus error salesforce iud outside proc edure error stating fatal canbus error lost communication node displayed iud additional recurring breaker tripping circuit electrical panel needed replaced iud new iud new iud new iud new iud new iud new iud new kit svc compressor kit new | NAV2047 | NAV2054 | NAV3263 | NAV3262 | NAV2057 | NAV2047NAV2054NAV3263NAV3262NAV2057 | False
5179 | 703612095 | XSE - X-eye Camera | 2020-02-13T00:04:45.000000+0000 | MazorX Stealth Edition | NAV4110 | NO TRACKING | 30 | ASM0362-01S | KIT SVC ASM0362-01S X-EYE FOR SSY 2 | XSE - X-eye Camera 2020-02-12 00683995 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the camera was not functional. The representative un/reinstalled the camera and it became operational. There was no patient involvement. 2020-02-13 IUD (Other) Additional information was received from the field service engineer (FSE). The FSE reported that he found the firewire connection at base of the camera was loose. 2020-02-19 IUD (Other): No new information. KIT SVC ASM0362-01S X-EYE FOR SSY 2 | 2020-02-12 00683995 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the camera was not functional. The representative un/reinstalled the camera and it became operational. There was no patient involvement. 2020-02-13 IUD (Other) Additional information was received from the field service engineer (FSE). The FSE reported that he found the firewire connection at base of the camera was loose. 2020-02-19 IUD (Other): No new information. | xse xeye camera outside procedure camera functional unreinstalled camera operational iud additional field service engineer fse firewire connection base camera loose iud new kit svc xeye ssy | NAV4109 | NAV2057 | NAV2054 | NAV3244 | NAV4110 | NAV4109NAV2057NAV2054NAV3244NAV4110 | True
1479 | 703202250 | X - Surgical Arm Issue | 2019-06-12T13:19:47.000000+0000 | MazorX | NAV4093 | INACCURACY DURING ADVANCED SETUP | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - Surgical Arm Issue CS completed SA set up and advanced test prior to case with normal results. After leaving the room and coming back several minutes later, CS noted that the SA was now making noise and shaking in the clear up position. CS attempted a 10 pt which failed. CS powered system down and tried new outlets x 2 with same results. RTC then plugged into a red outlet, waited 5min, and powered on. A shoulder calibration was done, after which the arm moved normally, but was still very loud, particularly the joint zero area. SA set up and 10pt both passed. Intraop, arm would not move to AP with error message saying arm did not reach final destination. RTC and surgeon noted that it's position following 3Define completion appeared to be extended further backwards than normal. CS sent to clear up and arm completed the movement but again looked extended backwards further than normal. CS attempted AP position, but the same error message was given. Surgeon engaged the E-stop and manually moved arm. E-stop was released and arm was successfully sent to AP. FWBMC System #92/Arm #101 3/22/18 Yellow -SA movement issues PSF L4-L5 Prep CS completed SA set up and advanced test prior to case with normal results. After leaving the room and coming back several minutes later, CS noted that the SA was now making noise and shaking in the clear up position. CS attempted a 10 pt which failed. CS powered system down and tried new outlets x 2 with same results. RTC then plugged into a red outlet, waited 5min, and powered on. A shoulder calibration was done, after which the arm moved normally, but was still very loud, particularly the joint zero area. SA set up and 10pt both passed. Planning: 11:15-11:20; surgeon, RTC and CS reviewed plan prior to case. Arm attach: 11:40 (1 person; 30 sec) Draping: 12:01; surgeon assisted with draping, no issues. Platform 12:07 Shanz inserted in right PSIS and connected to 15 degree BMB via shanz bond. 3Define 12:09 Recording; towel placed over field and lights directed away. Resulting scan was accurate. Image Acquisition and Registration 12:10 Draw Spine: L4 12:10 Arm would not move to AP with error message saying arm did not reach final destination. RTC and surgeon noted that it's position following 3Define completion appeared to be extended further backwards than normal. CS sent to clear up and arm completed the movement but again looked extended backwards further than normal. CS attempted AP position, but the same error message was given. Surgeon engaged the E-stop and manually moved arm. E-stop was released and arm was successfully sent to AP. 2 images. 12:14 Arm sent to OBL: 2 images 12:16 Successful registration on first attempt off L4 with basic algorithm resulting in green values at both levels. Operate 12:17 L4 left 12:18 L5 left 12:19 L4 right 12:20 L5 right 12:21 Clear left; confirmatory images. Surgeon pleased with results. Visiting surgeon very impressed with speed and accuracy of the system. UNK MAZOR X ARM | CS completed SA set up and advanced test prior to case with normal results. After leaving the room and coming back several minutes later, CS noted that the SA was now making noise and shaking in the clear up position. CS attempted a 10 pt which failed. CS powered system down and tried new outlets x 2 with same results. RTC then plugged into a red outlet, waited 5min, and powered on. A shoulder calibration was done, after which the arm moved normally, but was still very loud, particularly the joint zero area. SA set up and 10pt both passed. Intraop, arm would not move to AP with error message saying arm did not reach final destination. RTC and surgeon noted that it's position following 3Define completion appeared to be extended further backwards than normal. CS sent to clear up and arm completed the movement but again looked extended backwards further than normal. CS attempted AP position, but the same error message was given. Surgeon engaged the E-stop and manually moved arm. E-stop was released and arm was successfully sent to AP. FWBMC System #92/Arm #101 3/22/18 Yellow -SA movement issues PSF L4-L5 Prep CS completed SA set up and advanced test prior to case with normal results. After leaving the room and coming back several minutes later, CS noted that the SA was now making noise and shaking in the clear up position. CS attempted a 10 pt which failed. CS powered system down and tried new outlets x 2 with same results. RTC then plugged into a red outlet, waited 5min, and powered on. A shoulder calibration was done, after which the arm moved normally, but was still very loud, particularly the joint zero area. SA set up and 10pt both passed. Planning: 11:15-11:20; surgeon, RTC and CS reviewed plan prior to case. Arm attach: 11:40 (1 person; 30 sec) Draping: 12:01; surgeon assisted with draping, no issues. Platform 12:07 Shanz inserted in right PSIS and connected to 15 degree BMB via shanz bond. 3Define 12:09 Recording; towel placed over field and lights directed away. Resulting scan was accurate. Image Acquisition and Registration 12:10 Draw Spine: L4 12:10 Arm would not move to AP with error message saying arm did not reach final destination. RTC and surgeon noted that it's position following 3Define completion appeared to be extended further backwards than normal. CS sent to clear up and arm completed the movement but again looked extended backwards further than normal. CS attempted AP position, but the same error message was given. Surgeon engaged the E-stop and manually moved arm. E-stop was released and arm was successfully sent to AP. 2 images. 12:14 Arm sent to OBL: 2 images 12:16 Successful registration on first attempt off L4 with basic algorithm resulting in green values at both levels. Operate 12:17 L4 left 12:18 L5 left 12:19 L4 right 12:20 L5 right 12:21 Clear left; confirmatory images. Surgeon pleased with results. Visiting surgeon very impressed with speed and accuracy of the system. | x surgical arm issue cs completed sa set advanced test prior case normal results leaving room coming minutes later cs noted sa making noise shaking clear position cs attempted pt failed cs powered tried new outlets x results rtc plugged red outlet waited powered shoulder calibration arm moved normally loud particularly joint zero area sa set passed intraop arm ap error message saying arm did reach final destination rtc surgeon noted position following completion appeared extended backwards normal cs sent clear arm completed movement looked extended backwards normal cs attempted ap position error message given surgeon engaged estop manually moved arm estop released arm successfully sent ap fwbmc yellow sa movement issues psf prep cs completed sa set advanced test prior case normal results leaving room coming minutes later cs noted sa making noise shaking clear position cs attempted pt failed cs powered tried new outlets x results rtc plugged red outlet waited powered shoulder calibration arm moved normally loud particularly joint zero area sa set passed planning surgeon rtc cs reviewed plan prior case arm attach draping surgeon assisted draping issues platform shanz inserted right psis connected degree bmb shanz bond recording towel placed field lights directed away resulting scan accurate image acquisition registration draw spine arm ap error message saying arm did reach final destination rtc surgeon noted position following completion appeared extended backwards normal cs sent clear arm completed movement looked extended backwards normal cs attempted ap position error message given surgeon engaged estop manually moved arm estop released arm successfully sent ap images arm sent obl images successful registration attempt basic algorithm resulting green values levels operate left left right right clear left confirmatory images surgeon pleased results visiting surgeon impressed speed accuracy unk mazor x arm | NAV4095 | NAV4101 | NAV3264 | NAV4098 | NAV2050 | NAV4095NAV4101NAV3264NAV4098NAV2050 | False
3932 | 703470269 | XSE - CANBUS Error While Using Bovie | 2019-11-15T16:09:49.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CANBUS Error While Using Bovie 2019-11-15 00678452 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when using a bovie during the procedure, a canbus error was seen. A reboot was done and the system was plugged into a different outlet from the outlet that the bovie was plugged into. This resolved the issue and the case was completed without this issue occurring again. There was no patient harm and the procedure was not delayed over an hour. 2019-12-05 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2019-11-15 00678452 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when using a bovie during the procedure, a canbus error was seen. A reboot was done and the system was plugged into a different outlet from the outlet that the bovie was plugged into. This resolved the issue and the case was completed without this issue occurring again. There was no patient harm and the procedure was not delayed over an hour. 2019-12-05 IUD (Other): No new information. | xse canbus error using bovie spinal procedure using bovie procedure canbus error seen reboot plugged different outlet outlet bovie plugged resolved issue case completed issue occurring procedure delayed hour iud new unk mazor x workstation | NAV2047 | NAV3261 | NAV2058 | NAV3295 | NAV3247 | NAV2047NAV3261NAV2058NAV3295NAV3247 | True
5106 | 703600201 | XSE - Software Installation Failure | 2020-02-06T04:13:30.000000+0000 | MazorX Stealth Edition | NAV2056 | SOFTWARE ISSUES | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Software Installation Failure 2020-01-31 00683513, e1 (Rep, HCP): Medtronic received information regarding installation of system software on a health care provider's PC. The surgeon was unable to open their older version of the software (SW) after Windows updates were done on the computer. First the SW was uninstalled, then SW 3.0 was installed in attempt to fix the database. Then SW 3.0 was uninstalled and the PC was restarted. SW 4.2 was installed and an error message appeared: "A version of Mazor is already installed on this computer" - See logs. The control panel and C:/ProgramFiles were checked and there was no trace of any system software. There was no patient involvement. 2020-02-20 IUD (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-01-31 00683513, e1 (Rep, HCP): Medtronic received information regarding installation of system software on a health care provider's PC. The surgeon was unable to open their older version of the software (SW) after Windows updates were done on the computer. First the SW was uninstalled, then SW 3.0 was installed in attempt to fix the database. Then SW 3.0 was uninstalled and the PC was restarted. SW 4.2 was installed and an error message appeared: "A version of Mazor is already installed on this computer" - See logs. The control panel and C:/ProgramFiles were checked and there was no trace of any system software. There was no patient involvement. 2020-02-20 IUD (Other): No new information. | xse software installation failure installation software health care providers pc surgeon unable open older version software windows updates computer sw uninstalled sw installed attempt fix database sw uninstalled pc restarted sw installed error message appeared version mazor installed computer logs control panel cprogramfiles checked trace software iud new mazor x software installation kit | NAV3246 | NAV2056 | NAV3247 | NAV3248 | NAV3262 | NAV3246NAV2056NAV3247NAV3248NAV3262 | True
1502 | 703202694 | X - 3define fails | 2019-06-12T18:12:39.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3define fails 2019-Jun-05 SCL1901676 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that When the rep checked the 3Define camera it was green, but once the system was mounted and 3Define scan was run, the camera failed. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-Jun-07 SCL1901689 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3 define was said to be not working properly. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X 3DEFINE CAMERA | 2019-Jun-05 SCL1901676 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that When the rep checked the 3Define camera it was green, but once the system was mounted and 3Define scan was run, the camera failed. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-Jun-07 SCL1901689 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3 define was said to be not working properly. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | x fails spinal procedure checked camera green mounted scan run camera failed additional complications anticipated spinal procedure define said working properly additional complications anticipated unk mazor x camera | NAV3244 | NAV2054 | NAV3295 | NAV3238 | NAV1006 | NAV3244NAV2054NAV3295NAV3238NAV1006 | True
7586 | 704233253 | XSE - PC Failure | 2021-02-16T21:02:52.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - PC Failure 2021-Feb-15 Salesforce C00221178 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the rep gets the error, “Failed to load Navigation Module. Navigation Mode will be disable”. She also gets the error message, “Medtronic database installation failed. Medtronic implants will not be supported Contact Technical Support”. They tried uninstalling and reinstalling the SW with many errors and issues. 2021-02-16 IUD x4 (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2021-Feb-15 Salesforce C00221178 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the rep gets the error, “Failed to load Navigation Module. Navigation Mode will be disable”. She also gets the error message, “Medtronic database installation failed. Medtronic implants will not be supported Contact Technical Support”. They tried uninstalling and reinstalling the SW with many errors and issues. 2021-02-16 IUD x4 (Other): No new information. | xse pc failure salesforce outside procedure gets error “failed load navigation module navigation mode disable” gets error message “medtronic database installation failed implants supported contact technical support” tried uninstalling reinstalling sw errors issues iud new mazor x pc station | NAV3246 | NAV3262 | NAV2054 | NAV2047 | NAV3235 | NAV3246NAV3262NAV2054NAV2047NAV3235 | True
9511 | 704660603 | XSE - Lit - Morse - Adverse Event | 2021-10-13T17:19:31.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Lit - Morse - Adverse Event Morse K, Heath M, Avrumova F, Defrancesco C, Fabricant P, Lebl D, Widmann R. Comprehensive Error Analysis for Robotic-assisted Placement of Pedicle Screws in Pediatric Spinal Deformity: The Initial Learning Curve. Journal of Pediatric Orthopaedics. 2021. DOI: 10.1 097/BPO.0000000000001842. Abstract: Surgical navigation improves pedicle screw insertion accuracy and reliability. Robotic-assisted spinal surgery and screw placement has not been fully assessed in pediatric patients with spine deformity undergoing posterior spinal fusion. The purpose of this study was to describe the learning curve for robotically assisted pedicle screw placement in pediatric patients. A retrospective review on a consecutive series of the first 19 pediatric patients who underwent posterior spinal fusion by a single surgeon using robotic navigation was performed. Demographics, curve parameters, pedicle diameter, vertebral rotation, and additional outcome measures were recorded. Screw position was assessed with calibrated intraoperative 3-dimensional fluoroscopic images. All complications of planned and placed robotically placed screws were recorded. A total of 194 left-sided screws were planned as robot assisted. One hundred sixty-eight of the robotically planned screws (86.6%) were placed with robot assistance; 29 robotically planned screws (15.0%) were abandoned or converted to freehand. The mean time per robotically placed screw was 3.6 ±2.4 minutes. Fifteen breaches (8.9%) and 1 anterior perforation occurred with 2 critical (> 2 to 4 mm) breaches, 1 was associated with a durotomy, and both occurred in the first case. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. Six breaches occurred in the first case. The odds ratio of obtaining a breach in screws with a matched trajectory was 0.275 (95% confidence interval: 0.089-0.848). Screw time and accuracy improved and the number of breaches decreased after 10 cases. This series had 2 critical breaches (between 2 and 4 mm) on the first case. Overall, excluding the 2 critical breaches, 98.8% of robotically executed screws were placed without a critical breach, which is comparable to previous pediatric deformity studies. Caution should be exercised during the initial training period to avoid complications as experience and training lead to an improved understanding of surgical planning, skive, and soft-tissue pressure on the end-effector all of which can impact accuracy of robotically assisted pedicle screw placement. Reported Events: 1. After use of the guidance system during spinal procedure to place screws, two medial breaches occurred in the one case, one of which was associated with a durotomy. Critically breached screws were removed and replaced with either a pedicle screw using freehand technique or a pedicle hook. The durotomy occurred during drilling due to operator error from inadequate tightening of the screw between the bone mount bridge and dual spinous process clamp. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. MAZOR X SYSTEM TPL0059 | Morse K, Heath M, Avrumova F, Defrancesco C, Fabricant P, Lebl D, Widmann R. Comprehensive Error Analysis for Robotic-assisted Placement of Pedicle Screws in Pediatric Spinal Deformity: The Initial Learning Curve. Journal of Pediatric Orthopaedics. 2021. DOI: 10.1 097/BPO.0000000000001842. Abstract: Surgical navigation improves pedicle screw insertion accuracy and reliability. Robotic-assisted spinal surgery and screw placement has not been fully assessed in pediatric patients with spine deformity undergoing posterior spinal fusion. The purpose of this study was to describe the learning curve for robotically assisted pedicle screw placement in pediatric patients. A retrospective review on a consecutive series of the first 19 pediatric patients who underwent posterior spinal fusion by a single surgeon using robotic navigation was performed. Demographics, curve parameters, pedicle diameter, vertebral rotation, and additional outcome measures were recorded. Screw position was assessed with calibrated intraoperative 3-dimensional fluoroscopic images. All complications of planned and placed robotically placed screws were recorded. A total of 194 left-sided screws were planned as robot assisted. One hundred sixty-eight of the robotically planned screws (86.6%) were placed with robot assistance; 29 robotically planned screws (15.0%) were abandoned or converted to freehand. The mean time per robotically placed screw was 3.6 ±2.4 minutes. Fifteen breaches (8.9%) and 1 anterior perforation occurred with 2 critical (> 2 to 4 mm) breaches, 1 was associated with a durotomy, and both occurred in the first case. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. Six breaches occurred in the first case. The odds ratio of obtaining a breach in screws with a matched trajectory was 0.275 (95% confidence interval: 0.089-0.848). Screw time and accuracy improved and the number of breaches decreased after 10 cases. This series had 2 critical breaches (between 2 and 4 mm) on the first case. Overall, excluding the 2 critical breaches, 98.8% of robotically executed screws were placed without a critical breach, which is comparable to previous pediatric deformity studies. Caution should be exercised during the initial training period to avoid complications as experience and training lead to an improved understanding of surgical planning, skive, and soft-tissue pressure on the end-effector all of which can impact accuracy of robotically assisted pedicle screw placement. Reported Events: 1. After use of the guidance system during spinal procedure to place screws, two medial breaches occurred in the one case, one of which was associated with a durotomy. Critically breached screws were removed and replaced with either a pedicle screw using freehand technique or a pedicle hook. The durotomy occurred during drilling due to operator error from inadequate tightening of the screw between the bone mount bridge and dual spinous process clamp. There were no intraoperative/postoperative neuromonitoring changes and no sequela from the durotomy. | xse lit morse adverse morse k heath m avrumova f defrancesco c fabricant p lebl d widmann r comprehensive error analysis roboticassisted placement pedicle screws pediatric spinal deformity initial learning curve journal pediatric orthopaedics doi abstract surgical navigation improves pedicle screw insertion accuracy reliability roboticassisted spinal surgery screw placement fully assessed pediatric patients spine deformity undergoing posterior spinal fusion purpose study learning curve robotically assisted pedicle screw placement pediatric patients retrospective review consecutive series pediatric patients underwent posterior spinal fusion single surgeon using robotic navigation performed demographics curve parameters pedicle diameter vertebral rotation additional outcome measures recorded screw position assessed calibrated intraoperative fluoroscopic images complications planned placed robotically placed screws recorded total leftsided screws planned robot assisted sixtyeight robotically planned screws placed robot assistance robotically planned screws abandoned converted freehand mean time robotically placed screw ± minutes breaches anterior perforation occurred critical breaches associated durotomy occurred case intraoperativepostoperative neuromonitoring changes sequela durotomy breaches occurred case odds ratio obtaining breach screws matched trajectory screw time accuracy improved number breaches decreased cases series critical breaches case overall excluding critical breaches robotically executed screws placed critical breach comparable previous pediatric deformity studies caution exercised initial training period avoid complications experience training lead improved understanding surgical planning skive softtissue pressure endeffector impact accuracy robotically assisted pedicle screw placement events use spinal procedure place screws medial breaches occurred case associated durotomy critically breached screws removed replaced pedicle screw using freehand technique pedicle hook durotomy occurred drilling operator error inadequate tightening screw bone mount bridge dual spinous process clamp intraoperativepostoperative neuromonitoring changes sequela durotomy mazor x | NAV5000 | NAV3244 | NAV3243 | NAV1006 | NAV3261 | NAV5000NAV3244NAV3243NAV1006NAV3261 | True
6097 | 703868556 | XSE - Black Screen | 2020-07-23T23:18:52.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - Black Screen 2020-07-23 00693345 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a procedure. It was reported that the main monitor goes black for about 15 seconds then the surgeon monitor goes black. The main monitor comes back up but the surgeon monitor sometimes comes back up and sometimes it does not. They slave out to external monitors and they stay on during these times. This is reoccurring issue. There was no delay to the procedure and no known impact to the patient outcome. 2020-08-09 IUD x2 (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2020-07-23 00693345 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a procedure. It was reported that the main monitor goes black for about 15 seconds then the surgeon monitor goes black. The main monitor comes back up but the surgeon monitor sometimes comes back up and sometimes it does not. They slave out to external monitors and they stay on during these times. This is reoccurring issue. There was no delay to the procedure and no known impact to the patient outcome. 2020-08-09 IUD x2 (Other): No new information. | xse black screen intraoperatively procedure main monitor goes black seconds surgeon monitor goes black main monitor comes surgeon monitor comes does slave external monitors stay times reoccurring issue delay procedure known impact outcome iud new kit svc surgical syst type | NAV3261 | NAV2054 | NAV3253 | NAV3247 | NAV3250 | NAV3261NAV2054NAV3253NAV3247NAV3250 | True
6531 | 703968747 | XSE ¦ X-eye Not Functioning | 2020-09-21T20:00:02.000000+0000 | MazorX Stealth Edition | NAV4110 | NO TRACKING | 20 | KIT1284-01 | SVC KIT1284-01 X-EYE CAMERA | XSE ¦ X-eye Not Functioning 2020-09-21 00697137 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the x-eye camera does not have a video feed and not functioning. The green power light seems to be blinking. No patient was present. 2020-09-29 IUD (Other): No new information. SVC KIT1284-01 X-EYE CAMERA | 2020-09-21 00697137 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the x-eye camera does not have a video feed and not functioning. The green power light seems to be blinking. No patient was present. 2020-09-29 IUD (Other): No new information. | xse xeye functioning outside procedure xeye camera does video feed functioning green power light blinking present iud new svc xeye camera | NAV4110 | NAV3254 | NAV2058 | NAV3299 | NAV1006 | NAV4110NAV3254NAV2058NAV3299NAV1006 | True
10396 | 704871253 | XSE ¦ Surgeon Monitor Fell Off Hinge | 2022-02-18T05:05:51.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0356-01S | KIT SVC ASM0356-01S SURGEON SCREEN | XSE ¦ Surgeon Monitor Fell Off Hinge 2022-Feb-17 Salesforce C00243870 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the su rgeon monitor fell of the hinges and needs to be repaired. There was no known impact on the patient outcome. 2022-02-22 IUD x2 (Other): No new information. 2022-02-28 email 2022-02-28 16:30 (Rep): No new information. 2022-03-03 IUD (Other): No new information. 2022-03-17 IUD (Other): No new information. KIT SVC ASM0356-01S SURGEON SCREEN | 2022-Feb-17 Salesforce C00243870 (Rep, HCP): Medtronic received information regarding a guidance system. It was reported that the su rgeon monitor fell of the hinges and needs to be repaired. There was no known impact on the patient outcome. 2022-02-22 IUD x2 (Other): No new information. 2022-02-28 email 2022-02-28 16:30 (Rep): No new information. 2022-03-03 IUD (Other): No new information. 2022-03-17 IUD (Other): No new information. | xse surgeon monitor fell hinge salesforce su rgeon monitor fell hinges needs repaired known impact outcome iud new email new iud new iud new kit svc surgeon screen | NAV3253 | NAV3261 | NAV2054 | NAV3249 | NAV4101 | NAV3253NAV3261NAV2054NAV3249NAV4101 | True
5299 | 703644530 | XSE - BICE, Arm Failed Accuracy Tests | 2020-03-02T16:14:38.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - BICE, Arm Failed Accuracy Tests 2020-03-02 00685433 ID (Rep): Medtronic received information regarding a guidance system being used at a medical facility. It was reported that the surgical arm failed the set-up test; the accuracy pointer was hitting "the white part of the arm" but 1-2mm. The arm also failed the advanced accuracy test at every point and was again "hitting the white part of the arm." No patient was involved. 2020-03-06 IUD x4 (Other): No new information. 2020-04-07 IUD x2 (Other): No new information. 2020-05-31 IUD (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-03-02 00685433 ID (Rep): Medtronic received information regarding a guidance system being used at a medical facility. It was reported that the surgical arm failed the set-up test; the accuracy pointer was hitting "the white part of the arm" but 1-2mm. The arm also failed the advanced accuracy test at every point and was again "hitting the white part of the arm." No patient was involved. 2020-03-06 IUD x4 (Other): No new information. 2020-04-07 IUD x2 (Other): No new information. 2020-05-31 IUD (Other): No new information. | xse bice arm failed accuracy tests id medical facility surgical arm failed setup test accuracy pointer hitting white arm arm failed advanced accuracy test point hitting white arm involved iud new iud new iud new kit svc surgical arm | NAV4093 | NAV4092 | NAV2054 | NAV3249 | NAV4096 | NAV4093NAV4092NAV2054NAV3249NAV4096 | True
4405 | 703541326 | X- X-eye issue | 2020-01-02T18:36:21.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X- X-eye issue 2018-4-23 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the x-eye was having trouble seeing the stylus during draw spine. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X X-EYE CAMERA | 2018-4-23 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the x-eye was having trouble seeing the stylus during draw spine. There was no patient harm and the procedure was not delayed over an hour. | x xeye issue ccr spinal procedure xeye having trouble seeing stylus draw spine procedure delayed hour unk mazor x xeye camera | NAV4110 | NAV3254 | NAV2058 | NAV4109 | NAV3239 | NAV4110NAV3254NAV2058NAV4109NAV3239 | True
7690 | 704248829 | XSE - CANBUS Error Message | 2021-02-25T05:07:47.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error Message 2021-02-24 00707250, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error message was displayed and the system was not communicating with nodes 1-8. Bovie was being used when the error message was displayed. A soft restart of the system did not resolve the issue. A hard reboot of the system was done and the issue was resolved and the system was confirmed to be accurate. There was no patient harm and the procedure was delayed less than an hour. 2021-Mar-15 IUD (other): No new information. 2021-03-23 IUD x2 (Other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2021-02-24 00707250, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error message was displayed and the system was not communicating with nodes 1-8. Bovie was being used when the error message was displayed. A soft restart of the system did not resolve the issue. A hard reboot of the system was done and the issue was resolved and the system was confirmed to be accurate. There was no patient harm and the procedure was delayed less than an hour. 2021-Mar-15 IUD (other): No new information. 2021-03-23 IUD x2 (Other): No new information. | xse canbus error message iud spinal procedure canbus error message displayed communicating nodes bovie error message displayed soft restart did resolve issue hard reboot issue resolved confirmed accurate procedure delayed hour iud new iud new svc main cable | NAV2047 | NAV4095 | NAV4098 | NAV3256 | NAV3249 | NAV2047NAV4095NAV4098NAV3256NAV3249 | True
8433 | 704413953 | XSE - Improvement Suggestions | 2021-05-25T19:41:50.000000+0000 | MazorX Stealth Edition | NAV1013 | MAZOR X/STEALTH EDITION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Improvement Suggestions 2021-04-27 Email (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that one thing the manufacturer representative stated that the system could do better was to streamline the work flow and reduce set up errors and registration errors. No patient was present. MAZOR X SYSTEM TPL0059 | 2021-04-27 Email (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that one thing the manufacturer representative stated that the system could do better was to streamline the work flow and reduce set up errors and registration errors. No patient was present. | xse improvement suggestions email outside procedure thing stated better streamline work flow reduce set errors registration errors present mazor x | NAV3242 | NAV3246 | NAV3264 | NAV2047 | NAV2054 | NAV3242NAV3246NAV3264NAV2047NAV2054 | False
5041 | 703592057 | XSE - CANBUS Error | 2020-01-31T20:11:14.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CANBUS Error 2020-01-31 00683174 ID (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during the case an error T95 was received: "Fatal Error. CANBUS lost connection to nodes 2, 4, 7, and POsitioner IO 8." Technical Services were contacted and the manufacturer representative present for the case was instructed to perform a hard reboot. This was completed and the issue was resolved. No patient harm was reported and surgery was delayed less than an hour. 2020-02-05 Email 2020-02-05 07:16, IUD (Rep): No new information. UNK MAZOR X WORKSTATION | 2020-01-31 00683174 ID (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during the case an error T95 was received: "Fatal Error. CANBUS lost connection to nodes 2, 4, 7, and POsitioner IO 8." Technical Services were contacted and the manufacturer representative present for the case was instructed to perform a hard reboot. This was completed and the issue was resolved. No patient harm was reported and surgery was delayed less than an hour. 2020-02-05 Email 2020-02-05 07:16, IUD (Rep): No new information. | xse canbus error id spinal procedure case error fatal error canbus lost connection nodes positioner io technical services contacted present case instructed perform hard reboot completed issue resolved surgery delayed hour email iud new unk mazor x workstation | NAV2047 | NAV4104 | NAV4095 | NAV3244 | NAV3257 | NAV2047NAV4104NAV4095NAV3244NAV3257 | True
9880 | 704740567 | XSE - Cross-Threaded Screw | 2021-11-30T23:08:24.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 20 | KIT0586-03 | REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | XSE - Cross-Threaded Screw 2021-Nov-30 Salesforce C00238935 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was shipped with a cross-threaded screw on the 12in image adapter. No patient was present. 2021-12-01 IUD x2 (Other): No new information. 2021-12-09 IUD (Other): No new information. 2021-12-20 email 2021-12-20 12:46 (Rep): Additional information was received stating that this was an out of box failure. When trying to convert 12” to 9”, the screw wouldn’t come out. 2022-01-11 email 2022-01-11 13:16, email 2022-01-11 14:16, e1, e2 (Rep): No new information. 2022-01-23 e3, e4, e5, e6 (Rep): Additional information received from a manufacturer representative reported that when converting the image adapter from 12" to 9" the screw that connected the adapter to the C-arm would not come out. The representative was not able to put in the longer screw with the washers. The mount from a new image adapter was used to fix the image adapter at the site. 2022-02-08 IUD (Other): No new information. 2022-02-17 IUD (Other): No new information. 2022-03-20 email 2022-03-20 06:11 (Other): No new information. REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | 2021-Nov-30 Salesforce C00238935 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was shipped with a cross-threaded screw on the 12in image adapter. No patient was present. 2021-12-01 IUD x2 (Other): No new information. 2021-12-09 IUD (Other): No new information. 2021-12-20 email 2021-12-20 12:46 (Rep): Additional information was received stating that this was an out of box failure. When trying to convert 12” to 9”, the screw wouldn’t come out. 2022-01-11 email 2022-01-11 13:16, email 2022-01-11 14:16, e1, e2 (Rep): No new information. 2022-01-23 e3, e4, e5, e6 (Rep): Additional information received from a manufacturer representative reported that when converting the image adapter from 12" to 9" the screw that connected the adapter to the C-arm would not come out. The representative was not able to put in the longer screw with the washers. The mount from a new image adapter was used to fix the image adapter at the site. 2022-02-08 IUD (Other): No new information. 2022-02-17 IUD (Other): No new information. 2022-03-20 email 2022-03-20 06:11 (Other): No new information. | xse crossthreaded screw salesforce outside procedure shipped crossthreaded screw image adapter present iud new iud new email additional stating box failure trying convert ” ” screw wouldn’t come email email new additional converting image adapter screw connected adapter carm come able longer screw washers mount new image adapter fix image adapter site iud new iud new email new replacement grip set image adaptor | NAV3258 | NAV2054 | NAV3299 | NAV3237 | NAV4092 | NAV3258NAV2054NAV3299NAV3237NAV4092 | True
2136 | 703226449 | X - SURGICAL SYSTEM ISSUE | 2019-06-25T18:30:40.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - SURGICAL SYSTEM ISSUE Major trouble removing the surgical arm from the bed. The software would be put into storage state, and the manipulator moved to home position. However, a red warning would appear saying that the manipulator was not in the home position. The software would have to be turned off storage state, then turned back on and attempted again. After many attempts, the positioner finally went to the home position. When unlocking the positioner and removing it from bed frame, the lights began to blink in what appeared to be in a random sequence (a video of this will uploaded to dropbox) . The positioner provided significant resistance to being removed from the bed. After some physical motivation, the positioner came loose from the bed. The second the positioner cleared the bed frame adapter, the positioner fell to very lowest position and leaking air. With the positioner being in this position, it was impossible to put the surgical arm back into the workstation. After speaking with Leafar, the CSR overrode the lock and manually lifted the positioner. This required a great deal of force, so the CSR had the circulating nurse run the workstation during this process. This allowed the CSR to get the surgical arm back into the workstation, but it was then noted that the lights on the positioner were again blinking in a random sequence. Once the surgical arm was back into the workstation, a warning appeared stating "Manipulator D5 is not in home position Surgical Arm is Not in Storage Position. Do you wish to close the cover?". Since the CSR could clearly see that the surgical arm was in storage position, the cover was closed./ THP -PHOP Workstation #60/Arm #71 Software 3.0 Red -Problems with the positioner DX: Revision. PLANNING: The planning was done by the surgeon prior to the start of the procedure. PREP: 12:20 -Workstation turned on. When turned on, it was found that shoulder was unlocked from the start and the light on the shoulder was blinking as though it was unlocked. However, the shoulder was actually locked, would not move, and the button for unlocking the shoulderon the arm would not respond. After returning to the workstation, and locking the shoulder on the screen, the arm came loose and moved freely on it's own. The arm then returned to full functionality. A video will be uploaded to dropbox depicting the issue. 13:29 -Surgical arm attached to the bed -one person 13:35 -Surgical arm draped -one person -30 seconds 13:43 -Incision 13:43 -Workstation turned off 15:12 -Workstation turned on 15:16 -Tablet paired with workstation. PLATFORM MOUNTING: 15:21 -Surgical arm setup performed by the surgeon 15:23 -Dual schanze pins mounted. The pins were mounted angled with a more caudal tilt, and straight up and down in the PSIS. The surgeon was pleased with this mount. 15:23 -Surgical arm locked. 3DEFINE: 15:24 -Recording; wet towel placed over incision and lights pointed away. Resulting scan was clear and detailed. 15:26 -Drawspine -failed. CSR pointed out that the surgeon pulled the stylus away before scrubtechpressed stop on drawspine. 15:26 -Drawspine -successful. IMAGE ACQUISITION AND REGISTRATION: 15:27 -Surgical arm se nt to the AP position of L5 15:31 -Surgical arm sent to the OBL position of L5 15:33 -Registration attempt of L5 successful with the basic algorithm 15:34 -Fluoro match successful with all green values. OPERATION: 15:36 -Left S2 -aborted **CSR had to take control because the tablet would not register touch from scrubtech 15:40 -Right S2 15:41 -Left S1 15:45 -Left L5 15:47 -Left S2 15:53 -Right L5 15:58 -Right S1. 16:10 -Major trouble removing the surgical arm from the bed. The software would be put into storagestate, and the manipulator moved to home position. However, a red warning would appear saying that the manipulator was not in the home position. The software would have to be turned off storage state, then turned back on and attempted again. After many attempts, the positioner finally went to the home position. When unlocking the positioner and removing it from bed frame, the lights began to blink in what appeared to be in a random sequence (a video of this will uploaded to dropbox) . The positioner provided significant resistance to being removed from the bed. After some physical motivation, the positioner came loose from the bed. The second the positioner cleared the bed frame adapter, the positioner fell to very lowest position and leaking air. With the positioner being in this position,it was impossible to putthe surgical arm back into the workstation. After speaking with Leafar, the CSR overrode the lock and manually lifted the positioner. This required a great deal of force, so the CSR had the circulating nurse run the workstation during this p rocess. This allowed the CSR to get the surgical arm back into theworkstation, but it was then noted that the lights on the positioner were again blinking in a random sequence. Once the surgical arm was back into the workstation, a warning appeared stating "Manipu lator D5 is not in home position Surgical Arm is Not in Storage Position. Do you wish to close the cover?". Since the CSR could clearly see that the surgical arm was in storage position, the cover was closed. Overall, the surgeon was very pleased with the results. The surgeon seemed annoyed with the trouble of getting the positionerinto the home position. UNK MAZOR X SHOULDER | Major trouble removing the surgical arm from the bed. The software would be put into storage state, and the manipulator moved to home position. However, a red warning would appear saying that the manipulator was not in the home position. The software would have to be turned off storage state, then turned back on and attempted again. After many attempts, the positioner finally went to the home position. When unlocking the positioner and removing it from bed frame, the lights began to blink in what appeared to be in a random sequence (a video of this will uploaded to dropbox) . The positioner provided significant resistance to being removed from the bed. After some physical motivation, the positioner came loose from the bed. The second the positioner cleared the bed frame adapter, the positioner fell to very lowest position and leaking air. With the positioner being in this position, it was impossible to put the surgical arm back into the workstation. After speaking with Leafar, the CSR overrode the lock and manually lifted the positioner. This required a great deal of force, so the CSR had the circulating nurse run the workstation during this process. This allowed the CSR to get the surgical arm back into the workstation, but it was then noted that the lights on the positioner were again blinking in a random sequence. Once the surgical arm was back into the workstation, a warning appeared stating "Manipulator D5 is not in home position Surgical Arm is Not in Storage Position. Do you wish to close the cover?". Since the CSR could clearly see that the surgical arm was in storage position, the cover was closed./ THP -PHOP Workstation #60/Arm #71 Software 3.0 Red -Problems with the positioner DX: Revision. PLANNING: The planning was done by the surgeon prior to the start of the procedure. PREP: 12:20 -Workstation turned on. When turned on, it was found that shoulder was unlocked from the start and the light on the shoulder was blinking as though it was unlocked. However, the shoulder was actually locked, would not move, and the button for unlocking the shoulderon the arm would not respond. After returning to the workstation, and locking the shoulder on the screen, the arm came loose and moved freely on it's own. The arm then returned to full functionality. A video will be uploaded to dropbox depicting the issue. 13:29 -Surgical arm attached to the bed -one person 13:35 -Surgical arm draped -one person -30 seconds 13:43 -Incision 13:43 -Workstation turned off 15:12 -Workstation turned on 15:16 -Tablet paired with workstation. PLATFORM MOUNTING: 15:21 -Surgical arm setup performed by the surgeon 15:23 -Dual schanze pins mounted. The pins were mounted angled with a more caudal tilt, and straight up and down in the PSIS. The surgeon was pleased with this mount. 15:23 -Surgical arm locked. 3DEFINE: 15:24 -Recording; wet towel placed over incision and lights pointed away. Resulting scan was clear and detailed. 15:26 -Drawspine -failed. CSR pointed out that the surgeon pulled the stylus away before scrubtechpressed stop on drawspine. 15:26 -Drawspine -successful. IMAGE ACQUISITION AND REGISTRATION: 15:27 -Surgical arm se nt to the AP position of L5 15:31 -Surgical arm sent to the OBL position of L5 15:33 -Registration attempt of L5 successful with the basic algorithm 15:34 -Fluoro match successful with all green values. OPERATION: 15:36 -Left S2 -aborted **CSR had to take control because the tablet would not register touch from scrubtech 15:40 -Right S2 15:41 -Left S1 15:45 -Left L5 15:47 -Left S2 15:53 -Right L5 15:58 -Right S1. 16:10 -Major trouble removing the surgical arm from the bed. The software would be put into storagestate, and the manipulator moved to home position. However, a red warning would appear saying that the manipulator was not in the home position. The software would have to be turned off storage state, then turned back on and attempted again. After many attempts, the positioner finally went to the home position. When unlocking the positioner and removing it from bed frame, the lights began to blink in what appeared to be in a random sequence (a video of this will uploaded to dropbox) . The positioner provided significant resistance to being removed from the bed. After some physical motivation, the positioner came loose from the bed. The second the positioner cleared the bed frame adapter, the positioner fell to very lowest position and leaking air. With the positioner being in this position,it was impossible to putthe surgical arm back into the workstation. After speaking with Leafar, the CSR overrode the lock and manually lifted the positioner. This required a great deal of force, so the CSR had the circulating nurse run the workstation during this p rocess. This allowed the CSR to get the surgical arm back into theworkstation, but it was then noted that the lights on the positioner were again blinking in a random sequence. Once the surgical arm was back into the workstation, a warning appeared stating "Manipu lator D5 is not in home position Surgical Arm is Not in Storage Position. Do you wish to close the cover?". Since the CSR could clearly see that the surgical arm was in storage position, the cover was closed. Overall, the surgeon was very pleased with the results. The surgeon seemed annoyed with the trouble of getting the positionerinto the home position. | x surgical issue major trouble removing surgical arm bed software storage state manipulator moved home position red warning appear saying manipulator home position software turned storage state turned attempted attempts positioner finally went home position unlocking positioner removing bed frame lights began blink appeared random sequence positioner provided significant resistance removed bed physical motivation positioner came loose bed second positioner cleared bed frame adapter positioner fell lowest position leaking air positioner position impossible surgical arm workstation speaking leafar csr overrode lock manually lifted positioner required great deal force csr circulating nurse run workstation process allowed csr surgical arm workstation noted lights positioner blinking random sequence surgical arm workstation warning appeared stating manipulator home position surgical arm storage position wish close cover csr clearly surgical arm storage position cover closed thp phop workstation software red problems positioner dx revision planning planning surgeon prior start procedure prep workstation turned turned shoulder unlocked start light shoulder blinking unlocked shoulder actually locked button unlocking shoulderon arm respond returning workstation locking shoulder screen arm came loose moved freely arm returned functionality video uploaded dropbox depicting issue surgical arm attached bed person surgical arm draped person seconds incision workstation turned workstation turned tablet paired workstation platform mounting surgical arm setup performed surgeon dual schanze pins mounted pins mounted angled caudal tilt straight psis surgeon pleased mount surgical arm locked recording wet towel placed incision lights pointed away resulting scan clear detailed drawspine failed csr pointed surgeon pulled stylus away scrubtechpressed stop drawspine drawspine successful image acquisition registration surgical arm se nt ap position surgical arm sent obl position registration attempt successful basic algorithm fluoro match successful green values operation left aborted csr control tablet register touch scrubtech right left left left right right major trouble removing surgical arm bed software storagestate manipulator moved home position red warning appear saying manipulator home position software turned storage state turned attempted attempts positioner finally went home position unlocking positioner removing bed frame lights began blink appeared random sequence positioner provided significant resistance removed bed physical motivation positioner came loose bed second positioner cleared bed frame adapter positioner fell lowest position leaking air positioner positionit impossible putthe surgical arm workstation speaking leafar csr overrode lock manually lifted positioner required great deal force csr circulating nurse run workstation p rocess allowed csr surgical arm theworkstation noted lights positioner blinking random sequence surgical arm workstation warning appeared stating manipu lator home position surgical arm storage position wish close cover csr clearly surgical arm storage position cover closed overall surgeon pleased results surgeon annoyed trouble getting positionerinto home position unk mazor x shoulder | NAV3259 | NAV4101 | NAV4102 | NAV3253 | NAV3250 | NAV3259NAV4101NAV4102NAV3253NAV3250 | False
3584 | 703381586 | X - Surgical Arm Failed Accuracy Test | 2019-09-26T03:09:41.000000+0000 | MazorX | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical Arm Failed Accuracy Test 2019-02-07 SCL1900219 (REP): Medtronic received information regarding a guidance system during preventative maintenance. It was reported that the surgical arm failed the accuracy test. No patient was present. SURGICAL ARM | 2019-02-07 SCL1900219 (REP): Medtronic received information regarding a guidance system during preventative maintenance. It was reported that the surgical arm failed the accuracy test. No patient was present. | x surgical arm failed accuracy test preventative maintenance surgical arm failed accuracy test present surgical arm | NAV4093 | NAV4092 | NAV3249 | NAV2054 | NAV4097 | NAV4093NAV4092NAV3249NAV2054NAV4097 | True
111 | 703000356 | X - Communication Issue | 2019-02-11T11:21:15.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0206-01 | SURGICAL ARM | X - Communication Issue 2019-Apr-23 _e1 (rep): Additional information received from the manufacturer representative (rep) indicated that the system software was version 3.0.2.1. The surgery was delayed almost an hour before it was cancelled. There was no patient harm due to the delay. Troubleshooting included a soft restart, hard restart, tech support called and directions were followed until it became apparent that a field service engineer was needed to fix the issue. 2019-02-05 SCL1900314, Intake Form, Attachment 2019-02-11 05:28 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system lost connection to nodes 1-6 and 8 and the system thought the emergency stop button was pressed when it was not. The patient was anesthetized and there was no patient harm; however, the surgery was delayed for more than an hour and the case was canceled due to the errors. SURGICAL ARM | 2019-Apr-23 _e1 (rep): Additional information received from the manufacturer representative (rep) indicated that the system software was version 3.0.2.1. The surgery was delayed almost an hour before it was cancelled. There was no patient harm due to the delay. Troubleshooting included a soft restart, hard restart, tech support called and directions were followed until it became apparent that a field service engineer was needed to fix the issue. 2019-02-05 SCL1900314, Intake Form, Attachment 2019-02-11 05:28 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system lost connection to nodes 1-6 and 8 and the system thought the emergency stop button was pressed when it was not. The patient was anesthetized and there was no patient harm; however, the surgery was delayed for more than an hour and the case was canceled due to the errors. | x communication issue additional indicated software version surgery delayed hour cancelled delay troubleshooting included soft restart hard restart tech support called directions followed apparent field service engineer needed fix issue attachment spinal procedure lost connection nodes thought emergency stop button pressed anesthetized surgery delayed hour case canceled errors surgical arm | NAV4095 | NAV2047 | NAV3248 | NAV3249 | NAV3242 | NAV4095NAV2047NAV3248NAV3249NAV3242 | True
6811 | 704037519 | XSE - Joint 4 Error | 2020-10-27T22:38:57.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Joint 4 Error 2020-10-27 00699653 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the arm displayed an error in joint 4. Soft restart did not clear the error. Hard restart did not clear the error. This was prior to a case, during set up. No patient present. 2020-10-28 IUD x7 (Other): No new information. 2020-Nov-19 IUD x2 (rep): No new information. 2021-01-27 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-10-27 00699653 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the arm displayed an error in joint 4. Soft restart did not clear the error. Hard restart did not clear the error. This was prior to a case, during set up. No patient present. 2020-10-28 IUD x7 (Other): No new information. 2020-Nov-19 IUD x2 (rep): No new information. 2021-01-27 IUD (Other): No new information. | xse joint error outside procedure arm displayed error joint soft restart did clear error hard restart did clear error prior case set present iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4098 | NAV2047 | NAV4095 | NAV2054 | NAV3249NAV4098NAV2047NAV4095NAV2054 | True
9963 | 704758646 | XSE ¦ software freezes | 2021-12-10T15:18:17.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE ¦ software freezes 2021-Dec-10 Salesforce C00239525 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was delayed 3Define, software keeps timing out. Spinning wheel, delaying planning process. There was no patient involved. 2021-12-15 IUD x2 (Other): No new information. 2022-03-23 IUD (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Dec-10 Salesforce C00239525 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was delayed 3Define, software keeps timing out. Spinning wheel, delaying planning process. There was no patient involved. 2021-12-15 IUD x2 (Other): No new information. 2022-03-23 IUD (Other): No new information. | xse software freezes salesforce outside procedure delayed software keeps timing spinning wheel delaying planning process involved iud new iud new mazorx sw installation kit | NAV3246 | NAV3247 | NAV3248 | NAV2056 | NAV3249 | NAV3246NAV3247NAV3248NAV2056NAV3249 | True
4675 | 703561811 | X - Registration issues | 2020-01-14T20:03:08.000000+0000 | MazorX | NAV3236 | CT IMAGES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration issues 2020-1-12 Intake Form, 00682025(Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the CT quality was degraded , and therefore it was almost impossible to approve registration. There was no patient harm and the procedure was not delayed over an hour. 2020-02-23 Email 2020-02-23 01:30 (Rep): Additional information received from a manufacturer representative reported that during the case, an attempt was made to change contrast and brightness but the function did not make any change to the quality of the image. Also, CT quality degradation was due to the Mazor X SW. The Same CT was used for planning. During planning the CT quality was Good. The degradation effect only happened during the verification step. MAZOR X SYSTEM TPL0059 | 2020-1-12 Intake Form, 00682025(Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the CT quality was degraded , and therefore it was almost impossible to approve registration. There was no patient harm and the procedure was not delayed over an hour. 2020-02-23 Email 2020-02-23 01:30 (Rep): Additional information received from a manufacturer representative reported that during the case, an attempt was made to change contrast and brightness but the function did not make any change to the quality of the image. Also, CT quality degradation was due to the Mazor X SW. The Same CT was used for planning. During planning the CT quality was Good. The degradation effect only happened during the verification step. | x registration issues spinal procedure ct quality degraded impossible approve registration procedure delayed hour email additional case attempt change contrast brightness function did make change quality image ct quality degradation mazor x sw ct planning planning ct quality good degradation effect happened verification step mazor x | NAV3242 | NAV4112 | NAV3246 | NAV1012 | NAV4091 | NAV3242NAV4112NAV3246NAV1012NAV4091 | False
5982 | 703852143 | XSE - Soft Tissue Pressure/Aborted Case | 2020-07-14T23:13:37.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Soft Tissue Pressure/Aborted Case 2020-07-14 00692709 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that an open case was done to remove instrumentation and place screws. The screws were planned lateral and an imaging spin was done. The surgical arm was sent to right L1 and the screw was placed accurately. The arm was then sent to left L1 and the trajectory was difficult to reach due to soft tissue pressure, but the scre was able to be placed. While attempting to place the screws at L3, the dilator was dropped when being passed to the surgeon and was no longer sterile. The final screws were placed with stealth navigation due to the guidance system being aborted due to the weight of the patient causing soft tissue pressure and the dilator being dropped. There was no patient harm and the procedure was delayed less than an hour. 2020-07-21 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-07-14 00692709 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that an open case was done to remove instrumentation and place screws. The screws were planned lateral and an imaging spin was done. The surgical arm was sent to right L1 and the screw was placed accurately. The arm was then sent to left L1 and the trajectory was difficult to reach due to soft tissue pressure, but the scre was able to be placed. While attempting to place the screws at L3, the dilator was dropped when being passed to the surgeon and was no longer sterile. The final screws were placed with stealth navigation due to the guidance system being aborted due to the weight of the patient causing soft tissue pressure and the dilator being dropped. There was no patient harm and the procedure was delayed less than an hour. 2020-07-21 IUD (Other): No new information. | xse soft tissue pressureaborted case spinal procedure open case remove instrumentation place screws screws planned lateral imaging spin surgical arm sent right screw placed accurately arm sent left trajectory difficult reach soft tissue pressure scre able placed attempting place screws dilator dropped passed surgeon longer sterile final screws placed stealth navigation aborted weight causing soft tissue pressure dilator dropped procedure delayed hour iud new mazor x | NAV5000 | NAV1012 | NAV4095 | NAV4101 | NAV4105 | NAV5000NAV1012NAV4095NAV4101NAV4105 | True
5073 | 703594218 | X - Deviation issues | 2020-02-03T20:12:59.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation issues 2018-03-30 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were lateral deviations at right T4 - T9. Case was aborted and there was no patient harm. MAZOR X SYSTEM TPL0059 | 2018-03-30 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were lateral deviations at right T4 - T9. Case was aborted and there was no patient harm. | x deviation issues ccr spinal procedure lateral deviations right case aborted mazor x | NAV4091 | NAV4112 | NAV5000 | NAV4095 | NAV3242 | NAV4091NAV4112NAV5000NAV4095NAV3242 | True
3595 | 703383142 | XSE - BMB Stripped | 2019-09-26T16:18:42.000000+0000 | MazorX Stealth Edition | NAV3298 | PLATFORM ISSUES | 20 | MAS2188-02 | BM BRIDGE | XSE - BMB Stripped 2019-09-25 SCL1902914, email, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the bone mount bridge was stripped where the screw was used to tighten down on the Schanz bond. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. BM BRIDGE | 2019-09-25 SCL1902914, email, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the bone mount bridge was stripped where the screw was used to tighten down on the Schanz bond. There was no patient harm and the procedure was not delayed more than an hour. The case was completed. | xse bmb stripped email spinal procedure bone mount bridge stripped screw tighten schanz bond procedure delayed hour case completed bm bridge | NAV3299 | NAV4090 | NAV3295 | NAV3298 | NAV4100 | NAV3299NAV4090NAV3295NAV3298NAV4100 | True
1049 | 703186286 | X - Surgical arm/monitor issues | 2019-06-03T18:56:48.000000+0000 | MazorX | NAV4100 | STRIPPED THREADS | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical arm/monitor issues Prior to beginning of the case I noticed slight bit of "fog" underneath the workstation screen. Original spot disappeared but then two more larger foggy spots appeared and were present during entire case. No issue with work station functionality was noted. During second case, the holes where schanz arms attach were stripped and tech was unable to attach schanz arm properly. Replacement may be needed SURGICAL ARM | Prior to beginning of the case I noticed slight bit of "fog" underneath the workstation screen. Original spot disappeared but then two more larger foggy spots appeared and were present during entire case. No issue with work station functionality was noted. During second case, the holes where schanz arms attach were stripped and tech was unable to attach schanz arm properly. Replacement may be needed | x surgical armmonitor issues prior beginning case noticed slight bit fog underneath workstation screen original spot disappeared larger foggy spots appeared present entire case issue work station functionality noted second case holes schanz arms attach stripped tech unable attach schanz arm properly replacement needed surgical arm | NAV3261 | NAV2054 | NAV3295 | NAV4100 | NAV4101 | NAV3261NAV2054NAV3295NAV4100NAV4101 | True
3674 | 703413852 | XSE - Surgical Arm Failed Stress Test | 2019-10-15T15:15:07.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Surgical Arm Failed Stress Test 2019-04-22 SCL1900752 (REP): Medtronic received information regarding a guidance system, outside of procedure. It was reported that upon Stealth upgrade to the system, the arm failed the stress test. The arm was replaced to completed to upgrade. No patient was present. SURGICAL ARM | 2019-04-22 SCL1900752 (REP): Medtronic received information regarding a guidance system, outside of procedure. It was reported that upon Stealth upgrade to the system, the arm failed the stress test. The arm was replaced to completed to upgrade. No patient was present. | xse surgical arm failed stress test outside procedure stealth upgrade arm failed stress test arm replaced completed upgrade present surgical arm | NAV3249 | NAV4092 | NAV4098 | NAV2054 | NAV4095 | NAV3249NAV4092NAV4098NAV2054NAV4095 | True
197 | 703037030 | X - error message/shoulder shift | 2019-03-05T01:57:52.000000+0000 | MazorX | NAV4108 | SHOULDER SHIFTS | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - error message/shoulder shift 2019-Feb-27 SCL1900476 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they registered the first time very quickly and then received an error 4072 shoulder shift detected when trying to send to a trajectory. Then when attempting to register the second time, the C Arm hit the surgical arm when moving to the OBL position and they received the same error again. The third time they did not receive the shoulder shift error but registration took a very long time and ultimately gave the rep a YELLOW for L5 and S1 and a RED for L3 and L4. The surgical arm was sent and pedicles were drilled at S1 on the left and right and wires were placed. The arm was cleared up and X Ray was taken to confirm placement and wires looked perfect. After the screws were placed a lateral X Ray was taken to confirm and screws looked perfect. The rep then moved the arm to clear left so that an X Ray could be moved to an AP. Nothing was near the arm and they received the same error 4072 front shoulder shift detected. AP image was taken and screws confirmed in perfect position. After closing, the rep went to remove the surgical arm from the bed and put the robot into storage state. The rep then tried to engage the positioner to the robot but was unable the first time. When they went to try it again, the robot suddenly detached from the bed without the positioner being connected. They had to manually remove the robot from the bed as it was leaning forward and laying on the Mayo stand. The rep laid it on the floor and asked the sales rep and HCP to pick it up so they could re attach the positioner. Once the positioner was attached, they proceeded to put the robot away into the housing. This was the hospitals very first case using the Mazor X system. It was noted that the case was delayed approximately 45 minutes and there was no harm. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X SHOULDER | 2019-Feb-27 SCL1900476 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they registered the first time very quickly and then received an error 4072 shoulder shift detected when trying to send to a trajectory. Then when attempting to register the second time, the C Arm hit the surgical arm when moving to the OBL position and they received the same error again. The third time they did not receive the shoulder shift error but registration took a very long time and ultimately gave the rep a YELLOW for L5 and S1 and a RED for L3 and L4. The surgical arm was sent and pedicles were drilled at S1 on the left and right and wires were placed. The arm was cleared up and X Ray was taken to confirm placement and wires looked perfect. After the screws were placed a lateral X Ray was taken to confirm and screws looked perfect. The rep then moved the arm to clear left so that an X Ray could be moved to an AP. Nothing was near the arm and they received the same error 4072 front shoulder shift detected. AP image was taken and screws confirmed in perfect position. After closing, the rep went to remove the surgical arm from the bed and put the robot into storage state. The rep then tried to engage the positioner to the robot but was unable the first time. When they went to try it again, the robot suddenly detached from the bed without the positioner being connected. They had to manually remove the robot from the bed as it was leaning forward and laying on the Mayo stand. The rep laid it on the floor and asked the sales rep and HCP to pick it up so they could re attach the positioner. Once the positioner was attached, they proceeded to put the robot away into the housing. This was the hospitals very first case using the Mazor X system. It was noted that the case was delayed approximately 45 minutes and there was no harm. There were no additional complications reported or anticipated as a result of the event. | x error messageshoulder shift spinal procedure registered time quickly error shoulder shift detected trying send trajectory attempting register second time c arm hit surgical arm moving obl position error time did receive shoulder shift error registration took long time ultimately gave yellow red surgical arm sent pedicles drilled left right wires placed arm cleared x ray taken confirm placement wires looked perfect screws placed lateral x ray taken confirm screws looked perfect moved arm clear left x ray moved ap near arm error shoulder shift detected ap image taken screws confirmed perfect position closing went remove surgical arm bed robot storage state tried engage positioner robot unable time went try robot suddenly detached bed positioner connected manually remove robot bed leaning forward laying mayo stand laid floor asked sales hcp pick attach positioner positioner attached proceeded robot away housing hospitals case using mazor x noted case delayed approximately minutes additional complications anticipated unk mazor x shoulder | NAV4108 | NAV4098 | NAV3250 | NAV2050 | NAV4107 | NAV4108NAV4098NAV3250NAV2050NAV4107 | True
9885 | 704742547 | XSE - System Unresponsive | 2021-12-01T19:09:10.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - System Unresponsive 2021-Dec-01 Salesforce C00238987 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system is not operational and unresponsive. The manufacturer representative was unable to get the positioner into the storage position and was also unable to get the manipulator to move at all. He tried multiple reboots and and Crt Alt O to override multiple times and it did not resolve the issue. The manufacturer representative had to pull the surgical system off the bed not connected to the manipulator. No patient was present. 2021-12-06 IUD x2 (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Dec-01 Salesforce C00238987 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system is not operational and unresponsive. The manufacturer representative was unable to get the positioner into the storage position and was also unable to get the manipulator to move at all. He tried multiple reboots and and Crt Alt O to override multiple times and it did not resolve the issue. The manufacturer representative had to pull the surgical system off the bed not connected to the manipulator. No patient was present. 2021-12-06 IUD x2 (Other): No new information. | xse unresponsive salesforce outside procedure operational unresponsive unable positioner storage position unable manipulator tried multiple reboots crt alt o override multiple times did resolve issue pull surgical bed connected manipulator present iud new mazorx sw installation kit | NAV3248 | NAV3260 | NAV3253 | NAV3246 | NAV3250 | NAV3248NAV3260NAV3253NAV3246NAV3250 | True
1748 | 703213047 | X- Deviation/Accessory Kit Issues | 2019-06-18T14:59:33.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059-04 | MAZOR X SYSTEM (REV 4) | X- Deviation/Accessory Kit Issues The long cannulas from both trays are bent and need to be replaced. I was sent the newer black stylus for one tray but there are two trays. I need another black stylus. Lastly, both reference markers are the original model and are very hard for the x eye to pick up. We have redesigned them, I need two of the latest reference markers. This has been a big issue for the surgeons. 2018-01-02 CCR (REP, HCP): Medtronic received additional information regarding a guidance system being used during a spinal procedure. It was reported that the case was for two vertebrae on level L4 – L5. 3-point test was successful. Shoulder values were within range and 3Define was successful. During registration Drawspine was successful with manipulation of lighting to get system to see reference marker. Segmentation was successful on first try. Wires were inaccurate at L4, L5 left and L4 right. The representative went into planning, but surgeon felt the plan was fine. Trajectories were resent. The surgeon held the bottom of the cannula when maleting the drill guide to prevent skiving. All wires were accurate and verified with fluoro. The representative felt there was something wrong with the instruments as the facets were flat and skiving wasn't an issue. The representative and the decontamination rep went over the instruments and found the long cannula to be slightly bent. This was determined by putting the drill guide into the cannula and feeling resistance which was not got with the short tools. Representative spoke to the surgeon and he agreed. No patient harm was reported. MAZOR X SYSTEM (REV 4) | The long cannulas from both trays are bent and need to be replaced. I was sent the newer black stylus for one tray but there are two trays. I need another black stylus. Lastly, both reference markers are the original model and are very hard for the x eye to pick up. We have redesigned them, I need two of the latest reference markers. This has been a big issue for the surgeons. 2018-01-02 CCR (REP, HCP): Medtronic received additional information regarding a guidance system being used during a spinal procedure. It was reported that the case was for two vertebrae on level L4 – L5. 3-point test was successful. Shoulder values were within range and 3Define was successful. During registration Drawspine was successful with manipulation of lighting to get system to see reference marker. Segmentation was successful on first try. Wires were inaccurate at L4, L5 left and L4 right. The representative went into planning, but surgeon felt the plan was fine. Trajectories were resent. The surgeon held the bottom of the cannula when maleting the drill guide to prevent skiving. All wires were accurate and verified with fluoro. The representative felt there was something wrong with the instruments as the facets were flat and skiving wasn't an issue. The representative and the decontamination rep went over the instruments and found the long cannula to be slightly bent. This was determined by putting the drill guide into the cannula and feeling resistance which was not got with the short tools. Representative spoke to the surgeon and he agreed. No patient harm was reported. | x deviationaccessory kit issues long cannulas trays bent need replaced sent newer black stylus tray trays need black stylus lastly reference markers original model hard x eye pick redesigned need latest reference markers big issue surgeons ccr additional spinal procedure case vertebrae level – test successful shoulder values range successful registration drawspine successful manipulation lighting reference marker segmentation successful try wires inaccurate left right went planning surgeon felt plan fine trajectories resent surgeon held cannula maleting drill guide prevent skiving wires accurate verified fluoro felt wrong instruments facets flat skiving wasnt issue decontamination went instruments long cannula slightly bent determined putting drill guide cannula feeling resistance got short tools spoke surgeon agreed mazor x | NAV2044 | NAV3242 | NAV4091 | NAV3299 | NAV5000 | NAV2044NAV3242NAV4091NAV3299NAV5000 | True
3312 | 703310818 | XSE - NavLock Binding/Deviation | 2019-08-14T15:07:22.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059-06 | MAZOR X SYSTEM | XSE - NavLock Binding/Deviation 2019-08-14 SCL1902383, Intake Form, PDF, C00193354, ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the black and blue NavLocks bound to the drill bit. While the NavLocks were binding a screw was misplaced at S2AI left. The surgeon's plan for the S2AI screw had a high chance of skive and the representative had expressed concern, but the surgeon choose to place the screw. The case was aborted. There was no patient harm, but the case was delayed over an hour. Information unrelated to this event was included in other PEs. 2019-08-27 email 2019-08-27 21:15 (Rep): Additional information received from a manufacturer representative reported that the procedure was delayed one hour due to the guidance system being aborted. Troubleshooting included placing the NavLocks on other drill bits to eliminate outliers. The screws were deviated less than 3.5 mm. The case was completed freehand and all screws were successfully placed. There was no patient harm. 2019-09-11 e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported a post operative scan was not done. The screws were found to be deviated via x-ray. The representative indicated that drilling trajectories had been difficult in the majority of cases, but was enhanced when trying to do S2AI trajectories and crossing the joing. Even trying to tap across the joint rarely worked unless excessive pressure is placed down the vector. MAZOR X SYSTEM | 2019-08-14 SCL1902383, Intake Form, PDF, C00193354, ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the black and blue NavLocks bound to the drill bit. While the NavLocks were binding a screw was misplaced at S2AI left. The surgeon's plan for the S2AI screw had a high chance of skive and the representative had expressed concern, but the surgeon choose to place the screw. The case was aborted. There was no patient harm, but the case was delayed over an hour. Information unrelated to this event was included in other PEs. 2019-08-27 email 2019-08-27 21:15 (Rep): Additional information received from a manufacturer representative reported that the procedure was delayed one hour due to the guidance system being aborted. Troubleshooting included placing the NavLocks on other drill bits to eliminate outliers. The screws were deviated less than 3.5 mm. The case was completed freehand and all screws were successfully placed. There was no patient harm. 2019-09-11 e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported a post operative scan was not done. The screws were found to be deviated via x-ray. The representative indicated that drilling trajectories had been difficult in the majority of cases, but was enhanced when trying to do S2AI trajectories and crossing the joing. Even trying to tap across the joint rarely worked unless excessive pressure is placed down the vector. | xse navlock bindingdeviation pdf id spinal procedure black blue navlocks bound drill bit navlocks binding screw misplaced left surgeons plan screw high chance skive expressed concern surgeon choose place screw case aborted case delayed hour unrelated included pes email additional procedure delayed hour aborted troubleshooting included placing navlocks drill bits eliminate outliers screws deviated mm case completed freehand screws successfully placed additional post operative scan screws deviated xray indicated drilling trajectories difficult majority cases enhanced trying trajectories crossing joing trying tap joint rarely worked unless excessive pressure placed vector mazor x | NAV5000 | NAV4112 | NAV1012 | NAV3264 | NAV3243 | NAV5000NAV4112NAV1012NAV3264NAV3243 | True
6280 | 703903011 | XSE - Screens Going Black/Froze System | 2020-08-13T19:31:24.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 20 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - Screens Going Black/Froze System 2020-08-12 00694563 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system was booted up with nothing plugged in. When the HDMI was plugged in, the screen did not flicker, but the picture did come up. The manufacturer representative noted that everything was working fine for the first spin of the Scan & Plan case. After a second spin, the system froze and the surgeon and monitor screens went black. The screen on the boom monitor was frozen. The representative tried pushing Ctrl-Alt-R, but nothing would respond. A hard restart of the system was done and the case was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-14 email 2020-08-14 09:01, IUD x3 (Rep): Additional information received from a manufacturer representative reported that the cause of the screens going black and the system freezing was not determined. 2020-08-21 IUD x4 (Other): No new information. 2020-09-03 IUD x3 (Other): No new information. 2020-10-05 IUD (Other): No new information. 2020-10-30 IUD (Other): No new information. 2021-06-20 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2020-08-12 00694563 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the guidance system was booted up with nothing plugged in. When the HDMI was plugged in, the screen did not flicker, but the picture did come up. The manufacturer representative noted that everything was working fine for the first spin of the Scan & Plan case. After a second spin, the system froze and the surgeon and monitor screens went black. The screen on the boom monitor was frozen. The representative tried pushing Ctrl-Alt-R, but nothing would respond. A hard restart of the system was done and the case was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-14 email 2020-08-14 09:01, IUD x3 (Rep): Additional information received from a manufacturer representative reported that the cause of the screens going black and the system freezing was not determined. 2020-08-21 IUD x4 (Other): No new information. 2020-09-03 IUD x3 (Other): No new information. 2020-10-05 IUD (Other): No new information. 2020-10-30 IUD (Other): No new information. 2021-06-20 IUD (Other): No new information. | xse screens going blackfroze spinal procedure booted plugged hdmi plugged screen did flicker picture did come noted working fine spin scan plan case second spin froze surgeon monitor screens went black screen boom monitor frozen tried pushing ctrlaltr respond hard restart case able completed procedure delayed hour email iud additional cause screens going black freezing determined iud new iud new iud new iud new iud new mazor x pc station refurbished | NAV3261 | NAV3248 | NAV5000 | NAV3262 | NAV3253 | NAV3261NAV3248NAV5000NAV3262NAV3253 | True
1823 | 703218350 | X- Tablet Issue | 2019-06-20T17:21:03.000000+0000 | MazorX | NAV3253 | TABLET ISSUES | 20 | UNK_MAZORX_TBLT | UNK MAZOR X TABLET | X- Tablet Issue -X-eye failure resulting in several soft restarts. Could probably be from bovie use. -The tablet has become very difficult to use. The scrub tech and the surgeon have a difficult time using the tablet because it will not register finger taps. -The drape ripped as it moved from trajectory to trajectory. Pictures of the rip attached -Repeated shoulder shifts -Squeaking arm during movement -Significant crossview errors for no discernible reason. -There is a major delay of new images being updated to the tablet after they are taken from the c-arm. -Attached is an extremely short clip of the sound that the drawspine is making. It is a different sound than what is typically heard. It is in the very first millisecond of the video. This sound is what is heard throughout the drawspine process. I apologize for not having a better video. UNK MAZOR X TABLET | -X-eye failure resulting in several soft restarts. Could probably be from bovie use. -The tablet has become very difficult to use. The scrub tech and the surgeon have a difficult time using the tablet because it will not register finger taps. -The drape ripped as it moved from trajectory to trajectory. Pictures of the rip attached -Repeated shoulder shifts -Squeaking arm during movement -Significant crossview errors for no discernible reason. -There is a major delay of new images being updated to the tablet after they are taken from the c-arm. -Attached is an extremely short clip of the sound that the drawspine is making. It is a different sound than what is typically heard. It is in the very first millisecond of the video. This sound is what is heard throughout the drawspine process. I apologize for not having a better video. | x tablet issue xeye failure resulting soft restarts probably bovie use tablet difficult use scrub tech surgeon difficult time using tablet register finger taps drape ripped moved trajectory trajectory pictures rip attached repeated shoulder shifts squeaking arm movement significant crossview errors discernible reason major delay new images updated tablet taken carm attached extremely short clip sound drawspine making different sound typically heard millisecond video sound heard drawspine process apologize having better video unk mazor x tablet | NAV3253 | NAV4108 | NAV3249 | NAV3246 | NAV3262 | NAV3253NAV4108NAV3249NAV3246NAV3262 | True
9078 | 704547659 | XSE - Software Unresponsive | 2021-08-09T21:37:33.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 40 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - Software Unresponsive 2021-Aug-09 Salesforce C00231912, IUD x4 (rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the software went unresponsive and was not able to track navigation instruments. No patient was present. 2021-08-11 IUD x4, email 2021-08-13 08:46 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined, but replacing the PC resolved all of the issues. 2021-08-26 IUD (Other): No new information. 2021-09-03 IUD x3 (Other): No new information. 2021-09-22 IUD x3 (Other): No new information. 2021-12-06 IUD (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2021-Aug-09 Salesforce C00231912, IUD x4 (rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the software went unresponsive and was not able to track navigation instruments. No patient was present. 2021-08-11 IUD x4, email 2021-08-13 08:46 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined, but replacing the PC resolved all of the issues. 2021-08-26 IUD (Other): No new information. 2021-09-03 IUD x3 (Other): No new information. 2021-09-22 IUD x3 (Other): No new information. 2021-12-06 IUD (Other): No new information. | xse software unresponsive salesforce iud outside p rocedure software went unresponsive able track navigation instruments present iud email additional cause determined replacing pc resolved issues iud new iud new iud new iud new mazor x pc station | NAV3262 | NAV3248 | NAV2047 | NAV3246 | NAV3256 | NAV3262NAV3248NAV2047NAV3246NAV3256 | True
8093 | 704349364 | XSE - Pneumatic Failure Message | 2021-04-20T00:17:20.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Pneumatic Failure Message 2021-Apr-19 Salesforce C00224933, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that when the guidance system was powered on after being brought into the room, the screens went black when attempting to plan. After the system was recovered, an error message stating failure to build pneumatic pressure, maintenance is required was displayed twice. The surgeon decided to use a different guidance system for the case. There was no patient involvement. 2021-04-23 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Apr-19 Salesforce C00224933, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that when the guidance system was powered on after being brought into the room, the screens went black when attempting to plan. After the system was recovered, an error message stating failure to build pneumatic pressure, maintenance is required was displayed twice. The surgeon decided to use a different guidance system for the case. There was no patient involvement. 2021-04-23 IUD (Other): No new information. | xse pneumatic failure message salesforce iud outside proc edure powered brought room screens went black attempting plan recovered error message stating failure build pneumatic pressure maintenance required displayed twice surgeon decided use different case iud new unk mazor x workstation | NAV3264 | NAV3261 | NAV4095 | NAV2047 | NAV3244 | NAV3264NAV3261NAV4095NAV2047NAV3244 | True
3938 | 703470879 | XSE - Tablet screen is split | 2019-11-16T00:10:02.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0337-01 | SURGEON&APOS;S SCREEN | XSE - Tablet screen is split 2019-11-15 0678500 (Rep): Medtronic received information regarding a guidance system. It was reported that the surgeon's wireless tablet screen image was split it two; both images were the same. No patient was involved. 2019-11-21 IUD x5, Email 2019-12-20 5:31 x4 (Rep): Additional information was received from analysis: Visual inspection of the tablet confirmed DC to DC converter abnormally overheated. Also, there was no data because of a wire conductivity problem. 2020-02-26 IUD (Other): No new information. SURGEON&APOS;S SCREEN | 2019-11-15 0678500 (Rep): Medtronic received information regarding a guidance system. It was reported that the surgeon's wireless tablet screen image was split it two; both images were the same. No patient was involved. 2019-11-21 IUD x5, Email 2019-12-20 5:31 x4 (Rep): Additional information was received from analysis: Visual inspection of the tablet confirmed DC to DC converter abnormally overheated. Also, there was no data because of a wire conductivity problem. 2020-02-26 IUD (Other): No new information. | xse tablet screen split surgeons wireless tablet screen image split images involved iud email additional analysis visual inspection tablet confirmed dc dc converter abnormally overheated data wire conductivity problem iud new surgeonaposs screen | NAV3253 | NAV2054 | NAV1006 | NAV2058 | NAV3246 | NAV3253NAV2054NAV1006NAV2058NAV3246 | True
10208 | 704822061 | XSE - Graphics Card Broken | 2022-01-21T08:54:19.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 30 | ASM0215 | ASM0215 MAZOR X COMPUTER KIT | XSE - Graphics Card Broken 2022-Jan-21 Salesforce C00242023 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the graphics card was broken. Both screens show artifacts when touching and the computer randomly reboots losing all surgery information. No patient was present. 2022-01-31 IUD (Other, For): No new information. 2022-02-08 IUD (Other): No new information. 2022-02-14 IUD (Other): No new information. ASM0215 MAZOR X COMPUTER KIT | 2022-Jan-21 Salesforce C00242023 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the graphics card was broken. Both screens show artifacts when touching and the computer randomly reboots losing all surgery information. No patient was present. 2022-01-31 IUD (Other, For): No new information. 2022-02-08 IUD (Other): No new information. 2022-02-14 IUD (Other): No new information. | xse graphics card broken salesforce outside proc edure graphics card broken screens artifacts touching computer randomly reboots losing surgery present iud new iud new iud new mazor x computer kit | NAV3262 | NAV3261 | NAV2058 | NAV2054 | NAV2047 | NAV3262NAV3261NAV2058NAV2054NAV2047 | True
4249 | 703517920 | X - Registration issues/Deviation | 2019-12-13T15:48:19.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration issues/Deviation 2018-06-21 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was planned prior to surgery for two vertebrae on levels L2 - L3. During the procedure, there were registration difficulties and trajectory deviations. L3 right was deviated. No patient harm was reported. MAZOR X SYSTEM TPL0059 | 2018-06-21 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was planned prior to surgery for two vertebrae on levels L2 - L3. During the procedure, there were registration difficulties and trajectory deviations. L3 right was deviated. No patient harm was reported. | x registration issuesdeviation ccr spinal procedure case planned prior surgery vertebrae levels procedure registration difficulties trajectory deviations right deviated mazor x | NAV4091 | NAV3242 | NAV4112 | NAV2047 | NAV3261 | NAV4091NAV3242NAV4112NAV2047NAV3261 | True
1011 | 703183891 | X - Workstation Issue | 2019-05-31T19:52:34.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-01 | RNX-WS-1040 | X - Workstation Issue Britteni Rice reported via phone compressor did not initiate during system start up (system restart performed, compressor initiated). 3Define and X-eye not performed, no reason given. Error 'arm unable to reach trajectory' given during AP/OBL. Estop was pressed and arm physically moved to achieve AP/OBL images. Trajectories sent. Reported during case SW gave message 'shoulder unlocked'. Operator unlocked (mount lock) through SW and was required to take new images. RNX-WS-1040 | Britteni Rice reported via phone compressor did not initiate during system start up (system restart performed, compressor initiated). 3Define and X-eye not performed, no reason given. Error 'arm unable to reach trajectory' given during AP/OBL. Estop was pressed and arm physically moved to achieve AP/OBL images. Trajectories sent. Reported during case SW gave message 'shoulder unlocked'. Operator unlocked (mount lock) through SW and was required to take new images. | x workstation issue britteni rice phone compressor did initiate start xeye performed reason given error arm unable reach trajectory given apobl estop pressed arm physically moved achieve apobl images trajectories sent case sw gave message shoulder unlocked operator unlocked sw required new images | NAV4095 | NAV2047 | NAV3246 | NAV3264 | NAV4101 | NAV4095NAV2047NAV3246NAV3264NAV4101 | True
6197 | 703887196 | X - Missing Schanz Bond Screw | 2020-08-04T23:24:58.000000+0000 | MazorX | NAV2053 | MISSING PARTS | 20 | MEC4530-02S | KIT SVC MEC4530-02S BALL ARM TRGT SCREW | X - Missing Schanz Bond Screw 2020-08-04 00694004 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site lost the Schanz bond screws. There was no patient involvement. 2020-08-05 IUD x3 (Other): No new information. KIT SVC MEC4530-02S BALL ARM TRGT SCREW | 2020-08-04 00694004 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site lost the Schanz bond screws. There was no patient involvement. 2020-08-05 IUD x3 (Other): No new information. | x missing schanz bond screw outside procedure site lost schanz bond screws iud new kit svc ball arm trgt screw | NAV2053 | NAV3299 | NAV3295 | NAV4090 | NAV3249 | NAV2053NAV3299NAV3295NAV4090NAV3249 | True
4037 | 703493862 | XSE - Surgeon Screen Issues | 2019-12-01T11:18:27.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357 | KIT SVC ASM0357 SRGN SCRN W/ARM 02S NEW | XSE - Surgeon Screen Issues 2019-11-27 00679296 (HCP, REP, FOR): Medtronic received information regarding a guidance system being used pre-operatively of a spinal procedure. It was reported that after starting the system, the surgeon screen touch was not working and then the site needed to calibrate. After every restart of the system, calibration of the touch screen was needed again. There was no reported delay to the procedure and no impact on the patient outcome. 2020-02-25 Email (Rep, For): Additional information was received stating that after a restart and performing touchscreen calibration the issue was resolved. 2021-05-09 IUD (Other, For): No new information. KIT SVC ASM0357 SRGN SCRN W/ARM 02S NEW | 2019-11-27 00679296 (HCP, REP, FOR): Medtronic received information regarding a guidance system being used pre-operatively of a spinal procedure. It was reported that after starting the system, the surgeon screen touch was not working and then the site needed to calibrate. After every restart of the system, calibration of the touch screen was needed again. There was no reported delay to the procedure and no impact on the patient outcome. 2020-02-25 Email (Rep, For): Additional information was received stating that after a restart and performing touchscreen calibration the issue was resolved. 2021-05-09 IUD (Other, For): No new information. | xse surgeon screen issues preoperatively spinal procedure starting surgeon screen touch working site needed calibrate restart calibration touch screen needed delay procedure impact outcome email additional stating restart performing touchscreen calibration issue resolved iud new kit svc srgn scrn warm new | NAV3253 | NAV3244 | NAV3249 | NAV2047 | NAV5000 | NAV3253NAV3244NAV3249NAV2047NAV5000 | True
5387 | 703671082 | XSE - Damaged Schanz Arms | 2020-03-17T18:21:11.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Damaged Schanz Arms 2020-02-06 00683584 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that when inspecting their trays, they found that four schanz arms were stripped. They have 6 cases to complete this week. They have 8 total arms including the damaged ones, and they use two arms per surgery for stability. No patient was present. 2020-03-17 IUD x3 (Other): No new information. 2020-03-31 IUD x4 (Other): No new information. 2020-04-19 IUD x2 (Other): No new information. KIT SVC MAS2351-04S SCHANZ ARM | 2020-02-06 00683584 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that when inspecting their trays, they found that four schanz arms were stripped. They have 6 cases to complete this week. They have 8 total arms including the damaged ones, and they use two arms per surgery for stability. No patient was present. 2020-03-17 IUD x3 (Other): No new information. 2020-03-31 IUD x4 (Other): No new information. 2020-04-19 IUD x2 (Other): No new information. | xse damaged schanz arms outside procedure inspecting trays schanz arms stripped cases complete week total arms including damaged ones use arms surgery stability present iud new iud new iud new kit svc schanz arm | NAV3296 | NAV3299 | NAV2053 | NAV3295 | NAV4093 | NAV3296NAV3299NAV2053NAV3295NAV4093 | True
3141 | 703266425 | XSE - arm shift | 2019-07-18T18:40:30.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | XSE - arm shift 2019-Jul-16 SCL1902078 intake form (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the arm received a Joint 1 shift. The rep reset movements and performed a hard restart but error was not resolved. There was no patient involved. There were no additional complications reported or anticipated as a result of the event. 2019-Jul-22 SCL1902078 (rep): Additional information received from the rep indicated that they ran an advanced accuracy test and the arm passed all tests. SURGICAL ARM | 2019-Jul-16 SCL1902078 intake form (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the arm received a Joint 1 shift. The rep reset movements and performed a hard restart but error was not resolved. There was no patient involved. There were no additional complications reported or anticipated as a result of the event. 2019-Jul-22 SCL1902078 (rep): Additional information received from the rep indicated that they ran an advanced accuracy test and the arm passed all tests. | xse arm shift outside procedure arm joint shift reset movements performed hard restart error resolved involved additional complications anticipated additional indicated ran advanced accuracy test arm passed tests surgical arm | NAV4098 | NAV4108 | NAV2054 | NAV4093 | NAV4101 | NAV4098NAV4108NAV2054NAV4093NAV4101 | True
7828 | 704288190 | XSE - Deviated Trajectories | 2021-03-17T16:00:50.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Trajectories 2021-Mar-15 Salesforce C00222793 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that they placed four screws and registered at L3-S1. They placed the left screws left side first and then the right side. The first two screws were lateral to plan. They looked like skives and it was an MIS case. The screws on the right were perfect. They believe it was surgeon technique when placing the screws. There was no patient harm and no delay to the procedure. 2021-03-29 IUD (Other): No new information. 2021-05-19 email 2021-05-19 10:02 (Rep): Additional information received from a manufacturer representative reported the screws were deviated less than 3.5 mm. The screw were repositioned freehand and there were no patient symptoms. MAZOR X SYSTEM TPL0059 | 2021-Mar-15 Salesforce C00222793 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that they placed four screws and registered at L3-S1. They placed the left screws left side first and then the right side. The first two screws were lateral to plan. They looked like skives and it was an MIS case. The screws on the right were perfect. They believe it was surgeon technique when placing the screws. There was no patient harm and no delay to the procedure. 2021-03-29 IUD (Other): No new information. 2021-05-19 email 2021-05-19 10:02 (Rep): Additional information received from a manufacturer representative reported the screws were deviated less than 3.5 mm. The screw were repositioned freehand and there were no patient symptoms. | xse deviated trajectories salesforce spinal p rocedure placed screws registered placed left screws left right screws lateral plan looked like skives mis case screws right perfect believe surgeon technique placing screws delay procedure iud new email additional screws deviated mm screw repositioned freehand symptoms mazor x | NAV5000 | NAV1012 | NAV4112 | NAV2054 | NAV3243 | NAV5000NAV1012NAV4112NAV2054NAV3243 | True
9784 | 704719123 | XSE - Time Out Error/3Define Issues | 2021-11-17T00:10:48.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 40 | KIT1211-01 | SVC KIT1211-01 CABLE USB 3.0 HIGH F | XSE - Time Out Error/3Define Issues 2021-Nov-16 Salesforce C00238258 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a time out error when using the 3Define camera. There was no patient involvement. 2021-11-19 IUD (Other): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-06 IUD x2 (Other): No new information. 2021-12-09 IUD (Rep): Additional information received from a manufacturer representative reported that intermittent issues still per sisted. 2021-12-16 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. 2022-01-18 IUD (Other): No new information. 2022-02-01 IUD (Other): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-14 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-07 IUD (Other): No new information. 2022-04-05 IUD (Other): No new information. SVC KIT1211-01 CABLE USB 3.0 HIGH F | 2021-Nov-16 Salesforce C00238258 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a time out error when using the 3Define camera. There was no patient involvement. 2021-11-19 IUD (Other): No new information. 2021-11-29 IUD (Other): No new information. 2021-12-06 IUD x2 (Other): No new information. 2021-12-09 IUD (Rep): Additional information received from a manufacturer representative reported that intermittent issues still per sisted. 2021-12-16 IUD x2 (Other): No new information. 2022-01-05 IUD (Other): No new information. 2022-01-18 IUD (Other): No new information. 2022-02-01 IUD (Other): No new information. 2022-02-09 IUD (Other): No new information. 2022-02-14 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-07 IUD (Other): No new information. 2022-04-05 IUD (Other): No new information. | xse time issues salesforce outside procedure time error using camera iud new iud new iud new iud additional intermittent issues sisted iud new iud new iud new iud new iud new iud new iud new iud new iud new svc cable usb high f | NAV2047 | NAV3244 | NAV2058 | NAV3249 | NAV1006 | NAV2047NAV3244NAV2058NAV3249NAV1006 | True
2092 | 703225936 | X - DEVIATION | 2019-06-25T14:58:22.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X - DEVIATION Right Trajectories were lateral. Both right trajectories were lateral. SA resent back to right side, still not to plan. Surgeon re-positioned both right trajectories under navigation./Mazor X CCR Dr. Uschold –4th Case L3-L4 MIS OLIF. Asante Rogue Regional Medical Center –Medford, OR (03-August-2017) Arm #64 System #58 8th case at Asante Rogue. Planning -Performed and reviewed before case by surgeon and XCSR. Axial angles brought in to have skin incision to be only 4cm off midline. Setup -Shoulder calibration, X-eyes, and 3Define cameras checked at start of case. SA setup and advanced setup performed. 15:20 -Arm attached to bed with 1 person in less than 1 minute. PA and scrub tech draped arm successfully. O-Arm spin. Mount -15:38 Right PSIS Schanz pin. Surgeon used navigation to insert Schanz pin successfully on 1st attempt with good bone purchase. Bone Mount Bridge positioned at 15 degrees. Registration -15:41 -3Define was clear. Stylus not found in software after making attempts at altering the lighting. Generic AP and OBL used with minor adjustment. 1 AP and 1 OBL needed for fluoro acquisition. Basic Algorithm provided successful segmentation labeling L3. All segm entation lines were in disc space. Approved by surgeon on all operative levels. L5 not registered but no trajectories were at that level. L3 Left Red screw. Due to Medtronic fiducial, screw adjusted in lateral view 1 degree and was reachable. Operation -All trajectories sent to mark skin. Surgeon pleased with location. Clear arm up. L3 Right sent, tools docked and drilled. K-wire placed. L4 Right sent, tools docked and drilled. K-wire placed. L3 Left sent, tools docked and drilled. K-wire placed. L4 Left sent, tools docked and drilled. K-wire placed. Clear arm right. AP fluoro showed left wires to be accurate but right side looked lateral. Arm resent to L3 Right and L4 Right to be executed again due to possible skiving. AP Fluoro taken and surgeon still wasn't pleased with location so he free-handed the right side under navigation. MAZOR X SOFTWARE KIT VER-2.3 | Right Trajectories were lateral. Both right trajectories were lateral. SA resent back to right side, still not to plan. Surgeon re-positioned both right trajectories under navigation./Mazor X CCR Dr. Uschold –4th Case L3-L4 MIS OLIF. Asante Rogue Regional Medical Center –Medford, OR (03-August-2017) Arm #64 System #58 8th case at Asante Rogue. Planning -Performed and reviewed before case by surgeon and XCSR. Axial angles brought in to have skin incision to be only 4cm off midline. Setup -Shoulder calibration, X-eyes, and 3Define cameras checked at start of case. SA setup and advanced setup performed. 15:20 -Arm attached to bed with 1 person in less than 1 minute. PA and scrub tech draped arm successfully. O-Arm spin. Mount -15:38 Right PSIS Schanz pin. Surgeon used navigation to insert Schanz pin successfully on 1st attempt with good bone purchase. Bone Mount Bridge positioned at 15 degrees. Registration -15:41 -3Define was clear. Stylus not found in software after making attempts at altering the lighting. Generic AP and OBL used with minor adjustment. 1 AP and 1 OBL needed for fluoro acquisition. Basic Algorithm provided successful segmentation labeling L3. All segm entation lines were in disc space. Approved by surgeon on all operative levels. L5 not registered but no trajectories were at that level. L3 Left Red screw. Due to Medtronic fiducial, screw adjusted in lateral view 1 degree and was reachable. Operation -All trajectories sent to mark skin. Surgeon pleased with location. Clear arm up. L3 Right sent, tools docked and drilled. K-wire placed. L4 Right sent, tools docked and drilled. K-wire placed. L3 Left sent, tools docked and drilled. K-wire placed. L4 Left sent, tools docked and drilled. K-wire placed. Clear arm right. AP fluoro showed left wires to be accurate but right side looked lateral. Arm resent to L3 Right and L4 Right to be executed again due to possible skiving. AP Fluoro taken and surgeon still wasn't pleased with location so he free-handed the right side under navigation. | x deviation right trajectories lateral right trajectories lateral sa resent right plan surgeon repositioned right trajectories navigationmazor x ccr dr uschold – case mis olif asante rogue regional medical center –medford arm case asante rogue planning performed reviewed case surgeon xcsr axial angles brought skin incision midline setup shoulder calibration xeyes cameras checked start case sa setup advanced setup performed arm attached bed person minute pa scrub tech draped arm successfully oarm spin mount right psis schanz pin surgeon navigation insert schanz pin successfully attempt good bone purchase bone mount bridge positioned degrees registration clear stylus software making attempts altering lighting generic ap obl minor adjustment ap obl needed fluoro acquisition basic algorithm provided successful segmentation labeling segm entation lines disc space approved surgeon operative levels registered trajectories level left red screw fiducial screw adjusted lateral view degree reachable operation trajectories sent mark skin surgeon pleased location clear arm right sent tools docked drilled kwire placed right sent tools docked drilled kwire placed left sent tools docked drilled kwire placed left sent tools docked drilled kwire placed clear arm right ap fluoro showed left wires accurate right looked lateral arm resent right right executed possible skiving ap fluoro taken surgeon wasnt pleased location freehanded right navigation mazor x software kit | NAV4091 | NAV3246 | NAV3242 | NAV3248 | NAV3299 | NAV4091NAV3246NAV3242NAV3248NAV3299 | True
6300 | 703906696 | XSE - Shoulder Mismatch Error | 2020-08-17T13:36:55.000000+0000 | MazorX Stealth Edition | NAV4106 | CALIBRATION ISSUES | 20 | ASM0207-02S | KIT SVC ASM0207-02S SHOULDER | XSE - Shoulder Mismatch Error 2020-08-17 00694676 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a shoulder mismatch error was displayed prior to the case. A shoulder calibration was done and the issue was resolved. There was no patient involvement. KIT SVC ASM0207-02S SHOULDER | 2020-08-17 00694676 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a shoulder mismatch error was displayed prior to the case. A shoulder calibration was done and the issue was resolved. There was no patient involvement. | xse shoulder mismatch error outside procedure shoulder mismatch error displayed prior case shoulder calibration issue resolved kit svc shoulder | NAV4106 | NAV3252 | NAV2054 | NAV4107 | NAV4108 | NAV4106NAV3252NAV2054NAV4107NAV4108 | True
2008 | 703223680 | XSE -no registration/3define not working | 2019-06-24T17:43:39.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE -no registration/3define not working 2019-Jun-19 SCL1901809 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that registration could not be achieved. The rep went through all possible steps but still could not get registration. Also, 3-define was not working. The cable was checked and plugged in at the bottom of Workstation. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jun-19 SCL1901809 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that registration could not be achieved. The rep went through all possible steps but still could not get registration. Also, 3-define was not working. The cable was checked and plugged in at the bottom of Workstation. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse working spinal procedure registration achieved went possible steps registration working cable checked plugged workstation additional complications anticipated mazor x software installation kit | NAV5000 | NAV3246 | NAV3242 | NAV3247 | NAV2047 | NAV5000NAV3246NAV3242NAV3247NAV2047 | False
3563 | 703376611 | XSE - Target Extender/Schanz Bond Issue | 2019-09-24T00:31:09.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2310-03 | TARGET EXTENDER | XSE - Target Extender/Schanz Bond Issue 2019-09-20 mPXR 663391, email (Rep, HCP): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the target extender and Schanz bond were stripped. The threads were gone on the target extender and the square drive on the Schanz bond was stripped. Accessories from another set were used for the procedure. There was no patient involvement. TARGET EXTENDER | 2019-09-20 mPXR 663391, email (Rep, HCP): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the target extender and Schanz bond were stripped. The threads were gone on the target extender and the square drive on the Schanz bond was stripped. Accessories from another set were used for the procedure. There was no patient involvement. | xse target extenderschanz bond issue mpxr email prior procedure target extender schanz bond stripped threads gone target extender square drive schanz bond stripped accessories set procedure target extender | NAV3295 | NAV3299 | NAV4090 | NAV3233 | NAV3298 | NAV3295NAV3299NAV4090NAV3233NAV3298 | True
5279 | 703637189 | XSE - Registration Issues, L2 Deviation | 2020-02-26T16:13:33.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Issues, L2 Deviation 2020-02-24 CCR, 00685201 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a juvenile global deformity with fixation at T2-L3. It was reported that planning for the case was completed by the surgeon prior to the procedure to completed in three segments. The dual clamp was selected to mount the platform to the patient and was used for the entire construct. Due to the patient's kyphosis, the clamp was placed higher than usual at T4/T6 Registration for the first segment was achieved on the first attempt under basic but showed all red values with one yellow. The manufacturer representative returned to planning and adjusted the segmentation lines manually but all red and yellow values remained. New images were taken to clear up the end plates but still achieved the same results. After multiple attempts, the representative adjusted segmentation lines again and was able to get yellow and green values for T2-T6. The surgeon visually verified no shift and proceeded with the operation. Red screws were present at T2 due to unreachability. T2 was re-planned and the surgeon sent to the left trajectory which skiv ed superior. T3 screws and right T2 were not red but instead should "no solutions found." The representative returned to planning to review the screws, then reviewed work volume in the operation screen and saw no issues. Once the representative had click back into operation mode, without changing anything, there were solutions fro those trajectories and T3 trajectories were placed. The second segment of surgery was for T7-T12. Registration was successful with no issue and all trajectories were placed accurately. The third segment of surgery was for L1-L3. Automatic registration failed multiple times. Manual registration yielded all red values. New images were taken with the same result. The C-Arm was recalibrated and more new images were taken but registration was still not successful. The AP image was cleared up to remove K-wires in the image and then automatic registration was successful on basic with all green values. A similar issue occurred with the third segment registration as had occurred in the first segment; during planning it showed "no solutions found" for L2 but when they went back into planning then returned to the operation screen, the problem was resolved. No patient harm was reported and surgery was delayed less than an hour. The representative noted the surgeon expressed frustration regarding the failed registrations and "no solutions" for the screws, however was satisfied with all screw placements which were verified accurate with 2D/3D portable CT spins. 2020-02-26 IUD x3, email 2020-03-03 12:30 (Rep): Additional information received from a manufacturer representative reported that the deviation was less than 3.5 mm and was unacceptable to the surgeon so the placed the screw freehand. MAZOR X SYSTEM TPL0059 | 2020-02-24 CCR, 00685201 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a juvenile global deformity with fixation at T2-L3. It was reported that planning for the case was completed by the surgeon prior to the procedure to completed in three segments. The dual clamp was selected to mount the platform to the patient and was used for the entire construct. Due to the patient's kyphosis, the clamp was placed higher than usual at T4/T6 Registration for the first segment was achieved on the first attempt under basic but showed all red values with one yellow. The manufacturer representative returned to planning and adjusted the segmentation lines manually but all red and yellow values remained. New images were taken to clear up the end plates but still achieved the same results. After multiple attempts, the representative adjusted segmentation lines again and was able to get yellow and green values for T2-T6. The surgeon visually verified no shift and proceeded with the operation. Red screws were present at T2 due to unreachability. T2 was re-planned and the surgeon sent to the left trajectory which skiv ed superior. T3 screws and right T2 were not red but instead should "no solutions found." The representative returned to planning to review the screws, then reviewed work volume in the operation screen and saw no issues. Once the representative had click back into operation mode, without changing anything, there were solutions fro those trajectories and T3 trajectories were placed. The second segment of surgery was for T7-T12. Registration was successful with no issue and all trajectories were placed accurately. The third segment of surgery was for L1-L3. Automatic registration failed multiple times. Manual registration yielded all red values. New images were taken with the same result. The C-Arm was recalibrated and more new images were taken but registration was still not successful. The AP image was cleared up to remove K-wires in the image and then automatic registration was successful on basic with all green values. A similar issue occurred with the third segment registration as had occurred in the first segment; during planning it showed "no solutions found" for L2 but when they went back into planning then returned to the operation screen, the problem was resolved. No patient harm was reported and surgery was delayed less than an hour. The representative noted the surgeon expressed frustration regarding the failed registrations and "no solutions" for the screws, however was satisfied with all screw placements which were verified accurate with 2D/3D portable CT spins. 2020-02-26 IUD x3, email 2020-03-03 12:30 (Rep): Additional information received from a manufacturer representative reported that the deviation was less than 3.5 mm and was unacceptable to the surgeon so the placed the screw freehand. | xse registration issues deviation ccr id spinal procedure treat juvenile global deformity fixation planning case completed surgeon prior procedure completed segments dual clamp selected mount platform entire construct patients kyphosis clamp placed higher usual registration segment achieved attempt basic showed red values yellow returned planning adjusted segmentation lines manually red yellow values remained new images taken clear end plates achieved results multiple attempts adjusted segmentation lines able yellow green values surgeon visually verified shift proceeded operation red screws present unreachability replanned surgeon sent left trajectory skiv ed superior screws right red instead solutions returned planning review screws reviewed work volume operation screen saw issues click operation mode changing solutions fro trajectories trajectories placed second segment surgery registration successful issue trajectories placed accurately segment surgery automatic registration failed multiple times manual registration yielded red values new images taken carm recalibrated new images taken registration successful ap image cleared remove kwires image automatic registration successful basic green values similar issue occurred segment registration occurred segment planning showed solutions went planning returned operation screen problem resolved surgery delayed hour noted surgeon expressed frustration failed registrations solutions screws satisfied screw placements verified accurate portable ct spins iud email additional deviation mm unacceptable surgeon placed screw freehand mazor x | NAV3242 | NAV4112 | NAV5000 | NAV3241 | NAV2051 | NAV3242NAV4112NAV5000NAV3241NAV2051 | True
10082 | 704793918 | XSE ¦ Surgical System Movement | 2022-01-06T00:42:19.000000+0000 | MazorX Stealth Edition | NAV4104 | UNABLE TO MOVE | 20 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE ¦ Surgical System Movement 2022-Jan-05 Salesforce C00240490 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the manufacturer representative exited the operate task and turned mount lock off the system, she was unable to move the shoulder or positioner. She rebooted the software and there was no change. She used the override panel to get the system off the bed. She was unable to restart the system and test further. No patient was present. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2022-Jan-05 Salesforce C00240490 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the manufacturer representative exited the operate task and turned mount lock off the system, she was unable to move the shoulder or positioner. She rebooted the software and there was no change. She used the override panel to get the system off the bed. She was unable to restart the system and test further. No patient was present. | xse surgical movement salesforce outside procedure exited operate task turned mount lock unable shoulder positioner rebooted software change override panel bed unable restart test present kit svc surgical new | NAV3250 | NAV4103 | NAV4104 | NAV3249 | NAV2047 | NAV3250NAV4103NAV4104NAV3249NAV2047 | True
9154 | 704560777 | X - 3D Marker Disassembled | 2021-08-17T09:42:04.000000+0000 | MazorX | NAV2054 | OUT OF BOX FAILURES | 20 | KIT0574 | MAZOR X SPINE DISPOSABLE KIT KIT0574 | X - 3D Marker Disassembled 2021-Aug-15 Salesforce C00233031 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that during the preparation of the sterile tables the scrubbed nurse found that the 3D marker was supplied disassembled (the screw wasn’t secured in the socket) and that this may cause a sterility issue if the screw would have been discarded by mistake tor if the screw would fall to the floor. No patient was present at this time. 2021-Aug-18 IUD (Other, For): No new information. 2021-09-19 email 2021-09-19 00:00 (Rep, For): No new information. 2022-05-11 IUD (Other, For): No new information. MAZOR X SPINE DISPOSABLE KIT KIT0574 | 2021-Aug-15 Salesforce C00233031 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that during the preparation of the sterile tables the scrubbed nurse found that the 3D marker was supplied disassembled (the screw wasn’t secured in the socket) and that this may cause a sterility issue if the screw would have been discarded by mistake tor if the screw would fall to the floor. No patient was present at this time. 2021-Aug-18 IUD (Other, For): No new information. 2021-09-19 email 2021-09-19 00:00 (Rep, For): No new information. 2022-05-11 IUD (Other, For): No new information. | x marker disassembled salesforce outside proc edure preparation sterile tables scrubbed nurse marker supplied disassembled cause sterility issue screw discarded mistake tor screw fall floor present time iud new email new iud new mazor x spine disposable kit | NAV3233 | NAV3299 | NAV3246 | NAV2054 | NAV3242 | NAV3233NAV3299NAV3246NAV2054NAV3242 | True
7267 | 704160275 | X ¦ Lost Instruments | 2021-01-07T17:25:20.000000+0000 | MazorX | NAV2053 | MISSING PARTS | 20 | MAS2330-03S | KIT SVC MAS2330-03S ARM GUIDE | X ¦ Lost Instruments 2021-01-07 00703835 (Rep): Medtronic received information regarding guidance system accessories, outside of procedure. It was reported that that the arm guide and universal screw driver were lost. No patient was present. 2021-01-14 IUD x3 (Other): No new information. KIT SVC MAS2330-03S ARM GUIDE | 2021-01-07 00703835 (Rep): Medtronic received information regarding guidance system accessories, outside of procedure. It was reported that that the arm guide and universal screw driver were lost. No patient was present. 2021-01-14 IUD x3 (Other): No new information. | x lost instruments accessories outside procedure arm guide universal screw driver lost present iud new kit svc arm guide | NAV2053 | NAV3299 | NAV3249 | NAV3294 | NAV2046 | NAV2053NAV3299NAV3249NAV3294NAV2046 | True
7672 | 704246548 | XSE - Inaccurate Navigation (Case 1) | 2021-02-24T09:06:56.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Inaccurate Navigation (Case 1) 2019-09-30 00707153, 00707071, IUD, email, PDF (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were technical issues during a right L4-S1 lumbar interbody fusion procedure. The guidance system was inaccurate when tilting the patient and surgical bed. The most probable cause was a patient shift when rotating the bed. There was no patient harm and the procedure was delayed less than an hour. 2021-02-25 e1, e2, e3, e4, e5, e6, e7, e8, e9, e10 (Rep, For): Additional information received from a manufacturer representative re ported the inaccuracy was between 5 to 10 mm. Troubleshooting involved doing a snapshot check on the arm guide point and checking landmarks. The use of the guidance system was aborted and the procedure was completed using stealth navigation. 2021-12-07 IUD (Other, For): No new information. MAZOR X SYSTEM TPL0059 | 2019-09-30 00707153, 00707071, IUD, email, PDF (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were technical issues during a right L4-S1 lumbar interbody fusion procedure. The guidance system was inaccurate when tilting the patient and surgical bed. The most probable cause was a patient shift when rotating the bed. There was no patient harm and the procedure was delayed less than an hour. 2021-02-25 e1, e2, e3, e4, e5, e6, e7, e8, e9, e10 (Rep, For): Additional information received from a manufacturer representative re ported the inaccuracy was between 5 to 10 mm. Troubleshooting involved doing a snapshot check on the arm guide point and checking landmarks. The use of the guidance system was aborted and the procedure was completed using stealth navigation. 2021-12-07 IUD (Other, For): No new information. | xse inaccurate navigation iud email pdf spinal procedure technical issues right lumbar interbody fusion procedure inaccurate tilting surgical bed probable cause shift rotating bed procedure delayed hour additional ported inaccuracy mm troubleshooting involved doing snapshot check arm guide point checking landmarks use aborted procedure completed using stealth navigation iud new mazor x | NAV5000 | NAV1012 | NAV5001 | NAV2058 | NAV1006 | NAV5000NAV1012NAV5001NAV2058NAV1006 | True
9087 | 704547667 | XSE - OOBF/PC Failed & Screens Flicker | 2021-08-09T21:37:44.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215 | ASM0215 MAZOR X COMPUTER KIT | XSE - OOBF/PC Failed & Screens Flicker 2021-Aug-09 Salesforce C00231919 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that both screens flickered to black, the arm worked. The day after install both screens went to black and neither came back on. No patient was present. 2021-08-13 IUD x6 (Rep): Additional information received from a manufacturer representative reported that the replacement PC failed out of the box. 2021-09-05 e1 (Rep): No new information. 2021-09-07 IUD x2, e2 (Rep): Additional information received from a manufacturer representative reported the replacement PC would not provide an image to the screens. 2021-09-22 IUD (Other): No new information. 2021-09-30 e3, e4, e5, IUD x2 (Other): No new information. ASM0215 MAZOR X COMPUTER KIT | 2021-Aug-09 Salesforce C00231919 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that both screens flickered to black, the arm worked. The day after install both screens went to black and neither came back on. No patient was present. 2021-08-13 IUD x6 (Rep): Additional information received from a manufacturer representative reported that the replacement PC failed out of the box. 2021-09-05 e1 (Rep): No new information. 2021-09-07 IUD x2, e2 (Rep): Additional information received from a manufacturer representative reported the replacement PC would not provide an image to the screens. 2021-09-22 IUD (Other): No new information. 2021-09-30 e3, e4, e5, IUD x2 (Other): No new information. | xse oobfpc failed screens flicker salesforce outside procedure screens flickered black arm worked day install screens went black came present iud additional replacement pc failed box new iud additional replacement pc provide image screens iud new iud new mazor x computer kit | NAV3262 | NAV2054 | NAV3247 | NAV3261 | NAV3253 | NAV3262NAV2054NAV3247NAV3261NAV3253 | True
7155 | 704131645 | XSE - Broken Mazor Dual Clamp | 2020-12-17T20:01:14.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1355-05S | KIT SVC MAS1355-05S DUAL CLAMP | XSE - Broken Mazor Dual Clamp 2020-12-17 mPXR 788785, image 2020-12-19 10:38, 00703070 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the dual spinous process clamp was broken when being tightened during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-Dec-29 IUDx2 (other): No new information. 2021-01-12 IUD x2 (Other): No new information. 2021-03-01 IUD (Other): No new information. KIT SVC MAS1355-05S DUAL CLAMP | 2020-12-17 mPXR 788785, image 2020-12-19 10:38, 00703070 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the dual spinous process clamp was broken when being tightened during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-Dec-29 IUDx2 (other): No new information. 2021-01-12 IUD x2 (Other): No new information. 2021-03-01 IUD (Other): No new information. | xse broken mazor dual clamp mpxr image spinal procedure dual spinous process clamp broken tightened procedure procedure delayed hour new iud new iud new kit svc dual clamp | NAV5000 | NAV3295 | NAV4090 | NAV2053 | NAV3251 | NAV5000NAV3295NAV4090NAV2053NAV3251 | False
6252 | 703898284 | XSE - Instrument Verification Issue | 2020-08-11T19:30:12.000000+0000 | MazorX Stealth Edition | NAV3235 | 3D SCANS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Instrument Verification Issue 2020-08-04 00694406 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a scan and plan procedure was done. The usual setup steps were taken to get the system ready. During the procedure, the O1 system was selected, a 3Define scan was completed, a snapshot was completed and the surgical arm was positioned with the 3D marker. The O-arm was brought in and a scan was completed. The scan failed to be pushed to the guidance system. In the guidance system utilities, O1000 was unchecked. The manufacturer representative reported that there was a software bug, where if two profiles were setup (one O1 and one O2), then checking or unchecking this box in utilities in on profile also makes the change in the second profile. Since the images could not be transferred, a second spin was done. The representative noted that the only way to change the profile to the O2 was to exit out and redo the setup of the guidance system. When the representative exited the software, they checked the instruments in the setup menu and they were all green so re-verification was not done. During the second setup, the surgeon was using the marker to mark the middle of the surgical site, but the marker would not work as the instrument was not verified. All of the instruments were the same even though they showed green. There was no patient harm and the procedure was able delayed less than an hour. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-08-04 00694406 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a scan and plan procedure was done. The usual setup steps were taken to get the system ready. During the procedure, the O1 system was selected, a 3Define scan was completed, a snapshot was completed and the surgical arm was positioned with the 3D marker. The O-arm was brought in and a scan was completed. The scan failed to be pushed to the guidance system. In the guidance system utilities, O1000 was unchecked. The manufacturer representative reported that there was a software bug, where if two profiles were setup (one O1 and one O2), then checking or unchecking this box in utilities in on profile also makes the change in the second profile. Since the images could not be transferred, a second spin was done. The representative noted that the only way to change the profile to the O2 was to exit out and redo the setup of the guidance system. When the representative exited the software, they checked the instruments in the setup menu and they were all green so re-verification was not done. During the second setup, the surgeon was using the marker to mark the middle of the surgical site, but the marker would not work as the instrument was not verified. All of the instruments were the same even though they showed green. There was no patient harm and the procedure was able delayed less than an hour. | xse instrument verification issue spinal procedure scan plan procedure usual setup steps taken ready procedure selected scan completed snapshot completed surgical arm positioned marker oarm brought scan completed scan failed pushed utilities unchecked software bug profiles setup checking unchecking box utilities profile makes change second profile images transferred second spin noted way change profile exit redo setup exited software checked instruments setup menu green reverification second setup surgeon using marker mark middle surgical site marker work instrument verified instruments showed green procedure able delayed hour mazor x software installation kit | NAV3246 | NAV3235 | NAV3242 | NAV5000 | NAV3247 | NAV3246NAV3235NAV3242NAV5000NAV3247 | True
1275 | 703195105 | X- 3Define Functionality | 2019-06-07T14:36:21.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X- 3Define Functionality Error during 3-Define scan – timed out – Error – use generic work volume 2nd attempt at 3-Define Successful UNK MAZOR X 3DEFINE CAMERA | Error during 3-Define scan – timed out – Error – use generic work volume 2nd attempt at 3-Define Successful | x functionality error scan – timed – error – use generic work volume attempt successful unk mazor x camera | NAV3244 | NAV2047 | NAV3250 | NAV4101 | NAV4110 | NAV3244NAV2047NAV3250NAV4101NAV4110 | True
5104 | 703599797 | XSE - Unable to Track Instruments | 2020-02-06T00:13:19.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 30 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - Unable to Track Instruments 2020-02-05 00683483 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system had trouble tracking the instrumentation. It would pick up the navigation tracker and the attached instrument for a second and then a message would appear on the system station "low system performance, please contact Tech Services." This occurred every time an instrument was visible to the camera. The case was continued just using the trajectories without navigation. The manufacturer representative also noted that the system was delayed when doing CT/Fluoro matching. No patient harm was reported and surgery was delayed less than an hour. 2020-02-06 IUD x3, email 2020-02-20 19:31 (Rep): No new information. 2020-03-05 IUD (Other): No new information. 2020-05-14 IUD (Other): No new information. 2020-06-07 IUD x2 (Other): No new information. 2020-06-08 IUD x2 (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2020-02-05 00683483 (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system had trouble tracking the instrumentation. It would pick up the navigation tracker and the attached instrument for a second and then a message would appear on the system station "low system performance, please contact Tech Services." This occurred every time an instrument was visible to the camera. The case was continued just using the trajectories without navigation. The manufacturer representative also noted that the system was delayed when doing CT/Fluoro matching. No patient harm was reported and surgery was delayed less than an hour. 2020-02-06 IUD x3, email 2020-02-20 19:31 (Rep): No new information. 2020-03-05 IUD (Other): No new information. 2020-05-14 IUD (Other): No new information. 2020-06-07 IUD x2 (Other): No new information. 2020-06-08 IUD x2 (Other): No new information. | xse unable track instruments spinal procedure trouble tracking instrumentation pick navigation tracker attached instrument second message appear station low performance contact tech services occurred time instrument visible camera case continued just using trajectories navigation noted delayed doing ctfluoro matching surgery delayed hour iud email new iud new iud new iud new iud new mazor x pc station refurbished | NAV3262 | NAV3246 | NAV3299 | NAV1012 | NAV4110 | NAV3262NAV3246NAV3299NAV1012NAV4110 | True
2398 | 703234188 | X- Unmounting Issues | 2019-06-28T21:20:03.000000+0000 | MazorX | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- Unmounting Issues "CSR has had issues removing the system from the bed frame adapter following the past three cases. The CSR and engineer struggled to remove the system multiple times and force was required to allow the system to disengage. Please address this issue" UNK MAZOR X WORKSTATION | "CSR has had issues removing the system from the bed frame adapter following the past three cases. The CSR and engineer struggled to remove the system multiple times and force was required to allow the system to disengage. Please address this issue" | x unmounting issues csr issues removing bed frame adapter following past cases csr engineer struggled remove multiple times force required allow disengage address issue unk mazor x workstation | NAV3250 | NAV2058 | NAV3255 | NAV3264 | NAV2047 | NAV3250NAV2058NAV3255NAV3264NAV2047 | True
4541 | 703549308 | X - Registration Issues | 2020-01-07T21:38:14.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issues 2018-07-16 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a degenerative deformity with fixation at L2-L5. It was reported that calibration and accuracy checks were done before the procedure. On the first registration attempt which was done off of L2, red values were seen at L2 and green at all other levels. On the second attempt, manual adjustment was done on the L2 segmentation line and labeled off of L3, green values were seen at all levels. All trajectories were safe and accurate. Screws were verified in AP and lateral views. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-07-16 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a degenerative deformity with fixation at L2-L5. It was reported that calibration and accuracy checks were done before the procedure. On the first registration attempt which was done off of L2, red values were seen at L2 and green at all other levels. On the second attempt, manual adjustment was done on the L2 segmentation line and labeled off of L3, green values were seen at all levels. All trajectories were safe and accurate. Screws were verified in AP and lateral views. There was no patient harm and the procedure was not delayed over an hour. | x registration issues ccr spinal procedure treat degenerative deformity fixation calibration accuracy checks procedure registration attempt red values seen green levels second attempt manual adjustment segmentation line labeled green values seen levels trajectories safe accurate screws verified ap lateral views procedure delayed hour mazor x | NAV3242 | NAV4091 | NAV3246 | NAV3247 | NAV2047 | NAV3242NAV4091NAV3246NAV3247NAV2047 | True
6285 | 703903526 | X - System Responding Slowly/3Def Failur | 2020-08-13T23:31:15.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 30 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - System Responding Slowly/3Def Failur 2020-08-13 00694578 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when clicking the button to send the surgical arm to the next trajectory, the system took a long time for the surgical arm to respond. The delay was long enough for the surgeon to question why the arm was not moving. The manufacturer representative also noted that the 3Define scan failed during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-08-17 IUD x2 (Other): No new information. 2020-08-24 IUD x4 (Other): No new information. 2020-09-09 email 2020-09-09 08:46, IUD x2 (Rep): Additional information received from a manufacturer representative reported that tr oubleshooting involved restarting the software and checking the USB plug for looseness. The issue was resolved by replacing the computer and updating the software. The case was able to be successfully completed. 2021-06-20 IUD (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2020-08-13 00694578 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when clicking the button to send the surgical arm to the next trajectory, the system took a long time for the surgical arm to respond. The delay was long enough for the surgeon to question why the arm was not moving. The manufacturer representative also noted that the 3Define scan failed during the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-08-17 IUD x2 (Other): No new information. 2020-08-24 IUD x4 (Other): No new information. 2020-09-09 email 2020-09-09 08:46, IUD x2 (Rep): Additional information received from a manufacturer representative reported that tr oubleshooting involved restarting the software and checking the USB plug for looseness. The issue was resolved by replacing the computer and updating the software. The case was able to be successfully completed. 2021-06-20 IUD (Other): No new information. | x responding failur spinal procedure clicking button send surgical arm trajectory took long time surgical arm respond delay long surgeon question arm moving noted scan failed procedure procedure delayed hour iud new iud new email iud additional tr oubleshooting involved restarting software checking usb plug looseness issue resolved replacing computer updating software case able successfully completed iud new unk mazor x camera | NAV3244 | NAV3249 | NAV4095 | NAV3248 | NAV1012 | NAV3244NAV3249NAV4095NAV3248NAV1012 | True
8273 | 704383999 | XSE - Joint Shift Error | 2021-05-07T23:48:07.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 30 | ASM0206-01 | SURGICAL ARM | XSE - Joint Shift Error 2021-May-07 Salesforce C00226017 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgeon had placed all her guide wires accurately. They then cleared the arm to the left and placed an L2 spine jack and performed laminectomy at L2 to which they received joint shift error at Joint 6 (the patient was still mounted to the bed). This caused no delays for they were finished using the system. There was no impact on the patient outcome. 2021-05-11 IUD x4 (Other): No new information. 2021-06-24 302495326, IUD (Other): No new information. 2021-07-18 IUD x7 (Other): No new information, SURGICAL ARM | 2021-May-07 Salesforce C00226017 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgeon had placed all her guide wires accurately. They then cleared the arm to the left and placed an L2 spine jack and performed laminectomy at L2 to which they received joint shift error at Joint 6 (the patient was still mounted to the bed). This caused no delays for they were finished using the system. There was no impact on the patient outcome. 2021-05-11 IUD x4 (Other): No new information. 2021-06-24 302495326, IUD (Other): No new information. 2021-07-18 IUD x7 (Other): No new information, | xse joint shift error salesforce spinal p rocedure surgeon placed guide wires accurately cleared arm left placed spine jack performed laminectomy joint shift error joint caused delays finished using impact outcome iud new iud new iud new surgical arm | NAV4098 | NAV4101 | NAV2054 | NAV3253 | NAV5000 | NAV4098NAV4101NAV2054NAV3253NAV5000 | True
9085 | 704547667 | XSE - OOBF/PC Failed & Screens Flicker | 2021-08-09T21:37:44.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0215 | ASM0215 MAZOR X COMPUTER KIT | XSE - OOBF/PC Failed & Screens Flicker 2021-Aug-09 Salesforce C00231919 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that both screens flickered to black, the arm worked. The day after install both screens went to black and neither came back on. No patient was present. 2021-08-13 IUD x6 (Rep): Additional information received from a manufacturer representative reported that the replacement PC failed out of the box. 2021-09-05 e1 (Rep): No new information. 2021-09-07 IUD x2, e2 (Rep): Additional information received from a manufacturer representative reported the replacement PC would not provide an image to the screens. 2021-09-22 IUD (Other): No new information. 2021-09-30 e3, e4, e5, IUD x2 (Other): No new information. ASM0215 MAZOR X COMPUTER KIT | 2021-Aug-09 Salesforce C00231919 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that both screens flickered to black, the arm worked. The day after install both screens went to black and neither came back on. No patient was present. 2021-08-13 IUD x6 (Rep): Additional information received from a manufacturer representative reported that the replacement PC failed out of the box. 2021-09-05 e1 (Rep): No new information. 2021-09-07 IUD x2, e2 (Rep): Additional information received from a manufacturer representative reported the replacement PC would not provide an image to the screens. 2021-09-22 IUD (Other): No new information. 2021-09-30 e3, e4, e5, IUD x2 (Other): No new information. | xse oobfpc failed screens flicker salesforce outside procedure screens flickered black arm worked day install screens went black came present iud additional replacement pc failed box new iud additional replacement pc provide image screens iud new iud new mazor x computer kit | NAV3262 | NAV2054 | NAV3247 | NAV3261 | NAV3253 | NAV3262NAV2054NAV3247NAV3261NAV3253 | True
9619 | 704684640 | XSE - Intermittent Surgeon Monitor | 2021-10-27T14:24:28.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 60 | CAR0078-01S | KIT SVC CAR0078-01S Z AXIS SPLITTER PCB | XSE - Intermittent Surgeon Monitor 2021-Oct-27 Salesforce C00236884 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while planning, the surgeon monitor would intermittently go black. No patient was present. 2021-11-02 IUD x3 (Other): No new information. 2021-11-10 IUD (Other): No new information. 2021-12-21 IUD x3 (Other): No new information. 2022-03-29 IUD (Other): No new information. KIT SVC CAR0078-01S Z AXIS SPLITTER PCB | 2021-Oct-27 Salesforce C00236884 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while planning, the surgeon monitor would intermittently go black. No patient was present. 2021-11-02 IUD x3 (Other): No new information. 2021-11-10 IUD (Other): No new information. 2021-12-21 IUD x3 (Other): No new information. 2022-03-29 IUD (Other): No new information. | xse intermittent surgeon monitor salesforce outside procedure planning surgeon monitor intermittently black present iud new iud new iud new iud new kit svc z axis splitter pcb | NAV3253 | NAV3261 | NAV2047 | NAV2054 | NAV3249 | NAV3253NAV3261NAV2047NAV2054NAV3249 | True
7939 | 704310259 | XSE - Elbow Joint & Image Quality Issues | 2021-03-29T20:07:15.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Elbow Joint & Image Quality Issues 2021-Mar-29 Salesforce C00223472 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when the surgeon would send the surgical arm to a trajectory and put the elbow to the right, the elbow would remain on the left. The manufacturer representative tried to clear the elbow to the right and then send it to the trajectory but the elbow would go to the left. The surgeon was able to continue but had to change his position around the surgical arm. Additionally, after completing the imaging system spin for the scan and plan the image quality was very poor on the guidance system. The image quality was fine on the imaging system's Mobile Viewing Station (MVS). The surgeon was able to complete planning but complained on image quality. The surgeon ended up aborting the use of the guidance system for unrelated reasons. The issues caused a less than 1 hour surgical delay. There was no impact on the patient outcome. 2021-03-30 IUD x5 (Other): No new information. 2021-03-31 email 2021-03-31 18:30 (Rep): Additional information was received stating that the incision was not adequate so to extend the incision they would have had to unmount the robot. At that time the surgeon did not want to re-mount and re-register so aborting the use of the guidance system was decided to save time. The incision was incorrect. The arm only sending to the left side was an inconvenience but not the cause of aborting. Troubleshooting involved trying to re-send to the trajectory from different positions and clearing the arm but it did not resolve. After they aborted the use of the guidance system, the procedure was completed successfully and accurately. 2021-04-16 x3 IUD (Other): Additional information was received stating that the cause of the issue might have been due to being off or not positioned midline. 2021-04-28 IUD x4 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2021-Mar-29 Salesforce C00223472 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when the surgeon would send the surgical arm to a trajectory and put the elbow to the right, the elbow would remain on the left. The manufacturer representative tried to clear the elbow to the right and then send it to the trajectory but the elbow would go to the left. The surgeon was able to continue but had to change his position around the surgical arm. Additionally, after completing the imaging system spin for the scan and plan the image quality was very poor on the guidance system. The image quality was fine on the imaging system's Mobile Viewing Station (MVS). The surgeon was able to complete planning but complained on image quality. The surgeon ended up aborting the use of the guidance system for unrelated reasons. The issues caused a less than 1 hour surgical delay. There was no impact on the patient outcome. 2021-03-30 IUD x5 (Other): No new information. 2021-03-31 email 2021-03-31 18:30 (Rep): Additional information was received stating that the incision was not adequate so to extend the incision they would have had to unmount the robot. At that time the surgeon did not want to re-mount and re-register so aborting the use of the guidance system was decided to save time. The incision was incorrect. The arm only sending to the left side was an inconvenience but not the cause of aborting. Troubleshooting involved trying to re-send to the trajectory from different positions and clearing the arm but it did not resolve. After they aborted the use of the guidance system, the procedure was completed successfully and accurately. 2021-04-16 x3 IUD (Other): Additional information was received stating that the cause of the issue might have been due to being off or not positioned midline. 2021-04-28 IUD x4 (Other): No new information. | xse elbow joint image quality issues salesforce spinal p rocedure surgeon send surgical arm trajectory elbow right elbow remain left tried clear elbow right send trajectory elbow left surgeon able continue change position surgical arm additionally completing imaging spin scan plan image quality poor image quality fine imaging systems mobile viewing station surgeon able complete planning complained image quality surgeon ended aborting use unrelated reasons issues caused hour surgical delay impact outcome iud new email additional stating incision adequate extend incision unmount robot time surgeon did want remount reregister aborting use decided save time incision incorrect arm sending left inconvenience cause aborting troubleshooting involved trying resend trajectory different positions clearing arm did resolve aborted use procedure completed successfully accurately iud additional stating cause issue positioned midline iud new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV3237 | NAV3243 | NAV4096 | NAV3249NAV4095NAV3237NAV3243NAV4096 | False
4484 | 703546247 | XSE - Missing Iron Man | 2020-01-06T19:59:50.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | MAS2497-01S | KIT SVC MAS2497-01S IRON MAN | XSE - Missing Iron Man 2020-01-06 00681410 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the iron man is missing from the system. No patient was present. 2020-01-17 IUD x2 (Other): No new information. 2020-02-17 IUD x2 (Other): No new information. 2020-03-19 IUD (Other): No new information. 2020-04-06 IUD x2 (Other): No new information. 2020-05-04 IUD (Other): No new information. KIT SVC MAS2497-01S IRON MAN | 2020-01-06 00681410 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the iron man is missing from the system. No patient was present. 2020-01-17 IUD x2 (Other): No new information. 2020-02-17 IUD x2 (Other): No new information. 2020-03-19 IUD (Other): No new information. 2020-04-06 IUD x2 (Other): No new information. 2020-05-04 IUD (Other): No new information. | xse missing iron man outside procedure iron man missing present iud new iud new iud new iud new iud new kit svc iron man | NAV2053 | NAV2054 | NAV3249 | NAV1006 | NAV4095 | NAV2053NAV2054NAV3249NAV1006NAV4095 | True
7600 | 704235687 | XSE - System Noise Complaint | 2021-02-18T01:03:16.000000+0000 | MazorX Stealth Edition | NAV1013 | MAZOR X/STEALTH EDITION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - System Noise Complaint 2021-02-17 00706779 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the site was incredibly unhappy with the noise of the system. Its noted that when the navlock is placed to close to the arm the system produces a noise and flashes the red light of caution on the system. The site does not want this as a feature and wants it removed. There was no delay to the procedure and no impact on the patient outcome. 2021-02-24 IUD (Other): No new information. 2021-03-17 email 2021-03-17 11:46 (Rep): Additional information received from a manufacturer representative reported that the surgeon used long tools so the Navlock was not as close to the surgical arm. The system was functioning as intended. The surgeon was just annoyed by the sound. MAZOR X SYSTEM TPL0059 | 2021-02-17 00706779 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the site was incredibly unhappy with the noise of the system. Its noted that when the navlock is placed to close to the arm the system produces a noise and flashes the red light of caution on the system. The site does not want this as a feature and wants it removed. There was no delay to the procedure and no impact on the patient outcome. 2021-02-24 IUD (Other): No new information. 2021-03-17 email 2021-03-17 11:46 (Rep): Additional information received from a manufacturer representative reported that the surgeon used long tools so the Navlock was not as close to the surgical arm. The system was functioning as intended. The surgeon was just annoyed by the sound. | xse noise complaint intraoperatively spinal procedure site incredibly unhappy noise noted navlock placed close arm produces noise flashes red light caution site does want feature wants removed delay procedure impact outcome iud new email additional surgeon long tools navlock close surgical arm functioning intended surgeon just annoyed sound mazor x | NAV3249 | NAV5000 | NAV1006 | NAV3242 | NAV3243 | NAV3249NAV5000NAV1006NAV3242NAV3243 | False
1634 | 703208230 | X - Deviation issue | 2019-06-14T17:54:01.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviation issue It was reported that one of their plans skived. MAZOR X SOFTWARE KIT VER-3.0.2 | It was reported that one of their plans skived. | x deviation issue plans skived mazor x software kit | NAV4091 | NAV3246 | NAV3242 | NAV5000 | NAV2047 | NAV4091NAV3246NAV3242NAV5000NAV2047 | True
2830 | 703252201 | X- Positioner/3Define Camera Issues | 2019-07-10T20:21:13.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 30 | MAS2272-01 | RNX-WS-1046 | X- Positioner/3Define Camera Issues the compressor panel has a faulty main valve, which was bypassed to allow the manipulator to function for the lab. Additionally, the 3Define camera is not functional and the arm will need replaced. RNX-WS-1046 | the compressor panel has a faulty main valve, which was bypassed to allow the manipulator to function for the lab. Additionally, the 3Define camera is not functional and the arm will need replaced. | x camera issues compressor panel faulty main valve bypassed allow manipulator function lab additionally camera functional arm need replaced | NAV3244 | NAV1006 | NAV2054 | NAV2047 | NAV4091 | NAV3244NAV1006NAV2054NAV2047NAV4091 | False
2683 | 703245909 | X - SURGICAL SYSTEM ISSUE | 2019-07-07T17:16:48.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 30 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - SURGICAL SYSTEM ISSUE After surgeon had malleted drill guide - error message "shoulder movement detected/lock mount, and reregister Immediately following the error message the robot arm began to violently shake (video attached). Emergency stop button pressed and cannulas were removed from arm guide UNK MAZOR X ARM | After surgeon had malleted drill guide - error message "shoulder movement detected/lock mount, and reregister Immediately following the error message the robot arm began to violently shake (video attached). Emergency stop button pressed and cannulas were removed from arm guide | x surgical issue surgeon malleted drill guide error message shoulder movement detectedlock mount reregister immediately following error message robot arm began violently shake emergency stop button pressed cannulas removed arm guide unk mazor x arm | NAV4101 | NAV4108 | NAV3244 | NAV3294 | NAV4098 | NAV4101NAV4108NAV3244NAV3294NAV4098 | True
353 | 703085958 | XSE - difficulty mounting to bed | 2019-04-02T14:37:42.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | ASM0293-01 | RNX-PO-119 | XSE - difficulty mounting to bed 2019-Mar-28 SCL1900769 intake form, mov (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there was difficulty mounting the system to the bed. One tooth engaged the other was not. They used a screw driver to get it off the bed. The next time the other tooth did not latch, and they had to remove the whole bed frame adapter. Then it happened a third time, so they attached the bed frame first and then but the bed frame adapter on to the bed. It was noted that it seemed like one tooth was lagging behind. The right handle of the positioner was broken off and the left was loose soon to fall off. It was noted that there was no patient harm and the issue delayed the surgical time very close to an hour. There were no additional complications reported or anticipated as a result of the event. RNX-PO-119 | 2019-Mar-28 SCL1900769 intake form, mov (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there was difficulty mounting the system to the bed. One tooth engaged the other was not. They used a screw driver to get it off the bed. The next time the other tooth did not latch, and they had to remove the whole bed frame adapter. Then it happened a third time, so they attached the bed frame first and then but the bed frame adapter on to the bed. It was noted that it seemed like one tooth was lagging behind. The right handle of the positioner was broken off and the left was loose soon to fall off. It was noted that there was no patient harm and the issue delayed the surgical time very close to an hour. There were no additional complications reported or anticipated as a result of the event. | xse difficulty mounting bed mov spinal procedure difficulty mounting bed tooth engaged screw driver bed time tooth did latch remove bed frame adapter happened time attached bed frame bed frame adapter bed noted like tooth lagging right handle positioner broken left loose soon fall noted issue delayed surgical time close hour additional complications anticipated | NAV3250 | NAV3255 | NAV4103 | NAV2054 | NAV5000 | NAV3250NAV3255NAV4103NAV2054NAV5000 | True
9252 | 704589422 | XSE - SRA Fell/No Sound | 2021-09-01T22:10:02.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 30 | ASM0206-01 | SURGICAL ARM | XSE - SRA Fell/No Sound 2021-Sep-01 Salesforce C00233828, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was tested prior to the procedure and passed the 3 and 10 point accuracy tests. When moving the surgical arm to the oblique position during the case, there was no sound and the software did not acknowledge the surgical arm reached the position. The surgical arm was sent to the clear up position, but the same thing occurred. The surgical arm was sent to the oblique position for a third time and the buttons were used to manually lower the arm, but the arm fell on the patient. A soft restart was done and the surgical arm was sent to the clear up position. The surgical arm was then able to be sent to the oblique position and had passed an accuracy check. The case was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2021-09-06 IUD, email 2021-09-07 14:15 (Rep): Additional information received from a manufacturer representative reported that they had opened the virtual keyboard and started a movement with the surgical arm towards the patient in the oblique image acquisition position and it fell 8-12 inches. No other actions were being taken when the surgical arm fell besides using the virtual keyboard. 2021-09-22 IUD x2 (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-12-13 302664740 (Other): No new information. 2021-12-20 IUD x3 (Other): No new information. SURGICAL ARM | 2021-Sep-01 Salesforce C00233828, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was tested prior to the procedure and passed the 3 and 10 point accuracy tests. When moving the surgical arm to the oblique position during the case, there was no sound and the software did not acknowledge the surgical arm reached the position. The surgical arm was sent to the clear up position, but the same thing occurred. The surgical arm was sent to the oblique position for a third time and the buttons were used to manually lower the arm, but the arm fell on the patient. A soft restart was done and the surgical arm was sent to the clear up position. The surgical arm was then able to be sent to the oblique position and had passed an accuracy check. The case was able to be completed. There was no patient harm and the procedure was delayed less than an hour. 2021-09-06 IUD, email 2021-09-07 14:15 (Rep): Additional information received from a manufacturer representative reported that they had opened the virtual keyboard and started a movement with the surgical arm towards the patient in the oblique image acquisition position and it fell 8-12 inches. No other actions were being taken when the surgical arm fell besides using the virtual keyboard. 2021-09-22 IUD x2 (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-12-13 302664740 (Other): No new information. 2021-12-20 IUD x3 (Other): No new information. | xse sra fellno sound salesforce iud spinal procedure surgical arm tested prior procedure passed point accuracy tests moving surgical arm oblique position case sound software did acknowledge surgical arm reached position surgical arm sent clear position thing occurred surgical arm sent oblique position time buttons manually lower arm arm fell soft restart surgical arm sent clear position surgical arm able sent oblique position passed accuracy check case able completed procedure delayed hour iud email additional opened virtual keyboard started movement surgical arm oblique image acquisition position fell inches actions taken surgical arm fell using virtual keyboard iud new iud new new iud new surgical arm | NAV4101 | NAV3249 | NAV4095 | NAV3239 | NAV2047 | NAV4101NAV3249NAV4095NAV3239NAV2047 | False
10060 | 704792564 | XSE - Boot Up Error | 2022-01-05T20:32:55.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215 | ASM0215 MAZOR X COMPUTER KIT | XSE - Boot Up Error 2021-Dec-29 Salesforce C00240283 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that an error message was displayed every time the guidance system was booted up. The error message was for failing components detected. There was no patient involvement. 2022-01-06 IUD x2 (Other): No new information. 2022-01-19 IUD (Other): No new information 2022-02-08 IUD (Other): No new information. 2022-03-17 IUD (Other): No new information. ASM0215 MAZOR X COMPUTER KIT | 2021-Dec-29 Salesforce C00240283 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that an error message was displayed every time the guidance system was booted up. The error message was for failing components detected. There was no patient involvement. 2022-01-06 IUD x2 (Other): No new information. 2022-01-19 IUD (Other): No new information 2022-02-08 IUD (Other): No new information. 2022-03-17 IUD (Other): No new information. | xse boot error salesforce outside procedure error message displayed time booted error message failing components detected iud new iud new iud new iud new mazor x computer kit | NAV3262 | NAV2047 | NAV3248 | NAV3246 | NAV3244 | NAV3262NAV2047NAV3248NAV3246NAV3244 | True
4627 | 703554312 | XSE - Schanz Arm Issues | 2020-01-09T20:01:32.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 20 | MAS2351-04 | SCHANZ ARM | XSE - Schanz Arm Issues 2020-01-09 00681706 (Rep, HCP): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the manufacturer representative was unable to fit the pins through the Schanz arm. This was found before a patient was present. 2020-01-20 IUD (Other): No new information. SCHANZ ARM | 2020-01-09 00681706 (Rep, HCP): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the manufacturer representative was unable to fit the pins through the Schanz arm. This was found before a patient was present. 2020-01-20 IUD (Other): No new information. | xse schanz arm issues prior procedure unable fit pins schanz arm present iud new schanz arm | NAV3299 | NAV3249 | NAV4095 | NAV3295 | NAV3298 | NAV3299NAV3249NAV4095NAV3295NAV3298 | True
7174 | 704131899 | XSE - Dark Intra-op Images | 2020-12-17T21:16:25.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Dark Intra-op Images 2020-12-16 00702946, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after registration, the lateral view of the fluoro image was very dark and the manufacturer representative could only see the outline of the cortical bone. The representative tried changing the view, but the issue was not resolved. The representative then went back into planning to adjust the brightness and contrast of the image. The brightness was adjusted to be increased. The image looked good in the planning view, but the issue was not resolved in the operate/navigate views. The surgeon decided to continue with the case as is. After the procedure, the representative loaded the export file onto their laptop and they were unable able to replicate the failure. There was no patient harm and the procedure was delayed less than an hour. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-12-16 00702946, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after registration, the lateral view of the fluoro image was very dark and the manufacturer representative could only see the outline of the cortical bone. The representative tried changing the view, but the issue was not resolved. The representative then went back into planning to adjust the brightness and contrast of the image. The brightness was adjusted to be increased. The image looked good in the planning view, but the issue was not resolved in the operate/navigate views. The surgeon decided to continue with the case as is. After the procedure, the representative loaded the export file onto their laptop and they were unable able to replicate the failure. There was no patient harm and the procedure was delayed less than an hour. | xse dark intraop images iud spinal procedure registration lateral view fluoro image dark outline cortical bone tried changing view issue resolved went planning adjust brightness contrast image brightness adjusted increased image looked good planning view issue resolved operatenavigate views surgeon decided continue case procedure loaded export file laptop unable able replicate failure procedure delayed hour mazor x software installation kit | NAV3246 | NAV4112 | NAV2047 | NAV3242 | NAV3237 | NAV3246NAV4112NAV2047NAV3242NAV3237 | True
9368 | 704619712 | XSE - Damaged Part\Stripped Screws | 2021-09-20T18:00:57.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 20 | MAS1226-02S | KIT SVC MAS1226-02S MAIN IMG ADPTR BRCKT | XSE - Damaged Part\Stripped Screws 2021-Sep-20 Salesforce C00234697 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter clamp was cross threaded. It was also reported that while inspecting this site's instruments, they found 3 bone mount bride screws were stripped. No patient was present. 2021-09-27 IUD x3 (Other): No new information. 2021-10-13 IUD (Other): No new information. 2021-11-09 email 2021-11-09 12:45 (Rep): Additional information received from a manufacturer representative reported that the cause of the image adapter being cross threaded was overtightening of the clamp since it needs to be tight to make sure it is securely attached. The suspected cause of the stripped screws was overtightening and loosening over time. KIT SVC MAS1226-02S MAIN IMG ADPTR BRCKT | 2021-Sep-20 Salesforce C00234697 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter clamp was cross threaded. It was also reported that while inspecting this site's instruments, they found 3 bone mount bride screws were stripped. No patient was present. 2021-09-27 IUD x3 (Other): No new information. 2021-10-13 IUD (Other): No new information. 2021-11-09 email 2021-11-09 12:45 (Rep): Additional information received from a manufacturer representative reported that the cause of the image adapter being cross threaded was overtightening of the clamp since it needs to be tight to make sure it is securely attached. The suspected cause of the stripped screws was overtightening and loosening over time. | xse damaged partstripped screws salesforce outside procedure image adapter clamp cross threaded inspecting sites instruments bone mount bride screws stripped present iud new iud new email additional cause image adapter cross threaded overtightening clamp needs tight make sure securely attached suspected cause stripped screws overtightening loosening time kit svc main img adptr brckt | NAV3258 | NAV3299 | NAV2054 | NAV3256 | NAV3252 | NAV3258NAV3299NAV2054NAV3256NAV3252 | True
218 | 703041270 | XSE - monitor won't rotate/X-eye won't 0 | 2019-03-06T20:37:52.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | ASM0275-01 | RNX-WS-1210 | XSE - monitor won't rotate/X-eye won't 0 2019-Mar-04 SCL1900519 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the monitor would not rotate fully and was missing necessary hardware. The X-Eye would not zero out when trying to calibrate. It was noted there was no patient invovlement. There were no additional complications reported or anticipated as a result of the event. RNX-WS-1210 | 2019-Mar-04 SCL1900519 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the monitor would not rotate fully and was missing necessary hardware. The X-Eye would not zero out when trying to calibrate. It was noted there was no patient invovlement. There were no additional complications reported or anticipated as a result of the event. | xse monitor wont rotatexeye wont spinal procedure monitor rotate fully missing necessary hardware xeye zero trying calibrate noted invovlement additional complications anticipated | NAV3261 | NAV2054 | NAV5000 | NAV3295 | NAV4101 | NAV3261NAV2054NAV5000NAV3295NAV4101 | True
6299 | 703906165 | XSE - SRA Fails Advanced Accuracy Test | 2020-08-17T05:36:49.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 50 | ASM0206-01 | SURGICAL ARM | XSE - SRA Fails Advanced Accuracy Test 2020-08-14 00693465, 00694660 (Rep): Medtronic received information regarding a guidance system during preventative maintenance. It was reported that the surgical arm failed the advanced accuracy due to a loose joint 6. Joint 6 could be moved left or right and had too much play in the joint to pass the advanced accuracy test. The left positioner button was broken and the handle was replaced. No patient was present. 2020-08-17 IUD x2 (Other): No new information. 2020-08-24 IUD x5 (Other): No new information. 2020-09-09 IUD x5 (Other): No new information. 2020-09-22 IUD x8 (Other): No new information. 2020-10-09 IUD x9 (Other): No new information. 2020-11-10 IUD x2 (Other): No new information. 2020-11-16 IUD x2 (Other): No new information. SURGICAL ARM | 2020-08-14 00693465, 00694660 (Rep): Medtronic received information regarding a guidance system during preventative maintenance. It was reported that the surgical arm failed the advanced accuracy due to a loose joint 6. Joint 6 could be moved left or right and had too much play in the joint to pass the advanced accuracy test. The left positioner button was broken and the handle was replaced. No patient was present. 2020-08-17 IUD x2 (Other): No new information. 2020-08-24 IUD x5 (Other): No new information. 2020-09-09 IUD x5 (Other): No new information. 2020-09-22 IUD x8 (Other): No new information. 2020-10-09 IUD x9 (Other): No new information. 2020-11-10 IUD x2 (Other): No new information. 2020-11-16 IUD x2 (Other): No new information. | xse sra fails advanced accuracy test preventative maintenance surgical arm failed advanced accuracy loose joint joint moved left right play joint pass advanced accuracy test left positioner button broken handle replaced present iud new iud new iud new iud new iud new iud new iud new surgical arm | NAV4093 | NAV3251 | NAV2054 | NAV4097 | NAV3249 | NAV4093NAV3251NAV2054NAV4097NAV3249 | False
6236 | 703893961 | XSE - SRA Hit Bone Mount Bridge | 2020-08-07T23:28:22.000000+0000 | MazorX Stealth Edition | NAV4096 | ARM HIT WORK VOLUME | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - SRA Hit Bone Mount Bridge 2020-08-06 00694241 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when sending the surgical arm to a trajectory, the arm hit the bone mount bridge. The patient was re-registered and the case was continued. The manufacturer representative believed that the surgeon did not put a blue towel over the mount when the 3Define scan was completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-14 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-08-06 00694241 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when sending the surgical arm to a trajectory, the arm hit the bone mount bridge. The patient was re-registered and the case was continued. The manufacturer representative believed that the surgeon did not put a blue towel over the mount when the 3Define scan was completed. There was no patient harm and the procedure was delayed less than an hour. 2020-08-14 IUD (Other): No new information. | xse sra hit bone mount bridge spinal procedure sending surgical arm trajectory arm hit bone mount bridge reregistered case continued believed surgeon did blue towel mount scan completed procedure delayed hour iud new mazor x | NAV5000 | NAV4096 | NAV3246 | NAV3242 | NAV3243 | NAV5000NAV4096NAV3246NAV3242NAV3243 | True
9306 | 704605841 | XSE - Noise/No Connection to Controller | 2021-09-12T13:55:49.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Noise/No Connection to Controller 2021-Sep-12 Salesforce C00234294 (rep, hcp, for): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while booting the system they heard a noise coming from the internal compartment sounds like a minor "explosion ". The screen was uploaded but they received an error "no connection to controller." They performed hard shutdown and waited a few minutes. but after rebooting same error message was displayed. They opened the cover of the controller and verified that the short cable is connected properly, but they saw that there is no indication of power supply to any of the parts of the electrical panel. They verified that all switches are pressed and all surgical system cables are connected properly. There was no patient harm and a delay of less than one hour. The surgery was aborted. 2021-09-29 IUD x5 (Other, For): No new information. 2021-10-05 IUD x3 (Other): No new information. 2021-10-21 IUD x6, 302611538 (Other): No new information. 2021-12-07 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-Sep-12 Salesforce C00234294 (rep, hcp, for): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while booting the system they heard a noise coming from the internal compartment sounds like a minor "explosion ". The screen was uploaded but they received an error "no connection to controller." They performed hard shutdown and waited a few minutes. but after rebooting same error message was displayed. They opened the cover of the controller and verified that the short cable is connected properly, but they saw that there is no indication of power supply to any of the parts of the electrical panel. They verified that all switches are pressed and all surgical system cables are connected properly. There was no patient harm and a delay of less than one hour. The surgery was aborted. 2021-09-29 IUD x5 (Other, For): No new information. 2021-10-05 IUD x3 (Other): No new information. 2021-10-21 IUD x6, 302611538 (Other): No new information. 2021-12-07 IUD (Other): No new information. | xse noiseno connection controller salesforce spinal procedure booting heard noise coming internal compartment sounds like minor explosion screen uploaded error connection controller performed hard shutdown waited minutes rebooting error message displayed opened cover controller verified short cable connected properly saw indication power supply parts electrical panel verified switches pressed surgical cables connected properly delay hour surgery aborted iud new iud new iud new iud new kit svc cart elec panel assy | NAV2047 | NAV3264 | NAV2057 | NAV3263 | NAV2055 | NAV2047NAV3264NAV2057NAV3263NAV2055 | False
255 | 703054801 | X - Deviation | 2019-03-14T15:43:03.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviation 2019-Mar-12 SCL1900568, Intake form Bethesda Intake form .msg (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the HCP drilled all pedicle holes with drill and then placed bone wax in the holes. The HCP did not want to use cannulated screws or k-wires. After all holes were drilled and pedicle probe checked for accuracy, screws were placed. Screws all stimmed greater than 10. The HCP did not want to take an x-ray until finals, and the HCP closed the patient and took finals. T12 left was lateral, so the HCP had to re-open, remove the rod, and reposition the screw. Since the HCP did not use k-wires it was thought that while they were free handing, the hole the screw deviated from its original path. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-Mar-12 SCL1900568, Intake form Bethesda Intake form .msg (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the HCP drilled all pedicle holes with drill and then placed bone wax in the holes. The HCP did not want to use cannulated screws or k-wires. After all holes were drilled and pedicle probe checked for accuracy, screws were placed. Screws all stimmed greater than 10. The HCP did not want to take an x-ray until finals, and the HCP closed the patient and took finals. T12 left was lateral, so the HCP had to re-open, remove the rod, and reposition the screw. Since the HCP did not use k-wires it was thought that while they were free handing, the hole the screw deviated from its original path. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x deviation bethesda msg spinal procedure hcp drilled pedicle holes drill placed bone wax holes hcp did want use cannulated screws kwires holes drilled pedicle probe checked accuracy screws placed screws stimmed greater hcp did want xray finals hcp closed took finals left lateral hcp reopen remove rod reposition screw hcp did use kwires thought free handing hole screw deviated original path noted additional complications anticipated mazor x software kit | NAV5000 | NAV4091 | NAV2054 | NAV1006 | NAV3242 | NAV5000NAV4091NAV2054NAV1006NAV3242 | True
1870 | 703220346 | X - Accessories Issue | 2019-06-21T12:18:13.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 30 | KIT0736-02 | LONG DRILL GUIDE AND LONG CANNULA | X - Accessories Issue Surgeon mentioned that the screw is starting to strip on the clamp accessory piece. LONG DRILL GUIDE AND LONG CANNULA | Surgeon mentioned that the screw is starting to strip on the clamp accessory piece. | x accessories issue surgeon mentioned screw starting strip clamp accessory piece long drill guide long cannula | NAV3299 | NAV3295 | NAV4091 | NAV3242 | NAV3253 | NAV3299NAV3295NAV4091NAV3242NAV3253 | True
2243 | 703230814 | X - SOFTWARE ISSUE | 2019-06-27T15:02:15.000000+0000 | MazorX | NAV3247 | EXITS | 20 | KIT0639-09 | MAZOR X SOFTWARE KIT VER-2.3.1 | X - SOFTWARE ISSUE Screen on monitor would turn completely black except the mouse would still be recognized. Tried soft restart and it erased c-arm setup, 3 point test and Kit Code. Then tried hard restart and changed outlet, same problem occurred this time CSR waited for the screen to come back on its own and it did not deleted the previous 3 checkpoints required to move forward./ Software crashed x2 outlet changed, software crashed again no soft/hard restart system restored properly MAZOR X SOFTWARE KIT VER-2.3.1 | Screen on monitor would turn completely black except the mouse would still be recognized. Tried soft restart and it erased c-arm setup, 3 point test and Kit Code. Then tried hard restart and changed outlet, same problem occurred this time CSR waited for the screen to come back on its own and it did not deleted the previous 3 checkpoints required to move forward./ Software crashed x2 outlet changed, software crashed again no soft/hard restart system restored properly | x software issue screen monitor turn completely black mouse recognized tried soft restart erased carm setup point test kit code tried hard restart changed outlet problem occurred time csr waited screen come did deleted previous checkpoints required forward software crashed outlet changed software crashed softhard restart restored properly mazor x software kit | NAV3246 | NAV3247 | NAV3261 | NAV3248 | NAV3262 | NAV3246NAV3247NAV3261NAV3248NAV3262 | True
1745 | 703213029 | X - shoulder not locking | 2019-06-18T15:30:23.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - shoulder not locking 2019-Jun-13 SCL1901762 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the shoulder button began blinking during the case and was unable to lock manually. The rep had to lock through the software. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X SHOULDER | 2019-Jun-13 SCL1901762 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the shoulder button began blinking during the case and was unable to lock manually. The rep had to lock through the software. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x shoulder locking spinal procedure shoulder button began blinking case unable lock manually lock software additional complications anticipated unk mazor x shoulder | NAV4107 | NAV3244 | NAV2054 | NAV2047 | NAV3250 | NAV4107NAV3244NAV2054NAV2047NAV3250 | True
1881 | 703220549 | X - Accessory Problem | 2019-06-21T14:44:46.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | KIT0737-02 | SHORT DRILL GUIDE AND SHORT CANNULA | X - Accessory Problem MAS2245 Mazor X short drill guide , the top collar broke off of drill guide SHORT DRILL GUIDE AND SHORT CANNULA | MAS2245 Mazor X short drill guide , the top collar broke off of drill guide | x accessory problem mazor x short drill guide collar broke drill guide short drill guide short cannula | NAV3299 | NAV3295 | NAV4091 | NAV3242 | NAV4090 | NAV3299NAV3295NAV4091NAV3242NAV4090 | True
10047 | 704782835 | XSE - Failed 3Define Scan | 2021-12-28T16:38:37.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - Failed 3Define Scan 2021-Dec-27 Salesforce C00240214, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera took pictures of the field, but when it was done and computing to construct the model it displays "scan failed" on the screen. This is a known, ongoing issue on this system. Whenever this system is tested outside clinical use, the 3Define camera works but during a clinical case it keeps failing. The manufacturer representative entered generic work volume and case proceeded. There was no impact on the patient outcome. 2022-01-05 IUD x2 (Other): No new information. 2022-01-20 IUD (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2021-Dec-27 Salesforce C00240214, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera took pictures of the field, but when it was done and computing to construct the model it displays "scan failed" on the screen. This is a known, ongoing issue on this system. Whenever this system is tested outside clinical use, the 3Define camera works but during a clinical case it keeps failing. The manufacturer representative entered generic work volume and case proceeded. There was no impact on the patient outcome. 2022-01-05 IUD x2 (Other): No new information. 2022-01-20 IUD (Other): No new information. | xse failed scan salesforce iud spinal procedure camera took pictures field computing construct model displays scan failed screen known ongoing issue tested outside clinical use camera works clinical case keeps failing entered generic work volume case proceeded impact outcome iud new iud new unk mazor x camera | NAV3244 | NAV3253 | NAV3261 | NAV2058 | NAV1013 | NAV3244NAV3253NAV3261NAV2058NAV1013 | True
5523 | 703744965 | XSE - Deviation of trajectory | 2020-05-08T19:45:50.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation of trajectory 2020-05-08 00688558 (HCP, Rep): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a L4 L5 navigated case there was a reported deviation of trajectory. The manufacturer representative noticed that the technique of instrumenting caused intense vibration of the arm. The registration was completed with no issue and a screw was placed at L5 (left) with no issue. When tooling at L4 left, the deviation was showing grossly superior (~8mm). The manufacturer representative notes that navigation did not capture the deviation, therefore he expects patient movement or weak arm joints attributing to the deviation. The issue extended the surgical time by 45 min. There was no impact on the patient outcome. 2020-05-11 IUD x2 (Other): No new information. 2020-05-15 IUD (Other): No new information. 2020-05-26 IUD x2 (Other): No new information. 2020-06-07 IUD (Other): No new information. 2020-07-01 IUD (Other): No new information. 2020-07-13 IUD (Other): No new information. 2020-07-21 CCR, Attachment: No new information. 2021-03-16 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-05-08 00688558 (HCP, Rep): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a L4 L5 navigated case there was a reported deviation of trajectory. The manufacturer representative noticed that the technique of instrumenting caused intense vibration of the arm. The registration was completed with no issue and a screw was placed at L5 (left) with no issue. When tooling at L4 left, the deviation was showing grossly superior (~8mm). The manufacturer representative notes that navigation did not capture the deviation, therefore he expects patient movement or weak arm joints attributing to the deviation. The issue extended the surgical time by 45 min. There was no impact on the patient outcome. 2020-05-11 IUD x2 (Other): No new information. 2020-05-15 IUD (Other): No new information. 2020-05-26 IUD x2 (Other): No new information. 2020-06-07 IUD (Other): No new information. 2020-07-01 IUD (Other): No new information. 2020-07-13 IUD (Other): No new information. 2020-07-21 CCR, Attachment: No new information. 2021-03-16 IUD (Other): No new information. | xse deviation trajectory intraoperatively spinal procedure navigated case deviation trajectory noticed technique instrumenting caused intense vibration arm registration completed issue screw placed issue tooling left deviation showing grossly superior notes navigation did capture deviation expects movement weak arm joints attributing deviation issue extended surgical time min impact outcome iud new iud new iud new iud new iud new iud new ccr attachment new iud new mazor x | NAV5000 | NAV4091 | NAV4101 | NAV2047 | NAV3261 | NAV5000NAV4091NAV4101NAV2047NAV3261 | True
864 | 703178840 | X-Bed Frame Adaptor Issues | 2019-05-29T18:44:01.000000+0000 | MazorX | NAV3255 | BED FRAME ISSUES | 20 | MAS1574-01 | BED FRAME ADAPTOR | X-Bed Frame Adaptor Issues 2018-10-22 SCL1804829(Rep): Medtronic received information regarding a guidance system being used. It was reported that the knob used to tighten the bed frame adaptor was broken. No patient involvement. BED FRAME ADAPTOR | 2018-10-22 SCL1804829(Rep): Medtronic received information regarding a guidance system being used. It was reported that the knob used to tighten the bed frame adaptor was broken. No patient involvement. | xbed frame adaptor issues knob tighten bed frame adaptor broken bed frame adaptor | NAV3255 | NAV3295 | NAV4090 | NAV3299 | NAV2054 | NAV3255NAV3295NAV4090NAV3299NAV2054 | True
4053 | 703495049 | X - Arm, Tablet, & PC issue | 2019-12-02T18:35:18.000000+0000 | MazorX | NAV3262 | PC ISSUES | 50 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | X - Arm, Tablet, & PC issue 2018-11-8 SCL1805091(Rep): Medtronic received information regarding a guidance system being used during system installation. It was reported that the software installation was not completed until all files were deleted. The surgical arm produced gains lost errors at driver four and five. The arm was replaced, but its replacement had scratches on it and did not look ideal. The manufacturer representative tried to fix the tablet, but the board inside the tablet was fried so it needed to be replaced. The second tablet also had a fired board, so a third tablet was sent. The computer needed to be replaced because the deletion of the files may have caused a problem with the system's software. A program had to be launched from the C drive and would not launch from the guidance system software. No patient involvement. MAZOR X PC STATION - REFURBISHED | 2018-11-8 SCL1805091(Rep): Medtronic received information regarding a guidance system being used during system installation. It was reported that the software installation was not completed until all files were deleted. The surgical arm produced gains lost errors at driver four and five. The arm was replaced, but its replacement had scratches on it and did not look ideal. The manufacturer representative tried to fix the tablet, but the board inside the tablet was fried so it needed to be replaced. The second tablet also had a fired board, so a third tablet was sent. The computer needed to be replaced because the deletion of the files may have caused a problem with the system's software. A program had to be launched from the C drive and would not launch from the guidance system software. No patient involvement. | x arm tablet pc issue installation software installation completed files deleted surgical arm produced gains lost errors driver arm replaced replacement scratches did look ideal tried fix tablet board inside tablet fried needed replaced second tablet fired board tablet sent computer needed replaced deletion files caused problem systems software program launched c drive launch software mazor x pc station refurbished | NAV2054 | NAV3262 | NAV3246 | NAV2056 | NAV3253 | NAV2054NAV3262NAV3246NAV2056NAV3253 | True
9171 | 704565813 | XSE - SW Defect/Low IPeak Value | 2021-08-19T09:43:36.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - SW Defect/Low IPeak Value 2021-Aug-16 Salesforce C00233161 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the default value of IPeak in a driver is very low sometimes and would fail when the representative tried to set the Imax on MC startup since it was not possible to set the Imax higher than the IPeak. The scenario involved incorrect Imax values, the log on the ControlStudio showed an error that was not in the logs, and there was no function that read all values from the driver to verify the writing. An error message 8047 popped up after releasing the EMB in the surgical setup procedure. There was no patient involvement. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Aug-16 Salesforce C00233161 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the default value of IPeak in a driver is very low sometimes and would fail when the representative tried to set the Imax on MC startup since it was not possible to set the Imax higher than the IPeak. The scenario involved incorrect Imax values, the log on the ControlStudio showed an error that was not in the logs, and there was no function that read all values from the driver to verify the writing. An error message 8047 popped up after releasing the EMB in the surgical setup procedure. There was no patient involvement. | xse sw defectlow ipeak value salesforce outside procedure default value ipeak driver low fail tried set imax mc startup possible set imax higher ipeak scenario involved incorrect imax values log controlstudio showed error logs function read values driver verify writing error message popped releasing emb surgical setup procedure mazorx sw installation kit | NAV3246 | NAV3249 | NAV2054 | NAV2056 | NAV3251 | NAV3246NAV3249NAV2054NAV2056NAV3251 | True
1890 | 703220714 | X- Arm Issues | 2019-06-21T16:18:36.000000+0000 | MazorX | NAV4097 | ARM JOINT LOOSE | 20 | ASM0206-01 | SURGICAL ARM | X- Arm Issues Aborted case do to failing accuracy test. Arm did not pass accuracy test 2 times, then received Error 3254 and the arm dropped down fast. Tried to do reset movement and reset gain with Alex and did not change the accuracy test result. I attach a video and image. Mazor X was aborted and the surgeon used the RNS instead - CCR RNS was attached. SURGICAL ARM | Aborted case do to failing accuracy test. Arm did not pass accuracy test 2 times, then received Error 3254 and the arm dropped down fast. Tried to do reset movement and reset gain with Alex and did not change the accuracy test result. I attach a video and image. Mazor X was aborted and the surgeon used the RNS instead - CCR RNS was attached. | x arm issues aborted case failing accuracy test arm did pass accuracy test times error arm dropped fast tried reset movement reset gain alex did change accuracy test attach video image mazor x aborted surgeon rns instead ccr rns attached surgical arm | NAV4092 | NAV4112 | NAV4091 | NAV3246 | NAV4093 | NAV4092NAV4112NAV4091NAV3246NAV4093 | False
2484 | 703236474 | X - WORKSTATION/ARM ISSUE | 2019-07-01T18:43:31.000000+0000 | MazorX | NAV2054 | OUT OF BOX FAILURES | 30 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - WORKSTATION/ARM ISSUE This SC has been opened to document the issues with system 21. Medtronic accepted the system as is however there were several issues with the system: - The front handle of the system was damaged while in transport. It needs to be replaced. - SRA23 (system 21's original arm) was removed from the system due to it having wiring problems UNK MAZOR X ARM | This SC has been opened to document the issues with system 21. Medtronic accepted the system as is however there were several issues with the system: - The front handle of the system was damaged while in transport. It needs to be replaced. - SRA23 (system 21's original arm) was removed from the system due to it having wiring problems | x workstationarm issue sc opened document issues accepted issues handle damaged transport needs replaced removed having wiring problems unk mazor x arm | NAV2058 | NAV2055 | NAV3244 | NAV3256 | NAV3251 | NAV2058NAV2055NAV3244NAV3256NAV3251 | False
799 | 703176218 | X - deviation | 2019-05-28T18:36:26.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - deviation 2019-May-22 SCL1901311 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that after all of the trajectories were sent, accuracy was checked using fluoroscopic imaging, and the L2 wire on the left was lateral to the plan. The surgeon and the rep believe that skiving caused the inaccuracy. The screw was placed by hand and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE KIT VER-4.0.0 | 2019-May-22 SCL1901311 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that after all of the trajectories were sent, accuracy was checked using fluoroscopic imaging, and the L2 wire on the left was lateral to the plan. The surgeon and the rep believe that skiving caused the inaccuracy. The screw was placed by hand and there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x deviation spinal procedure trajectories sent accuracy checked using fluoroscopic imaging wire left lateral plan surgeon believe skiving caused inaccuracy screw placed hand additional complications anticipated mazor x software kit | NAV5000 | NAV4091 | NAV2047 | NAV1006 | NAV3264 | NAV5000NAV4091NAV2047NAV1006NAV3264 | True
1943 | 703222722 | X - Accessory Issue | 2019-06-24T11:25:57.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2202-01 | CLAMP LINK | X - Accessory Issue Screw is stripped on Clamp Accessory piece MAS2202 CLAMP LINK | Screw is stripped on Clamp Accessory piece MAS2202 | x accessory issue screw stripped clamp accessory piece clamp link | NAV3299 | NAV3295 | NAV4090 | NAV4091 | NAV3242 | NAV3299NAV3295NAV4090NAV4091NAV3242 | True
3446 | 703351899 | XSE - loosing pressure/3Define not work | 2019-09-09T15:45:34.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 40 | ASM0206-01 | SURGICAL ARM | XSE - loosing pressure/3Define not work 2019-Sep-04 SCL1902597 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the compressor builds up to 7.2 and then it drops to 5.2 within 2 minutes and then the compressor comes on again. It was noted that the compressor came on far too often. 3Define camera was also not functioning. There was no patient harm or additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Sep-04 SCL1902597 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the compressor builds up to 7.2 and then it drops to 5.2 within 2 minutes and then the compressor comes on again. It was noted that the compressor came on far too often. 3Define camera was also not functioning. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse loosing work spinal procedure compressor builds drops minutes compressor comes noted compressor came far camera functioning additional complications anticipated surgical arm | NAV3264 | NAV3244 | NAV1006 | NAV3249 | NAV4093 | NAV3264NAV3244NAV1006NAV3249NAV4093 | False
8814 | 704491743 | XSE - Bed Frame Adapter Screws Lost | 2021-07-08T20:03:54.000000+0000 | MazorX Stealth Edition | NAV3255 | BED FRAME ISSUES | 20 | KIT1113 | BED FRAME ADAPTER UPGRADE KIT KIT1113 | XSE - Bed Frame Adapter Screws Lost 2021-Jul-08 Salesforce C00229819 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws for the bed frame adapter were missing. There was no patient involvement. 2021-07-20 IUD (Other): No new information. 2021-08-03 IUD (Other): No new information. BED FRAME ADAPTER UPGRADE KIT KIT1113 | 2021-Jul-08 Salesforce C00229819 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws for the bed frame adapter were missing. There was no patient involvement. 2021-07-20 IUD (Other): No new information. 2021-08-03 IUD (Other): No new information. | xse bed frame adapter screws lost salesforce outside procedure screws bed frame adapter missing iud new iud new bed frame adapter upgrade kit | NAV3255 | NAV2053 | NAV3246 | NAV3249 | NAV2054 | NAV3255NAV2053NAV3246NAV3249NAV2054 | True
8147 | 704361679 | XSE - SRA Deviated/Not on Trajectory | 2021-04-26T20:22:40.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - SRA Deviated/Not on Trajectory 2021-Apr-21 Salesforce C00225302, CCR, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the surgical arm was sent to left L5 after completing setup of the system. The surgeon compared the trajectory of the surgical arm to the images an the C-arm and the arm was too far from the plan. The trajectory was lower/inferior to the plan. The main monitor showed that the surgical arm was on trajectory. Prior to the case, the manufacturer representative completed the setup and 10 point advanced accuracy test and the system passed the tests. The representative suspected the deviation could be due to skiving or movement of the patient or surgical system. The representative also mentioned the deviation could have been caused by poor registration, but they did not see anything wrong and no errors were displayed. The system was used the following day without any issue. The surgeon decided to abort the use of the guidance system and place the screws freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-05-11 IUD (Other): No new information. 2021-06-03 email 2021-06-03 05:46 (Rep): Additional information received from a manufacturer representative reported that it was unknown how much the trajectories were deviated. MAZOR X SYSTEM TPL0059 | 2021-Apr-21 Salesforce C00225302, CCR, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the surgical arm was sent to left L5 after completing setup of the system. The surgeon compared the trajectory of the surgical arm to the images an the C-arm and the arm was too far from the plan. The trajectory was lower/inferior to the plan. The main monitor showed that the surgical arm was on trajectory. Prior to the case, the manufacturer representative completed the setup and 10 point advanced accuracy test and the system passed the tests. The representative suspected the deviation could be due to skiving or movement of the patient or surgical system. The representative also mentioned the deviation could have been caused by poor registration, but they did not see anything wrong and no errors were displayed. The system was used the following day without any issue. The surgeon decided to abort the use of the guidance system and place the screws freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-05-11 IUD (Other): No new information. 2021-06-03 email 2021-06-03 05:46 (Rep): Additional information received from a manufacturer representative reported that it was unknown how much the trajectories were deviated. | xse sra deviatednot trajectory salesforce ccr iud spinal procedure treat degenerative discs fixation surgical arm sent left completing setup surgeon compared trajectory surgical arm images carm arm far plan trajectory lowerinferior plan main monitor showed surgical arm trajectory prior case completed setup point advanced accuracy test passed tests suspected deviation skiving movement surgical mentioned deviation caused poor registration did wrong errors displayed following day issue surgeon decided abort use place screws freehand procedure delayed hour iud new email additional unknown trajectories deviated mazor x | NAV5000 | NAV3243 | NAV3249 | NAV4095 | NAV4097 | NAV5000NAV3243NAV3249NAV4095NAV4097 | True
1911 | 703220991 | X- WS Issues | 2019-06-21T18:59:21.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- WS Issues workstation handle is very lose and needs to be fixed UNK MAZOR X WORKSTATION | workstation handle is very lose and needs to be fixed | x ws issues workstation handle lose needs fixed unk mazor x workstation | NAV2058 | NAV3264 | NAV2047 | NAV3256 | NAV3261 | NAV2058NAV3264NAV2047NAV3256NAV3261 | True
1667 | 703208472 | X - ACCESSORY ISSUES/SW ISSUES | 2019-06-15T11:39:50.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 50 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - ACCESSORY ISSUES/SW ISSUES 2018-02-01 CCR (Rep, HCP) K-wire skived out of the pedicle due to the 30mm stop long drill not being able to break through the cortical wall. Surgeon really had to hammer the K-wire in and it finally gave way and breached the lateral wall. Surgeon requests a longer drill – 45mm long to break through the cortical wall./ 2.1.18 Dr. Skidmore L4-L5 TLIF. preop: Cameras checked, AA completed, shoulder checked Plan: RTC planned with surgeon, surgeon agreed with the plan, discussed the patient's rotation and difficulties of getting good images. Arm attached to bed -1 person Arm Draped -1 person Schanz pin placed in left PSIS BMU locked at 15 degrees and connected to shanz ball connector-3define completed on the first try -clear-Drawspine completed on the first try labeled L5. Registration: Registered L4-L5 on the first try with green values and visually verified by surgeon. Arm sent to mark the skin incisions. Operation: Arm sent to first trajectory and placed the cannula and disector down the arm tube, lateral skive was present, and tools seemed to have skived. lateral image was the tools looked way off, RTC said it was a lateral skive based on the plan. surgeon didn't think it was a skive, he thought that the registration may have been off. Surgeon wanted to register and start fresh -RTC agreed.-RTC recalibrated C-arm as mag1 was used after registration.-New AP and new OBL were taken -registered with green values and visually verified by surgeon. -Arm sent to same trajectory -same issue, the tools seemed superior and deeper than the plan.-Wig-wag was used to get clearer images.-RTC told surgeon that we should move to L5 and that L4 plan will get changed.-L5 right k-wire placed RTC changed the plan at L4 and arm resent -much better purchase this time and no skive.-Surgeon placed k-wire. -Left side trajectories were placed. -Patient had really hard, dense bone, -SUreon used 3x80 long drill to get through the cortical wall-Surgeon did not like that he had to use this long of a drill to get through the cortical wall, felt that this was unsafe. -surgeon requests longer positive stop drill of 45mm drill to get through the cortical wall, said when planning on the facets it is rare break through the wall. -Took AP image and L5 right appeared to be lateral. Surgeon thought this was odd since it felt likesuch a good purchase. -RTC pulled up the plan, and noticed that the drill depth was several mm short of the cortical wall.-Surgeon had the hardest time breaking through the cortical wall with the k-wire and finally after really hammering it in, the k-wire gave RTC discussed with surgeon that the K-wire likley skived off the cortical wall and went out lateral.-Su rgeon agreed.-Surgeon wanted to adjust the plan however and resend to be sure. -Surgeon used the long drill, followed by the 3x80 long drill to break through the cortical wall. (This same issues happened the day before with another surgeon at the hospital). -Surgeon was happy with the use of the robot, Surgeon said he had never seen a patient with such hard bone. After removing the robot from the field, RTC had a hard time getting the robot to storage position and under the hood. Every movement the workstation would beep with a warning message about D5. Surgeon thought that there was an issue with the robot because of the alerts and requested that I turn the volume down. *D5 warning message needs to be removed from the software, does nothing to help the rep, only makes you look foolish putting the system away because of all the warning messages. UNK MAZOR X ARM | 2018-02-01 CCR (Rep, HCP) K-wire skived out of the pedicle due to the 30mm stop long drill not being able to break through the cortical wall. Surgeon really had to hammer the K-wire in and it finally gave way and breached the lateral wall. Surgeon requests a longer drill – 45mm long to break through the cortical wall./ 2.1.18 Dr. Skidmore L4-L5 TLIF. preop: Cameras checked, AA completed, shoulder checked Plan: RTC planned with surgeon, surgeon agreed with the plan, discussed the patient's rotation and difficulties of getting good images. Arm attached to bed -1 person Arm Draped -1 person Schanz pin placed in left PSIS BMU locked at 15 degrees and connected to shanz ball connector-3define completed on the first try -clear-Drawspine completed on the first try labeled L5. Registration: Registered L4-L5 on the first try with green values and visually verified by surgeon. Arm sent to mark the skin incisions. Operation: Arm sent to first trajectory and placed the cannula and disector down the arm tube, lateral skive was present, and tools seemed to have skived. lateral image was the tools looked way off, RTC said it was a lateral skive based on the plan. surgeon didn't think it was a skive, he thought that the registration may have been off. Surgeon wanted to register and start fresh -RTC agreed.-RTC recalibrated C-arm as mag1 was used after registration.-New AP and new OBL were taken -registered with green values and visually verified by surgeon. -Arm sent to same trajectory -same issue, the tools seemed superior and deeper than the plan.-Wig-wag was used to get clearer images.-RTC told surgeon that we should move to L5 and that L4 plan will get changed.-L5 right k-wire placed RTC changed the plan at L4 and arm resent -much better purchase this time and no skive.-Surgeon placed k-wire. -Left side trajectories were placed. -Patient had really hard, dense bone, -SUreon used 3x80 long drill to get through the cortical wall-Surgeon did not like that he had to use this long of a drill to get through the cortical wall, felt that this was unsafe. -surgeon requests longer positive stop drill of 45mm drill to get through the cortical wall, said when planning on the facets it is rare break through the wall. -Took AP image and L5 right appeared to be lateral. Surgeon thought this was odd since it felt likesuch a good purchase. -RTC pulled up the plan, and noticed that the drill depth was several mm short of the cortical wall.-Surgeon had the hardest time breaking through the cortical wall with the k-wire and finally after really hammering it in, the k-wire gave RTC discussed with surgeon that the K-wire likley skived off the cortical wall and went out lateral.-Su rgeon agreed.-Surgeon wanted to adjust the plan however and resend to be sure. -Surgeon used the long drill, followed by the 3x80 long drill to break through the cortical wall. (This same issues happened the day before with another surgeon at the hospital). -Surgeon was happy with the use of the robot, Surgeon said he had never seen a patient with such hard bone. After removing the robot from the field, RTC had a hard time getting the robot to storage position and under the hood. Every movement the workstation would beep with a warning message about D5. Surgeon thought that there was an issue with the robot because of the alerts and requested that I turn the volume down. *D5 warning message needs to be removed from the software, does nothing to help the rep, only makes you look foolish putting the system away because of all the warning messages. | x accessory issuessw issues ccr kwire skived pedicle stop long drill able break cortical wall surgeon really hammer kwire finally gave way breached lateral wall surgeon requests longer drill – long break cortical wall dr skidmore tlif preop cameras checked completed shoulder checked plan rtc planned surgeon surgeon agreed plan discussed patients rotation difficulties getting good images arm attached bed person arm draped person schanz pin placed left psis bmu locked degrees connected shanz ball completed try cleardrawspine completed try labeled registration registered try green values visually verified surgeon arm sent mark skin incisions operation arm sent trajectory placed cannula disector arm tube lateral skive present tools skived lateral image tools looked way rtc said lateral skive based plan surgeon didnt think skive thought registration surgeon wanted register start fresh rtc agreedrtc recalibrated carm registrationnew ap new obl taken registered green values visually verified surgeon arm sent trajectory issue tools superior deeper planwigwag clearer imagesrtc told surgeon plan right kwire placed rtc changed plan arm resent better purchase time skivesurgeon placed kwire left trajectories placed really hard dense bone sureon long drill cortical wallsurgeon did like use long drill cortical wall felt unsafe surgeon requests longer positive stop drill drill cortical wall said planning facets rare break wall took ap image right appeared lateral surgeon thought odd felt likesuch good purchase rtc pulled plan noticed drill depth mm short cortical wallsurgeon hardest time breaking cortical wall kwire finally really hammering kwire gave rtc discussed surgeon kwire likley skived cortical wall went lateralsu rgeon agreedsurgeon wanted adjust plan resend sure surgeon long drill followed long drill break cortical wall surgeon happy use robot surgeon said seen hard bone removing robot field rtc hard time getting robot storage position hood movement workstation beep warning message surgeon thought issue robot alerts requested turn volume warning message needs removed software does help makes look foolish putting away warning messages unk mazor x arm | NAV3233 | NAV4095 | NAV3242 | NAV3241 | NAV3299 | NAV3233NAV4095NAV3242NAV3241NAV3299 | True
9058 | 704542816 | XSE - Manipulator Button Not Function | 2021-08-05T20:21:15.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 30 | KIT0927-01 | SVC KIT0927-01 MANIPULATOR BUTTON | XSE - Manipulator Button Not Function 2021-Aug-05 Salesforce C00231697, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the left manipulator button was not working properly. There was no patient involvement. 2021-08-10 IUD x3 (Other): No new information. 2021-09-12 IUD (Other): No new information. SVC KIT0927-01 MANIPULATOR BUTTON | 2021-Aug-05 Salesforce C00231697, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the left manipulator button was not working properly. There was no patient involvement. 2021-08-10 IUD x3 (Other): No new information. 2021-09-12 IUD (Other): No new information. | xse manipulator button function salesforce iud outside p rocedure left manipulator button working properly iud new iud new svc manipulator button | NAV3260 | NAV2054 | NAV3249 | NAV3264 | NAV3251 | NAV3260NAV2054NAV3249NAV3264NAV3251 | True
1278 | 703195216 | X - DEVIATION | 2019-06-07T14:47:58.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - DEVIATION Surgeon felt that S1 left was lateral, but it looked identical to plan. Surgeon didn’t want to trust the robot and decided to freehand that trajectory. All robotic k-wires stimmed at 20, his free hand trajectory stimmed at 6. /5.4.18 Bailey PLIF L5-S1. PLAN: Plan reviewed with surgeon prior to case. PREOP: AA completed, cameras and shoulder calibration checked. BED ATTACHMENT: 1 person-Robot attached to bed with no issues. PATIENT ATTACHMENT: Shanz pin in patient's right iliac crest attached to BMB. REGISTRATION: 2 AP images taken, 2 oblique images taken. Basic registration successful for levels L3-S2. OPERATION: Robot sent to mark all trajectories. Surgeon exposed on left side and arm was resent for instrumentation. AP image taken to verify K wire placement and surgeon questioning S1 on the left, however K wire trajectory matched that of the surgical plan. Removed K wire and chose tofree hand. Surgeon exposed on right side, arm resent for instrumentation. K wire placement again verified with AP image. All robot placed K wires stimmed at or above 20,k wire free handed by surgeon stimmed at 6. Surgeon placed screws and proceeded with the decompression. REMOVAL: Robot removed from the bed with no issues. MAZOR X SOFTWARE KIT VER-3.0.2 | Surgeon felt that S1 left was lateral, but it looked identical to plan. Surgeon didn’t want to trust the robot and decided to freehand that trajectory. All robotic k-wires stimmed at 20, his free hand trajectory stimmed at 6. /5.4.18 Bailey PLIF L5-S1. PLAN: Plan reviewed with surgeon prior to case. PREOP: AA completed, cameras and shoulder calibration checked. BED ATTACHMENT: 1 person-Robot attached to bed with no issues. PATIENT ATTACHMENT: Shanz pin in patient's right iliac crest attached to BMB. REGISTRATION: 2 AP images taken, 2 oblique images taken. Basic registration successful for levels L3-S2. OPERATION: Robot sent to mark all trajectories. Surgeon exposed on left side and arm was resent for instrumentation. AP image taken to verify K wire placement and surgeon questioning S1 on the left, however K wire trajectory matched that of the surgical plan. Removed K wire and chose tofree hand. Surgeon exposed on right side, arm resent for instrumentation. K wire placement again verified with AP image. All robot placed K wires stimmed at or above 20,k wire free handed by surgeon stimmed at 6. Surgeon placed screws and proceeded with the decompression. REMOVAL: Robot removed from the bed with no issues. | x deviation surgeon felt left lateral looked identical plan surgeon didn’t want trust robot decided freehand trajectory robotic kwires stimmed free hand trajectory stimmed bailey plif plan plan reviewed surgeon prior case preop completed cameras shoulder calibration checked bed attachment personrobot attached bed issues attachment shanz pin patients right iliac crest attached bmb registration ap images taken oblique images taken basic registration successful levels operation robot sent mark trajectories surgeon exposed left arm resent instrumentation ap image taken verify k wire placement surgeon questioning left k wire trajectory matched surgical plan removed k wire chose tofree hand surgeon exposed right arm resent instrumentation k wire placement verified ap image robot placed k wires stimmed wire free handed surgeon stimmed surgeon placed screws proceeded decompression removal robot removed bed issues mazor x software kit | NAV4091 | NAV3242 | NAV5000 | NAV4112 | NAV3250 | NAV4091NAV3242NAV5000NAV4112NAV3250 | True
7706 | 704252657 | XSE - Breakers Tripped | 2021-02-26T21:09:18.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Breakers Tripped 2021-Feb-26 Salesforce C00221783, IUD (rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the controller is not connecting on the system. The 4A and 6A breakers are tripping and needs repaired quickly for a demo this weekend. No patient present. 2021-03-16 IUD x2 (Other): No new information. 2021-03-31 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-Feb-26 Salesforce C00221783, IUD (rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the controller is not connecting on the system. The 4A and 6A breakers are tripping and needs repaired quickly for a demo this weekend. No patient present. 2021-03-16 IUD x2 (Other): No new information. 2021-03-31 IUD (Other): No new information. | xse breakers tripped salesforce iud outside proc edure controller connecting breakers tripping needs repaired quickly demo weekend present iud new iud new kit svc cart elec panel assy | NAV2058 | NAV2047 | NAV3264 | NAV3249 | NAV3252 | NAV2058NAV2047NAV3264NAV3249NAV3252 | True
5784 | 703802690 | XSE - Lost Blunt Trocar | 2020-06-15T19:53:47.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MAS2255-05S | KIT SVC MAS2255-05S BLUNT | XSE - Lost Blunt Trocar 2020-06-15 00690857 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the blunt trocar was lost by the site. There was no patient involvement. 2020-06-24 IUD (Other): Additional information was received stating that the site found the part. KIT SVC MAS2255-05S BLUNT | 2020-06-15 00690857 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the blunt trocar was lost by the site. There was no patient involvement. 2020-06-24 IUD (Other): Additional information was received stating that the site found the part. | xse lost blunt trocar outside procedure blunt trocar lost site iud additional stating site kit svc blunt | NAV2053 | NAV2054 | NAV3249 | NAV2047 | NAV1006 | NAV2053NAV2054NAV3249NAV2047NAV1006 | True
4097 | 703498693 | XSE - Joint Shift Errors | 2019-12-04T03:19:50.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | XSE - Joint Shift Errors 2019-Dec-03 00679516 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a L3-L5 MIS fusion, the system was experiencing joint shift errors. The rep had set the system up and performed the arm set up an put the system in the docked position in the hood with the hood closed. The surgeon was performing his inter body work so the rep had left the room. When he got back, the system was displaying a error message stating that joint shift had been detected. The rep logged into superuser and reset movement of the arm. He then sent the arm to clear up and the error popped up again. He repeated and sent the arm to storage position. The error popped up again. He then hard reset the system. Upon reboot, on the log in screen, the error showed again. At this point, the rep went and got the other system at the account to continue with the case. This caused no delay, as the surgeon was not ready to use the robot at this point. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-12-11 IUD x8 (Other): No new information. 2020-02-18 IUD x2 (Other): No new information. 2020-03-22 IUD x2 (Other): No new information. SURGICAL ARM | 2019-Dec-03 00679516 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a L3-L5 MIS fusion, the system was experiencing joint shift errors. The rep had set the system up and performed the arm set up an put the system in the docked position in the hood with the hood closed. The surgeon was performing his inter body work so the rep had left the room. When he got back, the system was displaying a error message stating that joint shift had been detected. The rep logged into superuser and reset movement of the arm. He then sent the arm to clear up and the error popped up again. He repeated and sent the arm to storage position. The error popped up again. He then hard reset the system. Upon reboot, on the log in screen, the error showed again. At this point, the rep went and got the other system at the account to continue with the case. This caused no delay, as the surgeon was not ready to use the robot at this point. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-12-11 IUD x8 (Other): No new information. 2020-02-18 IUD x2 (Other): No new information. 2020-03-22 IUD x2 (Other): No new information. | xse joint shift errors spinal procedure mis fusion experiencing joint shift errors set performed arm set docked position hood hood closed surgeon performing inter body work left room got displaying error message stating joint shift detected logged superuser reset movement arm sent arm clear error popped repeated sent arm storage position error popped hard reset reboot log screen error showed point went got account continue case caused delay surgeon ready use robot point additional complications anticipated iud new iud new iud new surgical arm | NAV4098 | NAV4095 | NAV2058 | NAV4092 | NAV4108 | NAV4098NAV4095NAV2058NAV4092NAV4108 | True
4750 | 703573596 | XSE - Surgical Arm Issue | 2020-01-21T20:06:00.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | XSE - Surgical Arm Issue 2020-1-21 00682442(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system went to the OBL position for registration. The manufacturer representative tried to move the arm closer to the patient, but the system gave an error saying that the surgical arm could not reach the desired location. The representative took the OBL image from where the arm was, and the case proceeded. There was no patient harm and the procedure was not delayed over an hour. 2020-01-21 IUD x3 (Other): No new information. SURGICAL ARM | 2020-1-21 00682442(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system went to the OBL position for registration. The manufacturer representative tried to move the arm closer to the patient, but the system gave an error saying that the surgical arm could not reach the desired location. The representative took the OBL image from where the arm was, and the case proceeded. There was no patient harm and the procedure was not delayed over an hour. 2020-01-21 IUD x3 (Other): No new information. | xse surgical arm issue spinal procedure went obl position registration tried arm closer gave error saying surgical arm reach desired location took obl image arm case proceeded procedure delayed hour iud new surgical arm | NAV4095 | NAV3249 | NAV4101 | NAV4098 | NAV3242 | NAV4095NAV3249NAV4101NAV4098NAV3242 | True
9660 | 704691999 | XSE - OOBF/Faulty Positioner | 2021-11-01T18:27:32.000000+0000 | MazorX Stealth Edition | NAV4105 | UNEXPECTED POSITIONER MOVEMENT | 20 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - OOBF/Faulty Positioner 2021-Oct-26 Salesforce C00237108, e1, e2, e3, e4, e5 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner kept going up by itself at install due to a suspected faulty spring configuration. the manufacturer representative stated the position would not stay in a low position and was "jumping." The holes for the screws in the USB connection at the bottom of the positioner were not completely clean or empty, which made getting the cable firmly connected difficult. There was no patient involvement. 2021-11-10 e6, e7 (Rep, For): No new information. 2021-12-01 e8 (Other, For): No new information. 2021-12-06 IUD, e9, e10 (Other, For): No new information. 2021-12-24 IUD (Other, For): No new information. 2022-01-10 e11 (Other, For): No new information. 2022-01-25 IUD x2 (Other, For): No new information. 2022-04-03 e12, e13, e14, e15 (Other, For): No new information. 2022-04-19 email 2022-04-19 12:45, email 2022-04-20 06:45 (Rep, For): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-Oct-26 Salesforce C00237108, e1, e2, e3, e4, e5 (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner kept going up by itself at install due to a suspected faulty spring configuration. the manufacturer representative stated the position would not stay in a low position and was "jumping." The holes for the screws in the USB connection at the bottom of the positioner were not completely clean or empty, which made getting the cable firmly connected difficult. There was no patient involvement. 2021-11-10 e6, e7 (Rep, For): No new information. 2021-12-01 e8 (Other, For): No new information. 2021-12-06 IUD, e9, e10 (Other, For): No new information. 2021-12-24 IUD (Other, For): No new information. 2022-01-10 e11 (Other, For): No new information. 2022-01-25 IUD x2 (Other, For): No new information. 2022-04-03 e12, e13, e14, e15 (Other, For): No new information. 2022-04-19 email 2022-04-19 12:45, email 2022-04-20 06:45 (Rep, For): No new information. | xse oobffaulty positioner salesforce outside procedure positioner kept going install suspected faulty spring configuration stated position stay low position jumping holes screws usb connection positioner completely clean getting cable firmly connected difficult new new iud new iud new new iud new new email email new kit svc surgical syst type | NAV3251 | NAV4104 | NAV2054 | NAV3250 | NAV2047 | NAV3251NAV4104NAV2054NAV3250NAV2047 | False
6634 | 703994640 | X - Joint 6 Error | 2020-10-05T20:09:44.000000+0000 | MazorX | NAV3249 | ARM ISSUES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | X - Joint 6 Error 2020-10-05 00698133 (HCP, Rep): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a single level L4-L5 fusion revision there was an error at joint 6 when attempting to move to their first predefined target. Registration went smoothly and they were success in moving to both the left and right of each the L4 and L5. It was after doing this that the experienced the error. Several reboots were attempted, logged into super user to reset with no resolution. The surgeon decided to abort the use of the guidance system and do the procedure by hand. Patient present, 25 minute delay. There was no impact on the patient outcome. 2020-10-08 IUD x5 (Other): No new information. 2020-11-10 IUD x3 (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-10-05 00698133 (HCP, Rep): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that during a single level L4-L5 fusion revision there was an error at joint 6 when attempting to move to their first predefined target. Registration went smoothly and they were success in moving to both the left and right of each the L4 and L5. It was after doing this that the experienced the error. Several reboots were attempted, logged into super user to reset with no resolution. The surgeon decided to abort the use of the guidance system and do the procedure by hand. Patient present, 25 minute delay. There was no impact on the patient outcome. 2020-10-08 IUD x5 (Other): No new information. 2020-11-10 IUD x3 (Other): No new information. | x joint error intraoperatively spinal procedure single level fusion revision error joint attempting predefined target registration went smoothly success moving left right doing experienced error reboots attempted logged super user reset resolution surgeon decided abort use procedure hand present minute delay impact outcome iud new iud new kit svc surgical arm refurb | NAV3249 | NAV1006 | NAV4098 | NAV4095 | NAV2047 | NAV3249NAV1006NAV4098NAV4095NAV2047 | True
2415 | 703235109 | X - Pressure issue | 2019-06-29T18:38:16.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Pressure issue Pressure not building in system. Compressor turns on as normal. No leaks or indications of leaks found by the customer onsite. While inspecting the valves on the back of the pneumatic panel an abnormal grinding sound was heard coming from one of the valves. C ustomer tapped on the valve in an attempt to dislodge it but was unsuccessful in correcting the issue. UNK MAZOR X WORKSTATION | Pressure not building in system. Compressor turns on as normal. No leaks or indications of leaks found by the customer onsite. While inspecting the valves on the back of the pneumatic panel an abnormal grinding sound was heard coming from one of the valves. C ustomer tapped on the valve in an attempt to dislodge it but was unsuccessful in correcting the issue. | x pressure issue pressure building compressor turns normal leaks indications leaks customer onsite inspecting valves pneumatic panel abnormal grinding sound heard coming valves c ustomer tapped valve attempt dislodge unsuccessful correcting issue unk mazor x workstation | NAV3264 | NAV3250 | NAV3244 | NAV3261 | NAV2054 | NAV3264NAV3250NAV3244NAV3261NAV2054 | True
2561 | 703241285 | X - surgical arm issue | 2019-07-03T16:50:18.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - surgical arm issue During a staff in service, we experienced the following: 1. When sending arm to AP, it kept stopping after 2-3 seconds and we had to keep resending 2. 3 Define timed out twice in a row - went to default. 3. Arm reached clear up, but kept chiming like it was moving - SW froze and had to restart. 4. Hit E-stop and arm continued to move. UNK MAZOR X ARM | During a staff in service, we experienced the following: 1. When sending arm to AP, it kept stopping after 2-3 seconds and we had to keep resending 2. 3 Define timed out twice in a row - went to default. 3. Arm reached clear up, but kept chiming like it was moving - SW froze and had to restart. 4. Hit E-stop and arm continued to move. | x surgical arm issue staff service experienced following sending arm ap kept stopping seconds resending define timed twice row went default arm reached clear kept chiming like moving sw froze restart hit estop arm continued unk mazor x arm | NAV4101 | NAV4096 | NAV4095 | NAV3244 | NAV2058 | NAV4101NAV4096NAV4095NAV3244NAV2058 | True
3498 | 703361723 | XSE - Broken NavLock | 2019-09-13T22:05:29.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | 9734335 | ADAPTER 9734335 POWEREASE | XSE - Broken NavLock 2019-08-20 MPXR 653940, 00674296 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when the surgeon was placing the third screw they could hear a friction sound as the drill turned. A few seconds later the NavLock and drill bit locked together. The NavLock broke and the adapter began to spin as the drill bit did. There was no patient harm. 2019-09-20 IUD (Other): No new information. 2019-09-23 email 2019-09-23 08:15, IUD (Rep): Additional information received from a manufacturer representative reported the drill bit had to be malleted off of the tracker after the procedure. The surgeon was able to switch to long tools to complete the procedure. The representative thought the binding of the tracker was caused by the surgeon having to provide additional force to start the drill bit and the force probably not being co-linear to the trajectory leading to additional friction. ADAPTER 9734335 POWEREASE | 2019-08-20 MPXR 653940, 00674296 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that when the surgeon was placing the third screw they could hear a friction sound as the drill turned. A few seconds later the NavLock and drill bit locked together. The NavLock broke and the adapter began to spin as the drill bit did. There was no patient harm. 2019-09-20 IUD (Other): No new information. 2019-09-23 email 2019-09-23 08:15, IUD (Rep): Additional information received from a manufacturer representative reported the drill bit had to be malleted off of the tracker after the procedure. The surgeon was able to switch to long tools to complete the procedure. The representative thought the binding of the tracker was caused by the surgeon having to provide additional force to start the drill bit and the force probably not being co-linear to the trajectory leading to additional friction. | xse broken navlock mpxr spinal p rocedure surgeon placing screw hear friction sound drill turned seconds later navlock drill bit locked navlock broke adapter began spin drill bit did iud new email iud additional drill bit malleted tracker procedure surgeon able switch long tools complete procedure thought binding tracker caused surgeon having provide additional force start drill bit force probably colinear trajectory leading additional friction adapter powerease | NAV5000 | NAV1006 | NAV2054 | NAV5002 | NAV3249 | NAV5000NAV1006NAV2054NAV5002NAV3249 | True
4015 | 703488573 | XSE - Mult. Issues Surg System Replaced | 2019-11-26T20:46:11.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 50 | ASM0206-01 | SURGICAL ARM | XSE - Mult. Issues Surg System Replaced 2019-07-12 SCL1801616 (Rep) Medtronic received information regarding a guidance system being serviced. It was reported that the system had previously been upgraded for navigation and field service was installing additional upgrade of the surgical system. Then observed that the system was having trouble maintaining pressure when the positioner teeth were open. Manufacturer representatives did some trouble shooting: assessed the shoulder to ensure the pneumatic line wasn't pinched or kinked and it was not. The pneumatic line was isolated pressure stabilized but when reconnected to the positioner and teeth were opened, the pressure dropped again and stabilized when teeth were closed. The main cart handle was broken and was replaced it. During troubleshooting, the positioner handle buttons stopped working. Also during troubleshooting and servicing, it was suspected that the surgical screen rotation sensor might be cracked though though the screen did remain functional. The decision was made to replace the surgical system and this was completed. Troubleshooting and all servicing took place between 2019-07-12 through 07-23. Verification of the new system was completed with the exception of testing the arm as there was not an arm installed on 2019-08-23. The event occurred during service of the system and, therefore, no patient was involved. SURGICAL ARM | 2019-07-12 SCL1801616 (Rep) Medtronic received information regarding a guidance system being serviced. It was reported that the system had previously been upgraded for navigation and field service was installing additional upgrade of the surgical system. Then observed that the system was having trouble maintaining pressure when the positioner teeth were open. Manufacturer representatives did some trouble shooting: assessed the shoulder to ensure the pneumatic line wasn't pinched or kinked and it was not. The pneumatic line was isolated pressure stabilized but when reconnected to the positioner and teeth were opened, the pressure dropped again and stabilized when teeth were closed. The main cart handle was broken and was replaced it. During troubleshooting, the positioner handle buttons stopped working. Also during troubleshooting and servicing, it was suspected that the surgical screen rotation sensor might be cracked though though the screen did remain functional. The decision was made to replace the surgical system and this was completed. Troubleshooting and all servicing took place between 2019-07-12 through 07-23. Verification of the new system was completed with the exception of testing the arm as there was not an arm installed on 2019-08-23. The event occurred during service of the system and, therefore, no patient was involved. | xse mult issues surg replaced serviced previously upgraded navigation field service installing additional upgrade surgical observed having trouble maintaining pressure positioner teeth open representatives did trouble shooting assessed shoulder ensure pneumatic line wasnt pinched kinked pneumatic line isolated pressure stabilized reconnected positioner teeth opened pressure dropped stabilized teeth closed main cart handle broken replaced troubleshooting positioner handle buttons stopped working troubleshooting servicing suspected surgical screen rotation sensor cracked screen did remain functional decision replace surgical completed troubleshooting servicing took place verification new completed exception testing arm arm installed occurred service involved surgical arm | NAV3251 | NAV3264 | NAV3249 | NAV2054 | NAV2058 | NAV3251NAV3264NAV3249NAV2054NAV2058 | False
8896 | 704509095 | XSE - Damaged BNC Cable | 2021-07-19T16:10:13.000000+0000 | MazorX Stealth Edition | NAV3237 | FLUORO IMAGES | 30 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | XSE - Damaged BNC Cable 2021-Jul-19 Salesforce C00230372 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the the C-arm was not communicating with the guidance system. The BNC cable was replaced and the procedure was continued. There was no patient harm and the procedure was delayed less than an hour. 2021-Jul-21 IUDx5 (Rep): Additional information received from a manufacturer representative reported that after replacing the cable, the port on the system was found to be loose and they were still having issues transferring images. 2021-09-22 email 2021-09-22 07:31 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined. UNK MAZOR X ACCESSORIES | 2021-Jul-19 Salesforce C00230372 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the the C-arm was not communicating with the guidance system. The BNC cable was replaced and the procedure was continued. There was no patient harm and the procedure was delayed less than an hour. 2021-Jul-21 IUDx5 (Rep): Additional information received from a manufacturer representative reported that after replacing the cable, the port on the system was found to be loose and they were still having issues transferring images. 2021-09-22 email 2021-09-22 07:31 (Rep): Additional information received from a manufacturer representative reported that the cause was not determined. | xse damaged bnc cable salesforce spinal procedure carm communicating bnc cable replaced procedure continued procedure delayed hour additional replacing cable port loose having issues transferring images email additional cause determined unk mazor x accessories | NAV3256 | NAV2058 | NAV2044 | NAV3235 | NAV3299 | NAV3256NAV2058NAV2044NAV3235NAV3299 | False
3416 | 703342707 | XSE - lost connection | 2019-09-03T19:58:39.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - lost connection 2019-Aug-28 SCL1902500 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that while removing the positioner from the bed, the rep lost all control to the positioner and the system was turned off. The rep returned later that day, turned the system on, and received a failed to connect to the controller. Tech support had the rep remove the right bottom cover to confirm none of the breakers were tripped on the controller and they were not. System was turned off for about 10 mins but before system was turned off received the message "Your PC ran into a problem and needs to restart. We're just collecting some error info, and then we'll restart for you". Once the system shut off and was turned on it worked as designed. There was no patient harm or additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM TPL0059 | 2019-Aug-28 SCL1902500 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that while removing the positioner from the bed, the rep lost all control to the positioner and the system was turned off. The rep returned later that day, turned the system on, and received a failed to connect to the controller. Tech support had the rep remove the right bottom cover to confirm none of the breakers were tripped on the controller and they were not. System was turned off for about 10 mins but before system was turned off received the message "Your PC ran into a problem and needs to restart. We're just collecting some error info, and then we'll restart for you". Once the system shut off and was turned on it worked as designed. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse lost connection spinal procedure removing positioner bed lost control positioner turned returned later day turned failed connect controller tech support remove right cover confirm breakers tripped controller turned mins turned message pc ran problem needs restart just collecting error info restart shut turned worked designed additional complications anticipated mazor x | NAV3251 | NAV2047 | NAV1006 | NAV5000 | NAV3250 | NAV3251NAV2047NAV1006NAV5000NAV3250 | True
4638 | 703555462 | X - Intraop/Registration issues | 2020-01-10T13:47:31.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Intraop/Registration issues 2018-08-20 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on levels L2 – L5. Planning was done prior to the case. Surgeon did three level OLIF with Titanium cages at L2-L3 and L4-L5 and peek at L3-L4. This was to mitigate massive radio-opaque markings in the fluoro imaging. Single schanz arm was used with 120mm schanz in right-ilium. 3Define scan was used and draw spine as well labeling off L4. Registration first attempt segmented all levels through basic with good values at all levels excluding red at L5. Manual adjustment tried twice with no success. Proceeded to do bi-lateral screws at L2-L4. Once completed registration was attempted again focusing only on L4-L5 and L5-S1. Red value was got each time. New images were taken and still failed. Registration was attempted off of L5 and still could not get acceptable values. Finally, after new AP image and manual registration a yellow value was received at L5 with a slight shift. It was adjusted slightly and a yellow value with no shift was got. All trajectories were sent and were placed accurately and verified via fluoro and neuro monitoring. L4 right was sent twice as wire fell out to replace it. No patient harm was reported MAZOR X SYSTEM TPL0059 | 2018-08-20 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on levels L2 – L5. Planning was done prior to the case. Surgeon did three level OLIF with Titanium cages at L2-L3 and L4-L5 and peek at L3-L4. This was to mitigate massive radio-opaque markings in the fluoro imaging. Single schanz arm was used with 120mm schanz in right-ilium. 3Define scan was used and draw spine as well labeling off L4. Registration first attempt segmented all levels through basic with good values at all levels excluding red at L5. Manual adjustment tried twice with no success. Proceeded to do bi-lateral screws at L2-L4. Once completed registration was attempted again focusing only on L4-L5 and L5-S1. Red value was got each time. New images were taken and still failed. Registration was attempted off of L5 and still could not get acceptable values. Finally, after new AP image and manual registration a yellow value was received at L5 with a slight shift. It was adjusted slightly and a yellow value with no shift was got. All trajectories were sent and were placed accurately and verified via fluoro and neuro monitoring. L4 right was sent twice as wire fell out to replace it. No patient harm was reported | x intraopregistration issues ccr spinal procedure case vertebrae levels – planning prior case surgeon did level olif titanium cages peek mitigate massive radioopaque markings fluoro imaging single schanz arm schanz rightilium scan draw spine labeling registration attempt segmented levels basic good values levels excluding red manual adjustment tried twice success proceeded bilateral screws completed registration attempted focusing red value got time new images taken failed registration attempted acceptable values finally new ap image manual registration yellow value slight shift adjusted slightly yellow value shift got trajectories sent placed accurately verified fluoro neuro monitoring right sent twice wire fell replace mazor x | NAV3242 | NAV2051 | NAV4091 | NAV4112 | NAV3244 | NAV3242NAV2051NAV4091NAV4112NAV3244 | True
9199 | 704573762 | XSE ¦ Pneumatic Line Caught | 2021-08-24T17:47:09.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE ¦ Pneumatic Line Caught 2021-Aug-24 Salesforce C00233405 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the pneumatic line by the D5 sensor was getting caught in the flange. No patient was present. 2021-09-08 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Aug-24 Salesforce C00233405 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the pneumatic line by the D5 sensor was getting caught in the flange. No patient was present. 2021-09-08 IUD (Other): No new information. | xse pneumatic line caught salesforce outside procedure pneumatic line sensor getting caught flange present iud new unk mazor x workstation | NAV3264 | NAV2058 | NAV4110 | NAV3261 | NAV2047 | NAV3264NAV2058NAV4110NAV3261NAV2047 | True
9574 | 704675340 | XSE - SRA Set Up Not Recognized | 2021-10-21T18:20:46.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - SRA Set Up Not Recognized 2021-Oct-21 Salesforce C00236606, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the surgical system was attached to the bed and and the surgical arm setup was completed. When the manufacturer representative went to continue setting up the system, the system showed the surgical arm setup was not complete. The representative had to redo the surgical arm setup to resolve the issue. There was no patient harm and the procedure was delayed less than an hour. 2021-10-25 email 2021-10-25 21:30, IUD (Rep): No new information. 2021-11-05 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2021-Oct-21 Salesforce C00236606, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the surgical system was attached to the bed and and the surgical arm setup was completed. When the manufacturer representative went to continue setting up the system, the system showed the surgical arm setup was not complete. The representative had to redo the surgical arm setup to resolve the issue. There was no patient harm and the procedure was delayed less than an hour. 2021-10-25 email 2021-10-25 21:30, IUD (Rep): No new information. 2021-11-05 IUD (Other): No new information. | xse sra set recognized salesforce ccr spinal procedure treat degenerative discs fixation surgical attached bed surgical arm setup completed went continue setting showed surgical arm setup complete redo surgical arm setup resolve issue procedure delayed hour email iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV4098 | NAV4092 | NAV4101 | NAV3249NAV4095NAV4098NAV4092NAV4101 | False
4348 | 703536596 | XSE - BICE/Teeth Wont Unlock | 2019-12-27T19:57:40.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 30 | SLD0506-01 | PNEUMATIC MANIFULD | XSE - BICE/Teeth Wont Unlock 2019-12-27 00681114 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the teeth will not unlock from the bed frame adapter. The teeth lock onto the bed frame adapter without issues. Verified the manual lock was in the unlock position and still it would not unlock. Also tried the override panel and that did not unlock the teeth. No patient was present. 2019-12-27 IUD x9 (Rep): No new information. 2020-03-19 IUD (Other): No new information. 2020-04-22 IUD (Other): No new information. PNEUMATIC MANIFULD | 2019-12-27 00681114 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the teeth will not unlock from the bed frame adapter. The teeth lock onto the bed frame adapter without issues. Verified the manual lock was in the unlock position and still it would not unlock. Also tried the override panel and that did not unlock the teeth. No patient was present. 2019-12-27 IUD x9 (Rep): No new information. 2020-03-19 IUD (Other): No new information. 2020-04-22 IUD (Other): No new information. | xse biceteeth wont unlock outside procedure teeth unlock bed frame adapter teeth lock bed frame adapter issues verified manual lock unlock position unlock tried override panel did unlock teeth present iud new iud new iud new pneumatic manifuld | NAV4103 | NAV4107 | NAV2054 | NAV3255 | NAV2047 | NAV4103NAV4107NAV2054NAV3255NAV2047 | False
3106 | 703265954 | X - Bent Accuracy Pointer | 2019-07-18T14:58:11.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | MAS2285-01 | ACCURACY POINTER | X - Bent Accuracy Pointer The FSE found the accuracy pointer was bent and needed to be replaced. ACCURACY POINTER | The FSE found the accuracy pointer was bent and needed to be replaced. | x bent accuracy pointer fse accuracy pointer bent needed replaced accuracy pointer | NAV3299 | NAV3295 | NAV4091 | NAV2054 | NAV1006 | NAV3299NAV3295NAV4091NAV2054NAV1006 | True
912 | 703180766 | X- Registration Issue | 2019-05-30T13:43:45.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X- Registration Issue Hartford BJI System 27 Arm 33 Dr Esmende L2-3 L3-4 XLIF existing hardware L4-L5 Setup-Shoulder calibration checked Advanced accuracy completed Surgical arm setup completed While attempting 3Define scan there was no image, cleared up and resent 3Define with same results. Unlocked mount and went into utilities, no firmware present for 3Define camera Did soft restart, re-locked mount and completed 3Define Reg #1-Flouro seg w/ complex During matching it was noted that the AP was very poor quality Software matching gave green values at L2 and yellow at L3 Visual verification showed no shift in matching Sent L2 left, accurate Sent L3 Left and trajectory appeared superior on lateral xray Sent L2 right, accurate sent L3 right superior It was decided then to re-register for L3 body new images takes, AP again poor quality, attempted to get better AP, but unable to get much better Flouro seg w/basic yellow value at L3, during visual verification of matching it was noted that there was a major shiftin matching of L3 Took all new images, OBL and AP, this time we were able to get a slightly improved AP Flouro seg w/ basic L3 yellow value Visual ver ification showed no shift at L3 Sent L3 left, accurate Sent L3 right, accurate MAZOR X SOFTWARE KIT VER-3.0.2 | Hartford BJI System 27 Arm 33 Dr Esmende L2-3 L3-4 XLIF existing hardware L4-L5 Setup-Shoulder calibration checked Advanced accuracy completed Surgical arm setup completed While attempting 3Define scan there was no image, cleared up and resent 3Define with same results. Unlocked mount and went into utilities, no firmware present for 3Define camera Did soft restart, re-locked mount and completed 3Define Reg #1-Flouro seg w/ complex During matching it was noted that the AP was very poor quality Software matching gave green values at L2 and yellow at L3 Visual verification showed no shift in matching Sent L2 left, accurate Sent L3 Left and trajectory appeared superior on lateral xray Sent L2 right, accurate sent L3 right superior It was decided then to re-register for L3 body new images takes, AP again poor quality, attempted to get better AP, but unable to get much better Flouro seg w/basic yellow value at L3, during visual verification of matching it was noted that there was a major shiftin matching of L3 Took all new images, OBL and AP, this time we were able to get a slightly improved AP Flouro seg w/ basic L3 yellow value Visual ver ification showed no shift at L3 Sent L3 left, accurate Sent L3 right, accurate | x registration issue hartford bji arm dr esmende xlif existing hardware setupshoulder calibration checked advanced accuracy completed surgical arm setup completed attempting scan image cleared resent results unlocked mount went utilities firmware present camera did soft restart relocked mount completed reg seg w complex matching noted ap poor quality software matching gave green values yellow visual verification showed shift matching sent left accurate sent left trajectory appeared superior lateral xray sent right accurate sent right superior decided reregister body new images takes ap poor quality attempted better ap unable better flouro seg wbasic yellow value visual verification matching noted major shiftin matching took new images obl ap time able slightly improved ap flouro seg w basic yellow value visual ver ification showed shift sent left accurate sent right accurate mazor x software kit | NAV4091 | NAV3246 | NAV3242 | NAV3248 | NAV3247 | NAV4091NAV3246NAV3242NAV3248NAV3247 | False
3496 | 703361556 | X - Water Condensation in WS Screen | 2019-09-13T19:58:21.000000+0000 | MazorX | NAV3261 | MONITOR ISSUES | 20 | ASM0275-01 | RNX-WS-1210 | X - Water Condensation in WS Screen 2019-09-09 SCL1902729, Intake Form (Rep, For): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that there was water condensation on the inside of the workstation screen. There was no patient involved. RNX-WS-1210 | 2019-09-09 SCL1902729, Intake Form (Rep, For): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that there was water condensation on the inside of the workstation screen. There was no patient involved. | x water condensation ws screen outside pr ocedure water condensation inside workstation screen involved | NAV2058 | NAV3253 | NAV3261 | NAV3295 | NAV2054 | NAV2058NAV3253NAV3261NAV3295NAV2054 | True
6933 | 704067670 | Instrument Issues | 2020-11-13T00:31:24.000000+0000 | MazorX Stealth Edition | NAV4090 | BROKEN SCREW | 20 | MAS2188-04S | KIT SVC MAS2188-04S BM BRIDGE | Instrument Issues 2020-11-12 mPXR Report 776639, 00705527 (Rep, HCP): Medtronic received information regarding spinal and navigation instruments intra -operatively of a spinal procedure. It was reported that the square screw that locks the bone mount bridge to the ball/Shanz pin has stripped and needs to be replaced. The gray navlock has consistently displayed inferior and lateral on navigation. The actual instrument used with the Navlock is in the correct place, but the navigation is consistently off with this one Navlock. Also, the screwdriver tip broke off due to hard bone during a case. The broken piece was retrieved and trashed. There were no patient issues related to these events. 2020-11-18 IUD (Other): No new information. 2020-12-04 Attachment 2020-12-04 11:16 (Rep): No new information. 2020-12-23 email 2020-12-23 16:45 (Rep): No new information. 2020-12-28 email 2020-12-28 10:15 (Rep): No new information. 2021-02-01 Attachment 2021-02-01 01:59 (Other): No new information. 2021-02-05 IUD (Other): No new information. 2021-03-10 IUD x3 (Other): No new information. KIT SVC MAS2188-04S BM BRIDGE | 2020-11-12 mPXR Report 776639, 00705527 (Rep, HCP): Medtronic received information regarding spinal and navigation instruments intra -operatively of a spinal procedure. It was reported that the square screw that locks the bone mount bridge to the ball/Shanz pin has stripped and needs to be replaced. The gray navlock has consistently displayed inferior and lateral on navigation. The actual instrument used with the Navlock is in the correct place, but the navigation is consistently off with this one Navlock. Also, the screwdriver tip broke off due to hard bone during a case. The broken piece was retrieved and trashed. There were no patient issues related to these events. 2020-11-18 IUD (Other): No new information. 2020-12-04 Attachment 2020-12-04 11:16 (Rep): No new information. 2020-12-23 email 2020-12-23 16:45 (Rep): No new information. 2020-12-28 email 2020-12-28 10:15 (Rep): No new information. 2021-02-01 Attachment 2021-02-01 01:59 (Other): No new information. 2021-02-05 IUD (Other): No new information. 2021-03-10 IUD x3 (Other): No new information. | instrument issues mpxr report spinal navigation instruments intra operatively spinal procedure square screw locks bone mount bridge ballshanz pin stripped needs replaced gray navlock consistently displayed inferior lateral navigation actual instrument navlock correct place navigation consistently navlock screwdriver tip broke hard bone case broken piece retrieved trashed issues related events iud new attachment new email new email new attachment new iud new iud new kit svc bm bridge | NAV3299 | NAV1006 | NAV5002 | NAV3295 | NAV2047 | NAV3299NAV1006NAV5002NAV3295NAV2047 | False
652 | 703169573 | X - 3D marker connection difficulty | 2019-05-23T14:23:50.000000+0000 | MazorX | NAV3233 | UNUSABLE DISPOSABLES | 20 | KIT0574 | MAZOR X SPINE DISPOSABLE KIT KIT0574 | X - 3D marker connection difficulty 2019-Apr-02 SCL1901258 intake form, msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3D marker would not thread onto the Target Extender, they tried both Target Extenders before opening a new kit. The new 3D marker threaded without trouble and case was without issue after this. The image clamp adaptor would not tighten or loosen properly. There was no patient harm and the delay was less than 5 minutes. There were no additional complications reported or anticipated as a result of the event. MAZOR X SPINE DISPOSABLE KIT KIT0574 | 2019-Apr-02 SCL1901258 intake form, msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the 3D marker would not thread onto the Target Extender, they tried both Target Extenders before opening a new kit. The new 3D marker threaded without trouble and case was without issue after this. The image clamp adaptor would not tighten or loosen properly. There was no patient harm and the delay was less than 5 minutes. There were no additional complications reported or anticipated as a result of the event. | x marker connection difficulty msg spinal procedure marker thread target extender tried target extenders opening new kit new marker threaded trouble case issue image clamp adaptor tighten loosen properly delay minutes additional complications anticipated mazor x spine disposable kit | NAV2054 | NAV3233 | NAV3242 | NAV3299 | NAV3246 | NAV2054NAV3233NAV3242NAV3299NAV3246 | True
6902 | 704061159 | XSE - Dim Surgeon Monitor | 2020-11-10T14:46:08.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE - Dim Surgeon Monitor 2020-Nov-10 00700555 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon monitor was very dim and hard to see. No patient was present. 2020-11-11 IUD x6 (Other): No new information. 2021-02-16 IUD (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2020-Nov-10 00700555 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon monitor was very dim and hard to see. No patient was present. 2020-11-11 IUD x6 (Other): No new information. 2021-02-16 IUD (Other): No new information. | xse dim surgeon monitor outside procedure surgeon monitor dim hard present iud new iud new kit svc srgn screen arm | NAV3253 | NAV2054 | NAV2047 | NAV3249 | NAV4101 | NAV3253NAV2054NAV2047NAV3249NAV4101 | False
6582 | 703980917 | XSE - Damaged Pressure Tubing | 2020-09-28T16:04:06.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 20 | SLD0399-01 | OTHER, PLASTIC TUBING 3MM | XSE - Damaged Pressure Tubing 2020-09-28 00697578, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when putting the surgical system away, the pressure gauge on the manipulator tubing was separated and needed to be repaired. There was no patient involvement. 2020-10-05 IUD (Other): No new information. OTHER, PLASTIC TUBING 3MM | 2020-09-28 00697578, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when putting the surgical system away, the pressure gauge on the manipulator tubing was separated and needed to be repaired. There was no patient involvement. 2020-10-05 IUD (Other): No new information. | xse damaged pressure tubing iud outside procedure putting surgical away pressure gauge manipulator tubing separated needed repaired iud new plastic tubing | NAV3264 | NAV1006 | NAV3249 | NAV2054 | NAV4100 | NAV3264NAV1006NAV3249NAV2054NAV4100 | False
9413 | 704632635 | XSE - Prongs Are Not Sitting Correctly | 2021-09-28T02:05:02.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 20 | HRS0327-01S | KIT SVC HRS0327-01S TEETH LOCKING SENSOR | XSE - Prongs Are Not Sitting Correctly 2021-Sep-27 Salesforce C00235245 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner was catching onto the bedrail and not sitting properly. The prongs were not sitting correctly. No patient was present. 2021-09-30 IUD x2 (Other): No new information. 2021-10-08 IUD x4 (Other): No new information. 2021-10-20 IUD x2 (Other): No new information. 2021-10-29 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 2021-11-21 IUD (Other): No new information. 2021-12-12 IUD (Other): No new information. 2021-12-28 IUD (Other): No new information. 2022-01-09 IUD (Other): No new information. KIT SVC HRS0327-01S TEETH LOCKING SENSOR | 2021-Sep-27 Salesforce C00235245 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the positioner was catching onto the bedrail and not sitting properly. The prongs were not sitting correctly. No patient was present. 2021-09-30 IUD x2 (Other): No new information. 2021-10-08 IUD x4 (Other): No new information. 2021-10-20 IUD x2 (Other): No new information. 2021-10-29 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 2021-11-21 IUD (Other): No new information. 2021-12-12 IUD (Other): No new information. 2021-12-28 IUD (Other): No new information. 2022-01-09 IUD (Other): No new information. | xse prongs sitting correctly salesforce outside procedure positioner catching bedrail sitting properly prongs sitting correctly present iud new iud new iud new iud new iud new iud new iud new iud new iud new kit svc teeth locking sensor | NAV3251 | NAV3250 | NAV3249 | NAV2047 | NAV3299 | NAV3251NAV3250NAV3249NAV2047NAV3299 | True
7843 | 704291439 | XSE - Unexpected Reboot | 2021-03-18T20:01:36.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Unexpected Reboot 2021-Mar-17 Salesforce C00222888 (HCP, Rep): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during operation mode the manufacturer representative switched back to planning and went to plan an interbody, while planning that the system did a soft restart by itself, it took 30 seconds and rebooted up, it did eventually return to same screen. This occurred after executing several screws. This caused about a 30 second delay. There was no impact on the patient outcome. 2021-03-26 IUD x3 (Other): No new information. 2021-03-26 email 2021-03-26 10:49 (Rep): Additional information was received stating that the case was completed successfully once the system rebooted itself. The cause for the issue has not not been determined. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Mar-17 Salesforce C00222888 (HCP, Rep): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that during operation mode the manufacturer representative switched back to planning and went to plan an interbody, while planning that the system did a soft restart by itself, it took 30 seconds and rebooted up, it did eventually return to same screen. This occurred after executing several screws. This caused about a 30 second delay. There was no impact on the patient outcome. 2021-03-26 IUD x3 (Other): No new information. 2021-03-26 email 2021-03-26 10:49 (Rep): Additional information was received stating that the case was completed successfully once the system rebooted itself. The cause for the issue has not not been determined. | xse unexpected reboot salesforce spinal p rocedure operation mode switched planning went plan interbody planning did soft restart took seconds rebooted did eventually return screen occurred executing screws caused second delay impact outcome iud new email additional stating case completed successfully rebooted cause issue determined mazorx sw installation kit | NAV3247 | NAV2056 | NAV3246 | NAV2058 | NAV2047 | NAV3247NAV2056NAV3246NAV2058NAV2047 | True
71 | 702989092 | X - Schanz Bond Stripped Screw | 2019-02-02T18:43:50.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC4530-01 | SLOTTED BALL ARM TARGET SCREW | X - Schanz Bond Stripped Screw 2019-01-28 SCL1900236, Intake 2019-02-02 12:47 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the Schanz bond screws were stripped out. There was no patient harm, the surgery was not delayed more than an hour and the case was completed. SLOTTED BALL ARM TARGET SCREW | 2019-01-28 SCL1900236, Intake 2019-02-02 12:47 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the Schanz bond screws were stripped out. There was no patient harm, the surgery was not delayed more than an hour and the case was completed. | x schanz bond stripped screw spinal procedure schanz bond screws stripped surgery delayed hour case completed slotted ball arm target screw | NAV3299 | NAV3295 | NAV4090 | NAV4100 | NAV4091 | NAV3299NAV3295NAV4090NAV4100NAV4091 | True
2879 | 703258546 | X - Accessory issues | 2019-07-15T12:57:53.000000+0000 | MazorX | NAV3298 | PLATFORM ISSUES | 20 | MAS2188-01 | BM BRIDGE | X - Accessory issues As reported by Dany, THP has issues with their BM bridge and arm guide. BM BRIDGE | As reported by Dany, THP has issues with their BM bridge and arm guide. | x accessory issues dany thp issues bm bridge arm guide bm bridge | NAV3299 | NAV3295 | NAV3298 | NAV3249 | NAV3242 | NAV3299NAV3295NAV3298NAV3249NAV3242 | True
3197 | 703282505 | X - Stripped Screw/Instrument Issues | 2019-07-29T12:37:46.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2271-03 | SCHANZ BOND | X - Stripped Screw/Instrument Issues 2019-07-19 SCL1902153, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a few parts in the instrument trays were in poor condition. The Schanz bond had a stripped out screw. The target extender had difficulty being attached to the 3D marker and required the use of a wrench to attach the marker. The Schanz arm was not able to get a Schanz pin through it freely and easily. There was no patient harm and the procedure was not delayed by more than an hour. The case was able to be completed. 2019-08-27 e1 (Rep): No new information. SCHANZ BOND | 2019-07-19 SCL1902153, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a few parts in the instrument trays were in poor condition. The Schanz bond had a stripped out screw. The target extender had difficulty being attached to the 3D marker and required the use of a wrench to attach the marker. The Schanz arm was not able to get a Schanz pin through it freely and easily. There was no patient harm and the procedure was not delayed by more than an hour. The case was able to be completed. 2019-08-27 e1 (Rep): No new information. | x stripped screwinstrument issues spinal procedure parts instrument trays poor condition schanz bond stripped screw target extender difficulty attached marker required use wrench attach marker schanz arm able schanz pin freely easily procedure delayed hour case able completed new schanz bond | NAV3299 | NAV3295 | NAV4090 | NAV4100 | NAV3246 | NAV3299NAV3295NAV4090NAV4100NAV3246 | True
10171 | 704815198 | XSE - Unable to Complete Registration | 2022-01-18T16:51:37.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Unable to Complete Registration 2022-Jan-17 Salesforce C00241807, C00241761 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the site's system is "having continued problems with failing registration." The surgeon was performing a L5-S1 CT to Fluoro fusion that required 10 registrations before being able to pass and proceed. There was a 45 minute delay. There was no impact on the patient outcome. 2022-01-18 IUD (Other): No new information. 2022-01-25 IUD (Other): No new information. 2022-02-03 IUD (Other): No new information. 2022-02-08 IUD (Other): No new information. 2022-02-28 IUD (Other): No new information. 2022-03-06 IUD x2 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-04-07 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2022-Jan-17 Salesforce C00241807, C00241761 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the site's system is "having continued problems with failing registration." The surgeon was performing a L5-S1 CT to Fluoro fusion that required 10 registrations before being able to pass and proceed. There was a 45 minute delay. There was no impact on the patient outcome. 2022-01-18 IUD (Other): No new information. 2022-01-25 IUD (Other): No new information. 2022-02-03 IUD (Other): No new information. 2022-02-08 IUD (Other): No new information. 2022-02-28 IUD (Other): No new information. 2022-03-06 IUD x2 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-04-07 IUD (Other): No new information. | xse unable complete registration salesforce spinal procedure sites having continued problems failing registration surgeon performing ct fluoro fusion required registrations able pass proceed minute delay impact outcome iud new iud new iud new iud new iud new iud new iud new iud new mazor x | NAV4112 | NAV2047 | NAV3244 | NAV3299 | NAV3243 | NAV4112NAV2047NAV3244NAV3299NAV3243 | False
3545 | 703370913 | XSE - Deviation/Arm Guide Issue | 2019-09-19T15:35:26.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation/Arm Guide Issue 2019-09-13 SCL1902816, SCL1902838, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using fixation. It was reported that a Scan & Plan procedure for a L4/L5 TLIF was planned. Schanz pin was placed in the left PSIS and mounted to the surgical system using the BMB. A 3Define scan and snapshot were successfully completed. After the surgical arm was sent to the star marker position, the arm guide screw was difficult to remove. During the procedure, the first O-arm spin was successful and the star marker was found, but during import of the scan to the workstation an error occurred and the software was running very slowly. The error stated "System moved after scan, choose to continue yes/no." The representative checked no and another O-arm spin was done to continue using navigation. After the surgical arm was resent to the star marker position, the arm guide screw was extremely difficult to remove. A second O-arm spin was successful and the star marker was found. The scan was uploaded to the system successfully and the procedure was continued. Left L4 and L5 trajectories were instrumented and appeared accurate with navigation. During movement to the right side, a "Joint 4 error" was displayed. The error was resolved after troubleshooting and the system passed a 3 point accuracy test using the cannula and dilator. To continue using the system, a third O-arm spin was completed. During review of the spin, it was noted that the left screws were consistently lateral. The representative suspected a platform shift and inspected the image. The image revealed the Schanz pin had possible haloing around the pin indicating movement. The surgeon decided to abort the use of the guidance system and freehand the screws on the left side. After unmounting the surgical system, the representative requested the Schanz arm be checked for stability. The Schanz pin was found to be loose. The arm guide was inspected and found to be scored on the inside, possibly from known issues with NavLocks binding. There was no patient harm, but the procedure was delayed more than an hour due to the system movement error and the surgeon having to freehand the screws. Information unrelated to this event was included in a separate PE; See PE 703361320 for NavLocks binding. 2019-09-20 email 2019-09-20 12:00 (Rep): Additional information received from a manufacturer representative reported the deviation was between 3.5 and 10 mm and were visible on the O-arm spin. Two error messages were displayed. When the system was not moving and the O-arm scan was being uploaded, a "System Movement Error" was displayed. While the surgical arm was moving, a "Joint 4 Error" was displayed. Troubleshooting included reseting movements and completing the surgical arm setup. The representative believed the Schanz pin came loose in the patient due to difficulty passing tools through the arm guide, difficulty removing the PEEK screw from the arm guide, and multiple O-arm spins being required. All of those added more stress to the Schanz pin. The screws were able to be successfully placed freehand and the patient did not experience any symptoms. 2019-09-26 email 2019-09-27 12:43 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2019-09-13 SCL1902816, SCL1902838, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using fixation. It was reported that a Scan & Plan procedure for a L4/L5 TLIF was planned. Schanz pin was placed in the left PSIS and mounted to the surgical system using the BMB. A 3Define scan and snapshot were successfully completed. After the surgical arm was sent to the star marker position, the arm guide screw was difficult to remove. During the procedure, the first O-arm spin was successful and the star marker was found, but during import of the scan to the workstation an error occurred and the software was running very slowly. The error stated "System moved after scan, choose to continue yes/no." The representative checked no and another O-arm spin was done to continue using navigation. After the surgical arm was resent to the star marker position, the arm guide screw was extremely difficult to remove. A second O-arm spin was successful and the star marker was found. The scan was uploaded to the system successfully and the procedure was continued. Left L4 and L5 trajectories were instrumented and appeared accurate with navigation. During movement to the right side, a "Joint 4 error" was displayed. The error was resolved after troubleshooting and the system passed a 3 point accuracy test using the cannula and dilator. To continue using the system, a third O-arm spin was completed. During review of the spin, it was noted that the left screws were consistently lateral. The representative suspected a platform shift and inspected the image. The image revealed the Schanz pin had possible haloing around the pin indicating movement. The surgeon decided to abort the use of the guidance system and freehand the screws on the left side. After unmounting the surgical system, the representative requested the Schanz arm be checked for stability. The Schanz pin was found to be loose. The arm guide was inspected and found to be scored on the inside, possibly from known issues with NavLocks binding. There was no patient harm, but the procedure was delayed more than an hour due to the system movement error and the surgeon having to freehand the screws. Information unrelated to this event was included in a separate PE; See PE 703361320 for NavLocks binding. 2019-09-20 email 2019-09-20 12:00 (Rep): Additional information received from a manufacturer representative reported the deviation was between 3.5 and 10 mm and were visible on the O-arm spin. Two error messages were displayed. When the system was not moving and the O-arm scan was being uploaded, a "System Movement Error" was displayed. While the surgical arm was moving, a "Joint 4 Error" was displayed. Troubleshooting included reseting movements and completing the surgical arm setup. The representative believed the Schanz pin came loose in the patient due to difficulty passing tools through the arm guide, difficulty removing the PEEK screw from the arm guide, and multiple O-arm spins being required. All of those added more stress to the Schanz pin. The screws were able to be successfully placed freehand and the patient did not experience any symptoms. 2019-09-26 email 2019-09-27 12:43 (Rep): No new information. | xse deviationarm guide issue ccr spinal procedure treat degenerative discs using fixation scan plan procedure tlif planned schanz pin placed left psis mounted surgical using bmb scan snapshot successfully completed surgical arm sent star marker position arm guide screw difficult remove procedure oarm spin successful star marker import scan workstation error occurred software running slowly error stated moved scan choose continue yesno checked oarm spin continue using navigation surgical arm resent star marker position arm guide screw extremely difficult remove second oarm spin successful star marker scan uploaded successfully procedure continued left trajectories instrumented appeared accurate navigation movement right joint error displayed error resolved troubleshooting passed point accuracy test using cannula dilator continue using oarm spin completed review spin noted left screws consistently lateral suspected platform shift inspected image image revealed schanz pin possible haloing pin indicating movement surgeon decided abort use freehand screws left unmounting surgical requested schanz arm checked stability schanz pin loose arm guide inspected scored inside possibly known issues navlocks binding procedure delayed hour movement error surgeon having freehand screws unrelated included separate pe pe navlocks binding email additional deviation mm visible oarm spin error messages displayed moving oarm scan uploaded movement error displayed surgical arm moving joint error displayed troubleshooting included reseting movements completing surgical arm setup believed schanz pin came loose difficulty passing tools arm guide difficulty removing peek screw arm guide multiple oarm spins required added stress schanz pin screws able successfully placed freehand did experience symptoms email new mazor x | NAV3294 | NAV3235 | NAV4098 | NAV3246 | NAV5000 | NAV3294NAV3235NAV4098NAV3246NAV5000 | True
6445 | 703941950 | XSE - Registration Issues | 2020-09-04T21:48:58.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Issues 2020-09-04 00695937 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that they attempted to achieve registration but were unable. They had two levels with yellow screws and one with red screws. They took additional images but they were unable to get registration achieved. The use of navigation was aborted. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-09-08 email 2020-09-08 14:15 x2, IUD (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2020-09-04 00695937 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that they attempted to achieve registration but were unable. They had two levels with yellow screws and one with red screws. They took additional images but they were unable to get registration achieved. The use of navigation was aborted. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-09-08 email 2020-09-08 14:15 x2, IUD (Rep): No new information. | xse registration issues intraoperatively spinal procedure attempted achieve registration unable levels yellow screws red screws took additional images unable registration achieved use navigation aborted issue extended surgical time hour impact outcome email iud new mazor x | NAV4112 | NAV1012 | NAV3250 | NAV3243 | NAV2050 | NAV4112NAV1012NAV3250NAV3243NAV2050 | True
4810 | 703578496 | X - Registration/X-eye/3Define issues | 2020-01-23T19:31:08.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration/X-eye/3Define issues 2018-03-06 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that trajectories were planned by the surgeon and representatives. Schanz pin was placed in the left iliac crest and attached to the schanz arm. Generic AP and oblique images were attained and registered successfully. No 3Define scan or draw spine was performed due to known camera issues. Registration was achieved off of L4 on automatic through basic segmentation. Red screws at all levels. Drastic changes were made to S2 screws in order for reachability. Both S2 screws were placed accurately and verified using c-arm and neuro monitoring. No patient harm was reported. MAZOR X SYSTEM TPL0059 | 2018-03-06 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that trajectories were planned by the surgeon and representatives. Schanz pin was placed in the left iliac crest and attached to the schanz arm. Generic AP and oblique images were attained and registered successfully. No 3Define scan or draw spine was performed due to known camera issues. Registration was achieved off of L4 on automatic through basic segmentation. Red screws at all levels. Drastic changes were made to S2 screws in order for reachability. Both S2 screws were placed accurately and verified using c-arm and neuro monitoring. No patient harm was reported. | x issues ccr spinal procedure trajectories planned surgeon representatives schanz pin placed left iliac crest attached schanz arm generic ap oblique images attained registered successfully scan draw spine performed known camera issues registration achieved automatic basic segmentation red screws levels drastic changes screws order reachability screws placed accurately verified using carm neuro monitoring mazor x | NAV3242 | NAV4091 | NAV3246 | NAV4112 | NAV3244 | NAV3242NAV4091NAV3246NAV4112NAV3244 | True
8947 | 704514696 | XSE - Issue Attaching RRF | 2021-07-21T20:11:45.000000+0000 | MazorX Stealth Edition | NAV4100 | STRIPPED THREADS | 30 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - Issue Attaching RRF 2021-Jul-21 Salesforce C00230556 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was an issue with the connection between the surgical arm and the robotic reference frame. The threads on the surgical arm seemed to be pushed down making the frame hard to attach with a screw. There was no residual drape stuck in the threads. There was no patient involvement. 2021-07-27 IUD x6 (Other): No new information. 2021-08-13 IUD x2 (Other): No new information. 2021-10-19 302601761, IUD (Other): No new information. 2021-10-20 IUD x3 (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Jul-21 Salesforce C00230556 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was an issue with the connection between the surgical arm and the robotic reference frame. The threads on the surgical arm seemed to be pushed down making the frame hard to attach with a screw. There was no residual drape stuck in the threads. There was no patient involvement. 2021-07-27 IUD x6 (Other): No new information. 2021-08-13 IUD x2 (Other): No new information. 2021-10-19 302601761, IUD (Other): No new information. 2021-10-20 IUD x3 (Other): No new information. | xse issue attaching rrf salesforce outside procedure issue connection surgical arm robotic reference frame threads surgical arm pushed making frame hard attach screw residual drape stuck threads iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV2054 | NAV4100 | NAV2047 | NAV3249NAV4095NAV2054NAV4100NAV2047 | True
8213 | 704371635 | XSE - CANBUS Error | 2021-04-30T20:24:58.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error 2021-Apr-30 Salesforce C00225600 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that after using the bovie the system encountered CANBUS errors. This occurred during a case but caused less than an hour delay. There was no impact on the patient outcome. 2021-05-03 IUD x3 (Other): No new information. 2021-06-25 IUD (Other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2021-Apr-30 Salesforce C00225600 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that after using the bovie the system encountered CANBUS errors. This occurred during a case but caused less than an hour delay. There was no impact on the patient outcome. 2021-05-03 IUD x3 (Other): No new information. 2021-06-25 IUD (Other): No new information. | xse canbus error salesforce spinal p rocedure using bovie encountered canbus errors occurred case caused hour delay impact outcome iud new iud new svc main cable | NAV2047 | NAV3261 | NAV5000 | NAV2054 | NAV2057 | NAV2047NAV3261NAV5000NAV2054NAV2057 | True
3381 | 703329542 | XSE - deviation/aborted | 2019-08-26T16:34:16.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - deviation/aborted 2019-Aug-23 SCL1902469 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the setup test and accuracy test were completed. While waiting for the surgeon to start the procedure, the rep received a Fatal Canbus error to the positioner IO. This was resolved with a hard restart. After 10 minutes, the rep received a Fatal Canbus error 1,2,3,4,5,7 and cart IO. They performed a soft restart which did not resolve the issue and performed a hard restart and they did not receive the errors again in the case. The arm was sent to Left L4 and Left L5 for the surgeon to make initial incision prior to inserting screws. The arm was then resent to Left L4 and Left L5 for the surgeon to drill and insert screws. Surgeon requested that the assisting surgeon complete Right L4 and Right L5, and the rep asked surgeon to come around the table for the remaining trajectories. The arm was sent to Right L4 and Right L5 for surgeon to ensure incision from interbody insertion was appropriate for Right L4-5 screws. Assisting Surgeon then drilled and inserted the Right L5 screw with guidance from the rep and main Surgeon. The arm was sent to to Right L4 and assisting surgeon drilled trajectory, possibly applying too much force. Then the assisting Surgeon placed screw at Right L4. Screws were then checked with intra op fluoro. Left L4, Left L5 and Right L5 were accurate, but Right L4 was clearly breaching superiorly. Main Surgeon requested that the rep adjust the plan and resend to Right L4 so that new trajectory could be drilled. The rep changed Right L4 plan and sent to new trajectory where main surgeon drilled and placed screw. New trajectory at Right L4 was acceptable under fluoro, but surgeon requested O-Arm spin to verify new trajectory. O-Arm spin revealed that new Right L4 trajectory was still slightly higher than planned and was breaching slightly laterally. Assisting surgeon and main surgeon agreed to remove the robotically guided screw and attempt to reposition with jamshidi nee dle. Surgeon was able to reposition trajectory and put new screw into Right L4 under flouroscopy. Surgeon completed another O-Arm spin and verified all screws were placed accurately. Following the surgery, the rep discussed the operation and misplaced screw at Right L4 with the surgeon. Main Surgeon felt that assisting surgeon applied too much force to drill and drill bit when drilling Right L4 initially, causing the trajectory to breach superiorly. Surgeon was pleased with final outcome and did not place fault on the robot for misplacement, as 3/4 screws were placed accurately. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-08-29 e1, e2, e3, e4 (Rep): Additional information received from a manufacturer representative reported that after the misplacement of right and left L4, the screw was rotated by about 4 degrees. Trajectories for L4 were originally planned with an inferior to superior path and then flattened out the screw in the sagittal view. In the axial view, the trajectory was rotated laterally by about 2-3 degrees. No error messages were displayed when instrumenting. The surgeon was positioned on the right side of the patient and they did not lean on the patient. A lateral image was taken and the screw was found to be superior. No AP image had been requested. MAZOR X SYSTEM TPL0059 | 2019-Aug-23 SCL1902469 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the setup test and accuracy test were completed. While waiting for the surgeon to start the procedure, the rep received a Fatal Canbus error to the positioner IO. This was resolved with a hard restart. After 10 minutes, the rep received a Fatal Canbus error 1,2,3,4,5,7 and cart IO. They performed a soft restart which did not resolve the issue and performed a hard restart and they did not receive the errors again in the case. The arm was sent to Left L4 and Left L5 for the surgeon to make initial incision prior to inserting screws. The arm was then resent to Left L4 and Left L5 for the surgeon to drill and insert screws. Surgeon requested that the assisting surgeon complete Right L4 and Right L5, and the rep asked surgeon to come around the table for the remaining trajectories. The arm was sent to Right L4 and Right L5 for surgeon to ensure incision from interbody insertion was appropriate for Right L4-5 screws. Assisting Surgeon then drilled and inserted the Right L5 screw with guidance from the rep and main Surgeon. The arm was sent to to Right L4 and assisting surgeon drilled trajectory, possibly applying too much force. Then the assisting Surgeon placed screw at Right L4. Screws were then checked with intra op fluoro. Left L4, Left L5 and Right L5 were accurate, but Right L4 was clearly breaching superiorly. Main Surgeon requested that the rep adjust the plan and resend to Right L4 so that new trajectory could be drilled. The rep changed Right L4 plan and sent to new trajectory where main surgeon drilled and placed screw. New trajectory at Right L4 was acceptable under fluoro, but surgeon requested O-Arm spin to verify new trajectory. O-Arm spin revealed that new Right L4 trajectory was still slightly higher than planned and was breaching slightly laterally. Assisting surgeon and main surgeon agreed to remove the robotically guided screw and attempt to reposition with jamshidi nee dle. Surgeon was able to reposition trajectory and put new screw into Right L4 under flouroscopy. Surgeon completed another O-Arm spin and verified all screws were placed accurately. Following the surgery, the rep discussed the operation and misplaced screw at Right L4 with the surgeon. Main Surgeon felt that assisting surgeon applied too much force to drill and drill bit when drilling Right L4 initially, causing the trajectory to breach superiorly. Surgeon was pleased with final outcome and did not place fault on the robot for misplacement, as 3/4 screws were placed accurately. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-08-29 e1, e2, e3, e4 (Rep): Additional information received from a manufacturer representative reported that after the misplacement of right and left L4, the screw was rotated by about 4 degrees. Trajectories for L4 were originally planned with an inferior to superior path and then flattened out the screw in the sagittal view. In the axial view, the trajectory was rotated laterally by about 2-3 degrees. No error messages were displayed when instrumenting. The surgeon was positioned on the right side of the patient and they did not lean on the patient. A lateral image was taken and the screw was found to be superior. No AP image had been requested. | xse deviationaborted spinal procedure setup test accuracy test completed waiting surgeon start procedure fatal canbus error positioner io resolved hard restart minutes fatal canbus error cart io performed soft restart did resolve issue performed hard restart did receive errors case arm sent left left surgeon make initial incision prior inserting screws arm resent left left surgeon drill insert screws surgeon requested assisting surgeon complete right right asked surgeon come table remaining trajectories arm sent right right surgeon ensure incision interbody insertion appropriate right screws assisting surgeon drilled inserted right screw main surgeon arm sent right assisting surgeon drilled trajectory possibly applying force assisting surgeon placed screw right screws checked intra op fluoro left left right accurate right clearly breaching superiorly main surgeon requested adjust plan resend right new trajectory drilled changed right plan sent new trajectory main surgeon drilled placed screw new trajectory right acceptable fluoro surgeon requested oarm spin verify new trajectory oarm spin revealed new right trajectory slightly higher planned breaching slightly laterally assisting surgeon main surgeon agreed remove robotically guided screw attempt reposition jamshidi nee dle surgeon able reposition trajectory new screw right flouroscopy surgeon completed oarm spin verified screws placed accurately following surgery discussed operation misplaced screw right surgeon main surgeon felt assisting surgeon applied force drill drill bit drilling right initially causing trajectory breach superiorly surgeon pleased final outcome did place fault robot misplacement screws placed accurately additional complications anticipated additional misplacement right left screw rotated degrees trajectories originally planned inferior superior path flattened screw sagittal view axial view trajectory rotated laterally degrees error messages displayed instrumenting surgeon positioned right did lean lateral image taken screw superior ap image requested mazor x | NAV2047 | NAV3253 | NAV5000 | NAV1007 | NAV4091 | NAV2047NAV3253NAV5000NAV1007NAV4091 | True
8827 | 704494036 | XSE - Joint Shift Error 495 | 2021-07-09T20:04:39.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - Joint Shift Error 495 2021-Jul-09 Salesforce C00229897 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when setting up the system the manufacturer representative got a Joint Shift Error. He tried a soft reboot and a hard reboot. He then reset the movement and did an advanced accuracy test. No patient was present. 2021-07-16 IUD (Other): No new information. 2021-08-13 email 2021-08-13 13:31 (Rep): Additional information was received stating the reset and 10 point accuracy test resolved the issue. 2021-09-08 email 2021-09-08 12:15 (Rep): No new information. 2021-09-21 email 2021-09-21 15:02 (Rep): No new information. 2021-10-19 302601982 (Other): No new information. 2021-10-21 IUD (Other): No new information. 2021-10-21 IUD x7 (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Jul-09 Salesforce C00229897 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when setting up the system the manufacturer representative got a Joint Shift Error. He tried a soft reboot and a hard reboot. He then reset the movement and did an advanced accuracy test. No patient was present. 2021-07-16 IUD (Other): No new information. 2021-08-13 email 2021-08-13 13:31 (Rep): Additional information was received stating the reset and 10 point accuracy test resolved the issue. 2021-09-08 email 2021-09-08 12:15 (Rep): No new information. 2021-09-21 email 2021-09-21 15:02 (Rep): No new information. 2021-10-19 302601982 (Other): No new information. 2021-10-21 IUD (Other): No new information. 2021-10-21 IUD x7 (Other): No new information. | xse joint shift error salesforce outside procedure setting got joint shift error tried soft reboot hard reboot reset movement did advanced accuracy test present iud new email additional stating reset point accuracy test resolved issue email new email new new iud new iud new kit svc surgical arm refurb | NAV4092 | NAV4098 | NAV4093 | NAV4095 | NAV3249 | NAV4092NAV4098NAV4093NAV4095NAV3249 | True
2212 | 703230548 | X - 3Define Functionality issue | 2019-06-27T11:23:23.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define Functionality issue During the 3define scan at a phase 2 training, I received multiple messages resulting in a failed 3define scan. the messages were "surface mapper scanning time out" and "surface scanner is not found/not available." a few soft restarts were done which finally stopped the messages, leading to a successful scan. This is the new system that was recently placed and no surgeries have been conducted yet. I can definitely see this being an issue in the future when we begin cases. UNK MAZOR X 3DEFINE CAMERA | During the 3define scan at a phase 2 training, I received multiple messages resulting in a failed 3define scan. the messages were "surface mapper scanning time out" and "surface scanner is not found/not available." a few soft restarts were done which finally stopped the messages, leading to a successful scan. This is the new system that was recently placed and no surgeries have been conducted yet. I can definitely see this being an issue in the future when we begin cases. | x functionality issue scan phase training multiple messages resulting failed scan messages surface mapper scanning time surface scanner foundnot available soft restarts finally stopped messages leading successful scan new recently placed surgeries conducted definitely issue future begin cases unk mazor x camera | NAV3244 | NAV3246 | NAV4095 | NAV4110 | NAV3250 | NAV3244NAV3246NAV4095NAV4110NAV3250 | True
5348 | 703655033 | X - Universal Screwdriver Damaged | 2020-03-07T00:18:38.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | MAS1650-03S | KIT SVC MAS1650-03S UNIV SCREWDRIVER | X - Universal Screwdriver Damaged 2020-03-06 00685858 (rep, hcp): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the tip of the universal screwdriver was broken off. The surgeon was using it for a non-guidance system related activity and over-torqued. There was no patient involved. 2020-03-09 Email 2020-03-09 16:15, IUD x4 (Rep, Other): No new information. 2020-04-01 IUD x3 (Other): No new information. 2020-04-23 IUD (Other): No new information. KIT SVC MAS1650-03S UNIV SCREWDRIVER | 2020-03-06 00685858 (rep, hcp): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the tip of the universal screwdriver was broken off. The surgeon was using it for a non-guidance system related activity and over-torqued. There was no patient involved. 2020-03-09 Email 2020-03-09 16:15, IUD x4 (Rep, Other): No new information. 2020-04-01 IUD x3 (Other): No new information. 2020-04-23 IUD (Other): No new information. | x universal screwdriver damaged outside procedure tip universal screwdriver broken surgeon using nonguidance related activity overtorqued involved email iud new iud new iud new kit svc univ screwdriver | NAV3295 | NAV3299 | NAV2055 | NAV3249 | NAV2058 | NAV3295NAV3299NAV2055NAV3249NAV2058 | True
1955 | 703223196 | X - WORKSTATION ISSUE | 2019-06-24T13:31:14.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - WORKSTATION ISSUE Prior to the case, hood on the Mazor X Workstation was stuck in the closed position. XCSR had to push up and down on it before it ag gressively popped open. Post case: Attempted to open hood to remove arm from bed. Hood was stuck in the closed position and would not open. XCSR had to push up and down and verify that the positioner was breaking the laser under the hood./Diagnosis: 62 year old female with scoliosis and lumbar stenosis. Planning: 6 pedicle screws were planned by surgeon prior to procedure paying additional attention to bilateral, inferior skive potential at L2-L3 as well as inferior skive potential at L4 right. Platform Mounting: Clamp was chosen per clinical indication and surgeon preference. Clamp was placed on L3. Registration: 1 AP and 1 Oblique high-dose images were taken. Automatic registration was achieved on first attempt using the basic algorithm, labeling off L3. Operation: Surgeon commenced by exposing patient's boney anatomy prior to using the Mazor X to place 6 accurate wires. Surgeon noted that L2 left had skived medially and inferiorly but confirmed that placement was still within the pedicle and was accurate. Wire placement was confirmed using fluoroscopy images. Screw placement was verified using fluoro and neuromonitoring, all screws stimmed >20. Times: 9:25 Bed frame adaptor on bed 9:30 Hood on the Workstation was stuck in the closed position. XCSR had to push up and down on it before it aggressi vely popped open. 10:14 Arm on bed 10:18 Draping commenced 10:19 Draping completed 10:20 Attachments on 10:24 Tablet draped and placed on tablet arm. (When scrub tech touched tablet screen, cursor on workstation was only visible on half of the WS screen, despite the WS having full accessibility. Soft restart did not fix the problem. Asked scrub tech to touch the tablet again which prompted WS to give access to tablet. Once clicked yes, then clicked back to give accessibility to WS, issue with the cursor was resolved.) 10:25 Incision/ start time 10:45 Arm to Clamp link 10:49 3Define 10:50 Draw spine, labeling off L3 10:51 Attached AP marker (captured 1 AP image) 10:53 Attached Oblique 3D marker (captured 1 Oblique image) 10:55 Registration (1st attempt, labeled off L3) 10:57 Sent L2 left trajectory 10:59 Sent L3 left trajectory 11:00 Sent L4 left trajectory 11:02 Sent L2 right trajectory 11:03 Sent L3 right trajectory 11:07 Sent L4 right trajectory 11:09 Cleared right 11:11 Confirmation AP and Oblique images 11:20 Attempted to open hood to remove arm from bed. Hood was stuck and would not open. XCSR had to push up and down and verify that the positioner was breaking the laser under the hood. 11:25 Removed arm from bed UNK MAZOR X WORKSTATION | Prior to the case, hood on the Mazor X Workstation was stuck in the closed position. XCSR had to push up and down on it before it ag gressively popped open. Post case: Attempted to open hood to remove arm from bed. Hood was stuck in the closed position and would not open. XCSR had to push up and down and verify that the positioner was breaking the laser under the hood./Diagnosis: 62 year old female with scoliosis and lumbar stenosis. Planning: 6 pedicle screws were planned by surgeon prior to procedure paying additional attention to bilateral, inferior skive potential at L2-L3 as well as inferior skive potential at L4 right. Platform Mounting: Clamp was chosen per clinical indication and surgeon preference. Clamp was placed on L3. Registration: 1 AP and 1 Oblique high-dose images were taken. Automatic registration was achieved on first attempt using the basic algorithm, labeling off L3. Operation: Surgeon commenced by exposing patient's boney anatomy prior to using the Mazor X to place 6 accurate wires. Surgeon noted that L2 left had skived medially and inferiorly but confirmed that placement was still within the pedicle and was accurate. Wire placement was confirmed using fluoroscopy images. Screw placement was verified using fluoro and neuromonitoring, all screws stimmed >20. Times: 9:25 Bed frame adaptor on bed 9:30 Hood on the Workstation was stuck in the closed position. XCSR had to push up and down on it before it aggressi vely popped open. 10:14 Arm on bed 10:18 Draping commenced 10:19 Draping completed 10:20 Attachments on 10:24 Tablet draped and placed on tablet arm. (When scrub tech touched tablet screen, cursor on workstation was only visible on half of the WS screen, despite the WS having full accessibility. Soft restart did not fix the problem. Asked scrub tech to touch the tablet again which prompted WS to give access to tablet. Once clicked yes, then clicked back to give accessibility to WS, issue with the cursor was resolved.) 10:25 Incision/ start time 10:45 Arm to Clamp link 10:49 3Define 10:50 Draw spine, labeling off L3 10:51 Attached AP marker (captured 1 AP image) 10:53 Attached Oblique 3D marker (captured 1 Oblique image) 10:55 Registration (1st attempt, labeled off L3) 10:57 Sent L2 left trajectory 10:59 Sent L3 left trajectory 11:00 Sent L4 left trajectory 11:02 Sent L2 right trajectory 11:03 Sent L3 right trajectory 11:07 Sent L4 right trajectory 11:09 Cleared right 11:11 Confirmation AP and Oblique images 11:20 Attempted to open hood to remove arm from bed. Hood was stuck and would not open. XCSR had to push up and down and verify that the positioner was breaking the laser under the hood. 11:25 Removed arm from bed | x workstation issue prior case hood mazor x workstation stuck closed position xcsr push ag gressively popped open post case attempted open hood remove arm bed hood stuck closed position open xcsr push verify positioner breaking laser hooddiagnosis year old female scoliosis lumbar stenosis planning pedicle screws planned surgeon prior procedure paying additional attention bilateral inferior skive potential inferior skive potential right platform mounting clamp chosen clinical indication surgeon preference clamp placed registration ap oblique highdose images taken automatic registration achieved attempt using basic algorithm labeling operation surgeon commenced exposing patients boney anatomy prior using mazor x place accurate wires surgeon noted left skived medially inferiorly confirmed placement pedicle accurate wire placement confirmed using fluoroscopy images screw placement verified using fluoro neuromonitoring screws stimmed times bed frame adaptor bed hood workstation stuck closed position xcsr push aggressi vely popped open arm bed draping commenced draping completed attachments tablet draped placed tablet arm incision start time arm clamp link draw spine labeling attached ap marker attached oblique marker registration sent left trajectory sent left trajectory sent left trajectory sent right trajectory sent right trajectory sent right trajectory cleared right confirmation ap oblique images attempted open hood remove arm bed hood stuck open xcsr push verify positioner breaking laser hood removed arm bed unk mazor x workstation | NAV2058 | NAV2047 | NAV3237 | NAV3236 | NAV3243 | NAV2058NAV2047NAV3237NAV3236NAV3243 | True
7375 | 704192574 | XSE - Positioner Not Moving | 2021-01-25T17:34:42.000000+0000 | MazorX Stealth Edition | NAV4104 | UNABLE TO MOVE | 20 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - Positioner Not Moving 2021-Jan-25 Salesforce C00219823, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after the case was completed, the manufacturer representative went to move the positioner up and down in the Z-axis and the system would not move. The representative could move the positioner left and right, but could not move it up and down. The surgical system was removed from the bed and a soft and hard restart of the system was done. The representative mounted the surgical system to another be, but they were still unable to move the positioner up and down. The representative also tried to force the Z-axis open using the override panel, but the issue was not resolved. There was no patient harm and the procedure was not delayed. 2021-01-28 IUD x4 (Other): No new information. 2021-02-22 IUD (Other): No new information. 2021-04-20 IUD (Other): No new information. 2021-05-04 IUD (Other): No new information. 2021-05-24 IUD (Other): No new information. 2021-Jul-20 IUD (other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-Jan-25 Salesforce C00219823, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after the case was completed, the manufacturer representative went to move the positioner up and down in the Z-axis and the system would not move. The representative could move the positioner left and right, but could not move it up and down. The surgical system was removed from the bed and a soft and hard restart of the system was done. The representative mounted the surgical system to another be, but they were still unable to move the positioner up and down. The representative also tried to force the Z-axis open using the override panel, but the issue was not resolved. There was no patient harm and the procedure was not delayed. 2021-01-28 IUD x4 (Other): No new information. 2021-02-22 IUD (Other): No new information. 2021-04-20 IUD (Other): No new information. 2021-05-04 IUD (Other): No new information. 2021-05-24 IUD (Other): No new information. 2021-Jul-20 IUD (other): No new information. | xse positioner moving salesforce iud spinal procedure case completed went positioner zaxis positioner left right surgical removed bed soft hard restart mounted surgical unable positioner tried force zaxis open using override panel issue resolved procedure delayed iud new iud new iud new iud new iud new iud new kit svc surgical syst type | NAV4104 | NAV3251 | NAV3250 | NAV4105 | NAV4103 | NAV4104NAV3251NAV3250NAV4105NAV4103 | True
3330 | 703315121 | XSE - registration/hit WV/navlock issues | 2019-08-16T15:34:07.000000+0000 | MazorX Stealth Edition | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - registration/hit WV/navlock issues 2019-08-13 SCL1902373, Intake Form, CCR (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat deformity with fixation. It was reported that a T2-L2 open procedure was planned with the surgeon. Due to some of the pedicles being too small to comfortably fit a sizable screw, some pedicles were left out of the plan. For the first segment, a dual spinous process clamp was attached to T10 and T12 and connected to the surgical system via a bone mounting bridge. A 3Define scan was performed and Draw Spine was completed with a navigated passive planar prove. AP and oblique trajectories were sent and both images were accepted; however, when trying to register a warning that the AP image had been auto rotated appeared. The representative went back in the software and a new AP image was taken. Registration was successful from L2-T8 with green values at each level. All trajectories were sent, drilled and screws were placed. When sending the arm to the L1 and L2 trajectories, the surgical arm ran into the bone mounting bridge. L1 and L2 on the right side were obstructed by the clamp and the bone mounting bridge so the surgeon decided that they would move the clamp and try those trajectories at the end of the case. For the second segment, the dual spinous process clam was moved to the T4 and T6 ve rtebrae. A 3Define scan was performed and Draw Spine was completed with a passive planar prove. AP and oblique images were taken. The AP image was accepted on the first try, but several oblique images had to be taken to best visualize the anatomy of interest. Registration was successful when laveling off of T6 using the basic algorithm. All values were green except for T3 which had red values. The surgeon decided to go ahead and execute all trajectories except for T3. The surgeon planned to try and register T3 by itself. After placing all of the screws for the second segment, registration was completed for T3. Green values were achieved so the trajectory was drilled and screws placed. For segment 3, the surgeon tried to go back to L1 and L2 on the right side. The mount was unlocked and the surgeon used a single clamp on the L4 spinous process. A 3Define scan was performed and Draw Spine was completed using a passive planar probe. AP and oblique images were taken and both were accepted on the first try. Registration was suc cessful off of L3 with all green values. Both trajectories were sent, drill and had screws placed. All screws were placed and looked acceptable in fluoroscopic images. Neuro monitoring was used during the procedure and all screws were stimulated with good values. The representative noted that the drill bit bound and locked in the NavLock tracker and the surgeon's glove got stuck in between the tracker and the drill bit at least two times and tore the surgeons clove causing them to re-glove. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-08-22 email 2019-08-22 18:45 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2019-08-13 SCL1902373, Intake Form, CCR (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat deformity with fixation. It was reported that a T2-L2 open procedure was planned with the surgeon. Due to some of the pedicles being too small to comfortably fit a sizable screw, some pedicles were left out of the plan. For the first segment, a dual spinous process clamp was attached to T10 and T12 and connected to the surgical system via a bone mounting bridge. A 3Define scan was performed and Draw Spine was completed with a navigated passive planar prove. AP and oblique trajectories were sent and both images were accepted; however, when trying to register a warning that the AP image had been auto rotated appeared. The representative went back in the software and a new AP image was taken. Registration was successful from L2-T8 with green values at each level. All trajectories were sent, drilled and screws were placed. When sending the arm to the L1 and L2 trajectories, the surgical arm ran into the bone mounting bridge. L1 and L2 on the right side were obstructed by the clamp and the bone mounting bridge so the surgeon decided that they would move the clamp and try those trajectories at the end of the case. For the second segment, the dual spinous process clam was moved to the T4 and T6 ve rtebrae. A 3Define scan was performed and Draw Spine was completed with a passive planar prove. AP and oblique images were taken. The AP image was accepted on the first try, but several oblique images had to be taken to best visualize the anatomy of interest. Registration was successful when laveling off of T6 using the basic algorithm. All values were green except for T3 which had red values. The surgeon decided to go ahead and execute all trajectories except for T3. The surgeon planned to try and register T3 by itself. After placing all of the screws for the second segment, registration was completed for T3. Green values were achieved so the trajectory was drilled and screws placed. For segment 3, the surgeon tried to go back to L1 and L2 on the right side. The mount was unlocked and the surgeon used a single clamp on the L4 spinous process. A 3Define scan was performed and Draw Spine was completed using a passive planar probe. AP and oblique images were taken and both were accepted on the first try. Registration was suc cessful off of L3 with all green values. Both trajectories were sent, drill and had screws placed. All screws were placed and looked acceptable in fluoroscopic images. Neuro monitoring was used during the procedure and all screws were stimulated with good values. The representative noted that the drill bit bound and locked in the NavLock tracker and the surgeon's glove got stuck in between the tracker and the drill bit at least two times and tore the surgeons clove causing them to re-glove. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-08-22 email 2019-08-22 18:45 (Rep): No new information. | xse registrationhit wvnavlock issues ccr spinal procedure treat deformity fixation open procedure planned surgeon pedicles small comfortably fit sizable screw pedicles left plan segment dual spinous process clamp attached connected surgical bone mounting bridge scan performed draw spine completed navigated passive planar prove ap oblique trajectories sent images accepted trying register warning ap image auto rotated appeared went software new ap image taken registration successful green values level trajectories sent drilled screws placed sending arm trajectories surgical arm ran bone mounting bridge right obstructed clamp bone mounting bridge surgeon decided clamp try trajectories end case second segment dual spinous process clam moved ve rtebrae scan performed draw spine completed passive planar prove ap oblique images taken ap image accepted try oblique images taken best visualize anatomy registration successful laveling using basic algorithm values green red values surgeon decided ahead execute trajectories surgeon planned try register placing screws second segment registration completed green values achieved trajectory drilled screws placed segment surgeon tried right mount unlocked surgeon single clamp spinous process scan performed draw spine completed using passive planar probe ap oblique images taken accepted try registration suc cessful green values trajectories sent drill screws placed screws placed looked acceptable fluoroscopic images neuro monitoring procedure screws stimulated good values noted drill bit bound locked navlock tracker surgeons glove got stuck tracker drill bit times tore surgeons clove causing reglove additional complications anticipated email new mazor x | NAV3243 | NAV3242 | NAV5000 | NAV3246 | NAV3239 | NAV3243NAV3242NAV5000NAV3246NAV3239 | True
383 | 703095005 | X - deviation/shudder | 2019-04-08T16:06:02.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 30 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - deviation/shudder 2019-Apr-03 SCL1900815 intake form, msg (HCP,rep): ***See 703083932 for clamp issue*** Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were several issues that were occurring during each utilization. The system passed the advanced accuracy setup, failed the first attempted 3 point check, but passed the next attempt. Then the shoulder failed to lock, and a hard restart resolved the issue. During registration, the shoulder was moving to the oblique position and it was shuddering during movement. The shudder was substantial, but only noted during this one movement. The second k wire placed (L4 right) was not accurate, this was thought to be a result of the amount of bone removed on that side by the hemi laminectomywhich was preformed prior to drilling. It was noted that there was no patient harm. There were no additional complications reported or anticipated as a result of the event. during initial setup, the controller failed to initiate. This error went away once the system took more time to boot up. Next, the system passed the advanced accuracy setup but failed the 3 point check. A soft restart was attempted and the system failed again. A hard restart resolved the issue and the 3 point test was successful and the implants were accurate MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-Apr-03 SCL1900815 intake form, msg (HCP,rep): ***See 703083932 for clamp issue*** Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were several issues that were occurring during each utilization. The system passed the advanced accuracy setup, failed the first attempted 3 point check, but passed the next attempt. Then the shoulder failed to lock, and a hard restart resolved the issue. During registration, the shoulder was moving to the oblique position and it was shuddering during movement. The shudder was substantial, but only noted during this one movement. The second k wire placed (L4 right) was not accurate, this was thought to be a result of the amount of bone removed on that side by the hemi laminectomywhich was preformed prior to drilling. It was noted that there was no patient harm. There were no additional complications reported or anticipated as a result of the event. during initial setup, the controller failed to initiate. This error went away once the system took more time to boot up. Next, the system passed the advanced accuracy setup but failed the 3 point check. A soft restart was attempted and the system failed again. A hard restart resolved the issue and the 3 point test was successful and the implants were accurate | x deviationshudder msg clamp issue spinal procedure issues occurring utilization passed advanced accuracy setup failed attempted point check passed attempt shoulder failed lock hard restart resolved issue registration shoulder moving oblique position shuddering movement shudder substantial noted movement second k wire placed accurate thought bone removed hemi laminectomywhich preformed prior drilling noted additional complications anticipated initial setup controller failed initiate error went away took time boot passed advanced accuracy setup failed point check soft restart attempted failed hard restart resolved issue point test successful implants accurate mazor x software kit | NAV3242 | NAV3246 | NAV3248 | NAV2047 | NAV3262 | NAV3242NAV3246NAV3248NAV2047NAV3262 | False
4183 | 703511013 | XSE - Target Extender Damaged | 2019-12-10T19:23:34.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | MAS2310-03 | TARGET EXTENDER | XSE - Target Extender Damaged 2019-12-10 00679980 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the silver piece on the target extender is broken. No patient was present. 2019-12-27 IUD x4 (Other): No new information. TARGET EXTENDER | 2019-12-10 00679980 (REP): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the silver piece on the target extender is broken. No patient was present. 2019-12-27 IUD x4 (Other): No new information. | xse target extender damaged outside procedure silver piece target extender broken present iud new target extender | NAV3295 | NAV3299 | NAV2053 | NAV2054 | NAV3251 | NAV3295NAV3299NAV2053NAV2054NAV3251 | True
8179 | 704366774 | XSE - Difficulties Loading CT from Disc | 2021-04-28T20:23:25.000000+0000 | MazorX Stealth Edition | NAV3236 | CT IMAGES | 20 | UNK_MAZORX_LPTP | UNK MAZOR X LAPTOP | XSE - Difficulties Loading CT from Disc 2021-Apr-28 Salesforce C00225448 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative had a difficult time loading exams onto the guidance system using a disc. After several attempts, the representative was able to load the exams. There was no patient involvement. 2021-04-29 email 2021-04-29 10:32 (Rep): Additional information was received stating that the difficulty was with loading the CT onto the manufacturer representative's new work laptop. Not the system. 2021-04-30 IUD x2 (Other): No new information. 2021-05-19 email 2021-05-19 15:31 (Rep): The cause of the issue was due to an old slow laptop according to the manufacturer representative. The old CPU ended up dying. UNK MAZOR X LAPTOP | 2021-Apr-28 Salesforce C00225448 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative had a difficult time loading exams onto the guidance system using a disc. After several attempts, the representative was able to load the exams. There was no patient involvement. 2021-04-29 email 2021-04-29 10:32 (Rep): Additional information was received stating that the difficulty was with loading the CT onto the manufacturer representative's new work laptop. Not the system. 2021-04-30 IUD x2 (Other): No new information. 2021-05-19 email 2021-05-19 15:31 (Rep): The cause of the issue was due to an old slow laptop according to the manufacturer representative. The old CPU ended up dying. | xse difficulties loading ct disc salesforce outside procedure difficult time loading exams using disc attempts able load exams email additional stating difficulty loading ct representatives new work laptop iud new email cause issue old slow laptop according old cpu ended dying unk mazor x laptop | NAV3236 | NAV2058 | NAV3244 | NAV3257 | NAV2052 | NAV3236NAV2058NAV3244NAV3257NAV2052 | True
8112 | 704355389 | XSE - SRA Failed BIST Test | 2021-04-22T08:19:19.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - SRA Failed BIST Test 2021-Apr-20 Salesforce C00225104, image, IUD (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the built in self test (BIST) during set up and an Error 293 was displayed. There was no patient involvement. 2021-05-23 IUD x2 (Other, For): No new information. 2021-05-27 email 2021-05-27 05:16 (Rep, For): No new information. 2021-06-15 IUD x5, email 2021-06-15 03:01 (Rep, For): No new information. 2021-10-27 IUD x2 (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2021-Apr-20 Salesforce C00225104, image, IUD (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the built in self test (BIST) during set up and an Error 293 was displayed. There was no patient involvement. 2021-05-23 IUD x2 (Other, For): No new information. 2021-05-27 email 2021-05-27 05:16 (Rep, For): No new information. 2021-06-15 IUD x5, email 2021-06-15 03:01 (Rep, For): No new information. 2021-10-27 IUD x2 (Other): No new information. | xse sra failed bist test salesforce image iud outside procedure surgical arm failed built self test set error displayed iud new email new iud email new iud new kit svc surgical arm | NAV3249 | NAV2054 | NAV2047 | NAV3264 | NAV4093 | NAV3249NAV2054NAV2047NAV3264NAV4093 | True
6815 | 704040201 | XSE - Cracked Surgeon Screen | 2020-10-28T22:39:56.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE - Cracked Surgeon Screen 2020-10-28 00699779, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon screen was starting to crack. There was no patient involvement. 2020-11-03 IUD x4 (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2020-10-28 00699779, IUD x3 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon screen was starting to crack. There was no patient involvement. 2020-11-03 IUD x4 (Other): No new information. | xse cracked surgeon screen iud outside procedure surgeon screen starting crack iud new kit svc srgn screen arm | NAV3253 | NAV2054 | NAV2047 | NAV1006 | NAV2058 | NAV3253NAV2054NAV2047NAV1006NAV2058 | True
6738 | 704019993 | XSE - Pneumatic Leak on Manipulator | 2020-10-19T16:17:10.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE - Pneumatic Leak on Manipulator 2020-10-18 00699042, IUD, PDF (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when putting the surgical system into storage, a pneumatic leak was heard coming from the manipulator. There was no patient involvement. UNK MAZOR X MANIPULATOR | 2020-10-18 00699042, IUD, PDF (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when putting the surgical system into storage, a pneumatic leak was heard coming from the manipulator. There was no patient involvement. | xse pneumatic leak manipulator iud pdf outside procedure putting surgical storage pneumatic leak heard coming manipulator unk mazor x manipulator | NAV3260 | NAV3264 | NAV2054 | NAV4110 | NAV3261 | NAV3260NAV3264NAV2054NAV4110NAV3261 | True
2930 | 703259021 | X - Screw Fell Off of Bed Frame Adaptor | 2019-07-15T17:37:32.000000+0000 | MazorX | NAV3255 | BED FRAME ISSUES | 20 | MAS1574-01 | BED FRAME ADAPTOR | X - Screw Fell Off of Bed Frame Adaptor As a follow up to system installation, replacement of the Bed Adaptor. BED FRAME ADAPTOR | As a follow up to system installation, replacement of the Bed Adaptor. | x screw fell bed frame adaptor follow installation replacement bed adaptor bed frame adaptor | NAV3255 | NAV3295 | NAV4090 | NAV2054 | NAV3250 | NAV3255NAV3295NAV4090NAV2054NAV3250 | True
3714 | 703426916 | XSE - deviation | 2019-10-22T19:35:56.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059-04 | MAZOR X SYSTEM (REV 4) | XSE - deviation 2019-Oct-18 SCL1903132 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they had a L2-S1 case and L2 left and right were superior to plan. L3 left was superior as well. All three trajectories were placed free hand. Surgeon believes the patient could have bucked and the trajectories at L2 were in-out trajectories. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-06 e1, e2, e3, e4 (Rep): Additional information received from a manufacturer representative reported the surgeon noted that they had enabled a reverse trendelenburg on the table after the patient was registered. The surgeon thought that may hav shifted the positioning of the patient on the table. During the procedure, no tape was used on the patient's legs and they were strapped down with a belt. The surgeon did not lean on the patient or apply significant force while instrumenting. No boney work was done to teh surface prior to drilling. 2019-11-11 e6 (Rep): Additional information received from a manufacturer representative reported the surgeon placed the patient for comfort in most cases. The surgeon was aware of the problems it may cause and would avoid it from now on. MAZOR X SYSTEM (REV 4) | 2019-Oct-18 SCL1903132 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they had a L2-S1 case and L2 left and right were superior to plan. L3 left was superior as well. All three trajectories were placed free hand. Surgeon believes the patient could have bucked and the trajectories at L2 were in-out trajectories. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-06 e1, e2, e3, e4 (Rep): Additional information received from a manufacturer representative reported the surgeon noted that they had enabled a reverse trendelenburg on the table after the patient was registered. The surgeon thought that may hav shifted the positioning of the patient on the table. During the procedure, no tape was used on the patient's legs and they were strapped down with a belt. The surgeon did not lean on the patient or apply significant force while instrumenting. No boney work was done to teh surface prior to drilling. 2019-11-11 e6 (Rep): Additional information received from a manufacturer representative reported the surgeon placed the patient for comfort in most cases. The surgeon was aware of the problems it may cause and would avoid it from now on. | xse deviation spinal procedure case left right superior plan left superior trajectories placed free hand surgeon believes bucked trajectories inout trajectories additional complications anticipated additional surgeon noted enabled reverse trendelenburg table registered surgeon thought hav shifted positioning table procedure tape patients legs strapped belt surgeon did lean apply significant force instrumenting boney work teh surface prior drilling additional surgeon placed comfort cases surgeon aware problems cause avoid mazor x | NAV5000 | NAV2054 | NAV3253 | NAV3243 | NAV2051 | NAV5000NAV2054NAV3253NAV3243NAV2051 | True
4724 | 703570238 | X - Arm & SW issue | 2020-01-19T12:05:14.000000+0000 | MazorX | NAV2056 | SOFTWARE ISSUES | 30 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - Arm & SW issue 2020-1-12 Intake Form, 00682348(Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was lowered to the short trajectory, and the arm touched the platform. The software did not let the manufacturer representative send the arm to the clear up position. All the software options were grayed out and the only option available was "run simulation". A software restart was done to overcome the issue. There was no patient harm and the procedure was not delayed over an hour. 2020-01-22 email 2020-01-22 04:15, Image, CCR (Rep): Additional information received from a manufacturer representative reported the cause of the collision was the representative lowering the arm to reach short tools. MAZOR X SOFTWARE KIT VER-4.0.0 | 2020-1-12 Intake Form, 00682348(Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgical arm was lowered to the short trajectory, and the arm touched the platform. The software did not let the manufacturer representative send the arm to the clear up position. All the software options were grayed out and the only option available was "run simulation". A software restart was done to overcome the issue. There was no patient harm and the procedure was not delayed over an hour. 2020-01-22 email 2020-01-22 04:15, Image, CCR (Rep): Additional information received from a manufacturer representative reported the cause of the collision was the representative lowering the arm to reach short tools. | x arm sw issue spinal procedure surgical arm lowered short trajectory arm touched platform software did let send arm clear position software options grayed option available run simulation software restart overcome issue procedure delayed hour email image ccr additional cause collision lowering arm reach short tools mazor x software kit | NAV3246 | NAV3248 | NAV4096 | NAV3247 | NAV4095 | NAV3246NAV3248NAV4096NAV3247NAV4095 | False
580 | 703152752 | X - CD drive intermttent | 2019-05-14T14:28:48.000000+0000 | MazorX | NAV3257 | CD DRIVE ISSUES | 20 | DSD0038-01 | DVD, PORTABLE DVD REWRITER, GP70NS50 | X - CD drive intermttent 2019-May-10 SCL1901161 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the CD drive only loaded images 50% of the time. Sometimes it worked and sometimes it didn’t. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. DVD, PORTABLE DVD REWRITER, GP70NS50 | 2019-May-10 SCL1901161 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the CD drive only loaded images 50% of the time. Sometimes it worked and sometimes it didn’t. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x cd drive intermttent spinal procedure cd drive loaded images time worked didn’t additional complications anticipated dvd portable dvd rewriter | NAV3257 | NAV3295 | NAV3249 | NAV4101 | NAV1006 | NAV3257NAV3295NAV3249NAV4101NAV1006 | True
7815 | 704283190 | XSE - Main Monitor Dimmed | 2021-03-15T19:59:59.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 30 | ASM0275-01S | KIT SVC ASM0275-01S MONITOR | XSE - Main Monitor Dimmed 2021-03-15 00708454 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that when they were putting the system away and folded the surgeon screen up into storage position the main monitor became dimmed. They could see the software just the backlight was dimmed. They attempted a hard restart and that did not resolve the issue. No patient was present. 2021-03-24 IUD x3 (Rep): Additional information received from a manufacturer representative reported that the issue was due to a short in the surgeon monitor. 2021-04-05 IUD x4 (Other): No new information. 2021-04-26 IUD x5 (Other): No new information. 2021-06-13 IUD (Other): No new information. 2021-06-16 IUD (Other): No new information. 2021-06-28 IUD (Other): No new information. 2021-07-14 IUD (Other): No new information. KIT SVC ASM0275-01S MONITOR | 2021-03-15 00708454 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that when they were putting the system away and folded the surgeon screen up into storage position the main monitor became dimmed. They could see the software just the backlight was dimmed. They attempted a hard restart and that did not resolve the issue. No patient was present. 2021-03-24 IUD x3 (Rep): Additional information received from a manufacturer representative reported that the issue was due to a short in the surgeon monitor. 2021-04-05 IUD x4 (Other): No new information. 2021-04-26 IUD x5 (Other): No new information. 2021-06-13 IUD (Other): No new information. 2021-06-16 IUD (Other): No new information. 2021-06-28 IUD (Other): No new information. 2021-07-14 IUD (Other): No new information. | xse main monitor dimmed outside procedure putting away folded surgeon screen storage position main monitor dimmed software just backlight dimmed attempted hard restart did resolve issue present iud additional issue short surgeon monitor iud new iud new iud new iud new iud new iud new kit svc monitor | NAV3261 | NAV3253 | NAV2047 | NAV3246 | NAV2054 | NAV3261NAV3253NAV2047NAV3246NAV2054 | True
1421 | 703200084 | X- Hood Cover Issue | 2019-06-11T16:23:18.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- Hood Cover Issue Hood would not open without assistance. Robot also got stuck on bed rail adapter, one tooth engaged the mount while the other did not. We had to remove the robot with bed rail mount attached. We then reattached the bedrail mount and were ableto successfully attach robot. Once robot was installed on bed registration was quick and efficient,no software issues. UNK MAZOR X WORKSTATION | Hood would not open without assistance. Robot also got stuck on bed rail adapter, one tooth engaged the mount while the other did not. We had to remove the robot with bed rail mount attached. We then reattached the bedrail mount and were ableto successfully attach robot. Once robot was installed on bed registration was quick and efficient,no software issues. | x hood cover issue hood open assistance robot got stuck bed rail adapter tooth engaged mount did remove robot bed rail mount attached reattached bedrail mount ableto successfully attach robot robot installed bed registration quick efficientno software issues unk mazor x workstation | NAV3250 | NAV3255 | NAV2058 | NAV3254 | NAV3264 | NAV3250NAV3255NAV2058NAV3254NAV3264 | True
327 | 703078553 | X - Failed Accuracy Test | 2019-03-28T14:32:49.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 40 | ASM0270 | KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | X - Failed Accuracy Test 2019-03-20 SCL1900696, SCL1900709, SCL1900729 (Rep): Medtronic received information regarding a guidance system. It was reported that the surgical arm failed the advanced accuracy test and the electrical unit was faulty when the system was being upgraded. There was no patient involved. 2021-09-21 302568001 (Other): No new information. KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | 2019-03-20 SCL1900696, SCL1900709, SCL1900729 (Rep): Medtronic received information regarding a guidance system. It was reported that the surgical arm failed the advanced accuracy test and the electrical unit was faulty when the system was being upgraded. There was no patient involved. 2021-09-21 302568001 (Other): No new information. | x failed accuracy test surgical arm failed advanced accuracy test electrical unit faulty upgraded involved new kit svc compressor kit new | NAV4093 | NAV2054 | NAV4092 | NAV4091 | NAV2047 | NAV4093NAV2054NAV4092NAV4091NAV2047 | False
6850 | 704049287 | XSE - Inaccurate Navigation | 2020-11-03T18:42:33.000000+0000 | MazorX Stealth Edition | NAV4100 | STRIPPED THREADS | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Inaccurate Navigation 2020-11-03 00700107, email 2020-11-05 09:01, IUD x5, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using screws and rods. It was reported that navigation was inaccurate by more than 3 mm and inferior on all levels during a procedure. This included the sagittal views as well. After taking confirmation images, all screws were found to be placed accurately. The surgeon took a second snapshot, checked the plan and checked the reference frame placement. The cause of navigation being off was found to be the fastener on the surgical arm where the robotic reference frame attaches was loose. The loose fastener caused the reference frame to be loose and the navigation to be off. This issue was able to be reproduced after the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-01 IUD x4 (Other): No new information. 2020-Dec-28 IUD (other): No new information. 2021-02-01 IUD (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-11-03 00700107, email 2020-11-05 09:01, IUD x5, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using screws and rods. It was reported that navigation was inaccurate by more than 3 mm and inferior on all levels during a procedure. This included the sagittal views as well. After taking confirmation images, all screws were found to be placed accurately. The surgeon took a second snapshot, checked the plan and checked the reference frame placement. The cause of navigation being off was found to be the fastener on the surgical arm where the robotic reference frame attaches was loose. The loose fastener caused the reference frame to be loose and the navigation to be off. This issue was able to be reproduced after the procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-12-01 IUD x4 (Other): No new information. 2020-Dec-28 IUD (other): No new information. 2021-02-01 IUD (Other): No new information. | xse inaccurate navigation email iud ccr spinal procedure treat degenerative discs using screws rods navigation inaccurate mm inferior levels procedure included sagittal views taking confirmation images screws placed accurately surgeon took second snapshot checked plan checked reference frame placement cause navigation fastener surgical arm robotic reference frame attaches loose loose fastener caused reference frame loose navigation issue able reproduced procedure procedure delayed hour iud new iud new iud new kit svc surgical arm | NAV4097 | NAV5000 | NAV5001 | NAV4100 | NAV4109 | NAV4097NAV5000NAV5001NAV4100NAV4109 | True
7381 | 704193971 | XSE - Bent Side Panel | 2021-01-26T01:36:42.000000+0000 | MazorX Stealth Edition | NAV2055 | PRODUCTS DAMAGED DURING SHIPPING | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Bent Side Panel 2021-01-25 00705139 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the bed frame adapter side panel was bent after shipping. No patient was present. 2021-01-28 IUD x2 (Other): No new information. 2021-02-15 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2021-01-25 00705139 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the bed frame adapter side panel was bent after shipping. No patient was present. 2021-01-28 IUD x2 (Other): No new information. 2021-02-15 IUD (Other): No new information. | xse bent panel outside procedure bed frame adapter panel bent shipping present iud new iud new unk mazor x workstation | NAV2058 | NAV2047 | NAV3264 | NAV3255 | NAV2054 | NAV2058NAV2047NAV3264NAV3255NAV2054 | False
2414 | 703235106 | X - 3Define camera issue | 2019-06-29T14:49:01.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define camera issue Turned the workstation on and the message in the image attached is what appeared. Attempted soft restart and hard restart with no solution. There was also no firmware version # either. UNK MAZOR X 3DEFINE CAMERA | Turned the workstation on and the message in the image attached is what appeared. Attempted soft restart and hard restart with no solution. There was also no firmware version # either. | x camera issue turned workstation message image attached appeared attempted soft restart hard restart solution firmware version unk mazor x camera | NAV3244 | NAV2058 | NAV3261 | NAV2047 | NAV4110 | NAV3244NAV2058NAV3261NAV2047NAV4110 | True
7557 | 704226356 | XSE - Unable to Track Instruments | 2021-02-11T21:46:52.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 40 | ASM0275-01 | RNX-WS-1210 | XSE - Unable to Track Instruments 2021-Feb-11 Salesforce C00220934 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that of the instruments were not tracking. This included all Navlocks trackers, tracker kit, ball probe, robotic frame, ball probe, reference frame, and Navlock kit. There was no patient involved. 2021-03-01 IUD x3 (Other): No new information. 2021-03-16 email 2021-03-16 20:05 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved changing the arm, moving the camera around and trying new instruments. The cause of the tracking error was the camera and system error. 2021-04-19 IUD x5 (Other): No new information. 2021-05-27 IUD x2 (Other): No new information. RNX-WS-1210 | 2021-Feb-11 Salesforce C00220934 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that of the instruments were not tracking. This included all Navlocks trackers, tracker kit, ball probe, robotic frame, ball probe, reference frame, and Navlock kit. There was no patient involved. 2021-03-01 IUD x3 (Other): No new information. 2021-03-16 email 2021-03-16 20:05 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved changing the arm, moving the camera around and trying new instruments. The cause of the tracking error was the camera and system error. 2021-04-19 IUD x5 (Other): No new information. 2021-05-27 IUD x2 (Other): No new information. | xse unable track instruments salesforce outside procedure instruments tracking included navlocks trackers tracker kit ball probe robotic frame ball probe reference frame navlock kit involved iud new email additional troubleshooting involved changing arm moving camera trying new instruments cause tracking error camera error iud new iud new | NAV3299 | NAV3246 | NAV2054 | NAV4110 | NAV3249 | NAV3299NAV3246NAV2054NAV4110NAV3249 | False
4169 | 703508322 | XSE - SRA Failed Advanced Accuracy Test | 2019-12-09T19:22:16.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 30 | ASM0206-01 | SURGICAL ARM | XSE - SRA Failed Advanced Accuracy Test 2019-12-09 00679860 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the setup and advanced accuracy test at the second point. The representative tested the surgical arm with a new arm guide and accuracy pointer, but the results were the same. There was no patient involved. 2019-12-10 IUD x5 (Rep): No new information. 2020-01-16 IUD (Other): No new information. 2020-02-26 IUD (Other): No new information. 2020-03-05 IUD (Other): No new information. SURGICAL ARM | 2019-12-09 00679860 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed the setup and advanced accuracy test at the second point. The representative tested the surgical arm with a new arm guide and accuracy pointer, but the results were the same. There was no patient involved. 2019-12-10 IUD x5 (Rep): No new information. 2020-01-16 IUD (Other): No new information. 2020-02-26 IUD (Other): No new information. 2020-03-05 IUD (Other): No new information. | xse sra failed advanced accuracy test outside procedure surgical arm failed setup advanced accuracy test second point tested surgical arm new arm guide accuracy pointer results involved iud new iud new iud new iud new surgical arm | NAV4093 | NAV3249 | NAV2054 | NAV4092 | NAV1006 | NAV4093NAV3249NAV2054NAV4092NAV1006 | True
6914 | 704064180 | XSE - Navigation Off | 2020-11-11T18:47:15.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Navigation Off 2020-11-11 00700675, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the surgeon was drilling during an L5 and S1 procedure, the lateral projection was superior by 3 mm on all four screws. The surgeon decided to go ahead and implant the screws. Post operation imaging showed that all screws were accurately placed. The surgeon wanted to know if tightening the surgical arm would help resolve the issue. There was no patient harm and the procedure was delayed less than an hour. 2020-12-02 IUD (Other): No new information. 2020-12-10 email 2020-12-10 13:00 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved checking navigation accuracy with the pointer in the divot and on the patient's anatomy. The surgeon was okay to move forward and trusted the guidance system was correctly positioned. Navigation was just off in the lateral view. The axial view was accurate. The cause of the navigation was being off was not determined. MAZOR X SYSTEM TPL0059 | 2020-11-11 00700675, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the surgeon was drilling during an L5 and S1 procedure, the lateral projection was superior by 3 mm on all four screws. The surgeon decided to go ahead and implant the screws. Post operation imaging showed that all screws were accurately placed. The surgeon wanted to know if tightening the surgical arm would help resolve the issue. There was no patient harm and the procedure was delayed less than an hour. 2020-12-02 IUD (Other): No new information. 2020-12-10 email 2020-12-10 13:00 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved checking navigation accuracy with the pointer in the divot and on the patient's anatomy. The surgeon was okay to move forward and trusted the guidance system was correctly positioned. Navigation was just off in the lateral view. The axial view was accurate. The cause of the navigation was being off was not determined. | xse navigation ccr spinal procedure surgeon drilling procedure lateral projection superior mm screws surgeon decided ahead implant screws post operation imaging showed screws accurately placed surgeon wanted know tightening surgical arm help resolve issue procedure delayed hour iud new email additional troubleshooting involved checking navigation accuracy pointer divot patients anatomy surgeon okay forward trusted correctly positioned navigation just lateral view axial view accurate cause navigation determined mazor x | NAV5000 | NAV5001 | NAV1012 | NAV3246 | NAV2047 | NAV5000NAV5001NAV1012NAV3246NAV2047 | True
3049 | 703263801 | X - Shoulder Driver Card | 2019-07-17T17:27:00.000000+0000 | MazorX | NAV4106 | CALIBRATION ISSUES | 20 | ELB0066-01 | OTHER CARD, STEPPER DRIVE MODULE | X - Shoulder Driver Card Per Israel instructions, replace shoulder driver PCB and observe findings. OTHER CARD, STEPPER DRIVE MODULE | Per Israel instructions, replace shoulder driver PCB and observe findings. | x shoulder driver card israel instructions replace shoulder driver pcb observe findings card stepper drive module | NAV4106 | NAV3299 | NAV2054 | NAV3246 | NAV4091 | NAV4106NAV3299NAV2054NAV3246NAV4091 | True
7131 | 704125809 | XSE - Metal Tang Broke | 2020-12-15T17:14:13.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 20 | HRS0327-01S | KIT SVC HRS0327-01S TEETH LOCKING SENSOR | XSE - Metal Tang Broke 2020-12-15 00702686 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the metal tang on the teeth broke off and had to be removed from the case. The manufacturer representative was able to reattach the metal tang, but it needed to be replaced. There was no patient harm and the procedure was not delayed. KIT SVC HRS0327-01S TEETH LOCKING SENSOR | 2020-12-15 00702686 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the metal tang on the teeth broke off and had to be removed from the case. The manufacturer representative was able to reattach the metal tang, but it needed to be replaced. There was no patient harm and the procedure was not delayed. | xse metal tang broke spinal procedure metal tang teeth broke removed case able reattach metal tang needed replaced procedure delayed kit svc teeth locking sensor | NAV3251 | NAV4090 | NAV2047 | NAV5000 | NAV3264 | NAV3251NAV4090NAV2047NAV5000NAV3264 | True
5678 | 703776811 | X - No Data Found on Import from | 2020-05-29T18:18:37.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - No Data Found on Import from 2020-05-29 00689788 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while attempting to import an O-arm scan to the system from USB, a "No Data Found" error was received immediately after hitting the import button. The same behavior occurred with two different USBs. After a USB splitter bypass, and after creating a new patient, the manufacturer representative ensured that the USBs were wiped of all files before attempting an import. No patient was present. 2020-06-01 IUD (Other): No new information. 2020-06-15 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2020-05-29 00689788 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that while attempting to import an O-arm scan to the system from USB, a "No Data Found" error was received immediately after hitting the import button. The same behavior occurred with two different USBs. After a USB splitter bypass, and after creating a new patient, the manufacturer representative ensured that the USBs were wiped of all files before attempting an import. No patient was present. 2020-06-01 IUD (Other): No new information. 2020-06-15 IUD (Other): No new information. | x data import outside procedure attempting import oarm scan usb data error immediately hitting import button behavior occurred different usbs usb splitter bypass creating new ensured usbs wiped files attempting import present iud new iud new unk mazor x workstation | NAV2058 | NAV3244 | NAV3261 | NAV2054 | NAV3264 | NAV2058NAV3244NAV3261NAV2054NAV3264 | True
4508 | 703548192 | X - Arm and 3Define camera issues | 2020-01-07T15:10:23.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 30 | ASM0206-01 | SURGICAL ARM | X - Arm and 3Define camera issues 2018-07-20 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on level L3 - S1. The surgeon planned the case day of surgery. The system was attached to the bed and draped. The schanz pin was Inserted into a stable spinous process. 3Define failed. 3Define went through the right side of the surface scan perfectly but when it looked like it was almost completing finishing the left side of the surface scan, the guidance system would beep with a couldn't reach trajectory beep but would have a blank image on the workstation. Surgeon wanted to continue with a generic workstation. Four Ap and three oblique images were taken. Registration was achieved and operation started. Skin was cut and fluoro images taken. The operation started from left L3 - S1 and then Right L3 - S1. Confirmation images were taken, and suture done. During the procedure; the Arm shook in a finished sent trajectory position. Arm would not finish sending to any predefine trajectories. 3Define would not finish and would have no images of the surface scan. L3 and S1 trajectories were accurate. No patient harm was reported. SURGICAL ARM | 2018-07-20 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on level L3 - S1. The surgeon planned the case day of surgery. The system was attached to the bed and draped. The schanz pin was Inserted into a stable spinous process. 3Define failed. 3Define went through the right side of the surface scan perfectly but when it looked like it was almost completing finishing the left side of the surface scan, the guidance system would beep with a couldn't reach trajectory beep but would have a blank image on the workstation. Surgeon wanted to continue with a generic workstation. Four Ap and three oblique images were taken. Registration was achieved and operation started. Skin was cut and fluoro images taken. The operation started from left L3 - S1 and then Right L3 - S1. Confirmation images were taken, and suture done. During the procedure; the Arm shook in a finished sent trajectory position. Arm would not finish sending to any predefine trajectories. 3Define would not finish and would have no images of the surface scan. L3 and S1 trajectories were accurate. No patient harm was reported. | x arm camera issues ccr spinal procedure case vertebrae level surgeon planned case day surgery attached bed draped schanz pin inserted stable spinous process failed went right surface scan perfectly looked like completing finishing left surface scan beep reach trajectory beep blank image workstation surgeon wanted continue generic workstation ap oblique images taken registration achieved operation started skin cut fluoro images taken operation started left right confirmation images taken suture procedure arm shook finished sent trajectory position arm finish sending predefine trajectories finish images surface scan trajectories accurate surgical arm | NAV4095 | NAV3244 | NAV2058 | NAV3246 | NAV4091 | NAV4095NAV3244NAV2058NAV3246NAV4091 | False
8084 | 704348985 | XSE - BMB Screw Stripped | 2021-04-19T20:17:06.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5196-01S | KIT SVC MEC5196-01S RNX-ACC-063 | XSE - BMB Screw Stripped 2021-Apr-19 Salesforce C00224901 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws on the bone mount bridge that clamp down on the Schanz bond were stripped. There was no patient involvement. 2021-04-22 IUD x3 (Other): No new information. 2021-07-28 email 2021-07-28 22:15 (Rep): Additional information received from a manufacturer representative reported the cause of the stripped screws was wear and tear from consistent use. The screws stripped a little bit every time the screw was tightened and they got to the point where they could no longer be used. KIT SVC MEC5196-01S RNX-ACC-063 | 2021-Apr-19 Salesforce C00224901 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws on the bone mount bridge that clamp down on the Schanz bond were stripped. There was no patient involvement. 2021-04-22 IUD x3 (Other): No new information. 2021-07-28 email 2021-07-28 22:15 (Rep): Additional information received from a manufacturer representative reported the cause of the stripped screws was wear and tear from consistent use. The screws stripped a little bit every time the screw was tightened and they got to the point where they could no longer be used. | xse bmb screw stripped salesforce outside procedure screws bone mount bridge clamp schanz bond stripped iud new email additional cause stripped screws wear tear consistent use screws stripped little bit time screw tightened got point longer kit svc | NAV3299 | NAV4090 | NAV4100 | NAV4093 | NAV3256 | NAV3299NAV4090NAV4100NAV4093NAV3256 | True
31 | 702971867 | MazorX ¦ The shoulder will not properly | 2019-01-23T16:54:07.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | MazorX ¦ The shoulder will not properly When the shoulder is moved all the way to the left the shoulder will not lock. MAZOR X SYSTEM TPL0059 | When the shoulder is moved all the way to the left the shoulder will not lock. | mazorx shoulder properly shoulder moved way left shoulder lock mazor x | NAV4107 | NAV4091 | NAV3242 | NAV3246 | NAV2054 | NAV4107NAV4091NAV3242NAV3246NAV2054 | True
929 | 703181017 | X - Deviation issue | 2019-05-30T16:06:30.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviation issue This complaint will be separated to two complaints. This complaint will handle the deviation issue. Procedure last night around 9pm. We were placing the robot onto the bed rail and only one of the two teeth were able to clamp on. The robot locked and it took some force to get the robot from the bedrail. We had to override and force teeth open. The robot would still not let go. We pulled the robot from the bedrail and there was some jostling of the robot. We then successfully were able to mount the robot to the bed. We did not re-assess if the robot shoulder parameters were within the correct amounts. We then began the surgery per usual. We got through all verification steps and put k-wires in. On our confirmation k-wire fluoro shots we noticed that the L4 k-wires were inferior and somewhat medial. We opted to re-register thinking this would help. The robot maybe bumped our trajectory up a millimeter. After this, we decided to bail on the screws and bring O-Arm in. We confirmed each previous trajectory with the Nav pointer in pictures (I can share if needed). Because the case was late we had to run tests this morning. All 3 point, 10 point and left/right tests were run. Before I ran the 10 point test I noticed the robot parameters at Front:96 and Back-4. After the test, each position checking out, the parameters were within the normal +/- 5. In my conclusion, I believe the jostling caused the robot to come out of the +/- 5 parameters and thus put our trajectories in the wrong spot even though we had green from each vertebral body/no shifts. Unaware that we needed to check the robot parameters, we went ahead with the case. I look forward to hearing your thoughts on this matter. I have sent in the entire aborted case file MAZOR X SOFTWARE KIT VER-3.0.2 | This complaint will be separated to two complaints. This complaint will handle the deviation issue. Procedure last night around 9pm. We were placing the robot onto the bed rail and only one of the two teeth were able to clamp on. The robot locked and it took some force to get the robot from the bedrail. We had to override and force teeth open. The robot would still not let go. We pulled the robot from the bedrail and there was some jostling of the robot. We then successfully were able to mount the robot to the bed. We did not re-assess if the robot shoulder parameters were within the correct amounts. We then began the surgery per usual. We got through all verification steps and put k-wires in. On our confirmation k-wire fluoro shots we noticed that the L4 k-wires were inferior and somewhat medial. We opted to re-register thinking this would help. The robot maybe bumped our trajectory up a millimeter. After this, we decided to bail on the screws and bring O-Arm in. We confirmed each previous trajectory with the Nav pointer in pictures (I can share if needed). Because the case was late we had to run tests this morning. All 3 point, 10 point and left/right tests were run. Before I ran the 10 point test I noticed the robot parameters at Front:96 and Back-4. After the test, each position checking out, the parameters were within the normal +/- 5. In my conclusion, I believe the jostling caused the robot to come out of the +/- 5 parameters and thus put our trajectories in the wrong spot even though we had green from each vertebral body/no shifts. Unaware that we needed to check the robot parameters, we went ahead with the case. I look forward to hearing your thoughts on this matter. I have sent in the entire aborted case file | x deviation issue complaint separated complaints complaint handle deviation issue procedure night placing robot bed rail teeth able clamp robot locked took force robot bedrail override force teeth open robot let pulled robot bedrail jostling robot successfully able mount robot bed did reassess robot shoulder parameters correct amounts began surgery usual got verification steps kwires confirmation kwire fluoro shots noticed kwires inferior somewhat medial opted reregister thinking help robot maybe bumped trajectory millimeter decided bail screws bring oarm confirmed previous trajectory nav pointer pictures case late run tests morning point point leftright tests run ran point test noticed robot parameters test position checking parameters normal conclusion believe jostling caused robot come parameters trajectories wrong spot green vertebral bodyno shifts unaware needed check robot parameters went ahead case look forward hearing thoughts matter sent entire aborted case file mazor x software kit | NAV3250 | NAV4091 | NAV2051 | NAV4093 | NAV3246 | NAV3250NAV4091NAV2051NAV4093NAV3246 | True
8395 | 704405852 | X - X-eye Screws Missing | 2021-05-20T15:55:13.000000+0000 | MazorX | NAV3254 | X-EYE ISSUES | 20 | SCW0784-01S | KIT SVC SCW0784-01S PHLPS PAN 1/4 X 3/8 | X - X-eye Screws Missing 2021-May-20 Salesforce C00226762 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws that hold the x-eye camera in place are missing. No patient was present 2021-05-25 IUD x2 (Other): No new information. KIT SVC SCW0784-01S PHLPS PAN 1/4 X 3/8 | 2021-May-20 Salesforce C00226762 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws that hold the x-eye camera in place are missing. No patient was present 2021-05-25 IUD x2 (Other): No new information. | x xeye screws missing salesforce outside procedure screws hold xeye camera place missing present iud new kit svc phlps pan x | NAV2053 | NAV3254 | NAV4110 | NAV4091 | NAV2054 | NAV2053NAV3254NAV4110NAV4091NAV2054 | True
6615 | 703988734 | XSE ¦ System Not Booting | 2020-10-01T16:06:10.000000+0000 | MazorX Stealth Edition | NAV3248 | UNRESPONSIVE | 20 | UNK_MXSE_SW | UNKNOWN MAZOR X SE SW | XSE ¦ System Not Booting 2020-10-01 00697859 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that they started the system up and received the dell splash logo. It was stuck on the splash logo for 10 to 15 mins and it never got past this stage. The manufacturer representative powered the system down using the green power switch. When the system was turned back on the screen was black but the PC was powered on. Attempted another hard restart to resolve the issue. No patient was present. 2020-10-12 IUD (Other): A hard restart resolved the issue. UNKNOWN MAZOR X SE SW | 2020-10-01 00697859 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that they started the system up and received the dell splash logo. It was stuck on the splash logo for 10 to 15 mins and it never got past this stage. The manufacturer representative powered the system down using the green power switch. When the system was turned back on the screen was black but the PC was powered on. Attempted another hard restart to resolve the issue. No patient was present. 2020-10-12 IUD (Other): A hard restart resolved the issue. | xse booting outside procedure started dell splash logo stuck splash logo mins got past stage powered using green power switch turned screen black pc powered attempted hard restart resolve issue present iud hard restart resolved issue unknown mazor x se sw | NAV3248 | NAV2047 | NAV3262 | NAV3246 | NAV3261 | NAV3248NAV2047NAV3262NAV3246NAV3261 | True
7 | 702969746 | MazorX ¦ Schanz Bond Breakage | 2019-01-22T20:48:40.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | MazorX ¦ Schanz Bond Breakage MAS2271 Shanz Bond is stripped MAZOR X SYSTEM TPL0059 | MAS2271 Shanz Bond is stripped | mazorx schanz bond breakage shanz bond stripped mazor x | NAV3295 | NAV4091 | NAV3246 | NAV3242 | NAV3299 | NAV3295NAV4091NAV3246NAV3242NAV3299 | True
7428 | 704201139 | XSE ¦ Unable to Remove SRS From Bed | 2021-01-28T21:38:08.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | XSE ¦ Unable to Remove SRS From Bed 2021-01-28 00705386 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that after a case, when the manufacturer representative was trying to take the system off the bed, the manipulator would securely click into the back of the positioner column, but the positioner manipulator interface would no go green. She tried to go into the override panel to open the teeth, but that did not work. She did a soft reboot and then a hard reboot. She was still unable to get the system off, even though the positioner manipulator interface did go green after the hard reboot. She ended up taking the surgical system off the bed still attached to the bed frame adapter. This caused about a 10 minute delay, but no patient impact. 2021-01-29 IUD x3 (Other): No new information. 2021-Feb-23 IUD (other): No new information. UNK MAZOR X POSITIONER | 2021-01-28 00705386 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that after a case, when the manufacturer representative was trying to take the system off the bed, the manipulator would securely click into the back of the positioner column, but the positioner manipulator interface would no go green. She tried to go into the override panel to open the teeth, but that did not work. She did a soft reboot and then a hard reboot. She was still unable to get the system off, even though the positioner manipulator interface did go green after the hard reboot. She ended up taking the surgical system off the bed still attached to the bed frame adapter. This caused about a 10 minute delay, but no patient impact. 2021-01-29 IUD x3 (Other): No new information. 2021-Feb-23 IUD (other): No new information. | xse unable remove srs bed intraoperatively spinal procedure case trying bed manipulator securely click positioner column positioner manipulator interface green tried override panel open teeth did work did soft reboot hard reboot unable positioner manipulator interface did green hard reboot ended taking surgical bed attached bed frame adapter caused minute delay impact iud new iud new unk mazor x positioner | NAV3250 | NAV4103 | NAV3260 | NAV4104 | NAV3264 | NAV3250NAV4103NAV3260NAV4104NAV3264 | True
743 | 703175214 | X - x-eye loose and wobbling | 2019-05-28T12:15:24.000000+0000 | MazorX | NAV4109 | LOOSE | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - x-eye loose and wobbling The X-eye is loose and wobbling UNK MAZOR X X-EYE CAMERA | The X-eye is loose and wobbling | x xeye loose wobbling xeye loose wobbling unk mazor x xeye camera | NAV4109 | NAV4110 | NAV4097 | NAV3254 | NAV2054 | NAV4109NAV4110NAV4097NAV3254NAV2054 | True
6542 | 703971291 | XSE - Failed Registration/Inaccuracy | 2020-09-22T20:00:58.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 30 | M450001B020 | NAVIGATION SOFTWARE MODULE 1.1 | XSE - Failed Registration/Inaccuracy 2020-09-22 00697234, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that four registration attempts were made and green values were achieved, but when they sent the surgical arm to the trajectory, the surgeon did not believe the arm was accurate. After registration, the surgical arm was sent to the trajectory. The dilator was placed through the arm guide and the it sunk through the facet and into the pedicle on navigation. Navigation was checked and was accurate. The surgeon sued the drill bit and tap to see if the instrument depth was better, but it was not. The surgical arm was sent to another trajectory, but the issue was not resolved. Registration was done again and there were no issues with finding the levels. The dilator was placed down the arm guide and there was a small mismatch between the depth of the instrument and bone. The mismatch was not enough so the surgeon was ok with moving forward. The depth at L5 was off similar to the previous registration. The surgical arm was sent to right L5 and the same issue occurred. Right L4 was then drilled and tapped using fluoro. Left L4 was then drilled and tapped under fluoro, but neuromonitoring levels hit 16, 11 and 7. Registration was done again and a yellow solution was shown for L5. The surgeon confirmed the new images. The levels were adjusted, but L5 remained yellow. The surgeon continued the procedure and all screws were confirmed to be accurately placed when checked with a post op O-arm spin. The surgeon stated there was possible movement in the surgical arm. There was no patient harm and the procedure was delayed less than an hour. 2020-09-25 IUD x3 (Other): No new information. 2020-10-06 e1 (Rep): Additional information was received stating there was an inaccuracy of about 3.5-10mm and the cause of inaccuracy was either patient movement or a false green. 2020-10-08 IUD (Other): No new information. 2020-12-03 IUD (Other): No new information. 2021-01-26 IUD (Other): No new information. 2021-02-04 (Other): No new information. 2021-03-16 IUD (Other): No new information. NAVIGATION SOFTWARE MODULE 1.1 | 2020-09-22 00697234, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that four registration attempts were made and green values were achieved, but when they sent the surgical arm to the trajectory, the surgeon did not believe the arm was accurate. After registration, the surgical arm was sent to the trajectory. The dilator was placed through the arm guide and the it sunk through the facet and into the pedicle on navigation. Navigation was checked and was accurate. The surgeon sued the drill bit and tap to see if the instrument depth was better, but it was not. The surgical arm was sent to another trajectory, but the issue was not resolved. Registration was done again and there were no issues with finding the levels. The dilator was placed down the arm guide and there was a small mismatch between the depth of the instrument and bone. The mismatch was not enough so the surgeon was ok with moving forward. The depth at L5 was off similar to the previous registration. The surgical arm was sent to right L5 and the same issue occurred. Right L4 was then drilled and tapped using fluoro. Left L4 was then drilled and tapped under fluoro, but neuromonitoring levels hit 16, 11 and 7. Registration was done again and a yellow solution was shown for L5. The surgeon confirmed the new images. The levels were adjusted, but L5 remained yellow. The surgeon continued the procedure and all screws were confirmed to be accurately placed when checked with a post op O-arm spin. The surgeon stated there was possible movement in the surgical arm. There was no patient harm and the procedure was delayed less than an hour. 2020-09-25 IUD x3 (Other): No new information. 2020-10-06 e1 (Rep): Additional information was received stating there was an inaccuracy of about 3.5-10mm and the cause of inaccuracy was either patient movement or a false green. 2020-10-08 IUD (Other): No new information. 2020-12-03 IUD (Other): No new information. 2021-01-26 IUD (Other): No new information. 2021-02-04 (Other): No new information. 2021-03-16 IUD (Other): No new information. | xse failed registrationinaccuracy iud spinal procedure registration attempts green values achieved sent surgical arm trajectory surgeon did believe arm accurate registration surgical arm sent trajectory dilator placed arm guide sunk facet pedicle navigation navigation checked accurate surgeon sued drill bit tap instrument depth better surgical arm sent trajectory issue resolved registration issues finding levels dilator placed arm guide small mismatch depth instrument bone mismatch surgeon ok moving forward depth similar previous registration surgical arm sent right issue occurred right drilled tapped using fluoro left drilled tapped fluoro neuromonitoring levels hit registration yellow solution shown surgeon confirmed new images levels adjusted remained yellow surgeon continued procedure screws confirmed accurately placed checked post op oarm spin surgeon stated possible movement surgical arm procedure delayed hour iud new additional stating inaccuracy cause inaccuracy movement false green iud new iud new iud new new iud new navigation software module | NAV3242 | NAV5001 | NAV3249 | NAV4101 | NAV4095 | NAV3242NAV5001NAV3249NAV4101NAV4095 | True
1810 | 703217962 | X - Surgical Arm Issue | 2019-06-20T14:32:35.000000+0000 | MazorX | NAV3255 | BED FRAME ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Surgical Arm Issue Software issues: yes, 3 define failed once; also when mount button was hit got a message that said 3 define failed; software glitch when robot wouldn’t lock Mechanical issues: yes, only one side locked to bed, then wouldn’t unlock; after shutting down system and rebooting robot locked to bed adapter properly Ryan Kayvan was able to walk them through the mounting process. This was a user error. Still need to address the 3Define issues. UNK MAZOR X WORKSTATION | Software issues: yes, 3 define failed once; also when mount button was hit got a message that said 3 define failed; software glitch when robot wouldn’t lock Mechanical issues: yes, only one side locked to bed, then wouldn’t unlock; after shutting down system and rebooting robot locked to bed adapter properly Ryan Kayvan was able to walk them through the mounting process. This was a user error. Still need to address the 3Define issues. | x surgical arm issue software issues yes define failed once; mount button hit got message said define failed; software glitch robot wouldn’t lock mechanical issues yes locked bed wouldn’t unlock; shutting rebooting robot locked bed adapter properly ryan kayvan able walk mounting process user error need address issues unk mazor x workstation | NAV3244 | NAV3250 | NAV3264 | NAV4107 | NAV2047 | NAV3244NAV3250NAV3264NAV4107NAV2047 | False
8175 | 704364741 | XSE - Unable to Send to Trajectory | 2021-04-28T00:22:45.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - Unable to Send to Trajectory 2021-Apr-27 Salesforce C00225416 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that they had difficulties doing draw spine in a case today that ended up aborting the case. The manufacturer representative called in after the procedure was aborted already, so no troubleshooting was performed. They were doing draw spine and touched the head and foot positions, but when they sent the arm she said it was going too far lateral. They tried to re-do draw spine, but the same issue happened. They tried to send the arm to the next trajectory, but it failed to reach trajectory. They aborted the case at this point. There was no impact on the patient outcome. 2021-05-04 IUD x3 (Other): No new information. 2021-09-07 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-Apr-27 Salesforce C00225416 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that they had difficulties doing draw spine in a case today that ended up aborting the case. The manufacturer representative called in after the procedure was aborted already, so no troubleshooting was performed. They were doing draw spine and touched the head and foot positions, but when they sent the arm she said it was going too far lateral. They tried to re-do draw spine, but the same issue happened. They tried to send the arm to the next trajectory, but it failed to reach trajectory. They aborted the case at this point. There was no impact on the patient outcome. 2021-05-04 IUD x3 (Other): No new information. 2021-09-07 IUD (Other): No new information. | xse unable send trajectory salesforce spinal p rocedure difficulties doing draw spine case today ended aborting case called procedure aborted troubleshooting performed doing draw spine touched head foot positions sent arm said going far lateral tried redo draw spine issue happened tried send arm trajectory failed reach trajectory aborted case point impact outcome iud new iud new kit svc surgical arm refurb | NAV4095 | NAV4112 | NAV4107 | NAV3244 | NAV4104 | NAV4095NAV4112NAV4107NAV3244NAV4104 | False
4156 | 703505989 | XSE - Inaccurate Screw Placement | 2019-12-06T23:22:03.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Inaccurate Screw Placement 2019-12-05 00679807 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during the case from L2 to L5, all of the screws on the patient's left side were accurate, but all of the screws on the patient's right side were medially deviated. The screws on the left side were placed by one surgeon and the screws on the right side were placed by a second surgeon. At the beginning of the procedure, a successful 3Define scan was done and a successful snapshot was taken. A spin with the imaging system was done and the passive planar probe was used to send the surgical arm to the 3D marker position. Another spin was taken and then the snapshot was redone in case the drape pushed the reference frame. Planning was done and the surgeon started placing the screws. The surgeon started at L2 and ended at L5, alternating between left and right at each level. The screws on the right side were removed and repositioned. The representative suspected the deviations were caused by skin pressure and mechanical deviation from the tools. There was no patient harm and the procedure was delayed over an hour. 2019-12-10 Email 2019-12-10 16:00, email 2019-12-18 18:10, IUD x2, e1, e2 (REP): Additional information was received stating that the trajectories were deviated between 3.5 and 10mm. There was no known impact on the patient outcome. Neuro-monitoring reported no changes in motors. 2019-12-18 IUD x2 (Other): No new information. 2020-01-07 email 2020-01-07 17:45, e3 (Rep): No new information. 2020-02-19 Medwatch form, attachment: No new information. MAZOR X SYSTEM TPL0059 | 2019-12-05 00679807 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during the case from L2 to L5, all of the screws on the patient's left side were accurate, but all of the screws on the patient's right side were medially deviated. The screws on the left side were placed by one surgeon and the screws on the right side were placed by a second surgeon. At the beginning of the procedure, a successful 3Define scan was done and a successful snapshot was taken. A spin with the imaging system was done and the passive planar probe was used to send the surgical arm to the 3D marker position. Another spin was taken and then the snapshot was redone in case the drape pushed the reference frame. Planning was done and the surgeon started placing the screws. The surgeon started at L2 and ended at L5, alternating between left and right at each level. The screws on the right side were removed and repositioned. The representative suspected the deviations were caused by skin pressure and mechanical deviation from the tools. There was no patient harm and the procedure was delayed over an hour. 2019-12-10 Email 2019-12-10 16:00, email 2019-12-18 18:10, IUD x2, e1, e2 (REP): Additional information was received stating that the trajectories were deviated between 3.5 and 10mm. There was no known impact on the patient outcome. Neuro-monitoring reported no changes in motors. 2019-12-18 IUD x2 (Other): No new information. 2020-01-07 email 2020-01-07 17:45, e3 (Rep): No new information. 2020-02-19 Medwatch form, attachment: No new information. | xse inaccurate screw placement spinal procedure case screws patients left accurate screws patients right medially deviated screws left placed surgeon screws right placed second surgeon beginning procedure successful scan successful snapshot taken spin imaging passive planar probe send surgical arm marker position spin taken snapshot redone case drape pushed reference frame planning surgeon started placing screws surgeon started ended alternating left right level screws right removed repositioned suspected deviations caused skin pressure mechanical deviation tools procedure delayed hour email email iud additional stating trajectories deviated known impact outcome neuromonitoring changes motors iud new email new medwatch attachment new mazor x | NAV5000 | NAV4091 | NAV3243 | NAV1012 | NAV3299 | NAV5000NAV4091NAV3243NAV1012NAV3299 | True
5609 | 703764323 | XSE - Deviated Screws at L5 and L4 | 2020-05-21T14:14:52.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Screws at L5 and L4 2020-05-20 00689315, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a single clamp was placed at L4 and connected to the bone mount bridge to mount the surgical system to the patient for the L4-S1 procedure. The re-tractors were left in place during the entire procedure. Registration was achieved on the first try with green values and no observable shifts. The surgeon started with L4 right and placed the k-wire, tapped and placed the screw. The surgical arm was then sent to the clear up position and the surgeon moved to the left side. Imaging showed that L5 left and right were high and L4 left was medial when compared to plan. The surgeon removed screws and redirected them freehand. The surgeon believed the other screws to be off except for S1, but the surgeon decided not to redirect those screws. There was no patient harm and the procedure was not delayed over an hour. 2020-05-28 IUD x3 (Other): No new information. 2020-06-10 IUD (Other): No new information. 2020-06-12 email 2020-06-12 10:15 (Rep): Additional information received from a manufacturer representative reported the trajectories were deviated less than 3.5 mm. The suspected cause of the deviated screws was the way the clam was mounted and not all of the teeth being engaged. There was no patient harm from the deviated screws. MAZOR X SYSTEM TPL0059 | 2020-05-20 00689315, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a single clamp was placed at L4 and connected to the bone mount bridge to mount the surgical system to the patient for the L4-S1 procedure. The re-tractors were left in place during the entire procedure. Registration was achieved on the first try with green values and no observable shifts. The surgeon started with L4 right and placed the k-wire, tapped and placed the screw. The surgical arm was then sent to the clear up position and the surgeon moved to the left side. Imaging showed that L5 left and right were high and L4 left was medial when compared to plan. The surgeon removed screws and redirected them freehand. The surgeon believed the other screws to be off except for S1, but the surgeon decided not to redirect those screws. There was no patient harm and the procedure was not delayed over an hour. 2020-05-28 IUD x3 (Other): No new information. 2020-06-10 IUD (Other): No new information. 2020-06-12 email 2020-06-12 10:15 (Rep): Additional information received from a manufacturer representative reported the trajectories were deviated less than 3.5 mm. The suspected cause of the deviated screws was the way the clam was mounted and not all of the teeth being engaged. There was no patient harm from the deviated screws. | xse deviated screws spinal procedure single clamp placed connected bone mount bridge mount surgical procedure retractors left place entire procedure registration achieved try green values observable shifts surgeon started right placed kwire tapped placed screw surgical arm sent clear position surgeon moved left imaging showed left right high left medial compared plan surgeon removed screws redirected freehand surgeon believed screws surgeon decided redirect screws procedure delayed hour iud new iud new email additional trajectories deviated mm suspected cause deviated screws way clam mounted teeth engaged deviated screws mazor x | NAV5000 | NAV1012 | NAV3250 | NAV3253 | NAV2047 | NAV5000NAV1012NAV3250NAV3253NAV2047 | True
3836 | 703453595 | XSE - Bed Frame Adapter Loose | 2019-11-06T20:05:21.000000+0000 | MazorX Stealth Edition | NAV3255 | BED FRAME ISSUES | 30 | MAS1574-01 | BED FRAME ADAPTOR | XSE - Bed Frame Adapter Loose 2019-Nov-06 00677766 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the bed frame adaptor does not lock and the locking mechanism appears to be loose. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-18 e3 (Rep): Additional information received from a manufacturer representative reported the replacement bed frame adaptor was used and very beat up. 2019-11-26 e5, e7, e9, IUD x2 (Rep): No new information 2020-02-09 IUD x2 (Rep): No new information 2020-02-26 IUD (Other): No new information. BED FRAME ADAPTOR | 2019-Nov-06 00677766 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the bed frame adaptor does not lock and the locking mechanism appears to be loose. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-18 e3 (Rep): Additional information received from a manufacturer representative reported the replacement bed frame adaptor was used and very beat up. 2019-11-26 e5, e7, e9, IUD x2 (Rep): No new information 2020-02-09 IUD x2 (Rep): No new information 2020-02-26 IUD (Other): No new information. | xse bed frame adapter loose spinal procedure bed frame adaptor does lock locking mechanism appears loose additional complications anticipated additional replacement bed frame adaptor beat iud new iud new iud new bed frame adaptor | NAV3255 | NAV3299 | NAV2054 | NAV3295 | NAV3250 | NAV3255NAV3299NAV2054NAV3295NAV3250 | True
6860 | 704051740 | XSE - Compressor Not Turning On | 2020-11-04T18:43:33.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | ASM0263-02S | KIT SVC ASM0263-02S PNEUMATIC PANEL | XSE - Compressor Not Turning On 2020-11-03 00700203 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the compressor in the guidance system would not turn on. There did not seem to see to be a leak, but the pressure dropped when the compressor turned on. There was no patient involvement. 2021-02-04 (Other): No new information. KIT SVC ASM0263-02S PNEUMATIC PANEL | 2020-11-03 00700203 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the compressor in the guidance system would not turn on. There did not seem to see to be a leak, but the pressure dropped when the compressor turned on. There was no patient involvement. 2021-02-04 (Other): No new information. | xse compressor turning outside procedure compressor turn did leak pressure dropped compressor turned new kit svc pneumatic panel | NAV3264 | NAV2054 | NAV1006 | NAV4101 | NAV4093 | NAV3264NAV2054NAV1006NAV4101NAV4093 | True
3758 | 703436191 | XSE - Fatal Canbus lost connection | 2019-10-28T16:00:57.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Fatal Canbus lost connection 2019-Oct-28 00676974 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Mazor X lost connection to all nodes during the case. Hard restart performed and resolved the issue. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-06 IUD x5 (Rep): No new information. UNK MAZOR X WORKSTATION | 2019-Oct-28 00676974 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Mazor X lost connection to all nodes during the case. Hard restart performed and resolved the issue. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-06 IUD x5 (Rep): No new information. | xse fatal canbus lost connection spinal procedure mazor x lost connection nodes case hard restart performed resolved issue additional complications anticipated iud new unk mazor x workstation | NAV2047 | NAV3261 | NAV2054 | NAV2053 | NAV3262 | NAV2047NAV3261NAV2054NAV2053NAV3262 | True
7198 | 704137630 | XSE - Damaged Bed Frame Adaptor Screw | 2020-12-22T01:18:32.000000+0000 | MazorX Stealth Edition | NAV3255 | BED FRAME ISSUES | 20 | KIT1113 | BED FRAME ADAPTER UPGRADE KIT KIT1113 | XSE - Damaged Bed Frame Adaptor Screw 2020-Dec-21 00703181 (rep): Medtronic received information from a manufacturer representative (rep) regarding a device being used outside a procedure. It was reported that one of the screws on the bed frame adaptor is damaged and it wont tighten. There was no patient harm reported as a result of the event. 2020-12-28 IUD x5 (Other): No new information. 2021-Jan-22 IUD (other): No new information. 2021-02-15 IUD x4 (Other): No new information. BED FRAME ADAPTER UPGRADE KIT KIT1113 | 2020-Dec-21 00703181 (rep): Medtronic received information from a manufacturer representative (rep) regarding a device being used outside a procedure. It was reported that one of the screws on the bed frame adaptor is damaged and it wont tighten. There was no patient harm reported as a result of the event. 2020-12-28 IUD x5 (Other): No new information. 2021-Jan-22 IUD (other): No new information. 2021-02-15 IUD x4 (Other): No new information. | xse damaged bed frame adaptor screw outside procedure screws bed frame adaptor damaged wont tighten iud new iud new iud new bed frame adapter upgrade kit | NAV3255 | NAV3299 | NAV2054 | NAV3258 | NAV1006 | NAV3255NAV3299NAV2054NAV3258NAV1006 | True
1237 | 703193438 | X - Software Issue Bug | 2019-06-06T18:14:31.000000+0000 | MazorX | NAV2051 | INTRA-OP ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Software Issue Bug Event Description: After O-arm 3D scan, during planification, one image was shown on the axial view, as soon as we scrolled through the slice the image desapeared, even after stack up icon clicked, the axial view appeared and disappeared right after as soon as we scroll the mouse or if we click on the "play" icon to see the differents slice. We had to play with green lines on other view LT and AP to have at least on image on AX view to add implants. We never got to manage with this issue on the previous cases, scrolling through the slices was easy to do. Note: This time the 3D scan was performed in HD mode by surgeon decision, number of images for a HD scan are 750 Vs about 400 images for a normal 3D. 2018-05-24 CCR (REP, HCP): Medtronic received additional information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on level L4 - L5. Arm was attached to the bed and schanz pin inserted into left iliac. First 3D was transferred to the guidance system. There were difficulties to have the axial view even after stack up clicked. The axial view appeared and disappeared. Surgeon decided to redo 3D scan. Second 3D scan was transferred to the guidance system. Again, difficulties to scroll on the slices were experienced and there was black AX view during planification. L4 left was drilled. Surgeon was not able to remove the drill guide and used a mallet, so the shoulder got an error. Third 3D scan was sent to the guidance system. Previous patient scan not deleted so a new full scan had to be redone. Fourth 3D scan was transferred to the guidance system and four screws were placed accurately on L4 and L5. No patient harm was reported. MAZOR X SYSTEM TPL0059 | Event Description: After O-arm 3D scan, during planification, one image was shown on the axial view, as soon as we scrolled through the slice the image desapeared, even after stack up icon clicked, the axial view appeared and disappeared right after as soon as we scroll the mouse or if we click on the "play" icon to see the differents slice. We had to play with green lines on other view LT and AP to have at least on image on AX view to add implants. We never got to manage with this issue on the previous cases, scrolling through the slices was easy to do. Note: This time the 3D scan was performed in HD mode by surgeon decision, number of images for a HD scan are 750 Vs about 400 images for a normal 3D. 2018-05-24 CCR (REP, HCP): Medtronic received additional information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on level L4 - L5. Arm was attached to the bed and schanz pin inserted into left iliac. First 3D was transferred to the guidance system. There were difficulties to have the axial view even after stack up clicked. The axial view appeared and disappeared. Surgeon decided to redo 3D scan. Second 3D scan was transferred to the guidance system. Again, difficulties to scroll on the slices were experienced and there was black AX view during planification. L4 left was drilled. Surgeon was not able to remove the drill guide and used a mallet, so the shoulder got an error. Third 3D scan was sent to the guidance system. Previous patient scan not deleted so a new full scan had to be redone. Fourth 3D scan was transferred to the guidance system and four screws were placed accurately on L4 and L5. No patient harm was reported. | x software issue bug description oarm scan planification image shown axial view soon scrolled slice image desapeared stack icon clicked axial view appeared disappeared right soon scroll mouse click play icon differents slice play green lines view lt ap image ax view add implants got manage issue previous cases scrolling slices easy note time scan performed hd mode surgeon decision number images hd scan vs images normal ccr additional spinal procedure case vertebrae level arm attached bed schanz pin inserted left iliac transferred difficulties axial view stack clicked axial view appeared disappeared surgeon decided redo scan second scan transferred difficulties scroll slices experienced black ax view planification left drilled surgeon able remove drill guide mallet shoulder got error scan sent previous scan deleted new scan redone fourth scan transferred screws placed accurately mazor x | NAV3246 | NAV4091 | NAV3252 | NAV3299 | NAV3242 | NAV3246NAV4091NAV3252NAV3299NAV3242 | False
9789 | 704719127 | XSE - OOBF/Instruments Not Working | 2021-11-17T00:10:55.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | KIT2382 | KIT2382 BRIDGE ADAPTOR KIT | XSE - OOBF/Instruments Not Working 2021-Nov-16 Salesforce C00238262 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during a dry run with the guidance system, two pieces from the bone mounting tray were not working properly. The screw in knob that attached the two parts were extremely difficult to thread into one another. The manufacturer representative noted that the trays had not been used before. There was no patient involvement. 2021-11-17 email 2021-11-17 09:32, image 2021-11-17 09:32 x4 (Rep): No new information. 2021-11-24 e1, IUD (Other): No new information. 2021-12-05 e2, e3, e4 (Rep): No new information. 2021-12-07 IUD (Other): No new information. 2021-12-13 e5, IUD (Other): No new information. 2021-12-21 IUD x2 (Other): No new information. 2022-01-06 IUD (Other): No new information. KIT2382 BRIDGE ADAPTOR KIT | 2021-Nov-16 Salesforce C00238262 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during a dry run with the guidance system, two pieces from the bone mounting tray were not working properly. The screw in knob that attached the two parts were extremely difficult to thread into one another. The manufacturer representative noted that the trays had not been used before. There was no patient involvement. 2021-11-17 email 2021-11-17 09:32, image 2021-11-17 09:32 x4 (Rep): No new information. 2021-11-24 e1, IUD (Other): No new information. 2021-12-05 e2, e3, e4 (Rep): No new information. 2021-12-07 IUD (Other): No new information. 2021-12-13 e5, IUD (Other): No new information. 2021-12-21 IUD x2 (Other): No new information. 2022-01-06 IUD (Other): No new information. | xse oobfinstruments working salesforce outside procedure dry run pieces bone mounting tray working properly screw knob attached parts extremely difficult thread noted trays email image new iud new new iud new iud new iud new iud new bridge adaptor kit | NAV2054 | NAV3299 | NAV2053 | NAV3249 | NAV3244 | NAV2054NAV3299NAV2053NAV3249NAV3244 | True
7025 | 704102047 | XSE - OOBF/PC Replacement | 2020-12-03T03:00:05.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - OOBF/PC Replacement 2020-Dec-02 00701821 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that this system needs a new computer since being returned from the hospital. The computer was received and was bad on install. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-06 IUD (Other): No new information. 2020-12-21 IUD x6 (Other): No new information. 2021-01-20 IUD x3 (Other): No new information. 2021-06-20 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2020-Dec-02 00701821 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that this system needs a new computer since being returned from the hospital. The computer was received and was bad on install. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-06 IUD (Other): No new information. 2020-12-21 IUD x6 (Other): No new information. 2021-01-20 IUD x3 (Other): No new information. 2021-06-20 IUD (Other): No new information. | xse oobfpc replacement outside procedure needs new computer returned hospital computer bad install additional complications anticipated iud new iud new iud new iud new mazor x pc station refurbished | NAV3262 | NAV2054 | NAV3261 | NAV1006 | NAV2047 | NAV3262NAV2054NAV3261NAV1006NAV2047 | True
5451 | 703713251 | XSE - 3Define Camera Not Functioning | 2020-04-16T18:32:54.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 40 | ASM0357-01 | SURGEON SCREEN | XSE - 3Define Camera Not Functioning 2020-04-16 00687496 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that when the system was started the 3define camera was red. Performed a hard restart and that did not resolve the issue. The FW was incorrect and camera cover was missing. Troubleshooting involved verifying that all the connections were correct and performed a hard restart. No patient was present. 2020-04-22 IUD x5 (Other): No new information. 2020-05-14 IUD (Other): No new information. 2020-06-07 IUD (Other): No new information. SURGEON SCREEN | 2020-04-16 00687496 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that when the system was started the 3define camera was red. Performed a hard restart and that did not resolve the issue. The FW was incorrect and camera cover was missing. Troubleshooting involved verifying that all the connections were correct and performed a hard restart. No patient was present. 2020-04-22 IUD x5 (Other): No new information. 2020-05-14 IUD (Other): No new information. 2020-06-07 IUD (Other): No new information. | xse camera functioning outside procedure started camera red performed hard restart did resolve issue fw incorrect camera cover missing troubleshooting involved verifying connections correct performed hard restart present iud new iud new iud new surgeon screen | NAV3253 | NAV3244 | NAV3261 | NAV3249 | NAV4095 | NAV3253NAV3244NAV3261NAV3249NAV4095 | True
10101 | 704798498 | XSE - Monitor Loose & Missing Cover | 2022-01-07T20:43:34.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | MEC4274-01S | KIT SVC MEC4274-01S JOINT 3 LEFT COVER | XSE - Monitor Loose & Missing Cover 2022-Jan-07 Salesforce C00240709 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during preventative maintenance it was found that the surgeon monitor was loose and the joint 3's left cover on the arm was missing. No patient was present. 2022-01-07 IUD x6 (Other): No new information. 2022-01-21 IUD (Other): No new information. KIT SVC MEC4274-01S JOINT 3 LEFT COVER | 2022-Jan-07 Salesforce C00240709 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during preventative maintenance it was found that the surgeon monitor was loose and the joint 3's left cover on the arm was missing. No patient was present. 2022-01-07 IUD x6 (Other): No new information. 2022-01-21 IUD (Other): No new information. | xse monitor loose missing cover salesforce outside procedure preventative maintenance surgeon monitor loose joint left cover arm missing present iud new iud new kit svc joint left cover | NAV4099 | NAV2053 | NAV2054 | NAV3261 | NAV3253 | NAV4099NAV2053NAV2054NAV3261NAV3253 | True
2285 | 703231135 | X - Software issue | 2019-06-27T17:46:42.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - Software issue Issue was as follows: · Surgical arm failed setup and failed to reach predefined position upon restart in OR. ****CSR trouble shooting include: Pressing the emergency stop and manual moving surgical arm followed by another attempts to clear up and/or draping positon without success. WS was fully shutdown followed by the 5 min wait period which corrected the issues**** · Surgical system difficult to remove from bed. Cases on June 26th and June 22nd both required a full restart prior to the SRS releasing from the bed mount · CD got stuck in the CD-ROM drive. XCSR had to take drive out to get surgeon CD with MRI images on it. · Software issues - Registration difficulties even with good C-Arm images. Large cross view error · Start up issues – System would pause for 20+ mins at Dell screen upon CPU reboot UNK MAZOR X ARM | Issue was as follows: · Surgical arm failed setup and failed to reach predefined position upon restart in OR. ****CSR trouble shooting include: Pressing the emergency stop and manual moving surgical arm followed by another attempts to clear up and/or draping positon without success. WS was fully shutdown followed by the 5 min wait period which corrected the issues**** · Surgical system difficult to remove from bed. Cases on June 26th and June 22nd both required a full restart prior to the SRS releasing from the bed mount · CD got stuck in the CD-ROM drive. XCSR had to take drive out to get surgeon CD with MRI images on it. · Software issues - Registration difficulties even with good C-Arm images. Large cross view error · Start up issues – System would pause for 20+ mins at Dell screen upon CPU reboot | x software issue issue follows · surgical arm failed setup failed reach predefined position restart csr trouble shooting include pressing emergency stop manual moving surgical arm followed attempts clear andor draping positon success ws fully shutdown followed min wait period corrected issues · surgical difficult remove bed cases june june required restart prior srs releasing bed mount · cd got stuck cdrom drive xcsr drive surgeon cd mri images · software issues registration difficulties good carm images large cross view error · start issues – pause mins dell screen cpu reboot unk mazor x arm | NAV3250 | NAV3244 | NAV4095 | NAV3242 | NAV4101 | NAV3250NAV3244NAV4095NAV3242NAV4101 | True
6405 | 703935988 | XSE - Deviated Trajectory | 2020-09-02T13:47:03.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Trajectory 2020-09-02 00695700, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using fixation.. It was reported that planning was done by the surgeon for T8-T11 using an open incision. A dual spinous process clamp was attached to T9 and T11 spinous process. The clamp was connected to the bone mount bridge to mount the surgical system to the patient. A 3Define scan and Draw Spine were completed. Registration was successful on the first try with green values at every level. The surgical arm was sent to all trajectories and the surgeon drilled and placed k-wires in the pilot holes. Accuracy was checked with fluoro and the right T8 wire on the right side looked medial. The surgical arm was resent to the trajectory a few times, different techniques were applied, and the plan was adjusted slightly, but the wire fell into the same hole. The surgeon decided to free-hand right T8. The surgeon then decided to remove the right T9 and T10 screws as they felt the screws were medial too. There was no patient harm and the procedure was delayed less than an hour. 2020-09-09 IUD, email 2020-09-09 16:00 (Rep): Additional information received from a manufacturer representative reported that the t rajectories seemed to be deviated between 3.5 and 10 mm. Most of the trajectories were deviated 6 mm from the intended plan. There was no patient harm and the suspected cause was a medial skive potential. MAZOR X SYSTEM TPL0059 | 2020-09-02 00695700, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using fixation.. It was reported that planning was done by the surgeon for T8-T11 using an open incision. A dual spinous process clamp was attached to T9 and T11 spinous process. The clamp was connected to the bone mount bridge to mount the surgical system to the patient. A 3Define scan and Draw Spine were completed. Registration was successful on the first try with green values at every level. The surgical arm was sent to all trajectories and the surgeon drilled and placed k-wires in the pilot holes. Accuracy was checked with fluoro and the right T8 wire on the right side looked medial. The surgical arm was resent to the trajectory a few times, different techniques were applied, and the plan was adjusted slightly, but the wire fell into the same hole. The surgeon decided to free-hand right T8. The surgeon then decided to remove the right T9 and T10 screws as they felt the screws were medial too. There was no patient harm and the procedure was delayed less than an hour. 2020-09-09 IUD, email 2020-09-09 16:00 (Rep): Additional information received from a manufacturer representative reported that the t rajectories seemed to be deviated between 3.5 and 10 mm. Most of the trajectories were deviated 6 mm from the intended plan. There was no patient harm and the suspected cause was a medial skive potential. | xse deviated trajectory ccr spinal procedure treat degenerative discs using fixation planning surgeon using open incision dual spinous process clamp attached spinous process clamp connected bone mount bridge mount surgical scan draw spine completed registration successful try green values level surgical arm sent trajectories surgeon drilled placed kwires pilot holes accuracy checked fluoro right wire right looked medial surgical arm resent trajectory times different techniques applied plan adjusted slightly wire fell hole surgeon decided freehand right surgeon decided remove right screws felt screws medial procedure delayed hour iud email additional t rajectories deviated mm trajectories deviated mm intended plan suspected cause medial skive potential mazor x | NAV5000 | NAV3243 | NAV4097 | NAV3246 | NAV3247 | NAV5000NAV3243NAV4097NAV3246NAV3247 | True
2087 | 703225810 | XSE - OOB snapshot not recognized | 2019-06-25T13:30:17.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | KIT0898 | SNAPSHOT TRACKER | XSE - OOB snapshot not recognized 2019-Jun-21 SCL1901823 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the snap shot tracker was not recognized. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. SNAPSHOT TRACKER | 2019-Jun-21 SCL1901823 (rep): Medtronic received information from a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the snap shot tracker was not recognized. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. | xse oob snapshot recognized outside procedure snap shot tracker recognized additional complications anticipated snapshot tracker | NAV2054 | NAV3299 | NAV3295 | NAV3246 | NAV4092 | NAV2054NAV3299NAV3295NAV3246NAV4092 | True
10419 | 704877635 | XSE ¦ Stripped Screws | 2022-02-22T21:09:17.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MEC5196-01S | KIT SVC MEC5196-01S RNX-ACC-063 | XSE ¦ Stripped Screws 2022-Feb-22 Salesforce C00244027 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there are 2 BMB screws that are stripped in addition to two Schanz screws that are stripped. No patient was present. 2022-03-02 IUD x2 (Other): No new information. 2022-03-08 email 2022-03-08 14:01 (Rep): The cause of the damage to the screws was due to standard wear and tear. KIT SVC MEC5196-01S RNX-ACC-063 | 2022-Feb-22 Salesforce C00244027 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there are 2 BMB screws that are stripped in addition to two Schanz screws that are stripped. No patient was present. 2022-03-02 IUD x2 (Other): No new information. 2022-03-08 email 2022-03-08 14:01 (Rep): The cause of the damage to the screws was due to standard wear and tear. | xse stripped screws salesforce outside procedure bmb screws stripped addition schanz screws stripped present iud new email cause damage screws standard wear tear kit svc | NAV3299 | NAV4090 | NAV4100 | NAV3264 | NAV2058 | NAV3299NAV4090NAV4100NAV3264NAV2058 | True
2883 | 703258560 | X - Tablet Not Charging | 2019-07-15T13:01:55.000000+0000 | MazorX | NAV3253 | TABLET ISSUES | 20 | MON0030-01 | MONITOR_13.3, TABLET | X - Tablet Not Charging The tablet was not charging on the work station. Ryan was able to speak to CeCe and they discovered the tablet was not being docked correctly. Once the tablet was docked it began charging. Issue resolved over the phone. MONITOR_13.3, TABLET | The tablet was not charging on the work station. Ryan was able to speak to CeCe and they discovered the tablet was not being docked correctly. Once the tablet was docked it began charging. Issue resolved over the phone. | x tablet charging tablet charging work station ryan able speak cece discovered tablet docked correctly tablet docked began charging issue resolved phone tablet | NAV3253 | NAV3249 | NAV2054 | NAV4109 | NAV4091 | NAV3253NAV3249NAV2054NAV4109NAV4091 | True
6060 | 703862486 | X - RRF Does Not Connect to SRA | 2020-07-21T11:16:54.000000+0000 | MazorX | NAV4100 | STRIPPED THREADS | 20 | ASM0206-01 | SURGICAL ARM | X - RRF Does Not Connect to SRA 2020-07-19 00693069, CCR (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the reference frame did not connect to the thread of the shoulder on the surgical arm. There was no patient harm and the procedure was not delayed. 2020-09-03 email 2020-09-03 02:31 (Other, For): No new information. SURGICAL ARM | 2020-07-19 00693069, CCR (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the reference frame did not connect to the thread of the shoulder on the surgical arm. There was no patient harm and the procedure was not delayed. 2020-09-03 email 2020-09-03 02:31 (Other, For): No new information. | x rrf does connect sra ccr spinal procedure treat deformity fixation reference frame did connect thread shoulder surgical arm procedure delayed email new surgical arm | NAV4100 | NAV3249 | NAV4095 | NAV4101 | NAV1006 | NAV4100NAV3249NAV4095NAV4101NAV1006 | True
4783 | 703576249 | XSE- OOBF Arm Failed Stress Test | 2020-01-22T20:06:50.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE- OOBF Arm Failed Stress Test 2020-1-22 00682549(Rep): Medtronic received information regarding a guidance system being used. It was reported that the surgical arm failed a stress test so the manufacturer representative completed a field motor calibration. After the calibration, the stress test was completed. No patient involvement. 2020-01-22 IUD, Email 2020-01-23 08:32 (Rep): Additional information was received from the manufacturer's representative. The representative reported that there were errors in J3. Field service motor calibration was performed on J3 of the system and now it then performed as intended. 2020-02-11 IUD (Other): No new information. SURGICAL ARM | 2020-1-22 00682549(Rep): Medtronic received information regarding a guidance system being used. It was reported that the surgical arm failed a stress test so the manufacturer representative completed a field motor calibration. After the calibration, the stress test was completed. No patient involvement. 2020-01-22 IUD, Email 2020-01-23 08:32 (Rep): Additional information was received from the manufacturer's representative. The representative reported that there were errors in J3. Field service motor calibration was performed on J3 of the system and now it then performed as intended. 2020-02-11 IUD (Other): No new information. | xse oobf arm failed stress test surgical arm failed stress test completed field motor calibration calibration stress test completed iud email additional manufacturers errors field service motor calibration performed performed intended iud new surgical arm | NAV2054 | NAV3249 | NAV4092 | NAV4098 | NAV1006 | NAV2054NAV3249NAV4092NAV4098NAV1006 | True
10066 | 704792592 | XSE - CANBUS Error | 2022-01-05T20:34:39.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 80 | ASM0270 | KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | XSE - CANBUS Error 2021-Dec-30 Salesforce C00240316, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that an error stating "Fatal CANBUS error. Lost communication to node 7" was displayed. There was no patient involvement. 2022-01-05 IUD x3 (Rep): Additional information received from a manufacturer representative reported that there was a recurring breaker tripping at the 24V 4-amp circuit and the electrical panel needed to be replaced. 2022-01-16 IUD x2 (Other): No new information. 2022-01-26 IUD (Other): No new information. 2022-02-08 IUD x2 (Other): No new information. 2022-02-23 IUD x3 (Other): No new information. 2022-03-07 IUD x5 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-05-10 IUD (Other): No new information. KIT SVC ASM0270 COMPRESSOR KIT 02S NEW | 2021-Dec-30 Salesforce C00240316, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that an error stating "Fatal CANBUS error. Lost communication to node 7" was displayed. There was no patient involvement. 2022-01-05 IUD x3 (Rep): Additional information received from a manufacturer representative reported that there was a recurring breaker tripping at the 24V 4-amp circuit and the electrical panel needed to be replaced. 2022-01-16 IUD x2 (Other): No new information. 2022-01-26 IUD (Other): No new information. 2022-02-08 IUD x2 (Other): No new information. 2022-02-23 IUD x3 (Other): No new information. 2022-03-07 IUD x5 (Other): No new information. 2022-03-28 IUD (Other): No new information. 2022-05-10 IUD (Other): No new information. | xse canbus error salesforce iud outside proc edure error stating fatal canbus error lost communication node displayed iud additional recurring breaker tripping circuit electrical panel needed replaced iud new iud new iud new iud new iud new iud new iud new kit svc compressor kit new | NAV2047 | NAV2054 | NAV3263 | NAV3262 | NAV2057 | NAV2047NAV2054NAV3263NAV3262NAV2057 | False
853 | 703178615 | X - 3DEFINE CAMERA was red | 2019-05-29T18:28:09.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3DEFINE CAMERA was red CS turning on system and ran 10 point accuracy test. Went into operation setup and 3Define camera was red. CS checked FW and it was missing UNK MAZOR X 3DEFINE CAMERA | CS turning on system and ran 10 point accuracy test. Went into operation setup and 3Define camera was red. CS checked FW and it was missing | x camera red cs turning ran point accuracy test went operation setup camera red cs checked fw missing unk mazor x camera | NAV3244 | NAV2053 | NAV3246 | NAV4110 | NAV4109 | NAV3244NAV2053NAV3246NAV4110NAV4109 | True
3270 | 703301563 | XSE - Difficult Registration/Deviation | 2019-08-08T14:26:04.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059-06 | MAZOR X SYSTEM | XSE - Difficult Registration/Deviation 2019-Jul-05 email 7/10, Intake Form, CCR (HCP, Rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat degenerative discs using fixation. It was reported that there was a challenging registration possibly due to patient BMI of 49. The first attempts resulted in fluoro segmentation, but provided red values in approve registration. Fluoro segmentation lines were manually adjusted, and resulted in red and yellow values. A new AP image was obtained, fluoro segmentation on complex. L4 was a yellow value, and a superior shift was noted in AP verification. L5 was a green value, verified matching. Fluoro segmentation lines were manually adjusted, and L4 and L5 were both green, verified matching accurate for both bodies. It was also noted that L5 Right appeared superior compared to plan. They re-planned and re-instrumented, and was accurate per fluoro. There was no patient harm or additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM | 2019-Jul-05 email 7/10, Intake Form, CCR (HCP, Rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure to treat degenerative discs using fixation. It was reported that there was a challenging registration possibly due to patient BMI of 49. The first attempts resulted in fluoro segmentation, but provided red values in approve registration. Fluoro segmentation lines were manually adjusted, and resulted in red and yellow values. A new AP image was obtained, fluoro segmentation on complex. L4 was a yellow value, and a superior shift was noted in AP verification. L5 was a green value, verified matching. Fluoro segmentation lines were manually adjusted, and L4 and L5 were both green, verified matching accurate for both bodies. It was also noted that L5 Right appeared superior compared to plan. They re-planned and re-instrumented, and was accurate per fluoro. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse difficult registrationdeviation email ccr spinal procedure treat degenerative discs using fixation challenging registration possibly bmi attempts resulted fluoro segmentation provided red values approve registration fluoro segmentation lines manually adjusted resulted red yellow values new ap image obtained fluoro segmentation complex yellow value superior shift noted ap verification green value verified matching fluoro segmentation lines manually adjusted green verified matching accurate bodies noted right appeared superior compared plan replanned reinstrumented accurate fluoro additional complications anticipated mazor x | NAV5000 | NAV3242 | NAV2051 | NAV3295 | NAV4095 | NAV5000NAV3242NAV2051NAV3295NAV4095 | True
4291 | 703530516 | XSE - Damaged Short Drill Guide/Sch | 2019-12-20T19:54:25.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | KIT0737 | SHORT DRILL GUIDE AND CANNULA KIT0737 | XSE - Damaged Short Drill Guide/Sch 2019-12-20 00680904 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the short drill guide and cannula have become dull and schanz bond screws are stripped. No patient was present. 2019-12-30 IUD, Email 2020-01-27 (Rep): No additional information 2020-02-24 email 2020-02-24 13:45, email 2020-03-03 15:01, PDF 2020-03-03 15:01 (Rep): No new information. 2020-03-18 IUD (Other): No new information. 2020-03-19 x4 IUD (Other): No new information. SHORT DRILL GUIDE AND CANNULA KIT0737 | 2019-12-20 00680904 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the short drill guide and cannula have become dull and schanz bond screws are stripped. No patient was present. 2019-12-30 IUD, Email 2020-01-27 (Rep): No additional information 2020-02-24 email 2020-02-24 13:45, email 2020-03-03 15:01, PDF 2020-03-03 15:01 (Rep): No new information. 2020-03-18 IUD (Other): No new information. 2020-03-19 x4 IUD (Other): No new information. | xse damaged short drill guidesch outside procedure short drill guide cannula dull schanz bond screws stripped present iud email additional email email pdf new iud new iud new short drill guide cannula | NAV3299 | NAV3295 | NAV4090 | NAV2058 | NAV3264 | NAV3299NAV3295NAV4090NAV2058NAV3264 | True
1405 | 703199621 | X - Software issue | 2019-06-11T11:35:26.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Software issue It was reported that the software came back and brought us back to the home screen where we then attempted to enter operate mode when warning message asked used if we would like to unlock the mount we selected no but it wouldnt allow us to continue into operate mode we attempted this multiple times before unmounting. MAZOR X SOFTWARE KIT VER-3.0.2 | It was reported that the software came back and brought us back to the home screen where we then attempted to enter operate mode when warning message asked used if we would like to unlock the mount we selected no but it wouldnt allow us to continue into operate mode we attempted this multiple times before unmounting. | x software issue software came brought home screen attempted enter operate mode warning message asked like unlock mount selected wouldnt allow continue operate mode attempted multiple times unmounting mazor x software kit | NAV3246 | NAV4091 | NAV3247 | NAV3248 | NAV2047 | NAV3246NAV4091NAV3247NAV3248NAV2047 | True
10445 | 704882066 | XSE - Screw Deviation | 2022-02-24T17:11:46.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Screw Deviation 2022-Feb-24 Salesforce C00244157, IUD x3, email 2022-03-04 14:01 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was an inaccurate placement of a screw. The construct was L4-S1 open. All screws were placed accurately except for L5 right side which was placed superiorly by Dr. Lehman's fellow. The screw was deviated less than 3.5 mm. The patient was mounted via the Bone Mount Bridge in the PSIS. The screw was extracted and replaced freehand. There was about a 15 minute delay to the procedure. There was no impact on the patient outcome. 2022-03-09 email 2022-03-09 23:45 (Rep): No new information. 2022-03-18 email 2022-03-18 08:02 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2022-Feb-24 Salesforce C00244157, IUD x3, email 2022-03-04 14:01 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was an inaccurate placement of a screw. The construct was L4-S1 open. All screws were placed accurately except for L5 right side which was placed superiorly by Dr. Lehman's fellow. The screw was deviated less than 3.5 mm. The patient was mounted via the Bone Mount Bridge in the PSIS. The screw was extracted and replaced freehand. There was about a 15 minute delay to the procedure. There was no impact on the patient outcome. 2022-03-09 email 2022-03-09 23:45 (Rep): No new information. 2022-03-18 email 2022-03-18 08:02 (Rep): No new information. | xse screw deviation salesforce iud email spinal procedure inaccurate placement screw construct open screws placed accurately right placed superiorly dr lehmans fellow screw deviated mm mounted bone mount bridge psis screw extracted replaced freehand minute delay procedure impact outcome email new email new mazor x | NAV5000 | NAV4112 | NAV3243 | NAV3299 | NAV1012 | NAV5000NAV4112NAV3243NAV3299NAV1012 | True
4413 | 703541611 | X- Arm & Registration issues | 2020-01-02T20:28:28.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X- Arm & Registration issues 2018-2-12 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were registration issues because L3 and S1 always had poor values. Another AP image was taken and labeling was checked. Registration was achieved and the surgery was completed. The surgical arm aborted movement to the clear up position after set up and again after the surgery. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-2-12 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were registration issues because L3 and S1 always had poor values. Another AP image was taken and labeling was checked. Registration was achieved and the surgery was completed. The surgical arm aborted movement to the clear up position after set up and again after the surgery. There was no patient harm and the procedure was not delayed over an hour. | x arm registration issues ccr spinal procedure registration issues poor values ap image taken labeling checked registration achieved surgery completed surgical arm aborted movement clear position set surgery procedure delayed hour mazor x | NAV3242 | NAV4112 | NAV4095 | NAV3243 | NAV4091 | NAV3242NAV4112NAV4095NAV3243NAV4091 | True
2373 | 703234005 | X - SW, x-eye, WS, & 3Define issue | 2019-06-28T19:55:54.000000+0000 | MazorX | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 90 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | X - SW, x-eye, WS, & 3Define issue 2017-6-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on T5-T8. It was reported that the bed frame was mounted too low and the angle of the shoulder was too open. A 3Define scan was done, but it had poor quality so a second scan had to be done. The x-eye could not see the reference marker, so predefined AP and OBL trajectories were sent for the acquisition. The predefined trajectories were manually adjusted to get the optimal images. Registration was completed on the first attempt off of T8. Fluoro verification was accurate. All trajectories were sent and drilled, and the surgeon was pleased with the outcome. The manufacturer representative attempted to test the 3Define from utilities. A timeout pop-up appeared after 20 seconds, and the 3Define image froze. They refreshed the 3Define image and the timeout pop-up appeared again. While dismounting the surgical system from the bed the pneumatic system stopped working. The arm almost fell on the patient but the clinical engineer held it up. The representative performed a hard restart to try and fix the pressure issue, but the software crashed en dlessly. The representative had to return the surgical system to the workstation manually. During this process, the hood unexpectedly fell, so the clinical engineer inserted a drape near the hydraulic cylinders to keep the hood open. There was no patient harm a nd the procedure was not delayed over an hour. MAZOR X PC STATION - REFURBISHED | 2017-6-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on T5-T8. It was reported that the bed frame was mounted too low and the angle of the shoulder was too open. A 3Define scan was done, but it had poor quality so a second scan had to be done. The x-eye could not see the reference marker, so predefined AP and OBL trajectories were sent for the acquisition. The predefined trajectories were manually adjusted to get the optimal images. Registration was completed on the first attempt off of T8. Fluoro verification was accurate. All trajectories were sent and drilled, and the surgeon was pleased with the outcome. The manufacturer representative attempted to test the 3Define from utilities. A timeout pop-up appeared after 20 seconds, and the 3Define image froze. They refreshed the 3Define image and the timeout pop-up appeared again. While dismounting the surgical system from the bed the pneumatic system stopped working. The arm almost fell on the patient but the clinical engineer held it up. The representative performed a hard restart to try and fix the pressure issue, but the software crashed en dlessly. The representative had to return the surgical system to the workstation manually. During this process, the hood unexpectedly fell, so the clinical engineer inserted a drape near the hydraulic cylinders to keep the hood open. There was no patient harm a nd the procedure was not delayed over an hour. | x sw xeye ws issue ccr spinal procedure bed frame mounted low angle shoulder open scan poor quality second scan xeye reference marker predefined ap obl trajectories sent acquisition predefined trajectories manually adjusted optimal images registration completed attempt fluoro verification accurate trajectories sent drilled surgeon pleased outcome attempted test utilities timeout popup appeared seconds image froze refreshed image timeout popup appeared dismounting surgical bed pneumatic stopped working arm fell clinical engineer held performed hard restart try fix pressure issue software crashed en dlessly return surgical workstation manually process hood unexpectedly fell clinical engineer inserted drape near hydraulic cylinders hood open nd procedure delayed hour mazor x pc station refurbished | NAV3248 | NAV1006 | NAV3264 | NAV4110 | NAV3242 | NAV3248NAV1006NAV3264NAV4110NAV3242 | True
7112 | 704119868 | XSE - Failed BIST (Built in Seft Test) | 2020-12-11T15:05:07.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 30 | ASM0357-01 | SURGEON SCREEN | XSE - Failed BIST (Built in Seft Test) 2020-12-11 00702512, 00702527 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed during the built in self test. The surgical arm was sent to the clear up position and a hard reboot of the system was done. The surgical arm then passed the test after the system was booted up. While setting up the system for the case, the surgeon screen went blank, the surgical arm would not unlock when the unlock button on the arm was pressed, and the software would zoo in and and zoom out quickly correcting itself when they would move from the surgical arm setup to the operate mode screen. There was no patient involvement. 2020-12-23 IUD (Other): No new information. 2021-01-14 email 2021-01-14 13:15 (Rep): Additional information was received stating that the arm has been replaced and this issue has been fixed. The screen blacking out happens intermittently but has been much less frequent as of late. The arm not unlocking was thought to be due to the pneumatic tube supplying air pressure to the positioner being choked off by the positioner resting on it. This was stated to be the manufacturer representative's fault. 2021-Feb-03 email 2/3 x3 (rep): No new information. SURGEON SCREEN | 2020-12-11 00702512, 00702527 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed during the built in self test. The surgical arm was sent to the clear up position and a hard reboot of the system was done. The surgical arm then passed the test after the system was booted up. While setting up the system for the case, the surgeon screen went blank, the surgical arm would not unlock when the unlock button on the arm was pressed, and the software would zoo in and and zoom out quickly correcting itself when they would move from the surgical arm setup to the operate mode screen. There was no patient involvement. 2020-12-23 IUD (Other): No new information. 2021-01-14 email 2021-01-14 13:15 (Rep): Additional information was received stating that the arm has been replaced and this issue has been fixed. The screen blacking out happens intermittently but has been much less frequent as of late. The arm not unlocking was thought to be due to the pneumatic tube supplying air pressure to the positioner being choked off by the positioner resting on it. This was stated to be the manufacturer representative's fault. 2021-Feb-03 email 2/3 x3 (rep): No new information. | xse failed bist outside procedure surgical arm failed built self test surgical arm sent clear position hard reboot surgical arm passed test booted setting case surgeon screen went blank surgical arm unlock unlock button arm pressed software zoo zoom quickly correcting surgical arm setup operate mode screen iud new email additional stating arm replaced issue fixed screen blacking happens intermittently frequent late arm unlocking thought pneumatic tube supplying air pressure positioner choked positioner resting stated representatives fault email new surgeon screen | NAV3253 | NAV3249 | NAV4107 | NAV4101 | NAV2054 | NAV3253NAV3249NAV4107NAV4101NAV2054 | True
182 | 703027717 | X - unable to register | 2019-02-26T21:47:12.000000+0000 | MazorX | NAV4112 | NO REGISTRATION | 10 | TPL0059-04 | MAZOR X SYSTEM (REV 4) | X - unable to register 2019-Feb-21 SCL1900443 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they were unable to achieve registration for L5. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM (REV 4) | 2019-Feb-21 SCL1900443 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they were unable to achieve registration for L5. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x unable register spinal procedure unable achieve registration noted additional complications anticipated mazor x | NAV4112 | NAV3246 | NAV3250 | NAV2054 | NAV3242 | NAV4112NAV3246NAV3250NAV2054NAV3242 | True
5067 | 703593981 | XSE - Surgical Arm Covers | 2020-02-03T16:11:35.000000+0000 | MazorX Stealth Edition | NAV4099 | PLASTIC COVER ISSUES | 30 | MEC4654-01S | KIT SVC MEC4654-01S J6 - INTEL COVER | XSE - Surgical Arm Covers 2020-01-30 00683233 (REP, HCP): Medtronic received information regarding a guidance system in use outside of procedure. It was reported that the surgical arm covers are loose on the arm and need to be replaced. There was no patient involvement. 2020-02-03 IUD x3 Email 2020-02-04 10:45 (Rep): Additional information was received from the manufacturer's representative. The representative reported that two of the arm covers are loose and the shoulder button, not handle, doesn’t illuminate green like it should. The issue is attributed to wear and tear. 2020-02-25 IUD x3 (Other): No new information. KIT SVC MEC4654-01S J6 - INTEL COVER | 2020-01-30 00683233 (REP, HCP): Medtronic received information regarding a guidance system in use outside of procedure. It was reported that the surgical arm covers are loose on the arm and need to be replaced. There was no patient involvement. 2020-02-03 IUD x3 Email 2020-02-04 10:45 (Rep): Additional information was received from the manufacturer's representative. The representative reported that two of the arm covers are loose and the shoulder button, not handle, doesn’t illuminate green like it should. The issue is attributed to wear and tear. 2020-02-25 IUD x3 (Other): No new information. | xse surgical arm covers use outside procedure surgical arm covers loose arm need replaced iud email additional manufacturers arm covers loose shoulder button handle doesn’t illuminate green like issue attributed wear tear iud new kit svc intel cover | NAV4099 | NAV3252 | NAV4097 | NAV3249 | NAV2054 | NAV4099NAV3252NAV4097NAV3249NAV2054 | True
2329 | 703233180 | X - 3Define issue | 2019-06-28T14:03:33.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3Define issue 3Define camera error - See attached file for error screen capture. CSR trouble shooting included – soft restart, hard restart, power down with 5+ min wait and power back up, Power outlet move with power down with 5+ min wait and power back up, power down with USB connection test and power back up. None of these resolved the issue. UNK MAZOR X 3DEFINE CAMERA | 3Define camera error - See attached file for error screen capture. CSR trouble shooting included – soft restart, hard restart, power down with 5+ min wait and power back up, Power outlet move with power down with 5+ min wait and power back up, power down with USB connection test and power back up. None of these resolved the issue. | x issue camera error attached file error screen capture csr trouble shooting included – soft restart hard restart power min wait power power outlet power min wait power power usb connection test power resolved issue unk mazor x camera | NAV3263 | NAV3244 | NAV3261 | NAV2047 | NAV3248 | NAV3263NAV3244NAV3261NAV2047NAV3248 | True
4536 | 703549198 | X - Deviation Bilateral L5 | 2020-01-07T20:36:29.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation Bilateral L5 2018-01-18 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L4-L5. It was reported that the surgeon and manufacturer representative reviewed the plan for the surgery. A BMB with a Schanz pin was used to secure the system to the patient at the right PSIS. Registration was successful on the first attempt. All trajectories were drilled and placed. Confirmation images show that the K-wires for bilateral L5 were deviated laterally due to skiving. The arm was resent for left L5 without changing the plan and the surgeon edited the plan for right L5. Another set of confirmation images showed that all K-wires were accurate. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-01-18 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L4-L5. It was reported that the surgeon and manufacturer representative reviewed the plan for the surgery. A BMB with a Schanz pin was used to secure the system to the patient at the right PSIS. Registration was successful on the first attempt. All trajectories were drilled and placed. Confirmation images show that the K-wires for bilateral L5 were deviated laterally due to skiving. The arm was resent for left L5 without changing the plan and the surgeon edited the plan for right L5. Another set of confirmation images showed that all K-wires were accurate. There was no patient harm and the procedure was not delayed over an hour. | x deviation bilateral ccr spinal procedure treat degenerative disk disease fixation surgeon reviewed plan surgery bmb schanz pin secure right psis registration successful attempt trajectories drilled placed confirmation images kwires bilateral deviated laterally skiving arm resent left changing plan surgeon edited plan right set confirmation images showed kwires accurate procedure delayed hour mazor x | NAV4091 | NAV3242 | NAV5000 | NAV1012 | NAV3295 | NAV4091NAV3242NAV5000NAV1012NAV3295 | True
9063 | 704542829 | XSE - Missing BNC Cable | 2021-08-05T20:21:17.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | KIT1218-01 | SVC KIT1218-01 CBL BNC VIDEO | XSE - Missing BNC Cable 2021-Aug-05 Salesforce C00231704, IUD x4 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the BNC cable was missing. There was no patient involvement. SVC KIT1218-01 CBL BNC VIDEO | 2021-Aug-05 Salesforce C00231704, IUD x4 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the BNC cable was missing. There was no patient involvement. | xse missing bnc cable salesforce iud outside p rocedure bnc cable missing svc cbl bnc video | NAV2053 | NAV3256 | NAV3249 | NAV3237 | NAV3253 | NAV2053NAV3256NAV3249NAV3237NAV3253 | True
6331 | 703914376 | XSE - Bed Frame Adapter Damaged | 2020-08-20T17:39:49.000000+0000 | MazorX Stealth Edition | NAV3255 | BED FRAME ISSUES | 20 | MAS1574-01S | KIT SVC MAS1574-01S BED FRAME ADAPTOR | XSE - Bed Frame Adapter Damaged 2020-08-20 00694971 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the bed frame adapter screw was bent and needed to be replaced. There was no patient involvement. 2020-08-30 IUD x4 (Other): No new information. 2020-09-21 IUD x2 (Other): No new information. KIT SVC MAS1574-01S BED FRAME ADAPTOR | 2020-08-20 00694971 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the bed frame adapter screw was bent and needed to be replaced. There was no patient involvement. 2020-08-30 IUD x4 (Other): No new information. 2020-09-21 IUD x2 (Other): No new information. | xse bed frame adapter damaged outside procedure bed frame adapter screw bent needed replaced iud new iud new kit svc bed frame adaptor | NAV3255 | NAV3299 | NAV2054 | NAV1006 | NAV2058 | NAV3255NAV3299NAV2054NAV1006NAV2058 | True
1996 | 703223560 | X- Drape Rippped | 2019-06-24T17:39:19.000000+0000 | MazorX | NAV3232 | DRAPE ISSUES | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | X- Drape Rippped THP -PHOP Workstation #60/Arm #71 Software 3.0 Red -Surface scan issues, trajectory collisions, torn drape. DX: Scoliosis. PLANNING: -The planning was done by the surgeon prior to the start of the case on a personal computer. PREP: 08:14 -Surgical arm attached to the bed -one person 08:20 -Surgical arm draped -one person -30 seconds 08:24 -Workstation powered down 08:36 -Incision 12:09 -Workstation powered on. PLATFORM MOUNTING #1: 12:13 -Dual schanze pins mounted. The left schanze pin connector was mounted at the forward mounting point, while the right schanze pin connector wasmounted and the rearward mounting point 12:14 -Surgical arm setup performed by the surgeon 12:17 -Mount locked. 3DEFINE #1: ***12:18 -Recording; wet towel placed over the BMU and lights pointed away. Resulting scan was vague and missing an excessive amount of information. It was determined that too much blood on the 3Define camera. 3Define camera was cleaned. ***12:20 -Recording; wet towel placed over the BMU and lights pointed away. Resulting scan was missing an excessive amount of information. It was determined that the surgical arm was mounted to high above the patient. The surgical arm was adjusted closer. 12:22 -Surgical arm lowered 12:23 -Recording; wet towel placed over the BMU and lights pointed away. Resulting scan was clear and accurate. IMAGE ACQUISITION AND REGISTRATION #1: 12:25 -Drawspine successful 12:25 -Surgical arm sent to AP position of L5. ***12:26 -The surgical arm collided with the left schanze pin of the mount. Mount adjusted. 3DEFINE #1 CONTINUED: 12:27 -Recoding; wet towel placed over the BMU and lights pointed away. Resulting scan was clear and accurate . IMAGE ACQUISTION AND REGISTRATION #1 CONTINUED: 12:28 -Drawspine successful 12:29 -Surgical arm sent to AP position of L5 -adjusted. 1 image was taken. The CSR recognized that the c-arm tech had the magnification on. The magnification was removed and the image was retaken. 12:34 -Surgical arm sent to OBL of L5 ***12:34 -Drape tore on way to the OBL. The drape had to be repaired intraoperatively with the a tegaderm. 12:38 -Registration attempt of L5 successful with basic algorithm 12:40 -Fluoro match successful. OPERATION #1: 12:44 -Left S2 12:49 -Right S2 -too much pressure on the arm while lowering down. The surgeon decided to adjust the patient's soft tissue with retraction. 12 :50 -Cleared up 12:51 -Right S2 ***12:55 -Left S1 -The estimated movement time quoted 6 seconds, but actual movement time was closer to 45 seconds 13:00 -Left L5 13:04 -Left L4 13:08 -Left L3 13:13 -Right L3 13:16 -Right L4 13:22 -Right L5 13:26 -Right S1 13:29 -Cleared up. PLATFORM MOUNTING #2: 13:30 -Mount unlocked 13:35 -Dual clamp mounted at L1/T12 and T10/T9. Locked at 15 degrees 13:36 -Mount locked. 3DEFINE #2: 13:36 -Recording; wet towel placed over BMU and lights pointed away. Resulting scan was clear and accurate. IMAGE ACQUISITION AND REGISTRATION #2: 13:38 -Drawspine successful 13:40 -Surgical arm sent to AP position 13:45 -Surgical arm sent to OBL position 13:55 -Registration attempt successful with advanced algorithm ***13:59 -Fluoro match successful. 1 red screw found at L1. The surgeon chose to skip this screw and came back for it. When L1 was selected for fluoro matching, T12 was shown. OPERATION #2: 14:03 -Left L2 14:06 -Right L2 14:10 -Right T12 14:12 -Right T11 14:14 -Right T10 14:17 -Left T10 14:17 -The low battery indicator indicated that the tablet was about to loose power 14:20 -Left T11 14:22 -Left T12 14:22 -CSR went back to register L1 14:25 -Right L1 14:27 -Left L1 14:30 -Cleared up. PLATFORM MOUNTING #3: 14:31 -Mount unlocked 14:31 -Scrubtech attached charging plates to t ablet so that the tablet would charge 14:34 -Dual clamp clamped onto T6 and T8. BMU locked at 15 degrees. 14:34 -Tablet placed on charging arm. 3DEFINE #3: 14:35 -Recording; wet towel placed over BMU and lights pointed away. Resulting scan was clear and accurate. IMAGE ACQUISITION AND REGISTRATION #3: 14:37 -Drawspine successful 14:38 -Surgical arm sent to AP position 14:41 -Surgical arm sent to OBL position. 14:44 -Registration attempt of T6 successful with advanced algorithm. The resulting registration was missing T4. The surgeon decided to adjust the OBL image. 14:45 -New OBL image 14:47 -Registration attempt of T6 successful with advanced algorithm. 14:48 -Fluoro match successful with all green values. OPERATION #3: 14:52 -Left T9 14:54 -Left T7 14:57 -Left T5 15:00 -Left T4 15:02 -Right T4 15:06 -Right T5 15:09 -Right T6 15:10 -Right T8 15:13 -Surgical mount unlocked and removed from the bed. Overall, the surgeon was pleased with the results of the procedure. The surgeon requested that the schanze pin connectors have quick disconnect. This way, the surgeon would be able to remove the connectors when he moves to the dual clamp. UNK MAZOR X ACCESSORIES | THP -PHOP Workstation #60/Arm #71 Software 3.0 Red -Surface scan issues, trajectory collisions, torn drape. DX: Scoliosis. PLANNING: -The planning was done by the surgeon prior to the start of the case on a personal computer. PREP: 08:14 -Surgical arm attached to the bed -one person 08:20 -Surgical arm draped -one person -30 seconds 08:24 -Workstation powered down 08:36 -Incision 12:09 -Workstation powered on. PLATFORM MOUNTING #1: 12:13 -Dual schanze pins mounted. The left schanze pin connector was mounted at the forward mounting point, while the right schanze pin connector wasmounted and the rearward mounting point 12:14 -Surgical arm setup performed by the surgeon 12:17 -Mount locked. 3DEFINE #1: ***12:18 -Recording; wet towel placed over the BMU and lights pointed away. Resulting scan was vague and missing an excessive amount of information. It was determined that too much blood on the 3Define camera. 3Define camera was cleaned. ***12:20 -Recording; wet towel placed over the BMU and lights pointed away. Resulting scan was missing an excessive amount of information. It was determined that the surgical arm was mounted to high above the patient. The surgical arm was adjusted closer. 12:22 -Surgical arm lowered 12:23 -Recording; wet towel placed over the BMU and lights pointed away. Resulting scan was clear and accurate. IMAGE ACQUISITION AND REGISTRATION #1: 12:25 -Drawspine successful 12:25 -Surgical arm sent to AP position of L5. ***12:26 -The surgical arm collided with the left schanze pin of the mount. Mount adjusted. 3DEFINE #1 CONTINUED: 12:27 -Recoding; wet towel placed over the BMU and lights pointed away. Resulting scan was clear and accurate . IMAGE ACQUISTION AND REGISTRATION #1 CONTINUED: 12:28 -Drawspine successful 12:29 -Surgical arm sent to AP position of L5 -adjusted. 1 image was taken. The CSR recognized that the c-arm tech had the magnification on. The magnification was removed and the image was retaken. 12:34 -Surgical arm sent to OBL of L5 ***12:34 -Drape tore on way to the OBL. The drape had to be repaired intraoperatively with the a tegaderm. 12:38 -Registration attempt of L5 successful with basic algorithm 12:40 -Fluoro match successful. OPERATION #1: 12:44 -Left S2 12:49 -Right S2 -too much pressure on the arm while lowering down. The surgeon decided to adjust the patient's soft tissue with retraction. 12 :50 -Cleared up 12:51 -Right S2 ***12:55 -Left S1 -The estimated movement time quoted 6 seconds, but actual movement time was closer to 45 seconds 13:00 -Left L5 13:04 -Left L4 13:08 -Left L3 13:13 -Right L3 13:16 -Right L4 13:22 -Right L5 13:26 -Right S1 13:29 -Cleared up. PLATFORM MOUNTING #2: 13:30 -Mount unlocked 13:35 -Dual clamp mounted at L1/T12 and T10/T9. Locked at 15 degrees 13:36 -Mount locked. 3DEFINE #2: 13:36 -Recording; wet towel placed over BMU and lights pointed away. Resulting scan was clear and accurate. IMAGE ACQUISITION AND REGISTRATION #2: 13:38 -Drawspine successful 13:40 -Surgical arm sent to AP position 13:45 -Surgical arm sent to OBL position 13:55 -Registration attempt successful with advanced algorithm ***13:59 -Fluoro match successful. 1 red screw found at L1. The surgeon chose to skip this screw and came back for it. When L1 was selected for fluoro matching, T12 was shown. OPERATION #2: 14:03 -Left L2 14:06 -Right L2 14:10 -Right T12 14:12 -Right T11 14:14 -Right T10 14:17 -Left T10 14:17 -The low battery indicator indicated that the tablet was about to loose power 14:20 -Left T11 14:22 -Left T12 14:22 -CSR went back to register L1 14:25 -Right L1 14:27 -Left L1 14:30 -Cleared up. PLATFORM MOUNTING #3: 14:31 -Mount unlocked 14:31 -Scrubtech attached charging plates to t ablet so that the tablet would charge 14:34 -Dual clamp clamped onto T6 and T8. BMU locked at 15 degrees. 14:34 -Tablet placed on charging arm. 3DEFINE #3: 14:35 -Recording; wet towel placed over BMU and lights pointed away. Resulting scan was clear and accurate. IMAGE ACQUISITION AND REGISTRATION #3: 14:37 -Drawspine successful 14:38 -Surgical arm sent to AP position 14:41 -Surgical arm sent to OBL position. 14:44 -Registration attempt of T6 successful with advanced algorithm. The resulting registration was missing T4. The surgeon decided to adjust the OBL image. 14:45 -New OBL image 14:47 -Registration attempt of T6 successful with advanced algorithm. 14:48 -Fluoro match successful with all green values. OPERATION #3: 14:52 -Left T9 14:54 -Left T7 14:57 -Left T5 15:00 -Left T4 15:02 -Right T4 15:06 -Right T5 15:09 -Right T6 15:10 -Right T8 15:13 -Surgical mount unlocked and removed from the bed. Overall, the surgeon was pleased with the results of the procedure. The surgeon requested that the schanze pin connectors have quick disconnect. This way, the surgeon would be able to remove the connectors when he moves to the dual clamp. | x drape rippped thp phop workstation software red surface scan issues trajectory collisions torn drape dx scoliosis planning planning surgeon prior start case personal computer prep surgical arm attached bed person surgical arm draped person seconds workstation powered incision workstation powered platform mounting dual schanze pins mounted left schanze pin connector mounted forward mounting point right schanze pin connector wasmounted rearward mounting point surgical arm setup performed surgeon mount locked recording wet towel placed bmu lights pointed away resulting scan vague missing excessive determined blood camera camera cleaned recording wet towel placed bmu lights pointed away resulting scan missing excessive determined surgical arm mounted high surgical arm adjusted closer surgical arm lowered recording wet towel placed bmu lights pointed away resulting scan clear accurate image acquisition registration drawspine successful surgical arm sent ap position surgical arm collided left schanze pin mount mount adjusted continued recoding wet towel placed bmu lights pointed away resulting scan clear accurate image acquistion registration continued drawspine successful surgical arm sent ap position adjusted image taken csr recognized carm tech magnification magnification removed image retaken surgical arm sent obl drape tore way obl drape repaired intraoperatively tegaderm registration attempt successful basic algorithm fluoro match successful operation left right pressure arm lowering surgeon decided adjust patients soft tissue retraction cleared right left estimated movement time quoted seconds actual movement time closer seconds left left left right right right right cleared platform mounting mount unlocked dual clamp mounted locked degrees mount locked recording wet towel placed bmu lights pointed away resulting scan clear accurate image acquisition registration drawspine successful surgical arm sent ap position surgical arm sent obl position registration attempt successful advanced algorithm fluoro match successful red screw surgeon chose skip screw came selected fluoro matching shown operation left right right right right left low battery indicator indicated tablet loose power left left csr went register right left cleared platform mounting mount unlocked scrubtech attached charging plates t ablet tablet charge dual clamp clamped bmu locked degrees tablet placed charging arm recording wet towel placed bmu lights pointed away resulting scan clear accurate image acquisition registration drawspine successful surgical arm sent ap position surgical arm sent obl position registration attempt successful advanced algorithm resulting registration missing surgeon decided adjust obl image new obl image registration attempt successful advanced algorithm fluoro match successful green values operation left left left left right right right right surgical mount unlocked removed bed overall surgeon pleased results procedure surgeon requested schanze pin connectors quick disconnect way surgeon able remove connectors moves dual clamp unk mazor x accessories | NAV3246 | NAV3258 | NAV3259 | NAV4101 | NAV3253 | NAV3246NAV3258NAV3259NAV4101NAV3253 | False
6600 | 703986186 | XSE ¦ Failed to reach destination | 2020-09-30T16:05:52.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | XSE ¦ Failed to reach destination 2020-09-30 00697763 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that while sending the arm to clear up he received an message that "failed to reach destination". After rebooting the system the arm functioned as intended. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. UNK MAZOR X ARM | 2020-09-30 00697763 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that while sending the arm to clear up he received an message that "failed to reach destination". After rebooting the system the arm functioned as intended. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. | xse failed reach destination intraoperatively spinal procedure sending arm clear message failed reach destination rebooting arm functioned intended issue extended surgical time hour impact outcome unk mazor x arm | NAV4095 | NAV3244 | NAV3264 | NAV3261 | NAV2058 | NAV4095NAV3244NAV3264NAV3261NAV2058 | True
6690 | 704007642 | XSE - CANBUS Error | 2020-10-12T20:13:08.000000+0000 | MazorX Stealth Edition | NAV1006 | NO IDENTIFIED PRODUCT ISSUE | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - CANBUS Error 2020-10-12 00698596 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the site was having multiple CANBUS errors. The manufacturer representative also stated that the system showed gains lost and robot calibration not found on the driver. Error 4496 did not clear with reboots. The system positioner buttons then stayed depressed and he could not get them to function. They then did not use the system for the upcoming case. No patient was present. 2020-10-13 IUD x2 (Other): No new information. 2020-10-23 IUD x5 (Other): No new information. 2020-12-14 IUD x2 (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-10-12 00698596 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the site was having multiple CANBUS errors. The manufacturer representative also stated that the system showed gains lost and robot calibration not found on the driver. Error 4496 did not clear with reboots. The system positioner buttons then stayed depressed and he could not get them to function. They then did not use the system for the upcoming case. No patient was present. 2020-10-13 IUD x2 (Other): No new information. 2020-10-23 IUD x5 (Other): No new information. 2020-12-14 IUD x2 (Other): No new information. | xse canbus error outside procedure site having multiple canbus errors stated showed gains lost robot calibration driver error did clear reboots positioner buttons stayed depressed function did use upcoming case present iud new iud new iud new kit svc surgical arm | NAV2047 | NAV3249 | NAV4098 | NAV2054 | NAV4104 | NAV2047NAV3249NAV4098NAV2054NAV4104 | False
9641 | 704689727 | XSE - OOBF/3Define Camera Not Working | 2021-10-29T18:27:08.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | KIT1318 | SEN0088 SPARE PART KIT KIT1318 | XSE - OOBF/3Define Camera Not Working 2021-Oct-29 Salesforce C00237068, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the 3Define camera connection was lost during install. Troubleshooting did not resolve the issue. There was no patient involvement. 2021-11-02 IUD x3 (Other): No new information. 2021-11-10 IUD (Other): No new information. 2021-11-10 e1, e2, e3, e4 (Rep): Additional information received from a manufacturer representative reported that the original camera on the system was found to be working and the replacement camera was still in use since the issue was resolved by reseating the cables. 2021-11-17 IUD (Other): No new information. 2021-11-24 IUD x3 (Rep): No new information. 2021-12-03 IUD (Other): No new information. 2021-12-10 IUD x3 (Other): No new information. 2022-01-09 IUD x6, 302695109 (Other): No new information. 2022-02-08 IUD (Other): No new information. SEN0088 SPARE PART KIT KIT1318 | 2021-Oct-29 Salesforce C00237068, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the 3Define camera connection was lost during install. Troubleshooting did not resolve the issue. There was no patient involvement. 2021-11-02 IUD x3 (Other): No new information. 2021-11-10 IUD (Other): No new information. 2021-11-10 e1, e2, e3, e4 (Rep): Additional information received from a manufacturer representative reported that the original camera on the system was found to be working and the replacement camera was still in use since the issue was resolved by reseating the cables. 2021-11-17 IUD (Other): No new information. 2021-11-24 IUD x3 (Rep): No new information. 2021-12-03 IUD (Other): No new information. 2021-12-10 IUD x3 (Other): No new information. 2022-01-09 IUD x6, 302695109 (Other): No new information. 2022-02-08 IUD (Other): No new information. | xse camera working salesforce iud outside proc edure camera connection lost install troubleshooting did resolve issue iud new iud new additional original camera working replacement camera use issue resolved reseating cables iud new iud new iud new iud new iud new iud new spare kit | NAV3244 | NAV2054 | NAV3299 | NAV2047 | NAV3246 | NAV3244NAV2054NAV3299NAV2047NAV3246 | True
5273 | 703635364 | XSE - Inconsistent Accuracy | 2020-02-25T20:11:59.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 30 | ASM0206-01 | SURGICAL ARM | XSE - Inconsistent Accuracy 2020-02-25 00685115 ID (Rep): Medtronic received information regarding a guidance system being used at a medical facility. It was reported that the system had been getting inconsistent accuracy test results on both the 3 and 10 point tests. The manufacturer representative reported different instruments (e.g. accuracy pointer and trocar) yield different accuracy results. No patient was involved. 2020-02-25 UD x2, Email 2020-02-26 16:30 (Rep): Additional information received from a manufacturer representative reported that the arm was replaced. 2020-04-07 IUD (Other): No new information. SURGICAL ARM | 2020-02-25 00685115 ID (Rep): Medtronic received information regarding a guidance system being used at a medical facility. It was reported that the system had been getting inconsistent accuracy test results on both the 3 and 10 point tests. The manufacturer representative reported different instruments (e.g. accuracy pointer and trocar) yield different accuracy results. No patient was involved. 2020-02-25 UD x2, Email 2020-02-26 16:30 (Rep): Additional information received from a manufacturer representative reported that the arm was replaced. 2020-04-07 IUD (Other): No new information. | xse inconsistent accuracy id medical facility getting inconsistent accuracy test results point tests different instruments yield different accuracy results involved ud email additional arm replaced iud new surgical arm | NAV4093 | NAV4092 | NAV3299 | NAV2054 | NAV3246 | NAV4093NAV4092NAV3299NAV2054NAV3246 | True
3171 | 703268426 | XSE - target extender hole damaged | 2019-07-19T15:33:18.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 20 | MAS2310-03 | TARGET EXTENDER | XSE - target extender hole damaged 2019-Jul-18 SCL1902135 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the target extender Screw hole was damaged and required replacement. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. TARGET EXTENDER | 2019-Jul-18 SCL1902135 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the target extender Screw hole was damaged and required replacement. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | xse target extender hole damaged spinal procedure target extender screw hole damaged required replacement additional complications anticipated target extender | NAV3299 | NAV3295 | NAV2054 | NAV2053 | NAV3264 | NAV3299NAV3295NAV2054NAV2053NAV3264 | True
7666 | 704246547 | XSE - Failed Registration/SRA Low Moveme | 2021-02-24T09:06:55.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Failed Registration/SRA Low Moveme 2021-02-23 00707162, IUD, CCR, Intake Form, e1, e2, e3, e4 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the patient already had screws and rods placed from L4 to S1. Pre-op CT scans showed that the rods were too long and went up to L3 even though no screws were placed at L3. Planning was done by the surgeon's assistant from L3 to S1. Only screws at L3 would be placed. Prior to mounting the surgical system to the patient, the rods were removed. The Schanz pin was placed in the left PSIS and the surgical system was mounted to the patient using a Schanz arm. Registration failed at L3 and there was a red cross on the verify step. The failure was due to the rods being visible in the pre-op CT and then being removed. The surgeon went back to planning and re-defined the region of interest to exclude the rods from the region and narrowing the blue square. Both implant at L3 were then planned. When going to the Operate screen, the surgeon screen and workstation monitor went black for a minute. When sending the surgical arm to left L3, the surgical arm was really low so the surgeon pressed the emergency stop button. The surgical arm was sent to the clear up position, but the arm went down into the work volume instead of going up. The surgeon decided to abort the sue of the guidance system and place the L3 screws freehand. Before unmounting the surgical system from the patient, the emergency stop button was pressed and the surgical arm elbow was manually moved out of the work volume. The surgical arm was then able to be sent to the clear up position. There was no patient harm and the procedure was delayed less than an hour. 2021-02-25 e5, e6, e7, e8 (Rep, For): Additional information received from a manufacturer representative reported that going back to exclude the rods from the region of interest, made them go back to the first step of planning while the system setup was already done. The representative questioned if this was the cause of the software issue by not doing an unmount when going back that far in the planning. The surgical system was fixed to the patient using a Schanz screw in the left PSIS connected to the bone mount bridge. The surgeon decided to do this after the rods were removed. After the case, the accuracy and 3Define camera were checked and everything was fine. 2021-03-08 e9, e10, e11, e12, e13, e14, e15 (Rep, For): No new information. 2021-05-27 IUD (Rep, For): Additional information received from a manufacturer representative reported that the probable cause was the surgical arm being lowered down too much by the user to reach short tools and the arm entered the work volume. The representative indicated that the system had reversed the sense of up and down from the zero point based on the work volume limit, which made the arm go down even with the clear up command. 2021-06-15 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-02-23 00707162, IUD, CCR, Intake Form, e1, e2, e3, e4 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the patient already had screws and rods placed from L4 to S1. Pre-op CT scans showed that the rods were too long and went up to L3 even though no screws were placed at L3. Planning was done by the surgeon's assistant from L3 to S1. Only screws at L3 would be placed. Prior to mounting the surgical system to the patient, the rods were removed. The Schanz pin was placed in the left PSIS and the surgical system was mounted to the patient using a Schanz arm. Registration failed at L3 and there was a red cross on the verify step. The failure was due to the rods being visible in the pre-op CT and then being removed. The surgeon went back to planning and re-defined the region of interest to exclude the rods from the region and narrowing the blue square. Both implant at L3 were then planned. When going to the Operate screen, the surgeon screen and workstation monitor went black for a minute. When sending the surgical arm to left L3, the surgical arm was really low so the surgeon pressed the emergency stop button. The surgical arm was sent to the clear up position, but the arm went down into the work volume instead of going up. The surgeon decided to abort the sue of the guidance system and place the L3 screws freehand. Before unmounting the surgical system from the patient, the emergency stop button was pressed and the surgical arm elbow was manually moved out of the work volume. The surgical arm was then able to be sent to the clear up position. There was no patient harm and the procedure was delayed less than an hour. 2021-02-25 e5, e6, e7, e8 (Rep, For): Additional information received from a manufacturer representative reported that going back to exclude the rods from the region of interest, made them go back to the first step of planning while the system setup was already done. The representative questioned if this was the cause of the software issue by not doing an unmount when going back that far in the planning. The surgical system was fixed to the patient using a Schanz screw in the left PSIS connected to the bone mount bridge. The surgeon decided to do this after the rods were removed. After the case, the accuracy and 3Define camera were checked and everything was fine. 2021-03-08 e9, e10, e11, e12, e13, e14, e15 (Rep, For): No new information. 2021-05-27 IUD (Rep, For): Additional information received from a manufacturer representative reported that the probable cause was the surgical arm being lowered down too much by the user to reach short tools and the arm entered the work volume. The representative indicated that the system had reversed the sense of up and down from the zero point based on the work volume limit, which made the arm go down even with the clear up command. 2021-06-15 IUD (Other): No new information. | xse failed registrationsra low moveme iud ccr spinal procedure treat degenerative discs fixation screws rods placed preop ct scans showed rods long went screws placed planning surgeons assistant screws placed prior mounting surgical rods removed schanz pin placed left psis surgical mounted using schanz arm registration failed red cross verify step failure rods visible preop ct removed surgeon went planning redefined region exclude rods region narrowing blue square implant planned going operate screen surgeon screen workstation monitor went black minute sending surgical arm left surgical arm really low surgeon pressed emergency stop button surgical arm sent clear position arm went work volume instead going surgeon decided abort sue place screws freehand unmounting surgical emergency stop button pressed surgical arm elbow manually moved work volume surgical arm able sent clear position procedure delayed hour additional going exclude rods region step planning setup questioned cause software issue doing unmount going far planning surgical fixed using schanz screw left psis connected bone mount bridge surgeon decided rods removed case accuracy camera checked fine new iud additional probable cause surgical arm lowered user reach short tools arm entered work volume indicated reversed sense zero point based work volume limit arm clear command iud new mazor x | NAV3261 | NAV3253 | NAV4096 | NAV3246 | NAV5000 | NAV3261NAV3253NAV4096NAV3246NAV5000 | False
1091 | 703188398 | XSE - snapshot button greyed out/screen | 2019-06-04T16:02:01.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0356-01 | SURGEON SCREEN | XSE - snapshot button greyed out/screen 2019-May-29 SCL1901566 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon screen went black and had to soft restart to fix. It was also reported that they were unable to Aquire Snapshot because the button was greyed out. A soft restart also fixed the issue but the rep lost the Nav Code and KIT code which was a major delay. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. SURGEON SCREEN | 2019-May-29 SCL1901566 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgeon screen went black and had to soft restart to fix. It was also reported that they were unable to Aquire Snapshot because the button was greyed out. A soft restart also fixed the issue but the rep lost the Nav Code and KIT code which was a major delay. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | xse snapshot button greyed outscreen spinal procedure surgeon screen went black soft restart fix unable aquire snapshot button greyed soft restart fixed issue lost nav code kit code major delay additional complications anticipated surgeon screen | NAV3246 | NAV3253 | NAV3249 | NAV2047 | NAV3295 | NAV3246NAV3253NAV3249NAV2047NAV3295 | True
5259 | 703632072 | XSE - Deviation Superior at L4 | 2020-02-24T16:11:32.000000+0000 | MazorX Stealth Edition | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Deviation Superior at L4 2020-02-21 00684757 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a trajectory at L4 was superior to plan. The doctor freehanded the final placement. No patient harm was reported and surgery was delayed less than an hour. 2020-02-25 email 2020-02-25 11:30, IUD x2 (Rep): Additional information received from a manufacturer representative reported the surgical arm failed an accuracy test and the screws were deviated between 3.5 and 10 mm. KIT SVC ASM0206-01S SURGICAL ARM | 2020-02-21 00684757 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a trajectory at L4 was superior to plan. The doctor freehanded the final placement. No patient harm was reported and surgery was delayed less than an hour. 2020-02-25 email 2020-02-25 11:30, IUD x2 (Rep): Additional information received from a manufacturer representative reported the surgical arm failed an accuracy test and the screws were deviated between 3.5 and 10 mm. | xse deviation superior id spinal procedure trajectory superior plan doctor freehanded final placement surgery delayed hour email iud additional surgical arm failed accuracy test screws deviated mm kit svc surgical arm | NAV5000 | NAV4092 | NAV3249 | NAV4093 | NAV4097 | NAV5000NAV4092NAV3249NAV4093NAV4097 | True
2155 | 703226603 | X - Markers Issue | 2019-06-25T19:51:21.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Markers Issue -Stylus brought right in front of eyes, and it picked it up, but when placed on patient in the ROI, it could not pick it up. -X-eyes were moved to look more towards the stylus, still did not pick up. Registration Arm sent to AP -CSR moved arm more in the ROI -2 images Arm sent to OBL -CSR moved arm closer to patient and towards the foot -4 images Registration achieved on first try and visually verified by surgeon. Operation: Operation was quick and accurate, no skives present, and Dr. Nichols was very pleased with the overall outcome, speed, and accuracy of the robot, UNK MAZOR X X-EYE CAMERA | -Stylus brought right in front of eyes, and it picked it up, but when placed on patient in the ROI, it could not pick it up. -X-eyes were moved to look more towards the stylus, still did not pick up. Registration Arm sent to AP -CSR moved arm more in the ROI -2 images Arm sent to OBL -CSR moved arm closer to patient and towards the foot -4 images Registration achieved on first try and visually verified by surgeon. Operation: Operation was quick and accurate, no skives present, and Dr. Nichols was very pleased with the overall outcome, speed, and accuracy of the robot, | x markers issue stylus brought right eyes picked placed roi pick xeyes moved look stylus did pick registration arm sent ap csr moved arm roi images arm sent obl csr moved arm closer foot images registration achieved try visually verified surgeon operation operation quick accurate skives present dr nichols pleased overall outcome speed accuracy robot unk mazor x xeye camera | NAV4110 | NAV4101 | NAV3244 | NAV3295 | NAV3237 | NAV4110NAV4101NAV3244NAV3295NAV3237 | True
5789 | 703802704 | XSE - Workstation display cut out | 2020-06-15T19:54:07.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XSE - Workstation display cut out 2020-06-15 00690879, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a procedure to treat degenerative disc with fixation. It was reported planning was completed for a L1-L2 with a percutaneous incision plan. A Schanz screw was inserted into the right PSIS and then mounted to the surgical system using the Schanz arm. A 3Define scan and draw spine were completed successfully. AP and oblique images were taken. The AP image was accepted on the second try and the oblique image was accepted on the first try. Registration was successful with green values and no shifts. All trajectories were sent, drilled and had k-wires placed in the pilot holes. The left L1 wire was pulled prematurely so the the trajectory was resent and the wire was placed again. All wires were accurate when checked with fluoro and neuromonitoring. When the surgeon was finished with the guidance system, they wanted the manufacturer representative to unlock the mount. The representative pressed the unlock mount button on the screen, and the workstation display went black and the system did not unlock. After two minutes the display returned to normal and the r epresentative was able to unlock the mount. There was no patient harm and the procedure was not delayed. 2020-06-30 IUD (Other): No new information. UNK MAZOR X MONITOR | 2020-06-15 00690879, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a procedure to treat degenerative disc with fixation. It was reported planning was completed for a L1-L2 with a percutaneous incision plan. A Schanz screw was inserted into the right PSIS and then mounted to the surgical system using the Schanz arm. A 3Define scan and draw spine were completed successfully. AP and oblique images were taken. The AP image was accepted on the second try and the oblique image was accepted on the first try. Registration was successful with green values and no shifts. All trajectories were sent, drilled and had k-wires placed in the pilot holes. The left L1 wire was pulled prematurely so the the trajectory was resent and the wire was placed again. All wires were accurate when checked with fluoro and neuromonitoring. When the surgeon was finished with the guidance system, they wanted the manufacturer representative to unlock the mount. The representative pressed the unlock mount button on the screen, and the workstation display went black and the system did not unlock. After two minutes the display returned to normal and the r epresentative was able to unlock the mount. There was no patient harm and the procedure was not delayed. 2020-06-30 IUD (Other): No new information. | xse workstation display cut ccr procedure treat degenerative disc fixation planning completed percutaneous incision plan schanz screw inserted right psis mounted surgical using schanz arm scan draw spine completed successfully ap oblique images taken ap image accepted second try oblique image accepted try registration successful green values shifts trajectories sent drilled kwires placed pilot holes left wire pulled prematurely trajectory resent wire placed wires accurate checked fluoro neuromonitoring surgeon finished wanted unlock mount pressed unlock mount button screen workstation display went black did unlock minutes display returned normal r epresentative able unlock mount procedure delayed iud new unk mazor x monitor | NAV3261 | NAV2058 | NAV3260 | NAV3257 | NAV4110 | NAV3261NAV2058NAV3260NAV3257NAV4110 | True
349 | 703083458 | X - shoulder won't open or close | 2019-04-01T14:17:02.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - shoulder won't open or close 2019-Mar-27 SCL1900754 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the shoulder was in-between open and closed and couldn't open or close the shoulder through the software. They tried override panel and general tab in utilities. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X SHOULDER | 2019-Mar-27 SCL1900754 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the shoulder was in-between open and closed and couldn't open or close the shoulder through the software. They tried override panel and general tab in utilities. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x shoulder wont open close spinal procedure shoulder inbetween open closed open close shoulder software tried override panel general tab utilities noted additional complications anticipated unk mazor x shoulder | NAV4107 | NAV4108 | NAV4106 | NAV2058 | NAV2054 | NAV4107NAV4108NAV4106NAV2058NAV2054 | True
9354 | 704614877 | XSE - Unable to Lock Positioner | 2021-09-16T17:58:42.000000+0000 | MazorX Stealth Edition | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 50 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - Unable to Lock Positioner 2021-Sep-16 Salesforce C00234555 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the sensor on the jaws of the positioner were not being triggered and the positioner was unable to lock. The metal piece that triggers the sensor had fallen out and cannot be re-inserted because it was stepped on and the housing that it fit into seemed to be slightly damaged causing it to fall out. No patient was present. 2021-09-16 IUD x6 (Other): No new information. 2021-09-29 IUD x3 (Other): No new information. 2021-10-08 IUD x2 (Other): No new information. 2021-10-26 IUD x3 (Other): No new information. 2021-11-14 IUD (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-Sep-16 Salesforce C00234555 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the sensor on the jaws of the positioner were not being triggered and the positioner was unable to lock. The metal piece that triggers the sensor had fallen out and cannot be re-inserted because it was stepped on and the housing that it fit into seemed to be slightly damaged causing it to fall out. No patient was present. 2021-09-16 IUD x6 (Other): No new information. 2021-09-29 IUD x3 (Other): No new information. 2021-10-08 IUD x2 (Other): No new information. 2021-10-26 IUD x3 (Other): No new information. 2021-11-14 IUD (Other): No new information. | xse unable lock positioner salesforce outside procedure sensor jaws positioner triggered positioner unable lock metal piece triggers sensor fallen reinserted stepped housing fit slightly damaged causing fall present iud new iud new iud new iud new iud new kit svc surgical syst type | NAV3251 | NAV4103 | NAV4104 | NAV3250 | NAV3249 | NAV3251NAV4103NAV4104NAV3250NAV3249 | True
5971 | 703849701 | XSE - OOBF/SRA J4 High Current | 2020-07-13T23:12:48.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - OOBF/SRA J4 High Current 2020-07-13 00692626, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was flipped the images on the setup screen after install. There was no patient involvement. 2020-07-20 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-07-13 00692626, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was flipped the images on the setup screen after install. There was no patient involvement. 2020-07-20 IUD (Other): No new information. | xse oobfsra high current iud outside procedure surgical arm flipped images setup screen install iud new kit svc surgical arm refurb | NAV2054 | NAV3249 | NAV1006 | NAV4101 | NAV4095 | NAV2054NAV3249NAV1006NAV4101NAV4095 | True
4551 | 703550776 | X - X-eye Not Tracking Stylus | 2020-01-08T14:28:14.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - X-eye Not Tracking Stylus 2018-04-23 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the procedure was planned for fixation at L3-S1. The X-eye was having trouble seeing the stylus for Draw Spine. However, registration was able to be accomplished. The surgeon made the intraoperative decision not to place screws at L3. All trajectories for L4-S1 were placed accurately with left side trajectories placed first, then right side trajectories. No patient harm was reported and surgery was delayed less than an hour. UNK MAZOR X X-EYE CAMERA | 2018-04-23 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the procedure was planned for fixation at L3-S1. The X-eye was having trouble seeing the stylus for Draw Spine. However, registration was able to be accomplished. The surgeon made the intraoperative decision not to place screws at L3. All trajectories for L4-S1 were placed accurately with left side trajectories placed first, then right side trajectories. No patient harm was reported and surgery was delayed less than an hour. | x xeye tracking stylus ccr spinal procedure procedure planned fixation xeye having trouble seeing stylus draw spine registration able accomplished surgeon intraoperative decision place screws trajectories placed accurately left trajectories placed right trajectories surgery delayed hour unk mazor x xeye camera | NAV4110 | NAV3254 | NAV2058 | NAV3299 | NAV4095 | NAV4110NAV3254NAV2058NAV3299NAV4095 | True
7825 | 704285959 | XSE - CANBUS Error | 2021-03-16T20:00:32.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error 2021-Mar-16 Salesforce C00222746 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site had CANBUS no connection.to controller errors. There was no patient present. 2021-03-23 IUD x2 (Other): No new information 2021-04-02 IUD x3 (Other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2021-Mar-16 Salesforce C00222746 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site had CANBUS no connection.to controller errors. There was no patient present. 2021-03-23 IUD x2 (Other): No new information 2021-04-02 IUD x3 (Other): No new information. | xse canbus error salesforce outside procedure site canbus connectionto controller errors present iud new iud new svc main cable | NAV2047 | NAV2054 | NAV3256 | NAV3249 | NAV1006 | NAV2047NAV2054NAV3256NAV3249NAV1006 | True
939 | 703181370 | X- Surgical Arm Issues | 2019-05-30T18:55:24.000000+0000 | MazorX | NAV4100 | STRIPPED THREADS | 20 | ASM0206-01 | SURGICAL ARM | X- Surgical Arm Issues Reference marker hole stripped on the arm SURGICAL ARM | Reference marker hole stripped on the arm | x surgical arm issues reference marker hole stripped arm surgical arm | NAV4100 | NAV3249 | NAV4095 | NAV4101 | NAV4098 | NAV4100NAV3249NAV4095NAV4101NAV4098 | True
3351 | 703320647 | XSE - Screw Planning Issue | 2019-08-20T23:12:02.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Screw Planning Issue 2019-08-20 00672562, ID (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that when planning screws where the tulip heads collide, diagonal lines occur and the image looks streaky. The representative was planning the screws with the stealth edition software on their laptop. When the representative selected "show rods" the distortion went away. The issue continued to occur show full CT was selected. There was no patient involvement. 2019-09-08 email 2019-09-08 23:15 (Rep): No new information. 2019-09-26 email 2019-09-26 17:00 x3, video 2019-09-27 10:57 (Rep): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-08-20 00672562, ID (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that when planning screws where the tulip heads collide, diagonal lines occur and the image looks streaky. The representative was planning the screws with the stealth edition software on their laptop. When the representative selected "show rods" the distortion went away. The issue continued to occur show full CT was selected. There was no patient involvement. 2019-09-08 email 2019-09-08 23:15 (Rep): No new information. 2019-09-26 email 2019-09-26 17:00 x3, video 2019-09-27 10:57 (Rep): No new information. | xse screw planning issue id prior procedure planning screws tulip heads collide diagonal lines occur image looks streaky planning screws stealth edition software laptop selected rods distortion went away issue continued occur ct selected email new email video new mazor x software installation kit | NAV3246 | NAV5000 | NAV2047 | NAV3247 | NAV2058 | NAV3246NAV5000NAV2047NAV3247NAV2058 | True
6180 | 703882423 | XSE ¦ Non-Functioning Pneumatic Man | 2020-08-01T03:23:29.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | KIT1293-01 | SVC KIT1293-01 PNEUMATIC MANIFOLD | XSE ¦ Non-Functioning Pneumatic Man 2020-07-31 00693856 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that a valve on the pneumatic manifold is not operating and the teeth on the positioner will not release to be unmounted from the bed. No patient was present. 2020-08-07 IUD x2 (Other): No new information. 2020-08-30 IUD x4 (Other): No new information. SVC KIT1293-01 PNEUMATIC MANIFOLD | 2020-07-31 00693856 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that a valve on the pneumatic manifold is not operating and the teeth on the positioner will not release to be unmounted from the bed. No patient was present. 2020-08-07 IUD x2 (Other): No new information. 2020-08-30 IUD x4 (Other): No new information. | xse nonfunctioning pneumatic man outside procedure valve pneumatic manifold operating teeth positioner release unmounted bed present iud new iud new svc pneumatic manifold | NAV3264 | NAV2054 | NAV4103 | NAV3251 | NAV3250 | NAV3264NAV2054NAV4103NAV3251NAV3250 | True
9968 | 704761600 | XSE ¦ Intermittent Pneumatic issue | 2021-12-13T16:26:34.000000+0000 | MazorX Stealth Edition | NAV2057 | SURGICAL SYSTEM ISSUES | 20 | ASM0214-02 | SSY | XSE ¦ Intermittent Pneumatic issue 2021-Dec-13 Salesforce C00239585 (rep, hcp): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the rep was having problems with the pneumatic movement on this unit. Before the case they had a problem with getting the hood to close and open. After the case he attempted to place the positioner back in the hood and the manipulator a rm was stuck locked into place. He hasn't heard the compressor turn on recently. 2021-12-21 IUD (Other): No new information. SSY | 2021-Dec-13 Salesforce C00239585 (rep, hcp): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the rep was having problems with the pneumatic movement on this unit. Before the case they had a problem with getting the hood to close and open. After the case he attempted to place the positioner back in the hood and the manipulator a rm was stuck locked into place. He hasn't heard the compressor turn on recently. 2021-12-21 IUD (Other): No new information. | xse intermittent pneumatic issue salesforce outside proc edure having problems pneumatic movement unit case problem getting hood close open case attempted place positioner hood manipulator rm stuck locked place heard compressor turn recently iud new ssy | NAV3264 | NAV4103 | NAV2058 | NAV4105 | NAV3260 | NAV3264NAV4103NAV2058NAV4105NAV3260 | False
5358 | 703660341 | XSE - Leak in Pneumatic Line | 2020-03-10T23:05:22.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Leak in Pneumatic Line 2020-03-10 00686060 (rep, hcp): Medtronic received information regarding a guidance system being used for a spinal procedure. It was reported that when the manufacturer representative took the surgical system off the bed and hit the side positioner buttons, he was hearing a leak noise. Additionally, when hitting the matrix button, the leak sound was louder. This occurred at the end of the case when the system was being removed from the bed and thus caused no delay to the procedure. There was no impact to patient outcome. 2020-03-11 Email 2020-03-11 17:30, IUD x5 (Other): Additional information received from a manufacturer representative reported there was a leak in the positioner handle. The representative assumes that the pneumatic line got pinched which caused a little hole. The engineer cut out the leak in the pneumatic line and sealed the line with a pneumatic coupler. System was able to maintain pressure without any air leaks when manipulator was activated. System is fully functional. 2020-04-14 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2020-03-10 00686060 (rep, hcp): Medtronic received information regarding a guidance system being used for a spinal procedure. It was reported that when the manufacturer representative took the surgical system off the bed and hit the side positioner buttons, he was hearing a leak noise. Additionally, when hitting the matrix button, the leak sound was louder. This occurred at the end of the case when the system was being removed from the bed and thus caused no delay to the procedure. There was no impact to patient outcome. 2020-03-11 Email 2020-03-11 17:30, IUD x5 (Other): Additional information received from a manufacturer representative reported there was a leak in the positioner handle. The representative assumes that the pneumatic line got pinched which caused a little hole. The engineer cut out the leak in the pneumatic line and sealed the line with a pneumatic coupler. System was able to maintain pressure without any air leaks when manipulator was activated. System is fully functional. 2020-04-14 IUD (Other): No new information. | xse leak pneumatic line spinal procedure took surgical bed hit positioner buttons hearing leak noise additionally hitting matrix button leak sound louder occurred end case removed bed caused delay procedure impact outcome email iud additional leak positioner handle assumes pneumatic line got pinched caused little hole engineer cut leak pneumatic line sealed line pneumatic coupler able maintain pressure air leaks manipulator activated fully functional iud new unk mazor x workstation | NAV3264 | NAV1006 | NAV4096 | NAV3251 | NAV3244 | NAV3264NAV1006NAV4096NAV3251NAV3244 | True
6242 | 703895888 | XSE - Navigation Depth Off | 2020-08-10T19:29:02.000000+0000 | MazorX Stealth Edition | NAV5001 | MXSE - MISMATCH | 20 | M450001B020 | NAVIGATION SOFTWARE MODULE 1.1 | XSE - Navigation Depth Off 2020-08-10 00694297, email 2020-08-14 15:45, email 2020-08-19 15:15, email 2020-08-25 12:31, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that an L3-S1 OLIF was completed in a single position. At the first trajectory of right L3, the navigation was showing 3-5 mm deep in the bone when holding the drill right on top of the bone. The trajectory was accurate, but the navigation depth did not show correctly. The case was continued and when the surgeon went back to right L3, the navigation depth was accurate. Accuracy was checked with fluoro and all trajectories were accurate. The manufacturer representative thought there was a software issue with the depth. There was no patient harm and the procedure was delayed less than an hour. NAVIGATION SOFTWARE MODULE 1.1 | 2020-08-10 00694297, email 2020-08-14 15:45, email 2020-08-19 15:15, email 2020-08-25 12:31, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that an L3-S1 OLIF was completed in a single position. At the first trajectory of right L3, the navigation was showing 3-5 mm deep in the bone when holding the drill right on top of the bone. The trajectory was accurate, but the navigation depth did not show correctly. The case was continued and when the surgeon went back to right L3, the navigation depth was accurate. Accuracy was checked with fluoro and all trajectories were accurate. The manufacturer representative thought there was a software issue with the depth. There was no patient harm and the procedure was delayed less than an hour. | xse navigation depth email email email iud spinal procedure olif completed single position trajectory right navigation showing mm deep bone holding drill right bone trajectory accurate navigation depth did correctly case continued surgeon went right navigation depth accurate accuracy checked fluoro trajectories accurate thought software issue depth procedure delayed hour navigation software module | NAV5001 | NAV5000 | NAV5002 | NAV3249 | NAV3246 | NAV5001NAV5000NAV5002NAV3249NAV3246 | True
347 | 703083407 | X - x-eye noise/failed registration | 2019-04-01T13:10:28.000000+0000 | MazorX | NAV3254 | X-EYE ISSUES | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - x-eye noise/failed registration 2019-Mar-27 SCL1900750 intake form, msg, CCR (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the X Eye made a grinding / buzzing sound during most of the case and the only way to make it stop was to manually move it. Registration was successful for the placement on the T12-L3 wires and screws, but the second segment registration of L4-S1 failed. The surgeons used Ant buttress plates. The rep suggested they retake flouro shots but the HCP wanted to abort and move on without Mazor. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. Additional information received indicated that the issue was determined to be caused by a drape within the X-eye motor restricting the movement. The system was working as designed. UNK MAZOR X X-EYE CAMERA | 2019-Mar-27 SCL1900750 intake form, msg, CCR (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the X Eye made a grinding / buzzing sound during most of the case and the only way to make it stop was to manually move it. Registration was successful for the placement on the T12-L3 wires and screws, but the second segment registration of L4-S1 failed. The surgeons used Ant buttress plates. The rep suggested they retake flouro shots but the HCP wanted to abort and move on without Mazor. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. Additional information received indicated that the issue was determined to be caused by a drape within the X-eye motor restricting the movement. The system was working as designed. | x xeye noisefailed registration msg ccr spinal procedure x eye grinding buzzing sound case way make stop manually registration successful placement wires screws second segment registration failed surgeons ant buttress plates suggested retake flouro shots hcp wanted abort mazor additional complications anticipated additional indicated issue determined caused drape xeye motor restricting movement working designed unk mazor x xeye camera | NAV4110 | NAV4112 | NAV3261 | NAV3254 | NAV2054 | NAV4110NAV4112NAV3261NAV3254NAV2054 | True
1262 | 703194774 | X - DEVIATION | 2019-06-07T12:59:25.000000+0000 | MazorX | NAV4110 | NO TRACKING | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - DEVIATION This complaint is separated to two complaints:Medial deviation - this complaint. 3Define camera not working- SCL1801777. 3define camera was not working, soft restart done 2-3 times in order to get the camera functioning. L5 right was medial – surgeon requests an engineer come look at the robot./5.7.18 Dr. Francis L5-S1 TLIF Preop: Cameras checked, AA completed, shoulder checked. Plan: Done by Dr. Francis, axial angles were narrow and uniform from level to level. Schanz arm used and schanz pin placed in right PSIS. 3Define not completed on the firs try, RTC started 3define, but the 3define cameras were not working-RTC did a soft restart, and tried to start the 3define again, but the 3define camera was not responding again. '-RTC did another soft restart, this time it worked and 3Define was completed. -Drawspine completed on 2nd try -first try, the surgeon interfered with the the x-eye. Registration: Completed on the first try with green values and visually verified by surgeon.Operation: Operation seemed quick and accurate, and surgeon did not seem like anything seemed inaccurate. When final AP was taken surgeon said that the right L5 was medial. Surgeon removed the robot and said he would free hand the right L5. Surgeon requests that the engineers come and check the robots accuracy. UNK MAZOR X X-EYE CAMERA | This complaint is separated to two complaints:Medial deviation - this complaint. 3Define camera not working- SCL1801777. 3define camera was not working, soft restart done 2-3 times in order to get the camera functioning. L5 right was medial – surgeon requests an engineer come look at the robot./5.7.18 Dr. Francis L5-S1 TLIF Preop: Cameras checked, AA completed, shoulder checked. Plan: Done by Dr. Francis, axial angles were narrow and uniform from level to level. Schanz arm used and schanz pin placed in right PSIS. 3Define not completed on the firs try, RTC started 3define, but the 3define cameras were not working-RTC did a soft restart, and tried to start the 3define again, but the 3define camera was not responding again. '-RTC did another soft restart, this time it worked and 3Define was completed. -Drawspine completed on 2nd try -first try, the surgeon interfered with the the x-eye. Registration: Completed on the first try with green values and visually verified by surgeon.Operation: Operation seemed quick and accurate, and surgeon did not seem like anything seemed inaccurate. When final AP was taken surgeon said that the right L5 was medial. Surgeon removed the robot and said he would free hand the right L5. Surgeon requests that the engineers come and check the robots accuracy. | x deviation complaint separated complaintsmedial deviation complaint camera working camera working soft restart times order camera functioning right medial – surgeon requests engineer come look dr francis tlif preop cameras checked completed shoulder checked plan dr francis axial angles narrow uniform level level schanz arm schanz pin placed right psis completed firs try rtc started cameras workingrtc did soft restart tried start camera responding rtc did soft restart time worked completed drawspine completed try try surgeon interfered xeye registration completed try green values visually verified surgeonoperation operation quick accurate surgeon did like inaccurate final ap taken surgeon said right medial surgeon removed robot said free hand right surgeon requests engineers come check robots accuracy unk mazor x xeye camera | NAV4110 | NAV3244 | NAV4091 | NAV3233 | NAV3241 | NAV4110NAV3244NAV4091NAV3233NAV3241 | True
5360 | 703660344 | XSE - Damaged Image Adapter | 2020-03-10T23:05:40.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 40 | KIT0586-03 | REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | XSE - Damaged Image Adapter 2020-03-10 00686074 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the image adapter on the system is damaged and needs to be replaced. There was no patient involvement. 2020-03-11 IUD x4 (Other): No new information. 2020-03-31 IUD x7 (Other): No new information. 2020-04-23 IUD x2 (Other): No new information. REPLACEMENT GRIP SET FOR IMAGE ADAPTOR | 2020-03-10 00686074 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the image adapter on the system is damaged and needs to be replaced. There was no patient involvement. 2020-03-11 IUD x4 (Other): No new information. 2020-03-31 IUD x7 (Other): No new information. 2020-04-23 IUD x2 (Other): No new information. | xse damaged image adapter outside procedure image adapter damaged needs replaced iud new iud new iud new replacement grip set image adaptor | NAV3258 | NAV2054 | NAV3295 | NAV3299 | NAV3256 | NAV3258NAV2054NAV3295NAV3299NAV3256 | True
8744 | 704475122 | XSE - Unable to Transfer Exam | 2021-06-29T15:58:44.000000+0000 | MazorX Stealth Edition | NAV3236 | CT IMAGES | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - Unable to Transfer Exam 2021-Jun-29 Salesforce C00229248, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the manufacturer representative had trouble transferring exams to the guidance system from a C-am. The representative confirmed there was a green communication connection. A soft reboot was done, but the issue was not resolved. Troubleshooting found the system was not retaining the the "AE Title, IP Address, or Port Number." There was no patient involvement. 2021-Jul-06 IUDx5 (other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Jun-29 Salesforce C00229248, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the manufacturer representative had trouble transferring exams to the guidance system from a C-am. The representative confirmed there was a green communication connection. A soft reboot was done, but the issue was not resolved. Troubleshooting found the system was not retaining the the "AE Title, IP Address, or Port Number." There was no patient involvement. 2021-Jul-06 IUDx5 (other): No new information. | xse unable transfer exam salesforce iud outside proc edure trouble transferring exams cam confirmed green communication connection soft reboot issue resolved troubleshooting retaining ae title ip address port number new mazorx sw installation kit | NAV3246 | NAV2047 | NAV3235 | NAV2058 | NAV3247 | NAV3246NAV2047NAV3235NAV2058NAV3247 | False
2863 | 703256444 | X - Surgical arm Issue | 2019-07-12T17:54:12.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | X - Surgical arm Issue Completely non-functional. The RBT arm will not maneuver into any predefined locations or complete accuracy testing. SURGICAL ARM | Completely non-functional. The RBT arm will not maneuver into any predefined locations or complete accuracy testing. | x surgical arm issue completely nonfunctional rbt arm maneuver predefined locations complete accuracy testing surgical arm | NAV3249 | NAV4101 | NAV4095 | NAV2054 | NAV1006 | NAV3249NAV4101NAV4095NAV2054NAV1006 | True
9269 | 704594068 | XSE - Deviation | 2021-09-03T22:11:48.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviation 2021-Aug-06 Salesforce C00233991, email 2021-09-08 08:16, e1, e2 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a L4-L5 TLIF spinal procedure. It was reported that the L4 screws were not on trajectory during the procedure. There was a big difference between the planned trajectory and where the screw was placed. The amount of deviation was unknown. The manufacturer representative noted that the patient's anatomy was "weird." The L5 screws were accurately placed. The representative suspected the L4 screws had skived. The planning of L4 was adjusted and the screw was able to be placed with the guidance system. Prior to the procedure, the surgical arm passed the setup and advanced accuracy tests. There was no patient harm and the procedure was delayed less than an hour. 2021-11-08 e3 (Other, For): No new information. 2021-11-10 e4, e5 (Rep, For): Additional information received from a manufacturer representative reported that the surgeon followed their normal workflow. The surgeon's workflow included the following steps: planning the trajectory, send the surgical arm to the trajectory, mark the entry point on the patient's skin, make the incision, use the scalpel to gain access, place the cannula with the dilator down the arm guide, place the drill guide, drill the hole, place the reduction tube, insert the guide wire, tap and then place the screw. These steps were done with CT fluoro and the screw was found to be low during that time when compared to the plan. An interbody was noted to be placed. MAZOR X SYSTEM TPL0059 | 2021-Aug-06 Salesforce C00233991, email 2021-09-08 08:16, e1, e2 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a L4-L5 TLIF spinal procedure. It was reported that the L4 screws were not on trajectory during the procedure. There was a big difference between the planned trajectory and where the screw was placed. The amount of deviation was unknown. The manufacturer representative noted that the patient's anatomy was "weird." The L5 screws were accurately placed. The representative suspected the L4 screws had skived. The planning of L4 was adjusted and the screw was able to be placed with the guidance system. Prior to the procedure, the surgical arm passed the setup and advanced accuracy tests. There was no patient harm and the procedure was delayed less than an hour. 2021-11-08 e3 (Other, For): No new information. 2021-11-10 e4, e5 (Rep, For): Additional information received from a manufacturer representative reported that the surgeon followed their normal workflow. The surgeon's workflow included the following steps: planning the trajectory, send the surgical arm to the trajectory, mark the entry point on the patient's skin, make the incision, use the scalpel to gain access, place the cannula with the dilator down the arm guide, place the drill guide, drill the hole, place the reduction tube, insert the guide wire, tap and then place the screw. These steps were done with CT fluoro and the screw was found to be low during that time when compared to the plan. An interbody was noted to be placed. | xse deviation salesforce email tlif spinal procedure screws trajectory procedure big difference planned trajectory screw placed deviation unknown noted patients anatomy weird screws accurately placed suspected screws skived planning adjusted screw able placed prior procedure surgical arm passed setup advanced accuracy tests procedure delayed hour new additional surgeon followed normal workflow surgeons workflow included following steps planning trajectory send surgical arm trajectory mark entry point patients skin make incision use scalpel gain access place cannula dilator arm guide place drill guide drill hole place reduction tube insert guide wire tap place screw steps ct fluoro screw low time compared plan interbody noted placed mazor x | NAV5000 | NAV3299 | NAV3243 | NAV5001 | NAV3294 | NAV5000NAV3299NAV3243NAV5001NAV3294 | True
8339 | 704399008 | XSE - SRA Release Not Working | 2021-05-17T23:53:17.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Release Not Working 2021-May-17 Salesforce C00226574 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the release for the surgical arm stopped working. The manufacturer representative took the arm apart and thought the plug was loose. The light came back on, but the release to move the surgical arm did not work. There was no patient involvement. 2021-05-21 IUD x3 (Other): Additional information was received stating that at the time of the issue the connection to the shoulder button was secured, the system was then operational. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-May-17 Salesforce C00226574 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the release for the surgical arm stopped working. The manufacturer representative took the arm apart and thought the plug was loose. The light came back on, but the release to move the surgical arm did not work. There was no patient involvement. 2021-05-21 IUD x3 (Other): Additional information was received stating that at the time of the issue the connection to the shoulder button was secured, the system was then operational. | xse sra release working salesforce outside procedure release surgical arm stopped working took arm apart thought plug loose light came release surgical arm did work iud additional stating time issue connection shoulder button secured operational kit svc surgical arm refurb | NAV3249 | NAV4096 | NAV2054 | NAV4095 | NAV4099 | NAV3249NAV4096NAV2054NAV4095NAV4099 | True
6165 | 703877873 | XSE - Software Crash | 2020-07-29T23:22:07.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Software Crash 2020-07-29 00693680 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that once all the trajectories were placed the manufacturer representative attempted to press the back door in the top left to dismount the system. The system froze and the screen went black at the main monitor and surgeon screen. The system rebooted and brought to the same screen and they were able to click the back door and system acted as normal. There was no delay to the procedure and no impact on the patient outcome. 2020-07-30 email 2020-07-30 13:30, IUD x2 (Rep): No new information. 2020-08-12 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-07-29 00693680 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that once all the trajectories were placed the manufacturer representative attempted to press the back door in the top left to dismount the system. The system froze and the screen went black at the main monitor and surgeon screen. The system rebooted and brought to the same screen and they were able to click the back door and system acted as normal. There was no delay to the procedure and no impact on the patient outcome. 2020-07-30 email 2020-07-30 13:30, IUD x2 (Rep): No new information. 2020-08-12 IUD (Other): No new information. | xse software crash intraoperatively spinal procedure trajectories placed attempted press door left dismount froze screen went black main monitor surgeon screen rebooted brought screen able click door acted normal delay procedure impact outcome email iud new iud new mazor x | NAV1012 | NAV3247 | NAV3248 | NAV1006 | NAV5000 | NAV1012NAV3247NAV3248NAV1006NAV5000 | True
10271 | 704843463 | XSE - Frequent Low Pressure Warning | 2022-02-02T21:05:38.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Frequent Low Pressure Warning 2022-Feb-02 Salesforce C00242616 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that frequent low pressure warnings were displayed on the system. The manufacturer representative had to manually release pressure through the valve to get the compressor to turn on. When the low pressure messages occurred, the representative was unable to move the positioner. There was no patient involvement. 2022-02-08 IUD x2 (Other): No new information. 2022-03-01 IUD (Other): No new information. UNK MAZOR X WORKSTATION | 2022-Feb-02 Salesforce C00242616 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that frequent low pressure warnings were displayed on the system. The manufacturer representative had to manually release pressure through the valve to get the compressor to turn on. When the low pressure messages occurred, the representative was unable to move the positioner. There was no patient involvement. 2022-02-08 IUD x2 (Other): No new information. 2022-03-01 IUD (Other): No new information. | xse frequent low pressure warning salesforce outside procedure frequent low pressure warnings displayed manually release pressure valve compressor turn low pressure messages occurred unable positioner iud new iud new unk mazor x workstation | NAV3264 | NAV4104 | NAV1006 | NAV3261 | NAV3256 | NAV3264NAV4104NAV1006NAV3261NAV3256 | True
8061 | 704344780 | XBU ¦ Defective Screwdriver | 2021-04-16T07:01:20.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2202-04S | KIT SVC MAS2202-04S CLAMP LINK | XBU ¦ Defective Screwdriver 2021-Apr-16 MPXR 828963 (rep, for, uf): Information was received from multiple sources (Manufacturer Representative, Foreign, User Facility) regarding an instrument used for Spinal Therapy. It was reported that the part of the screwdriver where the Navlock sits on started to wear down, makes the screwdriver harder to turn. There was no patient involved in the event. 2021-Jun-23, Sukhia2: Additional info received from the rep. After the case the disposable spine clamp could not be taken of the link bridge, as the screw threads were sheered of. 2021-06-18 Attachments x5 (Other): No new information. 2021-08-05 email 2021-08-05 11:46 (Rep): No new information. KIT SVC MAS2202-04S CLAMP LINK | 2021-Apr-16 MPXR 828963 (rep, for, uf): Information was received from multiple sources (Manufacturer Representative, Foreign, User Facility) regarding an instrument used for Spinal Therapy. It was reported that the part of the screwdriver where the Navlock sits on started to wear down, makes the screwdriver harder to turn. There was no patient involved in the event. 2021-Jun-23, Sukhia2: Additional info received from the rep. After the case the disposable spine clamp could not be taken of the link bridge, as the screw threads were sheered of. 2021-06-18 Attachments x5 (Other): No new information. 2021-08-05 email 2021-08-05 11:46 (Rep): No new information. | xbu defective screwdriver mpxr multiple sources instrument spinal therapy screwdriver navlock sits started wear makes screwdriver harder turn involved additional info case disposable spine clamp taken link bridge screw threads sheered attachments new email new kit svc clamp link | NAV3299 | NAV3295 | NAV2054 | NAV3298 | NAV3250 | NAV3299NAV3295NAV2054NAV3298NAV3250 | True
2964 | 703261186 | X - Registration Issue | 2019-07-16T14:54:26.000000+0000 | MazorX | NAV4110 | NO TRACKING | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Registration Issue Significant Crossview error on second registration. First registration accurate. Horizontal line on AP image (when placing red reference point) appeared to be vertical as opposed to horizontal across the desired body during flouro segmentation./ Surgeon, Fellow, & RTC planned case prior to procedure and approved implant placement and trajectory plan. Single Arm Shanz pin agreed to be placed in left PSIS for the MIS procedure. Cage was placed in the lateral position at L3/4, Surgeon opted to add screws to L1/2 and L5/S1 for stability. System mounted and placed onto bed without any issue, arm draped, pin placed in Left PSIS. Robot mounted to patient without problem. While draping, scrub tech wrapped red rubber bands for X Eye cameras around the neck of the camera, causing an inability for the camera to move properly. 3Define executed properly as planned. X Eyes unable to consistently pick up the reference marker and stylus, RTC opted to move forward to generic AP & OB location for 3D Marker acquisition. Significant manual adjustments needed in order to place 3D marker in proper location. L1-S1 was unable to be achieved in the imaging due to larger patient, surgeon requested to register twice to achieve L5/S1 first before registering L1/2. Surgeon also requested to take the images for the upper portion of the registration ahead of time and save in case we needed them later, which RTC explained was not possible regardless. L5/S1 screws placed with accuracy bilaterally and checked with neuromonitering successfully. Surgeon requested to bring up the L1/2 images to try for registration, RTC explained why it was not possible & proceeded to accept the images to walk the surgeon through the process of why it will not work. (After accepting saved images and moving to segmentation) When placing reference marker in saved OB Image, the cross-view line was vertical to the vertebral bodies instead of horizontal as normal. RTC explained the cross-view problems and why registration could not be achieved. Surgeon agreed to take additional AP and OB image to capture L1/2. After achieving RBCs, RTC placed reference marker in OB image to achieve registration and again saw the cross-view lines running vertical to the spine as opposed to horizontal across a vertebral body. Surgeon opted to freehand the top screws due to lengthy duration of the procedure. UNK MAZOR X X-EYE CAMERA | Significant Crossview error on second registration. First registration accurate. Horizontal line on AP image (when placing red reference point) appeared to be vertical as opposed to horizontal across the desired body during flouro segmentation./ Surgeon, Fellow, & RTC planned case prior to procedure and approved implant placement and trajectory plan. Single Arm Shanz pin agreed to be placed in left PSIS for the MIS procedure. Cage was placed in the lateral position at L3/4, Surgeon opted to add screws to L1/2 and L5/S1 for stability. System mounted and placed onto bed without any issue, arm draped, pin placed in Left PSIS. Robot mounted to patient without problem. While draping, scrub tech wrapped red rubber bands for X Eye cameras around the neck of the camera, causing an inability for the camera to move properly. 3Define executed properly as planned. X Eyes unable to consistently pick up the reference marker and stylus, RTC opted to move forward to generic AP & OB location for 3D Marker acquisition. Significant manual adjustments needed in order to place 3D marker in proper location. L1-S1 was unable to be achieved in the imaging due to larger patient, surgeon requested to register twice to achieve L5/S1 first before registering L1/2. Surgeon also requested to take the images for the upper portion of the registration ahead of time and save in case we needed them later, which RTC explained was not possible regardless. L5/S1 screws placed with accuracy bilaterally and checked with neuromonitering successfully. Surgeon requested to bring up the L1/2 images to try for registration, RTC explained why it was not possible & proceeded to accept the images to walk the surgeon through the process of why it will not work. (After accepting saved images and moving to segmentation) When placing reference marker in saved OB Image, the cross-view line was vertical to the vertebral bodies instead of horizontal as normal. RTC explained the cross-view problems and why registration could not be achieved. Surgeon agreed to take additional AP and OB image to capture L1/2. After achieving RBCs, RTC placed reference marker in OB image to achieve registration and again saw the cross-view lines running vertical to the spine as opposed to horizontal across a vertebral body. Surgeon opted to freehand the top screws due to lengthy duration of the procedure. | x registration issue significant crossview error second registration registration accurate horizontal line ap image appeared vertical opposed horizontal desired body flouro segmentation surgeon fellow rtc planned case prior procedure approved implant placement trajectory plan single arm shanz pin agreed placed left psis mis procedure cage placed lateral position surgeon opted add screws stability mounted placed bed issue arm draped pin placed left psis robot mounted problem draping scrub tech wrapped red rubber bands x eye cameras neck camera causing inability camera properly executed properly planned x eyes unable consistently pick reference marker stylus rtc opted forward generic ap ob location marker acquisition significant manual adjustments needed order place marker proper location unable achieved imaging larger surgeon requested register twice achieve registering surgeon requested images upper portion registration ahead time save case needed later rtc explained possible regardless screws placed accuracy bilaterally checked neuromonitering successfully surgeon requested bring images try registration rtc explained possible proceeded accept images walk surgeon process work placing reference marker saved ob image crossview line vertical vertebral bodies instead horizontal normal rtc explained crossview problems registration achieved surgeon agreed additional ap ob image capture achieving rbcs rtc placed reference marker ob image achieve registration saw crossview lines running vertical spine opposed horizontal vertebral body surgeon opted freehand screws lengthy duration procedure unk mazor x xeye camera | NAV4110 | NAV3233 | NAV4112 | NAV3244 | NAV3237 | NAV4110NAV3233NAV4112NAV3244NAV3237 | True
8557 | 704438575 | XSE ¦ Difficulties with Set Up | 2021-06-08T23:05:21.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE ¦ Difficulties with Set Up 2021-Jun-08 Salesforce C00228051 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that before the case while completing the arm set up the arm was sent to the first position. The arm stopped at the first position but was beeping as if it was still moving. The arm was unable to move from this location. The manufacturer representative rebooted the system and successfully completed an arm set up. At this point they attempted to export video to the boom monitor but where unable to. Lastly, the surgeon tablet displayed images upside down. No patient was present. 2021-06-21 IUD x4 (Other): No new information. 2021-Jul-21 IUD (other): No new information. UNK MAZOR X WORKSTATION | 2021-Jun-08 Salesforce C00228051 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that before the case while completing the arm set up the arm was sent to the first position. The arm stopped at the first position but was beeping as if it was still moving. The arm was unable to move from this location. The manufacturer representative rebooted the system and successfully completed an arm set up. At this point they attempted to export video to the boom monitor but where unable to. Lastly, the surgeon tablet displayed images upside down. No patient was present. 2021-06-21 IUD x4 (Other): No new information. 2021-Jul-21 IUD (other): No new information. | xse difficulties set salesforce outside procedure case completing arm set arm sent position arm stopped position beeping moving arm unable location rebooted successfully completed arm set point attempted export video boom monitor unable lastly surgeon tablet displayed images upside present iud new iud new unk mazor x workstation | NAV2058 | NAV3253 | NAV4095 | NAV3264 | NAV3299 | NAV2058NAV3253NAV4095NAV3264NAV3299 | True
2548 | 703241060 | X - SOFTWARE ISSUE | 2019-07-03T13:50:53.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - SOFTWARE ISSUE Robotic arm came way too close to the patient during the fluoro registration OBL trajectory. A 3Define scan was performed and Draw Spine was also used. I’m unsure where the inaccuracy problem came from. MAZOR X SOFTWARE KIT VER-3.0.2 | Robotic arm came way too close to the patient during the fluoro registration OBL trajectory. A 3Define scan was performed and Draw Spine was also used. I’m unsure where the inaccuracy problem came from. | x software issue robotic arm came way close fluoro registration obl trajectory scan performed draw spine i’m unsure inaccuracy problem came mazor x software kit | NAV3246 | NAV4091 | NAV3247 | NAV2054 | NAV3242 | NAV3246NAV4091NAV3247NAV2054NAV3242 | True
1200 | 703193068 | X - Software Issue | 2019-06-06T15:03:11.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X - Software Issue After drawspine arm sent successfully to AP. Then arm lowered by a few clicks. Once AP acquired OBL couldn’t be reached, even though initial solution was achieved. Default WV resolved this issue (believed to be bad 3Dfine scan). Why doesn’t the SW know how to reverse out of a position it moved into? See image 2 MAZOR X SOFTWARE KIT VER-4.0.0 | After drawspine arm sent successfully to AP. Then arm lowered by a few clicks. Once AP acquired OBL couldn’t be reached, even though initial solution was achieved. Default WV resolved this issue (believed to be bad 3Dfine scan). Why doesn’t the SW know how to reverse out of a position it moved into? See image 2 | x software issue drawspine arm sent successfully ap arm lowered clicks ap acquired obl couldn’t reached initial solution achieved default wv resolved issue doesn’t sw know reverse position moved image mazor x software kit | NAV3246 | NAV3242 | NAV3248 | NAV3247 | NAV4091 | NAV3246NAV3242NAV3248NAV3247NAV4091 | True
133 | 703005936 | X - X-Eye Issues | 2019-02-13T20:58:40.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 30 | SEN0030-01 | SENSOR | X - X-Eye Issues 2019-02-07 SCL1900346, Intake Form 2019-02-13 15:05 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that X-eye showed green connection during setup, but was not recognizing the stylus during draw spine and there was no live feed from the utilities menu. There was no patient harm and the surgery was not delayed more than an hour. The case was completed. 2019-02-11 e1 (Rep, HCP): Additional information received from a manufacturer representative reported that a L4 to S1 transforaminal lumbar interbody fusion (TLIF) and posterior lateral spinal fusion (PSF) procedure was being done. During setup, the shoulder calibration was checked and then the X-Eye camera was checked, but there was no output. An advanced accuracy check was completed and the X-Eye was checked again and it worked. The rest of the system was setup successfully. After mounting the surgical system to the bed, the X-Eye camera was checked again and the screen remained black with no input. The cables were checked including unplugging and replugging in the lemo and X-Eye cables. At this point, the software started crashing and was not responding. When the software did respond a CANBUS and 4059 (Arm Shift Detected in Joint 2) error messages were displayed. A soft restart and hard restart were done. The manufacturer representative attempted to set up the surgical arm again with red lights for the X-Eye, but the 4059 error was displayed again. Technical services had the manufacturer representative manually enable arm movement in the software and complete an advanced accuracy check again. The surgical arm setup was able to be completed with there being a red light for the X-Eye. The surgeon was able to successfully place all of the screws in the correct position. Accuracy was verified visually and with lateral and AP images. SENSOR | 2019-02-07 SCL1900346, Intake Form 2019-02-13 15:05 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that X-eye showed green connection during setup, but was not recognizing the stylus during draw spine and there was no live feed from the utilities menu. There was no patient harm and the surgery was not delayed more than an hour. The case was completed. 2019-02-11 e1 (Rep, HCP): Additional information received from a manufacturer representative reported that a L4 to S1 transforaminal lumbar interbody fusion (TLIF) and posterior lateral spinal fusion (PSF) procedure was being done. During setup, the shoulder calibration was checked and then the X-Eye camera was checked, but there was no output. An advanced accuracy check was completed and the X-Eye was checked again and it worked. The rest of the system was setup successfully. After mounting the surgical system to the bed, the X-Eye camera was checked again and the screen remained black with no input. The cables were checked including unplugging and replugging in the lemo and X-Eye cables. At this point, the software started crashing and was not responding. When the software did respond a CANBUS and 4059 (Arm Shift Detected in Joint 2) error messages were displayed. A soft restart and hard restart were done. The manufacturer representative attempted to set up the surgical arm again with red lights for the X-Eye, but the 4059 error was displayed again. Technical services had the manufacturer representative manually enable arm movement in the software and complete an advanced accuracy check again. The surgical arm setup was able to be completed with there being a red light for the X-Eye. The surgeon was able to successfully place all of the screws in the correct position. Accuracy was verified visually and with lateral and AP images. | x xeye issues spinal procedure xeye showed green connection setup recognizing stylus draw spine live feed utilities menu surgery delayed hour case completed additional transforaminal lumbar interbody fusion posterior lateral spinal fusion procedure setup shoulder calibration checked xeye camera checked output advanced accuracy check completed xeye checked worked rest setup successfully mounting surgical bed xeye camera checked screen remained black input cables checked including unplugging replugging lemo xeye cables point software started crashing responding software did respond canbus error messages displayed soft restart hard restart attempted set surgical arm red lights xeye error displayed technical services manually enable arm movement software complete advanced accuracy check surgical arm setup able completed red light xeye surgeon able successfully place screws correct position accuracy verified visually lateral ap images sensor | NAV4110 | NAV4109 | NAV3254 | NAV3248 | NAV4095 | NAV4110NAV4109NAV3254NAV3248NAV4095 | False
8915 | 704512195 | XSE - Damaged Parts | 2021-07-20T20:10:41.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 70 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Damaged Parts 2021-Jul-20 Salesforce C00230460 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system had multiple damaged parts. There was no patient involvement. 2021-07-22 IUD x5 (Other): No new information. 2021-08-12 email 2021-08-12 07:45 (Rep): Additional information received from a manufacturer representative reported that the repair for the system was cancelled due to age of the system and cost. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2021-Jul-20 Salesforce C00230460 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system had multiple damaged parts. There was no patient involvement. 2021-07-22 IUD x5 (Other): No new information. 2021-08-12 email 2021-08-12 07:45 (Rep): Additional information received from a manufacturer representative reported that the repair for the system was cancelled due to age of the system and cost. | xse damaged parts salesforce outside procedure multiple damaged parts iud new email additional repair cancelled age cost kit svc cart elec panel assy | NAV2058 | NAV3264 | NAV2054 | NAV2053 | NAV3249 | NAV2058NAV3264NAV2054NAV2053NAV3249 | True
3958 | 703475413 | XSE - Workstation Cables Damaged | 2019-11-19T16:11:13.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 40 | HRS0350-01 | NAVIGATION CONNECTION CABLE | XSE - Workstation Cables Damaged 2019-11-19 00678633(Rep): Medtronic received information regarding a guidance system being used during training. It was reported that the cables inside the workstation were damaged and needed to be replaced. No patient involvement. 2019-12-05 IUD x4 (Other): No new information. NAVIGATION CONNECTION CABLE | 2019-11-19 00678633(Rep): Medtronic received information regarding a guidance system being used during training. It was reported that the cables inside the workstation were damaged and needed to be replaced. No patient involvement. 2019-12-05 IUD x4 (Other): No new information. | xse workstation cables damaged training cables inside workstation damaged needed replaced iud new navigation connection cable | NAV3256 | NAV3295 | NAV2058 | NAV2047 | NAV3264 | NAV3256NAV3295NAV2058NAV2047NAV3264 | True
8817 | 704492015 | XSE ¦ Shoulder Not Locking | 2021-07-09T00:04:04.000000+0000 | MazorX Stealth Edition | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | XSE ¦ Shoulder Not Locking 2021-Jul-08 Salesforce C00229834 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site was experiencing issues with the shoulder of the system not locking. No patient was present. 2021-07-15 IUD x2 (Other): No new information. UNK MAZOR X SHOULDER | 2021-Jul-08 Salesforce C00229834 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site was experiencing issues with the shoulder of the system not locking. No patient was present. 2021-07-15 IUD x2 (Other): No new information. | xse shoulder locking salesforce outside procedure site experiencing issues shoulder locking present iud new unk mazor x shoulder | NAV4107 | NAV3252 | NAV4108 | NAV2058 | NAV3264 | NAV4107NAV3252NAV4108NAV2058NAV3264 | True
6270 | 703902254 | XSE - CANBUS Error | 2020-08-13T15:31:00.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 40 | KIT1290-01 | SVC KIT1290-01 SWITCH EMERG STOP | XSE - CANBUS Error 2020-08-11 00694530, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error was displayed and the surgeon decided to abort the use of the guidance system. The surgeon screen was black and not working properly during the procedure. During the S1 screw on the left side, the surgeon was using bovie with the cannula just lifted up a little. The manufacturer representative noted that the bovie appeared to arch to the tool and and cause the CANBUS error. There was no patient harm and the procedure was delayed less than an hour. 2020-08-18 IUD x3 (Other): No new information. 2020-08-25 IUD (Other): No new information. 2020-08-28 IUD x3 (Other): No new information. 2020-09-09 email 2020-09-09 00:00, IUD x4 (Other): No new information. 2020-09-30 IUD x2 (Other): No new information. 2020-11-06 email 2020-11-06 13:15 (Rep): No new information. 2020-12-03 email 2020-12-03 11:00, email 2020-12-14 20:45 (Rep): Additional information received from a manufacturer representative reported that the use of the guidance system was aborted due to the CANBUS error and surgical arm shift with two screws left. The procedure was successfully completed freehand. All screws placed with the guidance system were accurate. SVC KIT1290-01 SWITCH EMERG STOP | 2020-08-11 00694530, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error was displayed and the surgeon decided to abort the use of the guidance system. The surgeon screen was black and not working properly during the procedure. During the S1 screw on the left side, the surgeon was using bovie with the cannula just lifted up a little. The manufacturer representative noted that the bovie appeared to arch to the tool and and cause the CANBUS error. There was no patient harm and the procedure was delayed less than an hour. 2020-08-18 IUD x3 (Other): No new information. 2020-08-25 IUD (Other): No new information. 2020-08-28 IUD x3 (Other): No new information. 2020-09-09 email 2020-09-09 00:00, IUD x4 (Other): No new information. 2020-09-30 IUD x2 (Other): No new information. 2020-11-06 email 2020-11-06 13:15 (Rep): No new information. 2020-12-03 email 2020-12-03 11:00, email 2020-12-14 20:45 (Rep): Additional information received from a manufacturer representative reported that the use of the guidance system was aborted due to the CANBUS error and surgical arm shift with two screws left. The procedure was successfully completed freehand. All screws placed with the guidance system were accurate. | xse canbus error iud spinal procedure canbus error displayed surgeon decided abort use surgeon screen black working properly procedure screw left surgeon using bovie cannula just lifted little noted bovie appeared arch tool cause canbus error procedure delayed hour iud new iud new iud new email iud new iud new email new email email additional use aborted canbus error surgical arm shift screws left procedure successfully completed freehand screws placed accurate svc switch emerg stop | NAV2047 | NAV4098 | NAV2054 | NAV3253 | NAV1012 | NAV2047NAV4098NAV2054NAV3253NAV1012 | True
5206 | 703619117 | XSE - Damage During Transport/WS Issue | 2020-02-17T16:06:16.000000+0000 | MazorX Stealth Edition | NAV2055 | PRODUCTS DAMAGED DURING SHIPPING | 60 | MEC4472-01S | KIT SVC MEC4472-01S J6 COVER LEFT | XSE - Damage During Transport/WS Issue 2020-02-12 00684214, Intake Form (REP, For): Medtronic received information regarding a guidance system outside of procedure. It was reported that the although the front side wheel was broken, after opening the wooden box they found that the remaining one front side wheel was bent. The monitor screen also got damaged. There was no patient involvement. 2020-04-17 email 2020-04-17 00:45 (Rep, For): Additional information received from a manufacturer representative reported that the issue was not resolved yet. 2021-03-10 IUD (Other, For): No new information. 2022-05-17 IUD (Other, For): No new information. KIT SVC MEC4472-01S J6 COVER LEFT | 2020-02-12 00684214, Intake Form (REP, For): Medtronic received information regarding a guidance system outside of procedure. It was reported that the although the front side wheel was broken, after opening the wooden box they found that the remaining one front side wheel was bent. The monitor screen also got damaged. There was no patient involvement. 2020-04-17 email 2020-04-17 00:45 (Rep, For): Additional information received from a manufacturer representative reported that the issue was not resolved yet. 2021-03-10 IUD (Other, For): No new information. 2022-05-17 IUD (Other, For): No new information. | xse damage transportws issue outside procedure wheel broken opening wooden box remaining wheel bent monitor screen got damaged email additional issue resolved iud new iud new kit svc cover left | NAV2055 | NAV2058 | NAV2054 | NAV3261 | NAV3299 | NAV2055NAV2058NAV2054NAV3261NAV3299 | True
6714 | 704012871 | XSE - OOBF/CANBUS Error | 2020-10-14T20:14:44.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 50 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - OOBF/CANBUS Error 2020-10-14 00698768 (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a Sacroiliac and Thoracolumbar procedure. It was reported ALL nodes had CANBUS error. Another error was received stated gains lost on driver one and stated requires maintenance error 8798 failed to reach destination and they could not clear the error and it would not move at all. They performed a soft reboot and that did not resolve the issue. They then went on to just use navigation and not the robot. There was no patient harm or additional complications reported or anticipated as a result of the event. There was a delay of less than 1 hour. 2020-10-27 00699658, 00699655, 00699656 (Rep): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that the wrong X-eye file was received with the surgical system type 2 at install. The manufacturer representative was able to get the correct calibration file for the camera to resolve the issue. The first encode for the new positioner was bad so an old one had to be used to resolve the issue. The 3define camera of the new surgical arm failed at install. There was no connection with the camera and the firmware would not pop up. The 3Define camera was changed in order to get the correct firmware version. Changing the camera resolved the issue. There was no patient involvement. 2020-10-27 00699658, 00699655, 00699656 (Rep): Additional information received from a manufacturer representative reported that the wrong X-eye file was received with the surgical system type 2 at install. The manufacturer representative was able to get the correct calibration file for the camera to resolve the issue. The first encode for the new positioner was bad so an old one had to be used to resolve the issue. The 3define camera of the new surgical arm failed at install. There was no connection with the camera and the firmware would not pop up. The 3Define camera was changed in order to get the correct firmware version. Changing the camera resolved the issue. 2020-11-10 email 11/10 (rep): Additional information received indicated that the surgical system positioner that was shipped showed to have a front shoulder encoder to be faulty. The rep replaced the encoder with the encoder from the system they had replaced and the old encoder ended up working. Testing passed the encoder. For the surgical arm, the 3define camera showed to be faulty. They switched the 3define camera from the new arm with the 3define camera from the arm that was replaced. The old 3define camera worked on the system after the change and tested with passing results. 2020-11-17 IUD x8 (Other): No new information. 2021-02-24 IUD x2 (Other): No new information. 2021-06-01 IUD x2 (Other): No new information. 2021-Jul-13 IUD (other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2020-10-14 00698768 (Rep, HCP): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a Sacroiliac and Thoracolumbar procedure. It was reported ALL nodes had CANBUS error. Another error was received stated gains lost on driver one and stated requires maintenance error 8798 failed to reach destination and they could not clear the error and it would not move at all. They performed a soft reboot and that did not resolve the issue. They then went on to just use navigation and not the robot. There was no patient harm or additional complications reported or anticipated as a result of the event. There was a delay of less than 1 hour. 2020-10-27 00699658, 00699655, 00699656 (Rep): Medtronic received information regarding a guidance system being used outside of a pr ocedure. It was reported that the wrong X-eye file was received with the surgical system type 2 at install. The manufacturer representative was able to get the correct calibration file for the camera to resolve the issue. The first encode for the new positioner was bad so an old one had to be used to resolve the issue. The 3define camera of the new surgical arm failed at install. There was no connection with the camera and the firmware would not pop up. The 3Define camera was changed in order to get the correct firmware version. Changing the camera resolved the issue. There was no patient involvement. 2020-10-27 00699658, 00699655, 00699656 (Rep): Additional information received from a manufacturer representative reported that the wrong X-eye file was received with the surgical system type 2 at install. The manufacturer representative was able to get the correct calibration file for the camera to resolve the issue. The first encode for the new positioner was bad so an old one had to be used to resolve the issue. The 3define camera of the new surgical arm failed at install. There was no connection with the camera and the firmware would not pop up. The 3Define camera was changed in order to get the correct firmware version. Changing the camera resolved the issue. 2020-11-10 email 11/10 (rep): Additional information received indicated that the surgical system positioner that was shipped showed to have a front shoulder encoder to be faulty. The rep replaced the encoder with the encoder from the system they had replaced and the old encoder ended up working. Testing passed the encoder. For the surgical arm, the 3define camera showed to be faulty. They switched the 3define camera from the new arm with the 3define camera from the arm that was replaced. The old 3define camera worked on the system after the change and tested with passing results. 2020-11-17 IUD x8 (Other): No new information. 2021-02-24 IUD x2 (Other): No new information. 2021-06-01 IUD x2 (Other): No new information. 2021-Jul-13 IUD (other): No new information. | xse oobfcanbus error sacroiliac thoracolumbar procedure nodes canbus error error stated gains lost driver stated requires maintenance error failed reach destination clear error performed soft reboot did resolve issue went just use navigation robot additional complications anticipated delay hour outside pr ocedure wrong xeye file surgical type install able correct calibration file camera resolve issue encode new positioner bad old resolve issue camera new surgical arm failed install connection camera firmware pop camera changed order correct firmware version changing camera resolved issue additional wrong xeye file surgical type install able correct calibration file camera resolve issue encode new positioner bad old resolve issue camera new surgical arm failed install connection camera firmware pop camera changed order correct firmware version changing camera resolved issue email additional indicated surgical positioner shipped showed shoulder encoder faulty replaced encoder encoder replaced old encoder ended working testing passed encoder surgical arm camera showed faulty switched camera new arm camera arm replaced old camera worked change tested passing results iud new iud new iud new iud new kit svc surgical syst type | NAV2054 | NAV2047 | NAV3244 | NAV4106 | NAV4095 | NAV2054NAV2047NAV3244NAV4106NAV4095 | True
3068 | 703264128 | X - Surgical System Unmounting Issue | 2019-07-17T19:06:49.000000+0000 | MazorX | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X - Surgical System Unmounting Issue Spoke to the CS on the phone. One cleat locked on the bed and the other was not when mounting the system. CS was trying to unmount the system to properly mountthe system. UNK MAZOR X POSITIONER | Spoke to the CS on the phone. One cleat locked on the bed and the other was not when mounting the system. CS was trying to unmount the system to properly mountthe system. | x surgical unmounting issue spoke cs phone cleat locked bed mounting cs trying unmount properly mountthe unk mazor x positioner | NAV3250 | NAV3244 | NAV3264 | NAV2054 | NAV3251 | NAV3250NAV3244NAV3264NAV2054NAV3251 | True
3289 | 703306062 | XSE - Error Message/Software Freeze | 2019-08-12T14:44:18.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Error Message/Software Freeze 2019-08-05 SCL1902329, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the surgeon was performing an L4-L5 percutaneous posterior instrumentation following a lateral interbody placement. Shoulder calibration and accuracy was checked prior to the procedure. The surgical system was mounted to the patient using a Schanz pin in the PSSIS with a Schanz bond and bone mounting bridge. After completing the 3Define scan, snapshot registration and draw spine, the system could not calculate the AP position for the 3D marker. When the AP button was clicked the system would begin to calculate, but would not pass 25%. After a bit of time the error message "Movement calculation is taking longer than usual. Do you with to abort?" was displayed. A generic AP position was attempted, but the same error message was displayed. The system was used to calculate the oblique so the oblique image was taken first. The AP position was then reattempted, but the same error message as before was displayed. The surgical arm was manually moved into the AP position and the AP image was taken. Registration was completed using the basic algorithm with good segmentation and green values for all levels. The surgeon then placed the left L4 screw without issue. The surgical arm was sent to the left L5 trajectory next. After drilling, blood was noticed coming out of the drill guide. All instruments were removed and the arm was sent to the clear up position. Fluoro was brought in while the surgeon worked to contain the bleeding. The surgeon reported the bleeding came from a facet artery and was due to the patient's anatomy and not the system. The surgeon then placed the right L4 screw and fluoro images were taken to confirm placement. While the x-ray technician brought in the fluoro, they bumped the surgical arm. No trajectory or shoulder shift errors were displayed. The confirmation images showed the right L4 screw was slightly superior to plan. The surgical arm was then sent to right L45 and fluoro was used to check the position of the drill guide. The trajectory was angled very superior to the vertebrae and would cause the screw to end up in the disc space. Reregistration was suggested and the surgeon was ok with that plan. Snapshot registration and draw spine were redone, but the system still would not calculate the AP position and froze twice while displaying a "Mazor X is not responding" error message. The system first froze when the draw spine AP was being calculated and once when the generic AP was being calculated. Both times, the system was left alone to continue processing the and system came back. Draw spine was then re-labeled as an adjacent body and the AP was calculated. Registration x-rays were taken and registration was successful on the sixth try using L3 vertebral body to register off of. Yellow values were given for L4 and green values were given for L5. When comparing fluoro to CT matching, the surgeon noticed that the superior end plate of L5 was imposed over the superior aspect of the lateral cage. The surgeon decide to proceed and approve the registration. The surgical arm was sent to the right L5 trajectory and fluoro shots of the drill guide were taken to confirm placement. The drill guide was pointed to the L4-L5 disc space so the surgeon decided to abort the guidance system and place the right side screws using a Jamshidi and K-wires. There was no patient harm, but the procedure was delayed more than an hour. 2019-08-20 email 2019-08-20 11:30 (Other): No new information. 2019-09-27 email 2019-09-27 09:00 (Rep): Additional information received from a manufacturer representative reported the surgeon was able to stop the bleeding by removing the drill guide and cannulas, moving the surgical arm, widening the incision, placing a retractor to expose the facet joint and then use bovie. The representative noted the trajectories were deviated between 3.5 and 10 mm. Due to the deviations, the procedure was delayed approximately one hour and 15 minutes. The delay was due to a combination of software issues extending the initial registration and needing to bring the C-arm back in for a second registration. There was no patient harm and all of the screws were able to be successfully placed suing the Jamshidi and K-wires. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-08-05 SCL1902329, Intake Form, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that the surgeon was performing an L4-L5 percutaneous posterior instrumentation following a lateral interbody placement. Shoulder calibration and accuracy was checked prior to the procedure. The surgical system was mounted to the patient using a Schanz pin in the PSSIS with a Schanz bond and bone mounting bridge. After completing the 3Define scan, snapshot registration and draw spine, the system could not calculate the AP position for the 3D marker. When the AP button was clicked the system would begin to calculate, but would not pass 25%. After a bit of time the error message "Movement calculation is taking longer than usual. Do you with to abort?" was displayed. A generic AP position was attempted, but the same error message was displayed. The system was used to calculate the oblique so the oblique image was taken first. The AP position was then reattempted, but the same error message as before was displayed. The surgical arm was manually moved into the AP position and the AP image was taken. Registration was completed using the basic algorithm with good segmentation and green values for all levels. The surgeon then placed the left L4 screw without issue. The surgical arm was sent to the left L5 trajectory next. After drilling, blood was noticed coming out of the drill guide. All instruments were removed and the arm was sent to the clear up position. Fluoro was brought in while the surgeon worked to contain the bleeding. The surgeon reported the bleeding came from a facet artery and was due to the patient's anatomy and not the system. The surgeon then placed the right L4 screw and fluoro images were taken to confirm placement. While the x-ray technician brought in the fluoro, they bumped the surgical arm. No trajectory or shoulder shift errors were displayed. The confirmation images showed the right L4 screw was slightly superior to plan. The surgical arm was then sent to right L45 and fluoro was used to check the position of the drill guide. The trajectory was angled very superior to the vertebrae and would cause the screw to end up in the disc space. Reregistration was suggested and the surgeon was ok with that plan. Snapshot registration and draw spine were redone, but the system still would not calculate the AP position and froze twice while displaying a "Mazor X is not responding" error message. The system first froze when the draw spine AP was being calculated and once when the generic AP was being calculated. Both times, the system was left alone to continue processing the and system came back. Draw spine was then re-labeled as an adjacent body and the AP was calculated. Registration x-rays were taken and registration was successful on the sixth try using L3 vertebral body to register off of. Yellow values were given for L4 and green values were given for L5. When comparing fluoro to CT matching, the surgeon noticed that the superior end plate of L5 was imposed over the superior aspect of the lateral cage. The surgeon decide to proceed and approve the registration. The surgical arm was sent to the right L5 trajectory and fluoro shots of the drill guide were taken to confirm placement. The drill guide was pointed to the L4-L5 disc space so the surgeon decided to abort the guidance system and place the right side screws using a Jamshidi and K-wires. There was no patient harm, but the procedure was delayed more than an hour. 2019-08-20 email 2019-08-20 11:30 (Other): No new information. 2019-09-27 email 2019-09-27 09:00 (Rep): Additional information received from a manufacturer representative reported the surgeon was able to stop the bleeding by removing the drill guide and cannulas, moving the surgical arm, widening the incision, placing a retractor to expose the facet joint and then use bovie. The representative noted the trajectories were deviated between 3.5 and 10 mm. Due to the deviations, the procedure was delayed approximately one hour and 15 minutes. The delay was due to a combination of software issues extending the initial registration and needing to bring the C-arm back in for a second registration. There was no patient harm and all of the screws were able to be successfully placed suing the Jamshidi and K-wires. | xse error messagesoftware freeze ccr spinal procedure treat degenerative discs fixation surgeon performing percutaneous posterior instrumentation following lateral interbody placement shoulder calibration accuracy checked prior procedure surgical mounted using schanz pin pssis schanz bond bone mounting bridge completing scan snapshot registration draw spine calculate ap position marker ap button clicked begin calculate pass bit time error message movement calculation taking longer usual abort displayed generic ap position attempted error message displayed calculate oblique oblique image taken ap position reattempted error message displayed surgical arm manually moved ap position ap image taken registration completed using basic algorithm good segmentation green values levels surgeon placed left screw issue surgical arm sent left trajectory drilling blood noticed coming drill guide instruments removed arm sent clear position fluoro brought surgeon worked contain bleeding surgeon bleeding came facet artery patients anatomy surgeon placed right screw fluoro images taken confirm placement xray technician brought fluoro bumped surgical arm trajectory shoulder shift errors displayed confirmation images showed right screw slightly superior plan surgical arm sent right fluoro check position drill guide trajectory angled superior vertebrae cause screw end disc space reregistration suggested surgeon ok plan snapshot registration draw spine redone calculate ap position froze twice displaying mazor x responding error message froze draw spine ap calculated generic ap calculated times left continue processing came draw spine relabeled adjacent body ap calculated registration xrays taken registration successful sixth try using vertebral body register yellow values given green values given comparing fluoro ct matching surgeon noticed superior end plate imposed superior aspect lateral cage surgeon decide proceed approve registration surgical arm sent right trajectory fluoro shots drill guide taken confirm placement drill guide pointed disc space surgeon decided abort place right screws using jamshidi kwires procedure delayed hour email new email additional surgeon able stop bleeding removing drill guide cannulas moving surgical arm widening incision placing retractor expose facet joint use bovie noted trajectories deviated mm deviations procedure delayed approximately hour minutes delay combination software issues extending initial registration needing bring carm second registration screws able successfully placed suing jamshidi kwires mazor x software installation kit | NAV3242 | NAV3248 | NAV5000 | NAV3246 | NAV2056 | NAV3242NAV3248NAV5000NAV3246NAV2056 | True
1345 | 703197348 | X - communication issue | 2019-06-10T12:49:07.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - communication issue Controller Error during system start up, clears after 5 sec. - Known software issue MAZOR X SOFTWARE KIT VER-3.0.2 | Controller Error during system start up, clears after 5 sec. - Known software issue | x communication issue controller error start clears sec known software issue mazor x software kit | NAV3246 | NAV2047 | NAV3248 | NAV3242 | NAV4091 | NAV3246NAV2047NAV3248NAV3242NAV4091 | True
4799 | 703578079 | X-No registration/Deviation/X-eye issues | 2020-01-23T16:57:48.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X-No registration/Deviation/X-eye issues 2018-05-17 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for L2-L3 DLIF with posterior pedicle screws. The case plan was reviewed with the surgeon prior to operation. Two plans were presented, and the surgeon chose to use the plan with screws planned on facets. Advanced accuracy test was completed. Shoulder calibration & camera function was checked. The guidance system was attached to the bed with no issues. The guidance system was attached to the patient with a single schanz pin into left PSIS. X-eyes were having trouble seeing the reference plate initially. The mount was unlocked and relocked and X-eyes were able to see. AP and oblique images were taken, and basic registration was successful for L1-L3 with a failure to register at L4. The arm was sent to mark all trajectories. The arm was resent to instrument, and X-ray images taken to verify placement. The trajectory looked too high. The arm was sent to remaining trajectories which looked equally high, inaccurate, and not in accordance with the plan. Surgeon chose not to re-register the patient but to free hand the screws instead. The system was removed from the bed with no issues. No patient harm was reported. UNK MAZOR X X-EYE CAMERA | 2018-05-17 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for L2-L3 DLIF with posterior pedicle screws. The case plan was reviewed with the surgeon prior to operation. Two plans were presented, and the surgeon chose to use the plan with screws planned on facets. Advanced accuracy test was completed. Shoulder calibration & camera function was checked. The guidance system was attached to the bed with no issues. The guidance system was attached to the patient with a single schanz pin into left PSIS. X-eyes were having trouble seeing the reference plate initially. The mount was unlocked and relocked and X-eyes were able to see. AP and oblique images were taken, and basic registration was successful for L1-L3 with a failure to register at L4. The arm was sent to mark all trajectories. The arm was resent to instrument, and X-ray images taken to verify placement. The trajectory looked too high. The arm was sent to remaining trajectories which looked equally high, inaccurate, and not in accordance with the plan. Surgeon chose not to re-register the patient but to free hand the screws instead. The system was removed from the bed with no issues. No patient harm was reported. | xno registrationdeviationxeye issues ccr spinal procedure case dlif posterior pedicle screws case plan reviewed surgeon prior operation plans presented surgeon chose use plan screws planned facets advanced accuracy test completed shoulder calibration camera function checked attached bed issues attached single schanz pin left psis xeyes having trouble seeing reference plate initially mount unlocked relocked xeyes able ap oblique images taken basic registration successful failure register arm sent mark trajectories arm resent instrument xray images taken verify placement trajectory looked high arm sent remaining trajectories looked equally high inaccurate accordance plan surgeon chose reregister free hand screws instead removed bed issues unk mazor x xeye camera | NAV4110 | NAV3244 | NAV2058 | NAV4091 | NAV4112 | NAV4110NAV3244NAV2058NAV4091NAV4112 | True
4113 | 703501112 | XSE - Nav and Planning Screw Issue | 2019-12-04T23:20:16.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 30 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Nav and Planning Screw Issue 2019-12-04 00679581, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when in operation mode of the software, the planned screws would not show along with the navigation tools and the screws would not display. Before entering the operation mode, the representative verified the tools. After entering operation mode, verification was lost. There was no patient harm and the procedure was delayed less than one hour. The case was completed. 2019-12-26 IUD x2 (Other): No new information. 2020-01-07 IUD x2 (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-12-04 00679581, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when in operation mode of the software, the planned screws would not show along with the navigation tools and the screws would not display. Before entering the operation mode, the representative verified the tools. After entering operation mode, verification was lost. There was no patient harm and the procedure was delayed less than one hour. The case was completed. 2019-12-26 IUD x2 (Other): No new information. 2020-01-07 IUD x2 (Other): No new information. | xse nav planning screw issue spinal procedure operation mode software planned screws navigation tools screws display entering operation mode verified tools entering operation mode verification lost procedure delayed hour case completed iud new iud new mazor x software installation kit | NAV3246 | NAV3247 | NAV5000 | NAV2056 | NAV3237 | NAV3246NAV3247NAV5000NAV2056NAV3237 | True
9181 | 704568196 | XSE - Intermittent C-Arm Import | 2021-08-20T09:44:16.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - Intermittent C-Arm Import 2021-Aug-06 Salesforce C00233235, IUD (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the C-arm was intermittently sending reference, AP and oblique images to the guidance system. The DICOM settings of the C-arm and guidance system were checked. The manufacturer representative found that changing the port or IP would resolve the issue, but the problem would come back during the next use. Two C-arms and an O2 were recently defined and the representative noted that a port can only be used once. During troubleshooting, the representative found that both DICOM C-arms were defined as two different entries when they should be defined as one. The configuration was updated and images were able to be transferred to the guidance system successfully. The system was working as intended. There was no patient involvement. ASM0215-06S MAZOR X 5820 PC STATION | 2021-Aug-06 Salesforce C00233235, IUD (Rep, For): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the C-arm was intermittently sending reference, AP and oblique images to the guidance system. The DICOM settings of the C-arm and guidance system were checked. The manufacturer representative found that changing the port or IP would resolve the issue, but the problem would come back during the next use. Two C-arms and an O2 were recently defined and the representative noted that a port can only be used once. During troubleshooting, the representative found that both DICOM C-arms were defined as two different entries when they should be defined as one. The configuration was updated and images were able to be transferred to the guidance system successfully. The system was working as intended. There was no patient involvement. | xse intermittent carm import salesforce iud outside procedure carm intermittently sending reference ap oblique images dicom settings carm checked changing port ip resolve issue problem come use carms recently defined noted port troubleshooting dicom carms defined different entries defined configuration updated images able transferred successfully working intended mazor x pc station | NAV3237 | NAV3262 | NAV2058 | NAV5000 | NAV3246 | NAV3237NAV3262NAV2058NAV5000NAV3246 | True
2565 | 703241463 | X - 3DEFINE CAMERA | 2019-07-03T18:31:55.000000+0000 | MazorX | NAV4094 | LATERAL APPROACH - 3DEFINE ISSUES | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - 3DEFINE CAMERA X Lat attempted 3Define camera would not function and when using generic work volume solutions for L4/L5 right trajectories were not acceptable they would give only unrealistic medial trajectories. Altered generic work volume about 15 different ways over a 20 minute period with no luck. Attempt Realsense DCM application restart, Software restart, and checked plug connection with no luck. Hard restart fixed it but by that time surgeon decided to go prone. (per Justin Brancaccio) UNK MAZOR X 3DEFINE CAMERA | X Lat attempted 3Define camera would not function and when using generic work volume solutions for L4/L5 right trajectories were not acceptable they would give only unrealistic medial trajectories. Altered generic work volume about 15 different ways over a 20 minute period with no luck. Attempt Realsense DCM application restart, Software restart, and checked plug connection with no luck. Hard restart fixed it but by that time surgeon decided to go prone. (per Justin Brancaccio) | x camera x lat attempted camera function using generic work volume solutions right trajectories acceptable unrealistic medial trajectories altered generic work volume different ways minute period luck attempt realsense dcm application restart software restart checked plug connection luck hard restart fixed time surgeon decided prone unk mazor x camera | NAV3244 | NAV2047 | NAV3261 | NAV4110 | NAV3233 | NAV3244NAV2047NAV3261NAV4110NAV3233 | False
2082 | 703225773 | X - Software issue | 2019-06-25T13:01:20.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X - Software issue It was reported that the system froze during draw spine and soft restart fixed the issue. MAZOR X SOFTWARE KIT VER-2.3 | It was reported that the system froze during draw spine and soft restart fixed the issue. | x software issue froze draw spine soft restart fixed issue mazor x software kit | NAV3248 | NAV3246 | NAV3242 | NAV4091 | NAV3247 | NAV3248NAV3246NAV3242NAV4091NAV3247 | True
8572 | 704443242 | XSE - Pressure Test Failure | 2021-06-10T20:41:31.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | ASM0263 | KIT SVC ASM0263 ELEC PANEL ASSY 03S NEW | XSE - Pressure Test Failure 2021-Jun-10 Salesforce C00228179, IUD x2 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the system failed a pressure test during servicing. There was no patient involvement. 2021-06-14 IUD x3, e1, e2 (Rep): Additional information received from a manufacturer representative reported that the pneumatic panel did not work correctly. The panel would not pressurize the actuator that lifts and supports the surgical system. The representative checked the hoses and everything seemed to hooked up correctly. 2021-06-16 IUD x3 (Other): No new information. 2021-Jul-27 IUDx2 (other): No new information. KIT SVC ASM0263 ELEC PANEL ASSY 03S NEW | 2021-Jun-10 Salesforce C00228179, IUD x2 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the system failed a pressure test during servicing. There was no patient involvement. 2021-06-14 IUD x3, e1, e2 (Rep): Additional information received from a manufacturer representative reported that the pneumatic panel did not work correctly. The panel would not pressurize the actuator that lifts and supports the surgical system. The representative checked the hoses and everything seemed to hooked up correctly. 2021-06-16 IUD x3 (Other): No new information. 2021-Jul-27 IUDx2 (other): No new information. | xse pressure test failure salesforce iud outside p rocedure failed pressure test servicing iud additional pneumatic panel did work correctly panel pressurize actuator lifts supports surgical checked hoses hooked correctly iud new new kit svc elec panel assy new | NAV3264 | NAV2054 | NAV3249 | NAV3251 | NAV2057 | NAV3264NAV2054NAV3249NAV3251NAV2057 | True
6392 | 703934616 | XSE - Screw Deviations | 2020-09-01T21:46:38.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Screw Deviations 2020-09-01 00695655, CCR, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that a S1-L4 percutaneous case was completed with cortical screws. Planning was completed and reviewed by the surgeon. Adjustments were made to try and line up the rods better. A Schanz pin was placed and connected to the bone mount bridge to mount the surgical system to the patient. Registration was completed successfully with 2 AP and 2 oblique images. Registration was completed using the complex algorithm with unacceptable values. The manufacturer representative adjusted the segmentation lines and was able to get green registration values. The surgeon verified that there were no shifts. Three of the six trajectories were completed successfully. The surgical arm was sent to the left side trajectories to mark the incisions and then the surgeon started at left S1 and then moved up to left L4. The same thing was then done on the right side. The manufacturer representative noted that the towers were in the way and the surgeon had to manually push the towers out of the way in order to do the L5 and L4 trajectories. The same issue occurred on the right side so the surgeon made the decision to skip right L5. While instrumenting right L4, the surgical arm started to shaking when the screw was being inserted. The surgeon stopped and looked at the screen. The surgical arm then stopped shaking. The surgeon proceed with the case and the surgical arm did not shake again. Imaging was taken, and right L4 was lateral to plan. L5 was also a bit lateral. The surgeon decided to remove the right L4 screw, but ended up placing it freehand in the same trajectory since it was still safe. The cause of the deviation was not determined. The manufacturer representative noted that navigation looked correct during the whole procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-09-10 e1, e2, e3, e4, e5, e6, e7 (Rep): Additional information received from a manufacturer representative reported that the surgeon was always on the left side of the patient during the procedure. The surgeon was not leaning on the patient and the patient was strapped to the bed around the thighs. The representative was unsure about a waterbed effect. 2020-09-22 e8 (Rep): No new information. MAZOR X SYSTEM TPL0059 | 2020-09-01 00695655, CCR, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that a S1-L4 percutaneous case was completed with cortical screws. Planning was completed and reviewed by the surgeon. Adjustments were made to try and line up the rods better. A Schanz pin was placed and connected to the bone mount bridge to mount the surgical system to the patient. Registration was completed successfully with 2 AP and 2 oblique images. Registration was completed using the complex algorithm with unacceptable values. The manufacturer representative adjusted the segmentation lines and was able to get green registration values. The surgeon verified that there were no shifts. Three of the six trajectories were completed successfully. The surgical arm was sent to the left side trajectories to mark the incisions and then the surgeon started at left S1 and then moved up to left L4. The same thing was then done on the right side. The manufacturer representative noted that the towers were in the way and the surgeon had to manually push the towers out of the way in order to do the L5 and L4 trajectories. The same issue occurred on the right side so the surgeon made the decision to skip right L5. While instrumenting right L4, the surgical arm started to shaking when the screw was being inserted. The surgeon stopped and looked at the screen. The surgical arm then stopped shaking. The surgeon proceed with the case and the surgical arm did not shake again. Imaging was taken, and right L4 was lateral to plan. L5 was also a bit lateral. The surgeon decided to remove the right L4 screw, but ended up placing it freehand in the same trajectory since it was still safe. The cause of the deviation was not determined. The manufacturer representative noted that navigation looked correct during the whole procedure. There was no patient harm and the procedure was delayed less than an hour. 2020-09-10 e1, e2, e3, e4, e5, e6, e7 (Rep): Additional information received from a manufacturer representative reported that the surgeon was always on the left side of the patient during the procedure. The surgeon was not leaning on the patient and the patient was strapped to the bed around the thighs. The representative was unsure about a waterbed effect. 2020-09-22 e8 (Rep): No new information. | xse screw deviations ccr iud spinal procedure treat degenerative discs fixation percutaneous case completed cortical screws planning completed reviewed surgeon adjustments try line rods better schanz pin placed connected bone mount bridge mount surgical registration completed successfully ap oblique images registration completed using complex algorithm unacceptable values adjusted segmentation lines able green registration values surgeon verified shifts trajectories completed successfully surgical arm sent left trajectories mark incisions surgeon started left moved left thing right noted towers way surgeon manually push towers way order trajectories issue occurred right surgeon decision skip right instrumenting right surgical arm started shaking screw inserted surgeon stopped looked screen surgical arm stopped shaking surgeon proceed case surgical arm did shake imaging taken right lateral plan bit lateral surgeon decided remove right screw ended placing freehand trajectory safe cause deviation determined noted navigation looked correct procedure procedure delayed hour additional surgeon left procedure surgeon leaning strapped bed thighs unsure waterbed effect new mazor x | NAV3242 | NAV4101 | NAV5000 | NAV3246 | NAV3243 | NAV3242NAV4101NAV5000NAV3246NAV3243 | True
8439 | 704414405 | XSE - Fan Not Functioning | 2021-05-25T22:56:40.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | FAN0023-01 | FAN-DC, 12V | XSE - Fan Not Functioning 2021-May-25 Salesforce C00227309, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that every time the guidance system booted up, a message was displayed stating the bottom right fan was not functioning. There was no patient involvement. 2021-05-28 IUD x5 (Other): No new information. 2021-Jun-29 IUDx3 (other): No new information. 2021-08-03 IUD x2 (Other): No new information. FAN-DC, 12V | 2021-May-25 Salesforce C00227309, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that every time the guidance system booted up, a message was displayed stating the bottom right fan was not functioning. There was no patient involvement. 2021-05-28 IUD x5 (Other): No new information. 2021-Jun-29 IUDx3 (other): No new information. 2021-08-03 IUD x2 (Other): No new information. | xse fan functioning salesforce iud outside proc edure time booted message displayed stating right fan functioning iud new new iud new fandc | NAV3299 | NAV2058 | NAV3249 | NAV3250 | NAV3253 | NAV3299NAV2058NAV3249NAV3250NAV3253 | True
5954 | 703846531 | X - Surgeon Monitor Blank | 2020-07-10T19:11:43.000000+0000 | MazorX | NAV3253 | TABLET ISSUES | 20 | ASM0337-01 | SURGEON&APOS;S SCREEN | X - Surgeon Monitor Blank 2020-07-10 00692493 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon monitor was working initially during the procedure. After registration, the screen went blank. A reboot of the system did not resolve the issue. There was no patient harm and the procedure was delayed less than an hour. 2020-07-23 IUD x2 (Rep, Other): It was stated that this is associated with a known issue, and will be addressed in the next software upgrade. 2020-08-05 IUD x2 (Other): No new information. SURGEON&APOS;S SCREEN | 2020-07-10 00692493 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon monitor was working initially during the procedure. After registration, the screen went blank. A reboot of the system did not resolve the issue. There was no patient harm and the procedure was delayed less than an hour. 2020-07-23 IUD x2 (Rep, Other): It was stated that this is associated with a known issue, and will be addressed in the next software upgrade. 2020-08-05 IUD x2 (Other): No new information. | x surgeon monitor blank spinal procedure surgeon monitor working initially procedure registration screen went blank reboot did resolve issue procedure delayed hour iud stated associated known issue addressed software upgrade iud new surgeonaposs screen | NAV3253 | NAV3247 | NAV3261 | NAV3246 | NAV3249 | NAV3253NAV3247NAV3261NAV3246NAV3249 | True
8227 | 704374270 | XSE - Screw Deviation | 2021-05-03T23:45:04.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Screw Deviation 2021-Apr-30 Salesforce C00225685, email 2021-05-10 09:46 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a Scan & Plan case for L3-L4 OLIF with perc screws was completed. The procedure started by taking a spin with the O-arm and sending the images to the guidance system for planning. While planning for the screws was done, the surgeon did the exposure for the OLIF. The surgeon then executed the screws, performed the OLIF, and then went back to place rods. Scout shots taken during the procedure showed the screws looked ok in the lateral view. Another lateral image was done and the screws looked to be in a different position. An AP image was taken and the screws on the left side were found to be lateral by 3.5-10 mm. The left L3 and L4 screws were removed during the procedure. A post operative CT scan was done of the right side and the screws were medial in the canal. A revision was done to remove the right side screws. No troubleshooting was done during the procedure since the surgeon thought the deviation may have been caused when trying to tighten the rods. Imaging was not done prior to that. The procedure was delayed less than an hour. 2021-05-17 email 2021-05-17 18:45, IUD (Rep): No new information. 2021-06-07 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2021-Apr-30 Salesforce C00225685, email 2021-05-10 09:46 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a Scan & Plan case for L3-L4 OLIF with perc screws was completed. The procedure started by taking a spin with the O-arm and sending the images to the guidance system for planning. While planning for the screws was done, the surgeon did the exposure for the OLIF. The surgeon then executed the screws, performed the OLIF, and then went back to place rods. Scout shots taken during the procedure showed the screws looked ok in the lateral view. Another lateral image was done and the screws looked to be in a different position. An AP image was taken and the screws on the left side were found to be lateral by 3.5-10 mm. The left L3 and L4 screws were removed during the procedure. A post operative CT scan was done of the right side and the screws were medial in the canal. A revision was done to remove the right side screws. No troubleshooting was done during the procedure since the surgeon thought the deviation may have been caused when trying to tighten the rods. Imaging was not done prior to that. The procedure was delayed less than an hour. 2021-05-17 email 2021-05-17 18:45, IUD (Rep): No new information. 2021-06-07 IUD (Other): No new information. | xse screw deviation salesforce email spinal procedure scan plan case olif perc screws completed procedure started taking spin oarm sending images planning planning screws surgeon did exposure olif surgeon executed screws performed olif went place rods scout shots taken procedure showed screws looked ok lateral view lateral image screws looked different position ap image taken screws left lateral mm left screws removed procedure post operative ct scan right screws medial canal revision remove right screws troubleshooting procedure surgeon thought deviation caused trying tighten rods imaging prior procedure delayed hour email iud new iud new mazor x | NAV5000 | NAV3299 | NAV2058 | NAV3242 | NAV1012 | NAV5000NAV3299NAV2058NAV3242NAV1012 | True
7556 | 704226356 | XSE - Unable to Track Instruments | 2021-02-11T21:46:52.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 30 | KIT1286-01 | SVC KIT1286-01 NDI VEGA CAMERA | XSE - Unable to Track Instruments 2021-Feb-11 Salesforce C00220934 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that of the instruments were not tracking. This included all Navlocks trackers, tracker kit, ball probe, robotic frame, ball probe, reference frame, and Navlock kit. There was no patient involved. 2021-03-01 IUD x3 (Other): No new information. 2021-03-16 email 2021-03-16 20:05 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved changing the arm, moving the camera around and trying new instruments. The cause of the tracking error was the camera and system error. 2021-04-19 IUD x5 (Other): No new information. 2021-05-27 IUD x2 (Other): No new information. SVC KIT1286-01 NDI VEGA CAMERA | 2021-Feb-11 Salesforce C00220934 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that of the instruments were not tracking. This included all Navlocks trackers, tracker kit, ball probe, robotic frame, ball probe, reference frame, and Navlock kit. There was no patient involved. 2021-03-01 IUD x3 (Other): No new information. 2021-03-16 email 2021-03-16 20:05 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved changing the arm, moving the camera around and trying new instruments. The cause of the tracking error was the camera and system error. 2021-04-19 IUD x5 (Other): No new information. 2021-05-27 IUD x2 (Other): No new information. | xse unable track instruments salesforce outside procedure instruments tracking included navlocks trackers tracker kit ball probe robotic frame ball probe reference frame navlock kit involved iud new email additional troubleshooting involved changing arm moving camera trying new instruments cause tracking error camera error iud new iud new svc ndi vega camera | NAV3299 | NAV3246 | NAV2054 | NAV4110 | NAV3256 | NAV3299NAV3246NAV2054NAV4110NAV3256 | False
9665 | 704692413 | XSE ¦ Disc Stuck & Unexpected Shutdown | 2021-11-01T22:27:38.000000+0000 | MazorX Stealth Edition | NAV3247 | EXITS | 30 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE ¦ Disc Stuck & Unexpected Shutdown 2021-Nov-01 Salesforce C00237144 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a disc got stuck in the external disc reader. The manufacturer representative believes there may be two discs in the drive that caused the issue even though they pressed the eject button before inserting the disc. It was also reported that the system shutdown on its own. It rebooted and they were able to use the system, but upon shutdown after a case the manufacturer representative got a blue screen with a sad face that stated the PC ran into issues and was collecting error information. No patient was present. 2021-11-02 IUD x5 (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Nov-01 Salesforce C00237144 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a disc got stuck in the external disc reader. The manufacturer representative believes there may be two discs in the drive that caused the issue even though they pressed the eject button before inserting the disc. It was also reported that the system shutdown on its own. It rebooted and they were able to use the system, but upon shutdown after a case the manufacturer representative got a blue screen with a sad face that stated the PC ran into issues and was collecting error information. No patient was present. 2021-11-02 IUD x5 (Other): No new information. | xse disc stuck unexpected shutdown salesforce outside procedure disc got stuck external disc reader believes discs drive caused issue pressed eject button inserting disc shutdown rebooted able use shutdown case got blue screen sad face stated pc ran issues collecting error present iud new mazorx sw installation kit | NAV3246 | NAV3247 | NAV3262 | NAV3248 | NAV3257 | NAV3246NAV3247NAV3262NAV3248NAV3257 | True
2457 | 703236295 | X - ARM/COMMUNICATION ERROR ISSUE | 2019-07-01T17:20:51.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 30 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X - ARM/COMMUNICATION ERROR ISSUE Arm Shift detected in Joint 3. Followed by "FATAL ERROR: CANBUS Connection Lost to Nodes # 1, 2, 3, 4, 5, 6, 7, 8, 10" UNKNOWN MAZOR X SW | Arm Shift detected in Joint 3. Followed by "FATAL ERROR: CANBUS Connection Lost to Nodes # 1, 2, 3, 4, 5, 6, 7, 8, 10" | x armcommunication error issue arm shift detected joint followed fatal error canbus connection lost nodes unknown mazor x sw | NAV2047 | NAV4098 | NAV3246 | NAV3242 | NAV2054 | NAV2047NAV4098NAV3246NAV3242NAV2054 | True
1794 | 703216096 | X - Shoulder issue | 2019-06-19T20:08:46.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Shoulder issue It was reported that before starting Advance set up, the Shoulder button was not working (light was on but no motion). I had to operate the shoulder from the software. Was fixed after I opened it on the software. UNK MAZOR X SHOULDER | It was reported that before starting Advance set up, the Shoulder button was not working (light was on but no motion). I had to operate the shoulder from the software. Was fixed after I opened it on the software. | x shoulder issue starting advance set shoulder button working operate shoulder software fixed opened software unk mazor x shoulder | NAV4107 | NAV3252 | NAV4108 | NAV2054 | NAV4106 | NAV4107NAV3252NAV4108NAV2054NAV4106 | True
4023 | 703491015 | X - Target Extenders Worn | 2019-11-27T21:21:29.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2310-03 | TARGET EXTENDER | X - Target Extenders Worn 2018-12-06 SCL1805189 (Rep) Medtronic received information regarding a guidance system being used during preventative maintenance. It was reported that the system was within acceptable operating specifications. Two target extenders were worn and due to be replaced as manufacturer recommends annual replacement. As the event occurred during maintenance, no patient was involved. TARGET EXTENDER | 2018-12-06 SCL1805189 (Rep) Medtronic received information regarding a guidance system being used during preventative maintenance. It was reported that the system was within acceptable operating specifications. Two target extenders were worn and due to be replaced as manufacturer recommends annual replacement. As the event occurred during maintenance, no patient was involved. | x target extenders worn preventative maintenance acceptable operating specifications target extenders worn replaced recommends annual replacement occurred maintenance involved target extender | NAV3299 | NAV3295 | NAV2054 | NAV4091 | NAV2058 | NAV3299NAV3295NAV2054NAV4091NAV2058 | True
5280 | 703637189 | XSE - Registration Issues, L2 Deviation | 2020-02-26T16:13:33.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Issues, L2 Deviation 2020-02-24 CCR, 00685201 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a juvenile global deformity with fixation at T2-L3. It was reported that planning for the case was completed by the surgeon prior to the procedure to completed in three segments. The dual clamp was selected to mount the platform to the patient and was used for the entire construct. Due to the patient's kyphosis, the clamp was placed higher than usual at T4/T6 Registration for the first segment was achieved on the first attempt under basic but showed all red values with one yellow. The manufacturer representative returned to planning and adjusted the segmentation lines manually but all red and yellow values remained. New images were taken to clear up the end plates but still achieved the same results. After multiple attempts, the representative adjusted segmentation lines again and was able to get yellow and green values for T2-T6. The surgeon visually verified no shift and proceeded with the operation. Red screws were present at T2 due to unreachability. T2 was re-planned and the surgeon sent to the left trajectory which skiv ed superior. T3 screws and right T2 were not red but instead should "no solutions found." The representative returned to planning to review the screws, then reviewed work volume in the operation screen and saw no issues. Once the representative had click back into operation mode, without changing anything, there were solutions fro those trajectories and T3 trajectories were placed. The second segment of surgery was for T7-T12. Registration was successful with no issue and all trajectories were placed accurately. The third segment of surgery was for L1-L3. Automatic registration failed multiple times. Manual registration yielded all red values. New images were taken with the same result. The C-Arm was recalibrated and more new images were taken but registration was still not successful. The AP image was cleared up to remove K-wires in the image and then automatic registration was successful on basic with all green values. A similar issue occurred with the third segment registration as had occurred in the first segment; during planning it showed "no solutions found" for L2 but when they went back into planning then returned to the operation screen, the problem was resolved. No patient harm was reported and surgery was delayed less than an hour. The representative noted the surgeon expressed frustration regarding the failed registrations and "no solutions" for the screws, however was satisfied with all screw placements which were verified accurate with 2D/3D portable CT spins. 2020-02-26 IUD x3, email 2020-03-03 12:30 (Rep): Additional information received from a manufacturer representative reported that the deviation was less than 3.5 mm and was unacceptable to the surgeon so the placed the screw freehand. MAZOR X SYSTEM TPL0059 | 2020-02-24 CCR, 00685201 ID (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a juvenile global deformity with fixation at T2-L3. It was reported that planning for the case was completed by the surgeon prior to the procedure to completed in three segments. The dual clamp was selected to mount the platform to the patient and was used for the entire construct. Due to the patient's kyphosis, the clamp was placed higher than usual at T4/T6 Registration for the first segment was achieved on the first attempt under basic but showed all red values with one yellow. The manufacturer representative returned to planning and adjusted the segmentation lines manually but all red and yellow values remained. New images were taken to clear up the end plates but still achieved the same results. After multiple attempts, the representative adjusted segmentation lines again and was able to get yellow and green values for T2-T6. The surgeon visually verified no shift and proceeded with the operation. Red screws were present at T2 due to unreachability. T2 was re-planned and the surgeon sent to the left trajectory which skiv ed superior. T3 screws and right T2 were not red but instead should "no solutions found." The representative returned to planning to review the screws, then reviewed work volume in the operation screen and saw no issues. Once the representative had click back into operation mode, without changing anything, there were solutions fro those trajectories and T3 trajectories were placed. The second segment of surgery was for T7-T12. Registration was successful with no issue and all trajectories were placed accurately. The third segment of surgery was for L1-L3. Automatic registration failed multiple times. Manual registration yielded all red values. New images were taken with the same result. The C-Arm was recalibrated and more new images were taken but registration was still not successful. The AP image was cleared up to remove K-wires in the image and then automatic registration was successful on basic with all green values. A similar issue occurred with the third segment registration as had occurred in the first segment; during planning it showed "no solutions found" for L2 but when they went back into planning then returned to the operation screen, the problem was resolved. No patient harm was reported and surgery was delayed less than an hour. The representative noted the surgeon expressed frustration regarding the failed registrations and "no solutions" for the screws, however was satisfied with all screw placements which were verified accurate with 2D/3D portable CT spins. 2020-02-26 IUD x3, email 2020-03-03 12:30 (Rep): Additional information received from a manufacturer representative reported that the deviation was less than 3.5 mm and was unacceptable to the surgeon so the placed the screw freehand. | xse registration issues deviation ccr id spinal procedure treat juvenile global deformity fixation planning case completed surgeon prior procedure completed segments dual clamp selected mount platform entire construct patients kyphosis clamp placed higher usual registration segment achieved attempt basic showed red values yellow returned planning adjusted segmentation lines manually red yellow values remained new images taken clear end plates achieved results multiple attempts adjusted segmentation lines able yellow green values surgeon visually verified shift proceeded operation red screws present unreachability replanned surgeon sent left trajectory skiv ed superior screws right red instead solutions returned planning review screws reviewed work volume operation screen saw issues click operation mode changing solutions fro trajectories trajectories placed second segment surgery registration successful issue trajectories placed accurately segment surgery automatic registration failed multiple times manual registration yielded red values new images taken carm recalibrated new images taken registration successful ap image cleared remove kwires image automatic registration successful basic green values similar issue occurred segment registration occurred segment planning showed solutions went planning returned operation screen problem resolved surgery delayed hour noted surgeon expressed frustration failed registrations solutions screws satisfied screw placements verified accurate portable ct spins iud email additional deviation mm unacceptable surgeon placed screw freehand mazor x | NAV3242 | NAV4112 | NAV5000 | NAV3241 | NAV2051 | NAV3242NAV4112NAV5000NAV3241NAV2051 | True
5590 | 703759911 | XSE - Lost Connection to Controller | 2020-05-19T14:13:44.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | ASM0270-01S | KIT SVC ASM0270-01S CART ELEC PANEL ASSY | XSE - Lost Connection to Controller 2020-05-18 00689135, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a communication issue with the guidance system during the procedure. The surgical arm was free to move, but all of the buttons on the guidance system were non-responsive. The manufacturer representative checked the indications tab and breakers, but the system still was not working. Two hard restarts of the system was completed and communication with the controller was restored. There was no patient harm and the procedure was delayed less than one hour. 2020-05-19 IUD x2 (Rep): No new information. 2020-07-20 IUD (Other): No new information. KIT SVC ASM0270-01S CART ELEC PANEL ASSY | 2020-05-18 00689135, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there was a communication issue with the guidance system during the procedure. The surgical arm was free to move, but all of the buttons on the guidance system were non-responsive. The manufacturer representative checked the indications tab and breakers, but the system still was not working. Two hard restarts of the system was completed and communication with the controller was restored. There was no patient harm and the procedure was delayed less than one hour. 2020-05-19 IUD x2 (Rep): No new information. 2020-07-20 IUD (Other): No new information. | xse lost connection controller iud spinal procedure communication issue procedure surgical arm free buttons nonresponsive checked indications tab breakers working hard restarts completed communication controller restored procedure delayed hour iud new iud new kit svc cart elec panel assy | NAV2047 | NAV3264 | NAV2058 | NAV5000 | NAV1006 | NAV2047NAV3264NAV2058NAV5000NAV1006 | True
4564 | 703551034 | XSE- X-eye connection issue | 2020-01-08T16:00:24.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | ASM0386-01 | X-EYE MOTOR ASSEMBLY | XSE- X-eye connection issue 2020-1-8 00681542(Rep): Medtronic received information regarding a guidance system being used during service. It was reported that the x-eye camera intermittently connected and a soft restart had to be performed to resolve the issue. No patient involved. 2020-1-13 IUD x2(Other): No new information. 2020-02-09 IUD (Other): No new information. 2020-02-26 IUD (Other): No new information. 2020-03-05 IUD (Other): No new information. X-EYE MOTOR ASSEMBLY | 2020-1-8 00681542(Rep): Medtronic received information regarding a guidance system being used during service. It was reported that the x-eye camera intermittently connected and a soft restart had to be performed to resolve the issue. No patient involved. 2020-1-13 IUD x2(Other): No new information. 2020-02-09 IUD (Other): No new information. 2020-02-26 IUD (Other): No new information. 2020-03-05 IUD (Other): No new information. | xse xeye connection issue service xeye camera intermittently connected soft restart performed resolve issue involved iud new iud new iud new iud new xeye motor assembly | NAV4110 | NAV4109 | NAV2047 | NAV2054 | NAV2058 | NAV4110NAV4109NAV2047NAV2054NAV2058 | True
5293 | 703641871 | XSE - Clamp Link Screw Stripped | 2020-02-28T16:14:06.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MEC5206-01S | KIT SVC MEC5206-01S RNX-ACC-021 | XSE - Clamp Link Screw Stripped 2020-02-28 00685366 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the clamp link screw was stripped and needs to be replaced. There was no patient involvement. 2020-03-02 Email 2020-03-02, IUD x3 (Rep): No new information 2020-03-31 IUD (Other): No new information. KIT SVC MEC5206-01S RNX-ACC-021 | 2020-02-28 00685366 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the clamp link screw was stripped and needs to be replaced. There was no patient involvement. 2020-03-02 Email 2020-03-02, IUD x3 (Rep): No new information 2020-03-31 IUD (Other): No new information. | xse clamp link screw stripped outside procedure clamp link screw stripped needs replaced email iud new iud new kit svc | NAV3299 | NAV2054 | NAV4090 | NAV4100 | NAV3251 | NAV3299NAV2054NAV4090NAV4100NAV3251 | True
2922 | 703258924 | X- Positioner Won't Lock/Tooth Stuck | 2019-07-15T17:53:02.000000+0000 | MazorX | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X- Positioner Won't Lock/Tooth Stuck system won't unlock - tooth stuck. Ryan K talked to Mackie Young on the phone anad they were able talk through the issue UNK MAZOR X POSITIONER | system won't unlock - tooth stuck. Ryan K talked to Mackie Young on the phone anad they were able talk through the issue | x positioner wont locktooth stuck wont unlock tooth stuck ryan k talked mackie young phone anad able talk issue unk mazor x positioner | NAV3244 | NAV2054 | NAV3250 | NAV2047 | NAV4105 | NAV3244NAV2054NAV3250NAV2047NAV4105 | False
10131 | 704804713 | XSE - Multiple Issues/Error Code | 2022-01-12T00:46:20.000000+0000 | MazorX Stealth Edition | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 20 | HRS0327-01S | KIT SVC HRS0327-01S TEETH LOCKING SENSOR | XSE - Multiple Issues/Error Code 2022-Jan-11 Salesforce C00241507 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the clip fell out of the locking mechanism of the positioner, the surgical arm covers were falling off, and an error code 4490 was being displayed. The manufacturer representative noted that they were unable to put the clip back in to complete the case. There was no patient harm and the procedure was delayed less than an hour. 2022-01-17 IUD (Other): No new information. 2022-01-25 IUD (Other): No new information. 2022-01-27 email 2022-01-27 10:16 (Rep): No new information. KIT SVC HRS0327-01S TEETH LOCKING SENSOR | 2022-Jan-11 Salesforce C00241507 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the clip fell out of the locking mechanism of the positioner, the surgical arm covers were falling off, and an error code 4490 was being displayed. The manufacturer representative noted that they were unable to put the clip back in to complete the case. There was no patient harm and the procedure was delayed less than an hour. 2022-01-17 IUD (Other): No new information. 2022-01-25 IUD (Other): No new information. 2022-01-27 email 2022-01-27 10:16 (Rep): No new information. | xse multiple issueserror code salesforce spinal p rocedure clip fell locking mechanism positioner surgical arm covers falling error code displayed noted unable clip complete case procedure delayed hour iud new iud new email new kit svc teeth locking sensor | NAV4099 | NAV3251 | NAV2047 | NAV4103 | NAV3250 | NAV4099NAV3251NAV2047NAV4103NAV3250 | True
7970 | 704322981 | XSE - Incorrect Date/Time On System | 2021-04-05T20:09:42.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Incorrect Date/Time On System 2021-Apr-02 Salesforce C00223832 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the time and date were not correct on the guidance system. The windows task manager was opening on the surgeon screen so the manufacturer representative was not able to correct the issue. There was no patient involvement. 2021-04-19 IUD x4 (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Apr-02 Salesforce C00223832 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the time and date were not correct on the guidance system. The windows task manager was opening on the surgeon screen so the manufacturer representative was not able to correct the issue. There was no patient involvement. 2021-04-19 IUD x4 (Other): No new information. | xse incorrect datetime salesforce outside procedure time date correct windows task manager opening surgeon screen able correct issue iud new unk mazor x workstation | NAV2058 | NAV3264 | NAV2047 | NAV3262 | NAV3253 | NAV2058NAV3264NAV2047NAV3262NAV3253 | True
666 | 703169788 | X-Accessory/WS/Tablet/SW/Arm Issues | 2019-05-23T17:11:01.000000+0000 | MazorX | NAV3247 | EXITS | 70 | KIT0639-05 | MAZOR X SOFTWARE KIT VER-2.2.1 | X-Accessory/WS/Tablet/SW/Arm Issues 2.2.1.14 build of the SW installed successfully and Site ATP initiated – completed successfully almost entirely. Surgical Arm, setup and advanced setup – Setup and advance setup were performed many times: SW Bug – there is no beep (arrived) sound when performing surgical arm setup (when the surgical arm arrives at the only point to be checked with the pointer). Accuracy was checked with the setup arm guide – very good but at times the pointer would brush against the side of the hole while remaining accurately in it. While simulating a 3D Marker images acquisition during surgery – sending to automatic AP and OB – the surgical arm reached the AP position and was adjusted manually with the 'Ctrl+Alt+K' – after about 10 adjustments (that may have caused the surgical arm to enter 'into the mapped surface') it got an 'arm shift detected in joint 2' error. The surgical arm wouldn’t move even after hard restarts and 'reset motion' was used. Advanced setup followed – surgical arm didn’t reach the first 3 points and moved forward – all points accurate. Another advanced setup was initiated as this one ended – all 10 points accurately reached. SW issue – after a 3Define scan auto AP position couldn’t be found by the SW. after a few attempts the SW crashed and restarted. Additional attempts but AP position not found. OB position was sent and successfully reached, and following this an acceptable position for the auto AP was available – sent and reached successfully. After a successful operation flow simulation was completed the advanced setup was initiated as the Site ATP requires – the surgical arm didn’t even reach the first point and 'arm shift detected in joint 4' error appeared. The surgical arm wouldn’t move even after hard restarts and 'reset motion' was used. The 'DualEncoderTolerance' value in the ini file was changed from 0.2 to 0.4 in the ini file with Eli's approval. Consequently advanced setup was successfully completed and the entire operation flow was repeated successfully. Tablet: The tablet is not charging from neither from the storage position nor from the intraoperative position – even before draping. It needs to be charged using a cable before the surgery to be usable intraoperatively. Accessories:'Older' version of the 5mm-shaft drill-bits are no longer available at this site - replaced with 'newer' version – long 80mm, short 80mm and short 2.4mm. New sets of long and short RN cannulas and drill guides were delivered. The interface of the 'new' drill bits with the RN drill-guides was manually checked and found to be satisfactory. Two surgical-arm arm-tubes for the SA tools were lef t on site. Their interface with the cannulas of one of the KIT0150 on site was tested and found to be not optimal – I believe there is too much play and recommend replacing the SA cannulas at this site as their tolerances seem to be loose interfacing with the the Mazor-X arm tube. The navigation stylus (the older version, not PEEK) is in the Mazor-X accessories tray. WS issues: Occasionally when the pneumatic compressor starts working it makes a loud popping sound. We couldn’t export a study from the DVD drive (couldn’t burn to a DVD from the WS) Frame grabber: for some reason the image from the C-Arm is slightly cut from the top and from the bottom when seen on the WS: MAZOR X SOFTWARE KIT VER-2.2.1 | 2.2.1.14 build of the SW installed successfully and Site ATP initiated – completed successfully almost entirely. Surgical Arm, setup and advanced setup – Setup and advance setup were performed many times: SW Bug – there is no beep (arrived) sound when performing surgical arm setup (when the surgical arm arrives at the only point to be checked with the pointer). Accuracy was checked with the setup arm guide – very good but at times the pointer would brush against the side of the hole while remaining accurately in it. While simulating a 3D Marker images acquisition during surgery – sending to automatic AP and OB – the surgical arm reached the AP position and was adjusted manually with the 'Ctrl+Alt+K' – after about 10 adjustments (that may have caused the surgical arm to enter 'into the mapped surface') it got an 'arm shift detected in joint 2' error. The surgical arm wouldn’t move even after hard restarts and 'reset motion' was used. Advanced setup followed – surgical arm didn’t reach the first 3 points and moved forward – all points accurate. Another advanced setup was initiated as this one ended – all 10 points accurately reached. SW issue – after a 3Define scan auto AP position couldn’t be found by the SW. after a few attempts the SW crashed and restarted. Additional attempts but AP position not found. OB position was sent and successfully reached, and following this an acceptable position for the auto AP was available – sent and reached successfully. After a successful operation flow simulation was completed the advanced setup was initiated as the Site ATP requires – the surgical arm didn’t even reach the first point and 'arm shift detected in joint 4' error appeared. The surgical arm wouldn’t move even after hard restarts and 'reset motion' was used. The 'DualEncoderTolerance' value in the ini file was changed from 0.2 to 0.4 in the ini file with Eli's approval. Consequently advanced setup was successfully completed and the entire operation flow was repeated successfully. Tablet: The tablet is not charging from neither from the storage position nor from the intraoperative position – even before draping. It needs to be charged using a cable before the surgery to be usable intraoperatively. Accessories:'Older' version of the 5mm-shaft drill-bits are no longer available at this site - replaced with 'newer' version – long 80mm, short 80mm and short 2.4mm. New sets of long and short RN cannulas and drill guides were delivered. The interface of the 'new' drill bits with the RN drill-guides was manually checked and found to be satisfactory. Two surgical-arm arm-tubes for the SA tools were lef t on site. Their interface with the cannulas of one of the KIT0150 on site was tested and found to be not optimal – I believe there is too much play and recommend replacing the SA cannulas at this site as their tolerances seem to be loose interfacing with the the Mazor-X arm tube. The navigation stylus (the older version, not PEEK) is in the Mazor-X accessories tray. WS issues: Occasionally when the pneumatic compressor starts working it makes a loud popping sound. We couldn’t export a study from the DVD drive (couldn’t burn to a DVD from the WS) Frame grabber: for some reason the image from the C-Arm is slightly cut from the top and from the bottom when seen on the WS: | xaccessorywstabletswarm issues build sw installed successfully site atp initiated – completed successfully entirely surgical arm setup advanced setup – setup advance setup performed times sw bug – beep sound performing surgical arm setup accuracy checked setup arm guide – good times pointer brush hole remaining accurately simulating marker images acquisition surgery – sending automatic ap ob – surgical arm reached ap position adjusted manually ctrlaltk – adjustments got arm shift detected joint error surgical arm wouldn’t hard restarts reset motion advanced setup followed – surgical arm didn’t reach points moved forward – points accurate advanced setup initiated ended – points accurately reached sw issue – scan auto ap position couldn’t sw attempts sw crashed restarted additional attempts ap position ob position sent successfully reached following acceptable position auto ap available – sent reached successfully successful operation flow simulation completed advanced setup initiated site atp requires – surgical arm didn’t reach point arm shift detected joint error appeared surgical arm wouldn’t hard restarts reset motion dualencodertolerance value ini file changed ini file elis approval consequently advanced setup successfully completed entire operation flow repeated successfully tablet tablet charging storage position intraoperative position – draping needs charged using cable surgery usable intraoperatively accessoriesolder version drillbits longer available site replaced newer version – long short short new sets long short rn cannulas drill guides delivered interface new drill bits rn drillguides manually checked satisfactory surgicalarm armtubes sa tools lef t site interface cannulas site tested optimal – believe play recommend replacing sa cannulas site tolerances loose interfacing mazorx arm tube navigation stylus mazorx accessories tray ws issues occasionally pneumatic compressor starts working makes loud popping sound couldn’t export study dvd drive frame grabber reason image carm slightly cut seen ws mazor x software kit | NAV3246 | NAV3242 | NAV3249 | NAV4092 | NAV2044 | NAV3246NAV3242NAV3249NAV4092NAV2044 | False
7873 | 704299214 | XSE - No Pressure | 2021-03-23T17:46:30.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 40 | KIT0759-01 | PNEUMATIC REPAIR KIT | XSE - No Pressure 2021-Mar-22 Salesforce C00223028 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system's pressure went down to 0 and the compressor was not turning on. The system had been working with the compressor cycling on and off like normal, but then just stopped. The manufacturer representative was performing a hard restart. This caused less than an hour delay. There was no impact on the patient outcome. 2021-03-24 IUD x2 (Other): No new information. 2021-04-26 IUD x3 (Other): No new information. 2021-06-20 IUD (Other): No new information. 2021-08-19 IUD x2 (Other): No new information. PNEUMATIC REPAIR KIT | 2021-Mar-22 Salesforce C00223028 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system's pressure went down to 0 and the compressor was not turning on. The system had been working with the compressor cycling on and off like normal, but then just stopped. The manufacturer representative was performing a hard restart. This caused less than an hour delay. There was no impact on the patient outcome. 2021-03-24 IUD x2 (Other): No new information. 2021-04-26 IUD x3 (Other): No new information. 2021-06-20 IUD (Other): No new information. 2021-08-19 IUD x2 (Other): No new information. | xse pressure salesforce spinal p rocedure systems pressure went compressor turning working compressor cycling like normal just stopped performing hard restart caused hour delay impact outcome iud new iud new iud new iud new pneumatic repair kit | NAV3264 | NAV4095 | NAV3299 | NAV2057 | NAV1006 | NAV3264NAV4095NAV3299NAV2057NAV1006 | True
4863 | 703582959 | X- Registration issues | 2020-01-27T18:10:39.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X- Registration issues 2018-3-8 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and an advanced accuracy test was competed. A schanz pin was inserted into the left PSIS and attached to a schanz arm. A 3Define scan and draw spine were completed. Fluoro segmentation failed on the first attempt, so a new OBL image was taken. The second attempt at segmentation was success using the complex algorithm, with green values at all levels. The trajectories were executed and verified with fluoro. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-3-8 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and an advanced accuracy test was competed. A schanz pin was inserted into the left PSIS and attached to a schanz arm. A 3Define scan and draw spine were completed. Fluoro segmentation failed on the first attempt, so a new OBL image was taken. The second attempt at segmentation was success using the complex algorithm, with green values at all levels. The trajectories were executed and verified with fluoro. There was no patient harm and the procedure was not delayed over an hour. | x registration issues ccr spinal procedure shoulder calibration checked advanced accuracy test competed schanz pin inserted left psis attached schanz arm scan draw spine completed fluoro segmentation failed attempt new obl image taken second attempt segmentation success using complex algorithm green values levels trajectories executed verified fluoro procedure delayed hour mazor x | NAV3242 | NAV4112 | NAV4091 | NAV3299 | NAV3244 | NAV3242NAV4112NAV4091NAV3299NAV3244 | True
4862 | 703582936 | X- 3Define issues | 2020-01-27T17:44:20.000000+0000 | MazorX | NAV3239 | INACCURATE WORK VOLUMES | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X- 3Define issues 2018-1-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and the surgical arm was setup. The schanz pin was inserted into the right PSIS, and attached to a BMB and schanz bond. A 3Define scan was completed, but it had very poor quality. The 3Define camera was checked to ensure the lenses were not covered. Another 3Define scan was completed, but the quality was still poor. The surgeon accepted the scan and moved forward with the procedure. Three AP and two OBL images were taken, and registration was achieved using the basic algorithm with green values at all levels. The trajectories were executed and verified with AP and lateral shots. After the procedure the manufacturer representative went into the camera settings , where the image was still blurry. A soft restart was completed, and the image became clear. It was also noted that the 3Define seemed to freeze if left inactive for a short period of time. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X 3DEFINE CAMERA | 2018-1-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked and the surgical arm was setup. The schanz pin was inserted into the right PSIS, and attached to a BMB and schanz bond. A 3Define scan was completed, but it had very poor quality. The 3Define camera was checked to ensure the lenses were not covered. Another 3Define scan was completed, but the quality was still poor. The surgeon accepted the scan and moved forward with the procedure. Three AP and two OBL images were taken, and registration was achieved using the basic algorithm with green values at all levels. The trajectories were executed and verified with AP and lateral shots. After the procedure the manufacturer representative went into the camera settings , where the image was still blurry. A soft restart was completed, and the image became clear. It was also noted that the 3Define seemed to freeze if left inactive for a short period of time. There was no patient harm and the procedure was not delayed over an hour. | x issues ccr spinal procedure shoulder calibration checked surgical arm setup schanz pin inserted right psis attached bmb schanz bond scan completed poor quality camera checked ensure lenses covered scan completed quality poor surgeon accepted scan moved forward procedure ap obl images taken registration achieved using basic algorithm green values levels trajectories executed verified ap lateral shots procedure went camera settings image blurry soft restart completed image clear noted freeze left inactive short period time procedure delayed hour unk mazor x camera | NAV3244 | NAV2057 | NAV3239 | NAV3235 | NAV3248 | NAV3244NAV2057NAV3239NAV3235NAV3248 | True
10162 | 704811757 | XSE - OOBF/SRA Failed Accuracy Test | 2022-01-14T20:49:40.000000+0000 | MazorX Stealth Edition | NAV4093 | INACCURACY DURING ADVANCED SETUP | 30 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - OOBF/SRA Failed Accuracy Test 2022-Jan-14 Salesforce C00241728 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed a stress test. There was no patient involvement. 2022-01-18 IUD (Other): No new information. 2022-01-24 IUD (Other): No new information. 2022-01-28 IUD (Other): No new information. 2022-02-01 email 2022-02-01 15:27 (Rep): No new information. 2022-02-02 e1, e2 (Rep): Additional information received from a manufacturer representative reported that the surgical arm failed an advanced accuracy test and not a stress test. 2022-02-11 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-02 IUD x2 (Other): No new information. 2022-03-21 IUD (Other): No new information. 2022-05-02 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2022-Jan-14 Salesforce C00241728 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm failed a stress test. There was no patient involvement. 2022-01-18 IUD (Other): No new information. 2022-01-24 IUD (Other): No new information. 2022-01-28 IUD (Other): No new information. 2022-02-01 email 2022-02-01 15:27 (Rep): No new information. 2022-02-02 e1, e2 (Rep): Additional information received from a manufacturer representative reported that the surgical arm failed an advanced accuracy test and not a stress test. 2022-02-11 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-02 IUD x2 (Other): No new information. 2022-03-21 IUD (Other): No new information. 2022-05-02 IUD (Other): No new information. | xse oobfsra failed accuracy test salesforce outside procedure surgical arm failed stress test iud new iud new iud new email new additional surgical arm failed advanced accuracy test stress test iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4093 | NAV2054 | NAV4092 | NAV4095 | NAV3249NAV4093NAV2054NAV4092NAV4095 | True
20 | 702971533 | X- Inaccuracy During Setup | 2019-01-23T13:31:38.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X- Inaccuracy During Setup during pre-operative setup ie “arm set up” the error 8047 popped up (Arm did not reach ) the arm did not make any movement other predefined location also return error 295 (Arm failed to reach final destination) . the error also appear for predefined location at other tab as well ie Storage Tab. I performed a soft restart but the issue was not resolved I then proceed to hard restart. The system is back to normal. The shoulder location was straight with seem to be accurate encoder reading (almost close to 0 on both shoulder MAZOR X SYSTEM TPL0059 | during pre-operative setup ie “arm set up” the error 8047 popped up (Arm did not reach ) the arm did not make any movement other predefined location also return error 295 (Arm failed to reach final destination) . the error also appear for predefined location at other tab as well ie Storage Tab. I performed a soft restart but the issue was not resolved I then proceed to hard restart. The system is back to normal. The shoulder location was straight with seem to be accurate encoder reading (almost close to 0 on both shoulder | x inaccuracy setup preoperative setup “arm set up” error popped arm did make movement predefined location return error error appear predefined location tab storage tab performed soft restart issue resolved proceed hard restart normal shoulder location straight accurate encoder reading close shoulder mazor x | NAV2047 | NAV1006 | NAV3242 | NAV4091 | NAV3249 | NAV2047NAV1006NAV3242NAV4091NAV3249 | False
8583 | 704443280 | XSE - C-arm Bracket Broken | 2021-06-10T20:41:52.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1226-02S | KIT SVC MAS1226-02S MAIN IMG ADPTR BRCKT | XSE - C-arm Bracket Broken 2021-Jun-10 Salesforce C00228202 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that that the C-arm bracket that allowed a 90 or 180 degree attachment was damaged. The push button on the bracket did not function so it could not be adjusted. There was no patient involvement. 2021-06-18 IUD x2 (Other): No new information. 2021-07-29 email 2021-07-29 19:30 (Rep): The cause or likely cause of the damage was due to the pin being jammed in the bracket. KIT SVC MAS1226-02S MAIN IMG ADPTR BRCKT | 2021-Jun-10 Salesforce C00228202 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that that the C-arm bracket that allowed a 90 or 180 degree attachment was damaged. The push button on the bracket did not function so it could not be adjusted. There was no patient involvement. 2021-06-18 IUD x2 (Other): No new information. 2021-07-29 email 2021-07-29 19:30 (Rep): The cause or likely cause of the damage was due to the pin being jammed in the bracket. | xse carm bracket broken salesforce outside procedure carm bracket allowed degree attachment damaged push button bracket did function adjusted iud new email cause likely cause damage pin jammed bracket kit svc main img adptr brckt | NAV3258 | NAV2058 | NAV3249 | NAV3299 | NAV2055 | NAV3258NAV2058NAV3249NAV3299NAV2055 | True
3583 | 703380941 | XSE - Joint Shift Error | 2019-09-25T18:10:26.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | XSE - Joint Shift Error 2019-09-25 SCL1902915, Intake Form x2, 00674880, ID, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the guidance system was booted up and after logging in, a joint 3 shift error was displayed when starting the setup process. The error message was closed and the representative attempted to proceed, but the error code "3836 Failed to Reach Defined Position" was displayed. A soft restart of the system was done, but the issue was not resolved. A hard restart of the system was done and the representative logged into the system as a supersuser. Once logged in, the movement was reset and the surgical arm was sent to a pre-defined location. Several pre-defined locations were tried, but there was failure at each location. In between sending the arm to each pre-defined location, the representative toggled through service tools and send surgical arm. A 3836 error continued to be displayed so a second hard restart was done. After booting up the syst em and logging in as a superuser again, the pre-defined draping position was selected. The surgical arm was sent to the location, but the joint 3 error appeared again and the arm did not fully stretch into position. The representative tried going into the advanced setup options to try and move each going separately to get the arm into draping position. All of the other joints were moved and then joint 3 was tried, but nothing happened. At this point the arm was manually moved into the draping position. The representative tried to send the arm to another pre-defined location, but the 3836 error code was displayed. The surgeon decided to abort the use of the guidance system and continue the procedure without the system. There was no patient harm and the procedure was not delayed more than an hour. 2019-10-03 IUD (Other): No new information. 2019-10-09 email 2019-10-09 08:30 (Rep): Additional information received from a manufacturer representative reported that the event occurred prior to the operation and the prior to the patient being in the operating room. The guidance system was aborted and the procedure was completed with navigation. 2022-02-19 IUD (Other): No new information. SURGICAL ARM | 2019-09-25 SCL1902915, Intake Form x2, 00674880, ID, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the guidance system was booted up and after logging in, a joint 3 shift error was displayed when starting the setup process. The error message was closed and the representative attempted to proceed, but the error code "3836 Failed to Reach Defined Position" was displayed. A soft restart of the system was done, but the issue was not resolved. A hard restart of the system was done and the representative logged into the system as a supersuser. Once logged in, the movement was reset and the surgical arm was sent to a pre-defined location. Several pre-defined locations were tried, but there was failure at each location. In between sending the arm to each pre-defined location, the representative toggled through service tools and send surgical arm. A 3836 error continued to be displayed so a second hard restart was done. After booting up the syst em and logging in as a superuser again, the pre-defined draping position was selected. The surgical arm was sent to the location, but the joint 3 error appeared again and the arm did not fully stretch into position. The representative tried going into the advanced setup options to try and move each going separately to get the arm into draping position. All of the other joints were moved and then joint 3 was tried, but nothing happened. At this point the arm was manually moved into the draping position. The representative tried to send the arm to another pre-defined location, but the 3836 error code was displayed. The surgeon decided to abort the use of the guidance system and continue the procedure without the system. There was no patient harm and the procedure was not delayed more than an hour. 2019-10-03 IUD (Other): No new information. 2019-10-09 email 2019-10-09 08:30 (Rep): Additional information received from a manufacturer representative reported that the event occurred prior to the operation and the prior to the patient being in the operating room. The guidance system was aborted and the procedure was completed with navigation. 2022-02-19 IUD (Other): No new information. | xse joint shift error id iud spinal procedure booted logging joint shift error displayed starting setup process error message closed attempted proceed error code failed reach defined position displayed soft restart issue resolved hard restart logged supersuser logged movement reset surgical arm sent predefined location predefined locations tried failure location sending arm predefined location toggled service tools send surgical arm error continued displayed second hard restart booting syst em logging superuser predefined draping position selected surgical arm sent location joint error appeared arm did fully stretch position tried going advanced setup options try going separately arm draping position joints moved joint tried happened point arm manually moved draping position tried send arm predefined location error code displayed surgeon decided abort use continue procedure procedure delayed hour iud new email additional occurred prior operation prior operating room aborted procedure completed navigation iud new surgical arm | NAV4095 | NAV4098 | NAV4101 | NAV2047 | NAV1006 | NAV4095NAV4098NAV4101NAV2047NAV1006 | True
4987 | 703589033 | X - Deviation Right L5 | 2020-01-30T17:02:46.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation Right L5 2018-02-13 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L3-L5. It was reported that All screws were placed according to plan except for Right L5. Right L5 was deviated inferior due to skiving. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-02-13 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L3-L5. It was reported that All screws were placed according to plan except for Right L5. Right L5 was deviated inferior due to skiving. There was no patient harm and the procedure was not delayed over an hour. | x deviation right ccr spinal procedure treat degenerative disk disease fixation screws placed according plan right right deviated inferior skiving procedure delayed hour mazor x | NAV4091 | NAV5000 | NAV3242 | NAV3247 | NAV2047 | NAV4091NAV5000NAV3242NAV3247NAV2047 | True
1633 | 703208225 | X -positioner locking button not working | 2019-06-14T18:06:31.000000+0000 | MazorX | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 20 | CAR0066-01 | POSITIONER CONNECTORS CARD | X -positioner locking button not working 2019-Jun-10 SCL1901705 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the only way to mount the system was to use Ctr-Alt-O to mount the system using the force tooth open and close. The shoulder would lock but after some time it lost its position and it was saying it neither open or closed.There was no patient harm. There were no additional complications reported or anticipated as a result of the event. POSITIONER CONNECTORS CARD | 2019-Jun-10 SCL1901705 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the only way to mount the system was to use Ctr-Alt-O to mount the system using the force tooth open and close. The shoulder would lock but after some time it lost its position and it was saying it neither open or closed.There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x positioner locking button working spinal procedure way mount use ctralto mount using force tooth open close shoulder lock time lost position saying open closedthere additional complications anticipated positioner connectors card | NAV4107 | NAV2053 | NAV3251 | NAV3250 | NAV4103 | NAV4107NAV2053NAV3251NAV3250NAV4103 | True
6617 | 703989503 | X - X-Eye Threads Stripped/Slow SW | 2020-10-01T20:06:05.000000+0000 | MazorX | NAV2056 | SOFTWARE ISSUES | 30 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | X - X-Eye Threads Stripped/Slow SW 2020-10-01 00697868 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the reference marker threads were stripped and the software was moving slowly. There was no patient involvement. 2020-10-12 IUD x8 (Rep): No new information. 2020-12-13 IUD x2 (Other): No new information. 2020-12-22 IUD x3 (Other): No new information. 2021-06-20 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2020-10-01 00697868 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the reference marker threads were stripped and the software was moving slowly. There was no patient involvement. 2020-10-12 IUD x8 (Rep): No new information. 2020-12-13 IUD x2 (Other): No new information. 2020-12-22 IUD x3 (Other): No new information. 2021-06-20 IUD (Other): No new information. | x xeye threads strippedslow sw outside procedure reference marker threads stripped software moving slowly iud new iud new iud new iud new mazor x pc station refurbished | NAV3262 | NAV3246 | NAV4100 | NAV4110 | NAV2054 | NAV3262NAV3246NAV4100NAV4110NAV2054 | False
5418 | 703677468 | XSE - Registration Difficulties | 2020-03-20T22:22:41.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Registration Difficulties 2020-03-20 00686676 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon and the representative planned the procedure the morning of the case on the workstation. Skin incisions and rod alignment were taken into consideration for the MIS procedure. Intraoperatively the surgeon chose to go open due to his decompression and inserting bone graft. The representative and surgeon discussed using two schanz arms for the mounting of the construct. One of the trays was missing and the loaner only contained one schanz arm so one schanz arm was used for the lower half of the construct and one schanz arm with a disposable clamp for the top portion of the construct. C-arm problems caused a delay in the registration process as well as a software crash while changing c-arms. Once the software came back, every step of registration was red one including taking AP images. When sending the arm to the oblique, the c-arm tech was not taking instruction and the representative hit the emergency button, but it was too late, and the c-arm hit the arm causing a shoulder shift. The third registration posed difficulties with segmentation after many attempts T10-L1 were the only bodies that were green in the first segment. All trajectories were instrumented and confirmed with fluoro. Schanz arm was extended straight so the arm had room to send to S2. Bone mount bridge was removed and a tegaderm was placed over the drape. The representative achieved registration and all values were green but L2. L2 had the L1 screw over the body in the fluoro shot and I believe that to be the reason for the yellow value. All trajectories were instrumented and confirmed with fluoro. There was no patient harm and the procedure was not delayed over an hour. 2020-03-26 IUD (Other): No new information. MAZOR X SYSTEM TPL0059 | 2020-03-20 00686676 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon and the representative planned the procedure the morning of the case on the workstation. Skin incisions and rod alignment were taken into consideration for the MIS procedure. Intraoperatively the surgeon chose to go open due to his decompression and inserting bone graft. The representative and surgeon discussed using two schanz arms for the mounting of the construct. One of the trays was missing and the loaner only contained one schanz arm so one schanz arm was used for the lower half of the construct and one schanz arm with a disposable clamp for the top portion of the construct. C-arm problems caused a delay in the registration process as well as a software crash while changing c-arms. Once the software came back, every step of registration was red one including taking AP images. When sending the arm to the oblique, the c-arm tech was not taking instruction and the representative hit the emergency button, but it was too late, and the c-arm hit the arm causing a shoulder shift. The third registration posed difficulties with segmentation after many attempts T10-L1 were the only bodies that were green in the first segment. All trajectories were instrumented and confirmed with fluoro. Schanz arm was extended straight so the arm had room to send to S2. Bone mount bridge was removed and a tegaderm was placed over the drape. The representative achieved registration and all values were green but L2. L2 had the L1 screw over the body in the fluoro shot and I believe that to be the reason for the yellow value. All trajectories were instrumented and confirmed with fluoro. There was no patient harm and the procedure was not delayed over an hour. 2020-03-26 IUD (Other): No new information. | xse registration difficulties spinal procedure surgeon planned procedure morning case workstation skin incisions rod alignment taken consideration mis procedure intraoperatively surgeon chose open decompression inserting bone graft surgeon discussed using schanz arms mounting construct trays missing loaner contained schanz arm schanz arm lower half construct schanz arm disposable clamp portion construct carm problems caused delay registration process software crash changing carms software came step registration red including taking ap images sending arm oblique carm tech taking instruction hit emergency button late carm hit arm causing shoulder shift registration posed difficulties segmentation attempts bodies green segment trajectories instrumented confirmed fluoro schanz arm extended straight arm room send bone mount bridge removed tegaderm placed drape achieved registration values green screw body fluoro shot believe reason yellow value trajectories instrumented confirmed fluoro procedure delayed hour iud new mazor x | NAV3242 | NAV3247 | NAV2053 | NAV4108 | NAV4112 | NAV3242NAV3247NAV2053NAV4108NAV4112 | True
4395 | 703540685 | XSE - Bed Frame Adapter Cannot Lock | 2020-01-02T11:57:30.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 90 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Bed Frame Adapter Cannot Lock 2019-12-21 00681221 (REP, FOR): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the bedframe adapter could not be properly locked. The surgical system was moving downward more than usual when pressing the manipulating button on the side of the locking plate and releasing it. The positioner was tilting to the heavier side (toward the arm’s weight) and could not be held firmly when the shoulder was manually moved to left or right. The software crashed a few times during the workshop. The 3Define camera unexpectedly stopped during taking the volume information. It worked properly after retrying. The software is lagging during making plans. The arm is oddly shaking after it reaches to planned trajectory. Two pin screws on the side of USB in the below cabinet (where the main cable is stored) are not properly going inside and therefore the connection is very loosened. The surgeon monitor is flipped in upside down direction. Attempted to fix it in Window’s Control Panel, but there is only one monitor (the main monitor) being recognized by the software. Flipping happened after trying to output video to the external monitor. No patient was present. 2020-01-05 IUD x3, email 2020-03-05 23:15, image x2, video x2 (Rep, For): Additional information received from a manufacturer representative reported that the bedframe adapter was replaced to resolve the issue. 2020-03-19 IUD, Email 2020-03-19 04:30 (Rep, Other): No new information. UNK MAZOR X WORKSTATION | 2019-12-21 00681221 (REP, FOR): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that the bedframe adapter could not be properly locked. The surgical system was moving downward more than usual when pressing the manipulating button on the side of the locking plate and releasing it. The positioner was tilting to the heavier side (toward the arm’s weight) and could not be held firmly when the shoulder was manually moved to left or right. The software crashed a few times during the workshop. The 3Define camera unexpectedly stopped during taking the volume information. It worked properly after retrying. The software is lagging during making plans. The arm is oddly shaking after it reaches to planned trajectory. Two pin screws on the side of USB in the below cabinet (where the main cable is stored) are not properly going inside and therefore the connection is very loosened. The surgeon monitor is flipped in upside down direction. Attempted to fix it in Window’s Control Panel, but there is only one monitor (the main monitor) being recognized by the software. Flipping happened after trying to output video to the external monitor. No patient was present. 2020-01-05 IUD x3, email 2020-03-05 23:15, image x2, video x2 (Rep, For): Additional information received from a manufacturer representative reported that the bedframe adapter was replaced to resolve the issue. 2020-03-19 IUD, Email 2020-03-19 04:30 (Rep, Other): No new information. | xse bed frame adapter lock outside procedure bedframe adapter properly locked surgical moving downward usual pressing manipulating button locking plate releasing positioner tilting heavier held firmly shoulder manually moved left right software crashed times workshop camera unexpectedly stopped taking volume worked properly retrying software lagging making plans arm oddly shaking reaches planned trajectory pin screws usb cabinet properly going inside connection loosened surgeon monitor flipped upside direction attempted fix window’s control panel monitor recognized software flipping happened trying output video external monitor present iud email image video additional bedframe adapter replaced resolve issue iud email new unk mazor x workstation | NAV3255 | NAV4101 | NAV3251 | NAV3247 | NAV3244 | NAV3255NAV4101NAV3251NAV3247NAV3244 | False
4670 | 703561289 | X - Registration issues | 2020-01-14T17:15:25.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration issues 2018-10-25 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked, an advanced accuracy test was completed, and the surgical arm was setup. The schanz pin was placed in the right PSIS and connected to a BMB. Registration was achieved on the second attempt with green values at all surgical levels. All screws were placed and verified with AP and lateral images. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-10-25 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the shoulder calibration was checked, an advanced accuracy test was completed, and the surgical arm was setup. The schanz pin was placed in the right PSIS and connected to a BMB. Registration was achieved on the second attempt with green values at all surgical levels. All screws were placed and verified with AP and lateral images. There was no patient harm and the procedure was not delayed over an hour. | x registration issues ccr spinal procedure shoulder calibration checked advanced accuracy test completed surgical arm setup schanz pin placed right psis connected bmb registration achieved second attempt green values surgical levels screws placed verified ap lateral images procedure delayed hour mazor x | NAV3242 | NAV4091 | NAV3246 | NAV2047 | NAV4112 | NAV3242NAV4091NAV3246NAV2047NAV4112 | True
5294 | 703641877 | XSE - Joint 5 Shift | 2020-02-28T16:14:00.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Joint 5 Shift 2020-02-28 00685362 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that there was a Joint 5 shift on the surgical arm. No patient was present. 2020-03-02 Email 2020-03-02 16:00 (Rep): Additional information received from a manufacturer representative reported the cause of the Joint 5 arm shift is presumed to be shipping as the error came up as soon as the workstation was powered on. The representative “reset movement” in Superuser mode/Utilities, then did an advanced accuracy test and this seemed to fix the problem. The representative wasn’t with the system for the entire demo event, but did not hear that this error message came up again. KIT SVC ASM0206-01S SURGICAL ARM | 2020-02-28 00685362 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that there was a Joint 5 shift on the surgical arm. No patient was present. 2020-03-02 Email 2020-03-02 16:00 (Rep): Additional information received from a manufacturer representative reported the cause of the Joint 5 arm shift is presumed to be shipping as the error came up as soon as the workstation was powered on. The representative “reset movement” in Superuser mode/Utilities, then did an advanced accuracy test and this seemed to fix the problem. The representative wasn’t with the system for the entire demo event, but did not hear that this error message came up again. | xse joint shift outside procedure joint shift surgical arm present email additional cause joint arm shift presumed shipping error came soon workstation powered “reset movement” superuser modeutilities did advanced accuracy test fix problem wasn’t entire demo did hear error message came kit svc surgical arm | NAV4098 | NAV2054 | NAV4108 | NAV4093 | NAV4101 | NAV4098NAV2054NAV4108NAV4093NAV4101 | True
8603 | 704448280 | XSE - 3Define Scan Failed/Pressure Issue | 2021-06-14T19:48:10.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - 3Define Scan Failed/Pressure Issue 2021-Jun-14 Salesforce C00228317, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define scan failed twice during a L5-S1 ALIF procedure. Troubleshooting involved two hard reboots of the system. After rebooting the system, the pneumatic system was no longer working. The manufacturer representative indicated there was a clicking noise, but the pneumatics were not fully turning on. The pneumatic system did turn on for a second, but it was not enough to move the system. A generic work volume was to continue the procedure. There was no patient harm and the procedure was delayed less than an hour. 2021-06-16 IUD x4 (Other): No new information. 2021-07-15 IUD (Other): No new information. 2021-07-20 IUD, e6, CCR (Rep): No new information. 2021-07-27 IUD x4 (Other): No new information. 2021-Aug-26 IUDx2 (other): No new information. UNK MAZOR X WORKSTATION | 2021-Jun-14 Salesforce C00228317, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define scan failed twice during a L5-S1 ALIF procedure. Troubleshooting involved two hard reboots of the system. After rebooting the system, the pneumatic system was no longer working. The manufacturer representative indicated there was a clicking noise, but the pneumatics were not fully turning on. The pneumatic system did turn on for a second, but it was not enough to move the system. A generic work volume was to continue the procedure. There was no patient harm and the procedure was delayed less than an hour. 2021-06-16 IUD x4 (Other): No new information. 2021-07-15 IUD (Other): No new information. 2021-07-20 IUD, e6, CCR (Rep): No new information. 2021-07-27 IUD x4 (Other): No new information. 2021-Aug-26 IUDx2 (other): No new information. | xse scan failedpressure issue salesforce iud spinal procedure scan failed twice alif procedure troubleshooting involved hard reboots rebooting pneumatic longer working indicated clicking noise pneumatics fully turning pneumatic did turn second generic work volume continue procedure procedure delayed hour iud new iud new iud ccr new iud new new unk mazor x workstation | NAV3264 | NAV3244 | NAV4095 | NAV4104 | NAV2058 | NAV3264NAV3244NAV4095NAV4104NAV2058 | True
5717 | 703787907 | XSE - Joint 1 Failure | 2020-06-05T14:21:52.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Joint 1 Failure 2020-06-05 00690193, IUD (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that when going to do a case it was noted that the system had failed at joint 1 and the case had to be cancelled. No patient was present. 2020-06-11 IUD x2 (Other): No new information. 2020-06-29 IUD x2 (Other): No new information. 2020-07-15 IUD (Other): No new information. 2020-07-29 IUD (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-06-05 00690193, IUD (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that when going to do a case it was noted that the system had failed at joint 1 and the case had to be cancelled. No patient was present. 2020-06-11 IUD x2 (Other): No new information. 2020-06-29 IUD x2 (Other): No new information. 2020-07-15 IUD (Other): No new information. 2020-07-29 IUD (Other): No new information. | xse joint failure iud outside procedure going case noted failed joint case cancelled present iud new iud new iud new iud new kit svc surgical arm | NAV3249 | NAV2054 | NAV3264 | NAV4098 | NAV2047 | NAV3249NAV2054NAV3264NAV4098NAV2047 | True
8195 | 704369162 | XSE - SRA Locked Up | 2021-04-29T20:24:13.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 30 | ASM0206-03S | SURGICAL ARM SPARE PART ASM0206-03S | XSE - SRA Locked Up 2021-Apr-29 Salesforce C00225527 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgical arm became locked during the procedure. An error message was displayed stating that the surgical arm was not in the oblique position and will not accept the image. The surgical arm would not move at all at this point. When moving from the oblique position to the AP position, there was a joint shift that pushed the system up an entire level in the AP view. The guidance system was pulled from the procedure and an accuracy test was done. Everything on the left side was on a trajectory that would hit the lateral wall. A stress test was started and the surgical arm would stop for 10 seconds and then move to the next position. There was no patient harm and the procedure was delayed less than an hour. 2021-04-30 IUD x4 (Other): No new information. 2021-05-20 IUD x2 (Other): No new information. 2021-06-24 IUD (Other): No new information. 2021-08-10 302523402 (Other): No new information. 2021-08-12 IUD (Other): No new information. 2021-08-15 IUD x2 (Other): No new information. 2021-09-20 IUD (Other): No new information. 2021-09-26 IUD x5 (Other): No new information. SURGICAL ARM SPARE PART ASM0206-03S | 2021-Apr-29 Salesforce C00225527 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgical arm became locked during the procedure. An error message was displayed stating that the surgical arm was not in the oblique position and will not accept the image. The surgical arm would not move at all at this point. When moving from the oblique position to the AP position, there was a joint shift that pushed the system up an entire level in the AP view. The guidance system was pulled from the procedure and an accuracy test was done. Everything on the left side was on a trajectory that would hit the lateral wall. A stress test was started and the surgical arm would stop for 10 seconds and then move to the next position. There was no patient harm and the procedure was delayed less than an hour. 2021-04-30 IUD x4 (Other): No new information. 2021-05-20 IUD x2 (Other): No new information. 2021-06-24 IUD (Other): No new information. 2021-08-10 302523402 (Other): No new information. 2021-08-12 IUD (Other): No new information. 2021-08-15 IUD x2 (Other): No new information. 2021-09-20 IUD (Other): No new information. 2021-09-26 IUD x5 (Other): No new information. | xse sra locked salesforce spinal p rocedure surgical arm locked procedure error message displayed stating surgical arm oblique position accept image surgical arm point moving oblique position ap position joint shift pushed entire level ap view pulled procedure accuracy test left trajectory hit lateral wall stress test started surgical arm stop seconds position procedure delayed hour iud new iud new iud new new iud new iud new iud new iud new surgical arm spare | NAV4101 | NAV4095 | NAV4092 | NAV3249 | NAV3243 | NAV4101NAV4095NAV4092NAV3249NAV3243 | False
9200 | 704574311 | XSE ¦ Worn Out Cover | 2021-08-24T21:47:12.000000+0000 | MazorX Stealth Edition | NAV4099 | PLASTIC COVER ISSUES | 20 | MEC4274-01S | KIT SVC MEC4274-01S JOINT 3 LEFT COVER | XSE ¦ Worn Out Cover 2021-Aug-24 Salesforce C00233426 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the cover on the left of joint is warn out and no longer stays in place. There was no patient involved. 2021-Aug-24 IUDx2 (other): No new information. KIT SVC MEC4274-01S JOINT 3 LEFT COVER | 2021-Aug-24 Salesforce C00233426 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the cover on the left of joint is warn out and no longer stays in place. There was no patient involved. 2021-Aug-24 IUDx2 (other): No new information. | xse worn cover salesforce outside procedure cover left joint warn longer stays place involved new kit svc joint left cover | NAV4099 | NAV2054 | NAV3299 | NAV2057 | NAV2058 | NAV4099NAV2054NAV3299NAV2057NAV2058 | True
139 | 703014763 | X - helicoils fell out | 2019-02-19T14:51:13.000000+0000 | MazorX | NAV4100 | STRIPPED THREADS | 20 | ASM0206-01 | SURGICAL ARM | X - helicoils fell out 2019-Feb-11 SCL1900358, intake form (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the helicoils on the side of the arm where the reference marker screws in fell out. It was noted there was no patient impact and no surgical delay. There were no additional complications reported or anticipated as a result of the event. SURGICAL ARM | 2019-Feb-11 SCL1900358, intake form (HCP, rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the helicoils on the side of the arm where the reference marker screws in fell out. It was noted there was no patient impact and no surgical delay. There were no additional complications reported or anticipated as a result of the event. | x helicoils fell spinal procedure helicoils arm reference marker screws fell noted impact surgical delay additional complications anticipated surgical arm | NAV3249 | NAV4101 | NAV2054 | NAV4100 | NAV5000 | NAV3249NAV4101NAV2054NAV4100NAV5000 | True
10110 | 704801451 | XSE - Detached Surgeon Screen | 2022-01-10T20:44:20.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 30 | ASM0357 | KIT SVC ASM0357 SRGN SCRN W/ARM 02S NEW | XSE - Detached Surgeon Screen 2022-Jan-10 Salesforce C00241032 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon screen detached from the mount arm due to normal wear and tear. The Medtronic team has tightened it down multiple times but it keeps coming loose. No patient was present. 2022-01-11 IUD x2 (Other): No new information. 2022-01-21 IUD (Other): No new information. 2022-02-24 IUD (Other): No new information. 2022-03-23 IUD (Other): No new information. KIT SVC ASM0357 SRGN SCRN W/ARM 02S NEW | 2022-Jan-10 Salesforce C00241032 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgeon screen detached from the mount arm due to normal wear and tear. The Medtronic team has tightened it down multiple times but it keeps coming loose. No patient was present. 2022-01-11 IUD x2 (Other): No new information. 2022-01-21 IUD (Other): No new information. 2022-02-24 IUD (Other): No new information. 2022-03-23 IUD (Other): No new information. | xse detached surgeon screen salesforce outside procedure surgeon screen detached mount arm normal wear tear team tightened multiple times keeps coming loose present iud new iud new iud new iud new kit svc srgn scrn warm new | NAV3253 | NAV2058 | NAV2054 | NAV2047 | NAV3299 | NAV3253NAV2058NAV2054NAV2047NAV3299 | True
9339 | 704612845 | XSE ¦ X-Eye Camera Not Functioning | 2021-09-15T21:58:11.000000+0000 | MazorX Stealth Edition | NAV3254 | X-EYE ISSUES | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | XSE ¦ X-Eye Camera Not Functioning 2021-Sep-15 Salesforce C00234510 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the x-eye camera is not functioning. Reboots do not resolve and he tried to reseat the firewire cable. No patient present. 2021-09-17 IUD x2 (Other): No new information. UNK MAZOR X X-EYE CAMERA | 2021-Sep-15 Salesforce C00234510 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the x-eye camera is not functioning. Reboots do not resolve and he tried to reseat the firewire cable. No patient present. 2021-09-17 IUD x2 (Other): No new information. | xse xeye camera functioning salesforce outside procedure xeye camera functioning reboots resolve tried reseat firewire cable present iud new unk mazor x xeye camera | NAV4110 | NAV3254 | NAV2058 | NAV3256 | NAV3261 | NAV4110NAV3254NAV2058NAV3256NAV3261 | True
464 | 703111346 | X - Inaccurate Trajectories (Case 2) | 2019-04-17T18:26:44.000000+0000 | MazorX | NAV4091 | DEVIATION | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Inaccurate Trajectories (Case 2) 2019-04-16 SCL1900969, Intake Form, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the trajectories were inaccurate and clearly low. The cause of the inaccurate trajectories was unknown. Registration was completed successfully in a quick manor with all green values and there was no contact with the patient prior to inserting instruments. There was no patient harm and the procedure was not delayed more than an hour. There was a change in the course of the procedure, but the case was able to be completed. Information unrelated to this event was included in a separate PE: See PE 703111355 for Case 1. 2020-01-29 e2, e3 (Rep): Additional information received from a manufacturer representative reported the issue was resolved. MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-04-16 SCL1900969, Intake Form, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the trajectories were inaccurate and clearly low. The cause of the inaccurate trajectories was unknown. Registration was completed successfully in a quick manor with all green values and there was no contact with the patient prior to inserting instruments. There was no patient harm and the procedure was not delayed more than an hour. There was a change in the course of the procedure, but the case was able to be completed. Information unrelated to this event was included in a separate PE: See PE 703111355 for Case 1. 2020-01-29 e2, e3 (Rep): Additional information received from a manufacturer representative reported the issue was resolved. | x inaccurate trajectories spinal procedure trajectories inaccurate clearly low cause inaccurate trajectories unknown registration completed successfully quick manor green values contact prior inserting instruments procedure delayed hour change course procedure case able completed unrelated included separate pe pe case additional issue resolved mazor x software kit | NAV4091 | NAV3247 | NAV3248 | NAV3242 | NAV3246 | NAV4091NAV3247NAV3248NAV3242NAV3246 | True
8475 | 704421639 | XSE - SRA Movement Issues | 2021-05-28T18:58:34.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - SRA Movement Issues 2021-May-27 Salesforce C00227489, IUD x2, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were arm movement issues in the case. When doing the surgical arm setup, the arm would get halfway through and stop. The monitor showed that it was still moving and was making the noise. A soft restart was completed and they were able to continue. They did another soft restart and same thing. They went to the main menu, sent to storage, system moved to storage, but did not give the confirmation noise that it completed movement. They did a hard restart, went to setup arm and it went through successfully. They mounted the system to the bed. They then went to do 3Define and the arm moved but the software never identified that the 3Define scan was completed. They did not have the option to try a generic default work volume. They aborted the use of the system. They ended up switching to a navigation system/imaging system The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-06-03 IUD x4 (Other): No new information. 2021-06-28 IUD (Other): No new information. 2021-07-27 IUD x3 (Other): No new information. 2021-08-30 302544137, IUD x3 (Other): No new information. 2022-01-05 IUD x2 (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2021-May-27 Salesforce C00227489, IUD x2, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that there were arm movement issues in the case. When doing the surgical arm setup, the arm would get halfway through and stop. The monitor showed that it was still moving and was making the noise. A soft restart was completed and they were able to continue. They did another soft restart and same thing. They went to the main menu, sent to storage, system moved to storage, but did not give the confirmation noise that it completed movement. They did a hard restart, went to setup arm and it went through successfully. They mounted the system to the bed. They then went to do 3Define and the arm moved but the software never identified that the 3Define scan was completed. They did not have the option to try a generic default work volume. They aborted the use of the system. They ended up switching to a navigation system/imaging system The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-06-03 IUD x4 (Other): No new information. 2021-06-28 IUD (Other): No new information. 2021-07-27 IUD x3 (Other): No new information. 2021-08-30 302544137, IUD x3 (Other): No new information. 2022-01-05 IUD x2 (Other): No new information. | xse sra movement issues salesforce iud ccr spinal procedure arm movement issues case doing surgical arm setup arm halfway stop monitor showed moving making noise soft restart completed able continue did soft restart thing went main menu sent storage moved storage did confirmation noise completed movement did hard restart went setup arm went successfully mounted bed went arm moved software identified scan completed did option try generic default work volume aborted use ended switching navigation systemimaging issue extended surgical time hour impact outcome iud new iud new iud new iud new iud new mazor x pc station refurbished | NAV4095 | NAV3262 | NAV3246 | NAV3248 | NAV3249 | NAV4095NAV3262NAV3246NAV3248NAV3249 | True
240 | 703048003 | XSE - Software Issues | 2019-03-12T01:31:05.000000+0000 | MazorX Stealth Edition | NAV4110 | NO TRACKING | 40 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | XSE - Software Issues 2019-Mar- 06 SCL1900538 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that upon startup, it was noticed there was no x-eye functionality. A soft re-start was attempted with no resolution. The rep called the engineer and it was noted that a hard re-start might be required. Hard re-start was attempted, and it appeared to resolve the issues. The software was freezing and moving slowly throughout setup. It was noticed by the rep that in the accuracy setup window there was no option to send the robot to storage. The case was started and upon sending for surface scan, it was noted that software wanted the navigation code. The rep had to move forward with generic work volume. Once spin was completed, the rep and HCP planned the screws on the software and all trajectories were presented as reachable via blue and yellow coloring. When attempting to send L2 and L1 right, the software presented an “unreachable solution.” It was noted that this was a large hassle because the rep had to continuously toggle between planning and operation and there was no indication that screws couldn’t be reached until attempting to send the trajectory. This slowed the rep and HCP down roughly ten minutes. When loading trajectories, it would take roughly 45 seconds to load the new pathway. It was noted that this was vastly slower than old software versions. Overall the software froze several times when toggling between screens. On four trajectories, RTC was presented with error message “movement calculation is taking longer than expected, do you wish to abort?” It was noted that the screw stripped out of the peek bone mount bridge that was noticed upon opening of trays. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-03-20 e1, e2 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a procedure using the k-wire approach, the system showed slow performance and had small bugs. The case was able to be completed successfully. Information unrelated to this event was included in three separate PEs; See PE 703076216 for XSE - Camera Issues, 703066185 for XSE - No Pressure, and 703076180 for XSE - Software Crash. 2019-03-06 e3, e4 (Rep): No new information. Information unrelated to this event was included in three separate PEs; See PE 703076216 for XSE - Camera Issues, 703066185 for XSE - No Pressure, and 703076180 for XSE - Software Crash. 2019-04-22 e7 (Rep): Additional information received from a manufacturer representative reported the issue had been resolved and no further information was available. UNK MAZOR X X-EYE CAMERA | 2019-Mar- 06 SCL1900538 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that upon startup, it was noticed there was no x-eye functionality. A soft re-start was attempted with no resolution. The rep called the engineer and it was noted that a hard re-start might be required. Hard re-start was attempted, and it appeared to resolve the issues. The software was freezing and moving slowly throughout setup. It was noticed by the rep that in the accuracy setup window there was no option to send the robot to storage. The case was started and upon sending for surface scan, it was noted that software wanted the navigation code. The rep had to move forward with generic work volume. Once spin was completed, the rep and HCP planned the screws on the software and all trajectories were presented as reachable via blue and yellow coloring. When attempting to send L2 and L1 right, the software presented an “unreachable solution.” It was noted that this was a large hassle because the rep had to continuously toggle between planning and operation and there was no indication that screws couldn’t be reached until attempting to send the trajectory. This slowed the rep and HCP down roughly ten minutes. When loading trajectories, it would take roughly 45 seconds to load the new pathway. It was noted that this was vastly slower than old software versions. Overall the software froze several times when toggling between screens. On four trajectories, RTC was presented with error message “movement calculation is taking longer than expected, do you wish to abort?” It was noted that the screw stripped out of the peek bone mount bridge that was noticed upon opening of trays. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-03-20 e1, e2 (Rep): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a procedure using the k-wire approach, the system showed slow performance and had small bugs. The case was able to be completed successfully. Information unrelated to this event was included in three separate PEs; See PE 703076216 for XSE - Camera Issues, 703066185 for XSE - No Pressure, and 703076180 for XSE - Software Crash. 2019-03-06 e3, e4 (Rep): No new information. Information unrelated to this event was included in three separate PEs; See PE 703076216 for XSE - Camera Issues, 703066185 for XSE - No Pressure, and 703076180 for XSE - Software Crash. 2019-04-22 e7 (Rep): Additional information received from a manufacturer representative reported the issue had been resolved and no further information was available. | xse software issues spinal procedure startup noticed xeye functionality soft restart attempted resolution called engineer noted hard restart required hard restart attempted appeared resolve issues software freezing moving slowly setup noticed accuracy setup window option send robot storage case started sending surface scan noted software wanted navigation code forward generic work volume spin completed hcp planned screws software trajectories presented reachable blue yellow coloring attempting send right software presented “unreachable solution” noted large hassle continuously toggle planning operation indication screws couldn’t reached attempting send trajectory slowed hcp roughly minutes loading trajectories roughly seconds load new pathway noted vastly slower old software versions overall software froze times toggling screens trajectories rtc presented error message “movement calculation taking longer expected wish abort” noted screw stripped peek bone mount bridge noticed opening trays noted additional complications anticipated spinal procedure procedure using kwire approach showed slow performance small bugs case able completed successfully unrelated included separate pes pe xse camera issues xse pressure xse software crash new unrelated included separate pes pe xse camera issues xse pressure xse software crash additional issue resolved available unk mazor x xeye camera | NAV4095 | NAV3299 | NAV4110 | NAV3248 | NAV3262 | NAV4095NAV3299NAV4110NAV3248NAV3262 | True
3792 | 703443873 | XSE - Failed to Reach Destination | 2019-10-31T16:02:27.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | ASM0206-01 | SURGICAL ARM | XSE - Failed to Reach Destination 2019-Oct-31 00677306 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a case, while sending the arm to clear up he received an message that "failed to reach destination" . After the 3rd attempt to clear the arm, it reached its destination. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-12 IUD x2 (Other): No new information. SURGICAL ARM | 2019-Oct-31 00677306 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that during a case, while sending the arm to clear up he received an message that "failed to reach destination" . After the 3rd attempt to clear the arm, it reached its destination. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-12 IUD x2 (Other): No new information. | xse failed reach destination spinal procedure case sending arm clear message failed reach destination attempt clear arm reached destination additional complications anticipated iud new surgical arm | NAV4095 | NAV3299 | NAV3295 | NAV3264 | NAV5000 | NAV4095NAV3299NAV3295NAV3264NAV5000 | True
7905 | 704304353 | XSE - Deviations | 2021-03-25T16:04:32.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviations 2021-Mar-24 Salesforce C00223232, Image 2021-03-28 13:30 x2, Image 2021-03-28 13:31 x2, CCR, e1, e2, e3, e4 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that screws were planned for L4, L5, and S1 during a spondylptosis case. A dual clamp was used to mount the system to the patient. There was one minor challenge with registration. The surgeon could only get both anatomy and marker around 75 degrees from the AP position and not the ideal 60 degrees. The surgical arm was sent to the first trajectory at left L4. The dilator was slowly placed down the arm guide and an image was taken to confirm the location once bone was hit. The surgical arm was inferior when compared to plan. The plan was adjusted and the entry point was made slightly superior and a superior angulation was given. The surgical was was still not able to reach the desired superior position. A second registration was done, but the same issue occurred. The manufacturer representative thought the issue was due to the way the C-arm image was taken and not being able to do a true lateral/oblique image so the end plates did not seem parallel. The trajectory was not drilled. There was a scope of skiving at the traje ctory and the surgeon was aware. The surgeon had planned to flatten the entry point surface before drilling, but the surgeon did not proceed. The case was an open procedure do avoid skiving. The medial angulation was clos to 30 degrees and soft tissue pressure could affect the trajectory, the surgeon created an additional incision to avoid soft tissue pressure. No force was exerted on the instruments when checking the entry point to avoid skiving. The surgeon moved onto right L4 and the position was confirmed with the C-arm. The trajectory was ok in the superior/inferior direction, but seemed lateral. The surgeon thought this was because of the AP image, but they were sure if it was a visualization issue or if the trajectory was lateral. The surgeon decided to abort the use of the guidance system and place the screws freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-04-07 email 2021-04-07 06:16 (Rep): Additional information received from a manufacturer representative reported that the trajectories were deviated less than 3.5 mm. 2021-05-09 IUD (Rep): Additional information received from a manufacturer representative reported the suspected cause was a planning issue and software functionality. MAZOR X SYSTEM TPL0059 | 2021-Mar-24 Salesforce C00223232, Image 2021-03-28 13:30 x2, Image 2021-03-28 13:31 x2, CCR, e1, e2, e3, e4 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with fixation. It was reported that screws were planned for L4, L5, and S1 during a spondylptosis case. A dual clamp was used to mount the system to the patient. There was one minor challenge with registration. The surgeon could only get both anatomy and marker around 75 degrees from the AP position and not the ideal 60 degrees. The surgical arm was sent to the first trajectory at left L4. The dilator was slowly placed down the arm guide and an image was taken to confirm the location once bone was hit. The surgical arm was inferior when compared to plan. The plan was adjusted and the entry point was made slightly superior and a superior angulation was given. The surgical was was still not able to reach the desired superior position. A second registration was done, but the same issue occurred. The manufacturer representative thought the issue was due to the way the C-arm image was taken and not being able to do a true lateral/oblique image so the end plates did not seem parallel. The trajectory was not drilled. There was a scope of skiving at the traje ctory and the surgeon was aware. The surgeon had planned to flatten the entry point surface before drilling, but the surgeon did not proceed. The case was an open procedure do avoid skiving. The medial angulation was clos to 30 degrees and soft tissue pressure could affect the trajectory, the surgeon created an additional incision to avoid soft tissue pressure. No force was exerted on the instruments when checking the entry point to avoid skiving. The surgeon moved onto right L4 and the position was confirmed with the C-arm. The trajectory was ok in the superior/inferior direction, but seemed lateral. The surgeon thought this was because of the AP image, but they were sure if it was a visualization issue or if the trajectory was lateral. The surgeon decided to abort the use of the guidance system and place the screws freehand. There was no patient harm and the procedure was delayed less than an hour. 2021-04-07 email 2021-04-07 06:16 (Rep): Additional information received from a manufacturer representative reported that the trajectories were deviated less than 3.5 mm. 2021-05-09 IUD (Rep): Additional information received from a manufacturer representative reported the suspected cause was a planning issue and software functionality. | xse deviations salesforce image image ccr spinal procedure treat degenerative discs fixation screws planned spondylptosis case dual clamp mount minor challenge registration surgeon anatomy marker degrees ap position ideal degrees surgical arm sent trajectory left dilator slowly placed arm guide image taken confirm location bone hit surgical arm inferior compared plan plan adjusted entry point slightly superior superior angulation given surgical able reach desired superior position second registration issue occurred thought issue way carm image taken able true lateraloblique image end plates did parallel trajectory drilled scope skiving traje ctory surgeon aware surgeon planned flatten entry point surface drilling surgeon did proceed case open procedure avoid skiving medial angulation clos degrees soft tissue pressure affect trajectory surgeon created additional incision avoid soft tissue pressure force exerted instruments checking entry point avoid skiving surgeon moved right position confirmed carm trajectory ok superiorinferior direction lateral surgeon thought ap image sure visualization issue trajectory lateral surgeon decided abort use place screws freehand procedure delayed hour email additional trajectories deviated mm iud additional suspected cause planning issue software functionality mazor x | NAV5000 | NAV4095 | NAV3246 | NAV3243 | NAV1012 | NAV5000NAV4095NAV3246NAV3243NAV1012 | False
3589 | 703383069 | XSE - SRA Failed Accuracy Tests | 2019-09-26T15:27:09.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 30 | ASM0206-01 | SURGICAL ARM | XSE - SRA Failed Accuracy Tests 2019-09-20 SCL1902849, SCL1902910 (Rep): Medtronic received information regarding a guidance system being used during testing. It was reported that the surgical arm failed the advanced accuracy test. A field service engineer reported five days later that the replacement arm failed an advanced accuracy test. There was no patient involvement in the event. SURGICAL ARM | 2019-09-20 SCL1902849, SCL1902910 (Rep): Medtronic received information regarding a guidance system being used during testing. It was reported that the surgical arm failed the advanced accuracy test. A field service engineer reported five days later that the replacement arm failed an advanced accuracy test. There was no patient involvement in the event. | xse sra failed accuracy tests testing surgical arm failed advanced accuracy test field service engineer days later replacement arm failed advanced accuracy test surgical arm | NAV4093 | NAV4092 | NAV2054 | NAV3249 | NAV3295 | NAV4093NAV4092NAV2054NAV3249NAV3295 | True
7156 | 704131868 | XSE - CANBUS Error/Inaccurate Model | 2020-12-17T21:15:56.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error/Inaccurate Model 2020-12-16 00702904, 00702906, 00702907 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error was displayed during the case and the monitor was unresponsive. During the procedure, the CAD model showed the surgical arm was moving when the arm was not actually moving. There was patient harm and the procedure was delayed less than an hour. 2020-12-18 IUD x9 (Other): No new information. 2021-01-21 IUD (Other): No new information. 2021-02-17 email 2021-02-17 11:30 (Rep): Additional information received from a manufacturer representative reported that the main cable was replaced during service and the issue was resolved. 2021-Mar-18 IUDx2 (other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2020-12-16 00702904, 00702906, 00702907 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that a CANBUS error was displayed during the case and the monitor was unresponsive. During the procedure, the CAD model showed the surgical arm was moving when the arm was not actually moving. There was patient harm and the procedure was delayed less than an hour. 2020-12-18 IUD x9 (Other): No new information. 2021-01-21 IUD (Other): No new information. 2021-02-17 email 2021-02-17 11:30 (Rep): Additional information received from a manufacturer representative reported that the main cable was replaced during service and the issue was resolved. 2021-Mar-18 IUDx2 (other): No new information. | xse canbus errorinaccurate model spinal procedure canbus error displayed case monitor unresponsive procedure cad model showed surgical arm moving arm actually moving procedure delayed hour iud new iud new email additional main cable replaced service issue resolved new svc main cable | NAV2047 | NAV3256 | NAV2054 | NAV3248 | NAV3243 | NAV2047NAV3256NAV2054NAV3248NAV3243 | True
8327 | 704398165 | XSE - OOBF/Surgical System & PC Failure | 2021-05-17T15:53:04.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 80 | ELB0057-01 | MAZOR X FIRE WIRE CARD | XSE - OOBF/Surgical System & PC Failure 2021-May-17 Salesforce C00226529, IUD x2 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the surgical system and PC failed at install. There was no patient involvement. 2021-05-19 IUD x4 (Other): No new information. 2021-05-31 e1 (Other): No new information. 2021-06-03 IUD, e2, e3, e4 (Rep): No new information. 2021-06-21 e5 (Other): No new information. 2021-06-24 e6 (Other): No new information. 2021-Jun-29 IUDx6 (other): No new information. 2021-07-30 IUD (Other): No new information. 2021-08-30 IUD x3 (Other): No new information. 2021-10-28 302612706, IUD x5 (Other): No new information. MAZOR X FIRE WIRE CARD | 2021-May-17 Salesforce C00226529, IUD x2 (Rep): Medtronic received information regarding a guidance system being used outside of a p rocedure. It was reported that the surgical system and PC failed at install. There was no patient involvement. 2021-05-19 IUD x4 (Other): No new information. 2021-05-31 e1 (Other): No new information. 2021-06-03 IUD, e2, e3, e4 (Rep): No new information. 2021-06-21 e5 (Other): No new information. 2021-06-24 e6 (Other): No new information. 2021-Jun-29 IUDx6 (other): No new information. 2021-07-30 IUD (Other): No new information. 2021-08-30 IUD x3 (Other): No new information. 2021-10-28 302612706, IUD x5 (Other): No new information. | xse oobfsurgical pc failure salesforce iud outside p rocedure surgical pc failed install iud new new iud new new new new iud new iud new iud new mazor x wire card | NAV2054 | NAV3262 | NAV2047 | NAV3246 | NAV2057 | NAV2054NAV3262NAV2047NAV3246NAV2057 | False
5505 | 703741354 | XSE - SRA Damaged/Shipped with System | 2020-05-06T23:44:44.000000+0000 | MazorX Stealth Edition | NAV2055 | PRODUCTS DAMAGED DURING SHIPPING | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - SRA Damaged/Shipped with System 2020-05-06 00688425 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was shipped while attached to the system and there was concern the arm would be damaged. There was no patient harm. 2020-05-13 IUD x2 (Other): No new information. 2020-05-20 IUD (Other): No new information. 2020-05-27 IUD x3 (Other): No new information. 2020-06-07 IUD (Other): No new information. 2020-06-30 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-05-06 00688425 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the surgical arm was shipped while attached to the system and there was concern the arm would be damaged. There was no patient harm. 2020-05-13 IUD x2 (Other): No new information. 2020-05-20 IUD (Other): No new information. 2020-05-27 IUD x3 (Other): No new information. 2020-06-07 IUD (Other): No new information. 2020-06-30 IUD (Other): No new information. | xse sra damagedshipped outside procedure surgical arm shipped attached concern arm damaged iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV2054 | NAV1006 | NAV2055 | NAV4098 | NAV3249NAV2054NAV1006NAV2055NAV4098 | True
2445 | 703236028 | X - SOFTWARE ISSUE | 2019-07-01T14:07:58.000000+0000 | MazorX | NAV3247 | EXITS | 20 | UNK_MAZORX_SW | UNKNOWN MAZOR X SW | X - SOFTWARE ISSUE SW crach was reported (happened 3 times during the procedure) UNKNOWN MAZOR X SW | SW crach was reported (happened 3 times during the procedure) | x software issue sw crach unknown mazor x sw | NAV3246 | NAV3242 | NAV4091 | NAV3247 | NAV3248 | NAV3246NAV3242NAV4091NAV3247NAV3248 | True
10315 | 704853384 | XSE - SW Flips Name on Images | 2022-02-08T21:56:40.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - SW Flips Name on Images 2022-Feb-07 Salesforce C00242838 (Rep, HCP): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the patient name would sometimes flip when loading images on the guidance system. The origin of the images did not matter. The manufacturer representative tried loading images from a disc burned from PACS, multiple USB sticks from different laptops, and while transferring images from an O-arm. The issue appeared to be related to the software and not the system. There was no patient involvement. 2022-02-15 IUD (Other): No new information. 2022-03-13 e1, e2, e3, e4, e5 (Rep): No new information. 2022-03-16 e7 2022-03-22 10:13 (Rep): When the surgeon imports a brand study into the software, the names get scrambled around between the first and last name. This behavior is reproducible across his all of his computers, the hospital’s workstation and even when importing o-arm scans via usb or Ethernet during a scan and plan case. It tends to happen pretty much during every import. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2022-Feb-07 Salesforce C00242838 (Rep, HCP): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the patient name would sometimes flip when loading images on the guidance system. The origin of the images did not matter. The manufacturer representative tried loading images from a disc burned from PACS, multiple USB sticks from different laptops, and while transferring images from an O-arm. The issue appeared to be related to the software and not the system. There was no patient involvement. 2022-02-15 IUD (Other): No new information. 2022-03-13 e1, e2, e3, e4, e5 (Rep): No new information. 2022-03-16 e7 2022-03-22 10:13 (Rep): When the surgeon imports a brand study into the software, the names get scrambled around between the first and last name. This behavior is reproducible across his all of his computers, the hospital’s workstation and even when importing o-arm scans via usb or Ethernet during a scan and plan case. It tends to happen pretty much during every import. | xse sw flips images salesforce outside proc edure flip loading images origin images did matter tried loading images disc burned pacs multiple usb sticks different laptops transferring images oarm issue appeared related software iud new new surgeon imports brand study software names scrambled behavior reproducible computers hospital’s workstation importing oarm scans usb ethernet scan plan case tends happen pretty import mazorx sw installation kit | NAV3235 | NAV3246 | NAV2058 | NAV3236 | NAV3248 | NAV3235NAV3246NAV2058NAV3236NAV3248 | True
3061 | 703264053 | X - Electric Panel Damaged | 2019-07-17T19:42:03.000000+0000 | MazorX | NAV2055 | PRODUCTS DAMAGED DURING SHIPPING | 20 | ASM0270-01 | CART ELECTRIC PANEL ASSEMBLY | X - Electric Panel Damaged Electrical Panel (ASM0270) was damaged during shipping. See attached images. CART ELECTRIC PANEL ASSEMBLY | Electrical Panel (ASM0270) was damaged during shipping. See attached images. | x electric panel damaged electrical panel damaged shipping attached images cart electric panel assembly | NAV2055 | NAV3295 | NAV2058 | NAV3249 | NAV4091 | NAV2055NAV3295NAV2058NAV3249NAV4091 | True
5169 | 703610125 | XSE - Monitors Going Black/Arm issues | 2020-02-12T16:04:01.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 50 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XSE - Monitors Going Black/Arm issues 2020-02-12 00683927 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that both the monitor and surgeon screen are going black. This happened four times and at one point it was black for upwards of 30 seconds. In between these screens going black they attempted to send the arm to clear up, but it moved downward instead. After an additional attempt the arm did move to clear up. After the final screen going black, they lost all of the information in the case and had to enter new codes, take new scans, and images. There was no patient harm. 2020-02-14 IUD x6 (Other): Additional information was received from the manufacturer's representative. The Field service engineer arrived on site to troubleshoot the issue. Upon inspection engineer noticed the cart fan was not spinning as fast as usual, potentially causing the PC to overheat and display the noted symptoms. Fan was ordered and the representative was instructed on how to provide proper ventilation to the PC during cases. 2020-03-12 IUD x5 (Other): Additional information received from a manufacturer representative reported that the engineer arrived on site and, before replacing the fan, connected the fan to the power source externally. At first start up the fan didn't move and exhibited the same symptoms as the installed fan. Engineer verified that there was a steady 12.03V source. The fault in the new fan was caused by a small exposed wire protruding from the bottom center of fan. This wire is standard on all fans the engineer has seen but the wire was getting stuck in the bearing causing the fan to not turn/twitch or not spin at the proper RPMs. The engineer tried to free the caught wire but could not position the wire adequately enough that it would not inhibit the operation of the fan. A new fan was then ordered and will be installed once the system is available. 2020-04-06 IUD x5 (Other): No new information. 2020-04-22 IUD x3 (Other): No new information. 2020-04-30 IUD x2 (Other): No new information. UNK MAZOR X MONITOR | 2020-02-12 00683927 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that both the monitor and surgeon screen are going black. This happened four times and at one point it was black for upwards of 30 seconds. In between these screens going black they attempted to send the arm to clear up, but it moved downward instead. After an additional attempt the arm did move to clear up. After the final screen going black, they lost all of the information in the case and had to enter new codes, take new scans, and images. There was no patient harm. 2020-02-14 IUD x6 (Other): Additional information was received from the manufacturer's representative. The Field service engineer arrived on site to troubleshoot the issue. Upon inspection engineer noticed the cart fan was not spinning as fast as usual, potentially causing the PC to overheat and display the noted symptoms. Fan was ordered and the representative was instructed on how to provide proper ventilation to the PC during cases. 2020-03-12 IUD x5 (Other): Additional information received from a manufacturer representative reported that the engineer arrived on site and, before replacing the fan, connected the fan to the power source externally. At first start up the fan didn't move and exhibited the same symptoms as the installed fan. Engineer verified that there was a steady 12.03V source. The fault in the new fan was caused by a small exposed wire protruding from the bottom center of fan. This wire is standard on all fans the engineer has seen but the wire was getting stuck in the bearing causing the fan to not turn/twitch or not spin at the proper RPMs. The engineer tried to free the caught wire but could not position the wire adequately enough that it would not inhibit the operation of the fan. A new fan was then ordered and will be installed once the system is available. 2020-04-06 IUD x5 (Other): No new information. 2020-04-22 IUD x3 (Other): No new information. 2020-04-30 IUD x2 (Other): No new information. | xse monitors going blackarm issues spinal procedure monitor surgeon screen going black happened times point black upwards seconds screens going black attempted send arm clear moved downward instead additional attempt arm did clear final screen going black lost case enter new codes new scans images iud additional manufacturers field service engineer arrived site troubleshoot issue inspection engineer noticed cart fan spinning fast usual potentially causing pc overheat display noted symptoms fan ordered instructed provide proper ventilation pc cases iud additional engineer arrived site replacing fan connected fan power source externally start fan didnt exhibited symptoms installed fan engineer verified steady source fault new fan caused small exposed wire protruding center fan wire standard fans engineer seen wire getting stuck bearing causing fan turntwitch spin proper rpms engineer tried free caught wire position wire adequately inhibit operation fan new fan ordered installed available iud new iud new iud new unk mazor x monitor | NAV2058 | NAV3261 | NAV4101 | NAV3253 | NAV3244 | NAV2058NAV3261NAV4101NAV3253NAV3244 | True
7711 | 704252668 | XSE - CANBUS Error | 2021-02-26T21:09:23.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | XSE - CANBUS Error 2021-Feb-26 Salesforce C00221792 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that while setting up for a case they performed a scan plan registration she got a CANBUS error that she could not clear. At that time the surgeon wanted the system off the table and she could not remove the system from the table due to the teeth not disengaging. She tried the override panel and tried to unmount the system manually without resolution . The emergency stop would not function. The button that releases the shoulder would not work. The foot pedal pump would not function as well to remove the system from the bed as well. She finally had to release the bed frame adaptor to get the system off the bed. The whole system then became unresponsive and nothing would function. Surgeon stated he did not want to use the system again. There was no patient harm reported but there was a delay of procedure of one hour or longer. 2021-03-01 IUD x3 (Other): No new information. 2021-03-17 IUD, e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the CANBUS error occurred with the old Lemo cable prior to installing the new software version. The lemo cable was replaced as part of troubleshooting. 2021-03-23 IUD x2 (Other): No new information. 2021-04-13 IUD x2 (Other): No new information. 2021-05-12 IUD (Other): No new information. 2021-05-20 IUD x2 (Other): No new information. 2021-08-04 IUD (Other): No new information. 2021-09-03 IUD (Other): No new information. UNK MAZOR X ACCESSORIES | 2021-Feb-26 Salesforce C00221792 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that while setting up for a case they performed a scan plan registration she got a CANBUS error that she could not clear. At that time the surgeon wanted the system off the table and she could not remove the system from the table due to the teeth not disengaging. She tried the override panel and tried to unmount the system manually without resolution . The emergency stop would not function. The button that releases the shoulder would not work. The foot pedal pump would not function as well to remove the system from the bed as well. She finally had to release the bed frame adaptor to get the system off the bed. The whole system then became unresponsive and nothing would function. Surgeon stated he did not want to use the system again. There was no patient harm reported but there was a delay of procedure of one hour or longer. 2021-03-01 IUD x3 (Other): No new information. 2021-03-17 IUD, e1, e2, e3 (Rep): Additional information received from a manufacturer representative reported that the CANBUS error occurred with the old Lemo cable prior to installing the new software version. The lemo cable was replaced as part of troubleshooting. 2021-03-23 IUD x2 (Other): No new information. 2021-04-13 IUD x2 (Other): No new information. 2021-05-12 IUD (Other): No new information. 2021-05-20 IUD x2 (Other): No new information. 2021-08-04 IUD (Other): No new information. 2021-09-03 IUD (Other): No new information. | xse canbus error salesforce spinal p rocedure setting case performed scan plan registration got canbus error clear time surgeon wanted table remove table teeth disengaging tried override panel tried unmount manually resolution emergency stop function button releases shoulder work foot pedal pump function remove bed finally release bed frame adaptor bed unresponsive function surgeon stated did want use delay procedure hour longer iud new iud additional canbus error occurred old lemo cable prior installing new software version lemo cable replaced troubleshooting iud new iud new iud new iud new iud new iud new unk mazor x accessories | NAV4107 | NAV2047 | NAV2054 | NAV2057 | NAV3250 | NAV4107NAV2047NAV2054NAV2057NAV3250 | True
5123 | 703605261 | XSE - Monitor & Tablet Screens Went Dark | 2020-02-10T13:59:32.000000+0000 | MazorX Stealth Edition | NAV3261 | MONITOR ISSUES | 20 | UNK_MAZORX_MON | UNK MAZOR X MONITOR | XSE - Monitor & Tablet Screens Went Dark 2020-02-06 email x2, CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while the software was in the operate process, the screen on the workstation dimmed. Shortly after, both the workstation screen and surgeon tablet screen blacked out completely for about 30 seconds before turning back on and resumed the operate screen. The workstation monitor remained dim for most of the procedure but returned to normal brightness at the end of the operate process. It was also noted by the manufacturer representative that when extending the workstation screen, a grinding noise was heard coming from the workstation. No patient harm was reported and surgery was delayed less than an hour. 2020-02-10 Email 2020-02-10 11:01, IUD x6 (Rep): No new information. UNK MAZOR X MONITOR | 2020-02-06 email x2, CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that while the software was in the operate process, the screen on the workstation dimmed. Shortly after, both the workstation screen and surgeon tablet screen blacked out completely for about 30 seconds before turning back on and resumed the operate screen. The workstation monitor remained dim for most of the procedure but returned to normal brightness at the end of the operate process. It was also noted by the manufacturer representative that when extending the workstation screen, a grinding noise was heard coming from the workstation. No patient harm was reported and surgery was delayed less than an hour. 2020-02-10 Email 2020-02-10 11:01, IUD x6 (Rep): No new information. | xse monitor tablet screens went dark email ccr spinal procedure software operate process screen workstation dimmed shortly workstation screen surgeon tablet screen blacked completely seconds turning resumed operate screen workstation monitor remained dim procedure returned normal brightness end operate process noted extending workstation screen grinding noise heard coming workstation surgery delayed hour email iud new unk mazor x monitor | NAV3253 | NAV3261 | NAV2058 | NAV3264 | NAV3257 | NAV3253NAV3261NAV2058NAV3264NAV3257 | True
4825 | 703579164 | XSE - Bed Frame Adaptor failing test | 2020-01-24T04:07:22.000000+0000 | MazorX Stealth Edition | NAV3255 | BED FRAME ISSUES | 40 | MAS1574-01 | BED FRAME ADAPTOR | XSE - Bed Frame Adaptor failing test 2020-01-23 00681956 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that the representative requested a replacement bed frame adaptor be sent to the field service engineer for testing. There was no patient involvement. 2020-01-24 IUD x5 (Rep): No new information. BED FRAME ADAPTOR | 2020-01-23 00681956 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that the representative requested a replacement bed frame adaptor be sent to the field service engineer for testing. There was no patient involvement. 2020-01-24 IUD x5 (Rep): No new information. | xse bed frame adaptor failing test use requested replacement bed frame adaptor sent field service engineer testing iud new bed frame adaptor | NAV3255 | NAV4090 | NAV2054 | NAV3295 | NAV1006 | NAV3255NAV4090NAV2054NAV3295NAV1006 | True
188 | 703029477 | X - Deviation | 2019-02-27T17:05:13.000000+0000 | MazorX | NAV5000 | MXSE - INACCURACY | 20 | KIT0639-11 | MAZOR X SOFTWARE KIT VER-3.0.2 | X - Deviation 2019-Feb-25 SCL1900461, SCL1900479, (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the patient was over 40 BMI. Everthing went well with registration and there were good flouro matches to vertebral bodies. After registration the HCP seemed to be doing osteotomies and was removing bone. The patient was moving so much that they got a shoulder shift. They then had to re-register and had good values. There was a lot of deep bi-lateral retraction. Two trajectories were missed inferior at L4. The HCP then placed those screws manually under flouro. After the case the rep did a ten point check and found the system to be accurate. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE KIT VER-3.0.2 | 2019-Feb-25 SCL1900461, SCL1900479, (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the patient was over 40 BMI. Everthing went well with registration and there were good flouro matches to vertebral bodies. After registration the HCP seemed to be doing osteotomies and was removing bone. The patient was moving so much that they got a shoulder shift. They then had to re-register and had good values. There was a lot of deep bi-lateral retraction. Two trajectories were missed inferior at L4. The HCP then placed those screws manually under flouro. After the case the rep did a ten point check and found the system to be accurate. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | x deviation spinal procedure bmi everthing went registration good flouro matches vertebral bodies registration hcp doing osteotomies removing bone moving got shoulder shift reregister good values lot deep bilateral retraction trajectories missed inferior hcp placed screws manually flouro case did point check accurate noted additional complications anticipated mazor x software kit | NAV5000 | NAV4091 | NAV1006 | NAV3242 | NAV4108 | NAV5000NAV4091NAV1006NAV3242NAV4108 | True
7444 | 704207891 | XSE - System Stopped Responding | 2021-02-02T13:39:50.000000+0000 | MazorX Stealth Edition | NAV3251 | POSITIONER ISSUES | 70 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - System Stopped Responding 2021-02-02 00705633 (Rep, FOR): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system stopped responding. The system was mounted to the table. The teeth had been engaged on the bed frame adapter. When they went to disconnect the manipulator, they were not able to. Diodes on manipulator where off. Diodes on the positioner where off. The shoulder switch was not working. The monitor up/down button was not working. The surgical arm was working. They used the emergency stop button and it was working, but did not resolve the issues. They performed a hard reboot three times, 5 minutes, with no resolution. After a reboot the system displayed CANBUS error 7. They were trying to use emergency tool to release the positioner, but only one of the teeth opened. They removed the positioner with the bed frame adapter from the table to give surgeon access to the patient. The surgery continued with guidewire. There was no patient harm and the procedure was delayed less than an hour. 2021-02-03 IUD x5, email 2021-02-21 12:15 (Rep, For): No new information. 2021-Feb-23 email 2/2 (Rep, For): It was found that the surgical system was damaged and had to be replaced. The electrical panel was working properly and there was no need to replace it. Bed frame adaptor and emergency pump were also replaced as recommended after the issue. 2021-04-08 IUD x2 (Other, For): No new information. 2021-06-8 e1, e2 (Other, For): No new information. 2021-06-16 IUD (Other, For): No new information. 2021-06-28 IUD x3 (Other), For): No new information. 2021-Aug-09 IUD (other): No new information. 2021-10-03 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-02-02 00705633 (Rep, FOR): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system stopped responding. The system was mounted to the table. The teeth had been engaged on the bed frame adapter. When they went to disconnect the manipulator, they were not able to. Diodes on manipulator where off. Diodes on the positioner where off. The shoulder switch was not working. The monitor up/down button was not working. The surgical arm was working. They used the emergency stop button and it was working, but did not resolve the issues. They performed a hard reboot three times, 5 minutes, with no resolution. After a reboot the system displayed CANBUS error 7. They were trying to use emergency tool to release the positioner, but only one of the teeth opened. They removed the positioner with the bed frame adapter from the table to give surgeon access to the patient. The surgery continued with guidewire. There was no patient harm and the procedure was delayed less than an hour. 2021-02-03 IUD x5, email 2021-02-21 12:15 (Rep, For): No new information. 2021-Feb-23 email 2/2 (Rep, For): It was found that the surgical system was damaged and had to be replaced. The electrical panel was working properly and there was no need to replace it. Bed frame adaptor and emergency pump were also replaced as recommended after the issue. 2021-04-08 IUD x2 (Other, For): No new information. 2021-06-8 e1, e2 (Other, For): No new information. 2021-06-16 IUD (Other, For): No new information. 2021-06-28 IUD x3 (Other), For): No new information. 2021-Aug-09 IUD (other): No new information. 2021-10-03 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse stopped responding spinal procedure stopped responding mounted table teeth engaged bed frame adapter went disconnect manipulator able diodes manipulator diodes positioner shoulder switch working monitor updown button working surgical arm working emergency stop button working did resolve issues performed hard reboot times minutes resolution reboot displayed canbus error trying use emergency tool release positioner teeth opened removed positioner bed frame adapter table surgeon access surgery continued guidewire procedure delayed hour iud email new email surgical damaged replaced electrical panel working properly need replace bed frame adaptor emergency pump replaced recommended issue iud new new iud new iud new iud new iud new iud new kit svc surgical syst type | NAV4103 | NAV3260 | NAV3250 | NAV2047 | NAV3251 | NAV4103NAV3260NAV3250NAV2047NAV3251 | True
2813 | 703251767 | X - 3DEFINE ISSUE | 2019-07-10T16:10:42.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | ASM0206-01 | SURGICAL ARM | X - 3DEFINE ISSUE Separated from Comp-0000319. This complaint will handle the 3-Define issue. Before registration there were complications with the 3 define scan. As the arm stopped at each position during the scan and took an image, multiple lines were noticed as if parts of the surface were missing. The scan was completed but less than half of the surface was visible. The tech examined the drape to make sure it was not blocking the 3define camera but there was nothing that seemed out of the ordinary. The scan was attempted again but the arm stopped on it's on and said that the 3define scan failed. The working volume was then deleted, and a generic volume was set to move forward with the procedure. SURGICAL ARM | Separated from Comp-0000319. This complaint will handle the 3-Define issue. Before registration there were complications with the 3 define scan. As the arm stopped at each position during the scan and took an image, multiple lines were noticed as if parts of the surface were missing. The scan was completed but less than half of the surface was visible. The tech examined the drape to make sure it was not blocking the 3define camera but there was nothing that seemed out of the ordinary. The scan was attempted again but the arm stopped on it's on and said that the 3define scan failed. The working volume was then deleted, and a generic volume was set to move forward with the procedure. | x issue separated complaint handle issue registration complications define scan arm stopped position scan took image multiple lines noticed parts surface missing scan completed half surface visible tech examined drape make sure blocking camera ordinary scan attempted arm stopped said scan failed working volume deleted generic volume set forward procedure surgical arm | NAV3244 | NAV4095 | NAV3239 | NAV4096 | NAV4091 | NAV3244NAV4095NAV3239NAV4096NAV4091 | True
777 | 703175999 | X - Clamp Link Screw Strip | 2019-05-28T17:33:13.000000+0000 | MazorX | NAV4090 | BROKEN SCREW | 20 | MAS2202-03 | CLAMP LINK | X - Clamp Link Screw Strip 2018-12-19 SCL1805319(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the clamp link screw stripped. There was no patient harm and the procedure was not delayed over an hour. CLAMP LINK | 2018-12-19 SCL1805319(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the clamp link screw stripped. There was no patient harm and the procedure was not delayed over an hour. | x clamp link screw strip spinal procedure clamp link screw stripped procedure delayed hour clamp link | NAV3299 | NAV4090 | NAV3295 | NAV3242 | NAV4112 | NAV3299NAV4090NAV3295NAV3242NAV4112 | True
9251 | 704588952 | XSE - Emergency Stop Error | 2021-09-01T18:09:34.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Emergency Stop Error 2021-Sep-01 Salesforce C00233813 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the guidance system was booted up ,the system showed the emergency stop button was pressed when it was not. The manufacturer representative powered down the system, waited 30 minutes, and then powered it back up, but the error was still there. There was no patient involvement. 2021-Sep-06 IUDx2 (other): No new information. UNK MAZOR X WORKSTATION | 2021-Sep-01 Salesforce C00233813 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the guidance system was booted up ,the system showed the emergency stop button was pressed when it was not. The manufacturer representative powered down the system, waited 30 minutes, and then powered it back up, but the error was still there. There was no patient involvement. 2021-Sep-06 IUDx2 (other): No new information. | xse emergency stop error salesforce outside procedure booted showed emergency stop button pressed powered waited minutes powered error new unk mazor x workstation | NAV2047 | NAV2058 | NAV3261 | NAV3244 | NAV4108 | NAV2047NAV2058NAV3261NAV3244NAV4108 | True
5885 | 703828895 | XSE - Image Adapter Latch Damaged | 2020-06-30T19:06:14.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1226-02S | KIT SVC MAS1226-02S MAIN IMG ADPTR BRCKT | XSE - Image Adapter Latch Damaged 2020-06-30 00691829 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter latch was very stiff, which made it difficult to tighten. There was no patient involvement. 2020-07-06 IUD x2 (Other): No new information. 2020-07-28 IUD x4 (Other): No new information. KIT SVC MAS1226-02S MAIN IMG ADPTR BRCKT | 2020-06-30 00691829 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter latch was very stiff, which made it difficult to tighten. There was no patient involvement. 2020-07-06 IUD x2 (Other): No new information. 2020-07-28 IUD x4 (Other): No new information. | xse image adapter latch damaged outside procedure image adapter latch stiff difficult tighten iud new iud new kit svc main img adptr brckt | NAV3258 | NAV2054 | NAV3299 | NAV2058 | NAV2047 | NAV3258NAV2054NAV3299NAV2058NAV2047 | True
5405 | 703673202 | X - X-Eye Motor/Surgeon Screen Issues | 2020-03-18T18:22:15.000000+0000 | MazorX | NAV3253 | TABLET ISSUES | 50 | UNK_MAZORX_TBLT | UNK MAZOR X TABLET | X - X-Eye Motor/Surgeon Screen Issues 2020-03-18 00686590 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the x-eye motor was not functioning, the surgeon screen was not displaying a picture but responding to touch and the software was not responding properly. The representative took the c-arm images and the system said the bodies were not labeled correctly and could not get past registration. The representative attempted to test it in demo but the c-arm images did not save. The representative attempted to perform a soft restart and received a blue screen saying the PC needs to restart. The system got stuck on the black screen and the representative had to perform a hard restart. There was no patient harm and the procedure was not delayed over an hour. 2020-03-19 Email 2020-03-19 12:45, IUD x6 (Rep, Other): Additional information received from a manufacturer representative reported that the engineer backed up the case files on the PC being replaced. Once backed up, the PC was upgraded to a new PC. Software 4.2.1 was put on the new PC. Once the software was installed per the installation documents, the case files were put back on the PC using the import feature of the software. The system was then checkout via the checklist ATP for the Mazor X surgical system type 2. Patient information was also updated. 2020-04-17 IUD x4 (Other): No new information. 2020-05-13 IUD x2 (Other): No new information. 2020-06-07 IUD x3 (Other): No new information. UNK MAZOR X TABLET | 2020-03-18 00686590 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the x-eye motor was not functioning, the surgeon screen was not displaying a picture but responding to touch and the software was not responding properly. The representative took the c-arm images and the system said the bodies were not labeled correctly and could not get past registration. The representative attempted to test it in demo but the c-arm images did not save. The representative attempted to perform a soft restart and received a blue screen saying the PC needs to restart. The system got stuck on the black screen and the representative had to perform a hard restart. There was no patient harm and the procedure was not delayed over an hour. 2020-03-19 Email 2020-03-19 12:45, IUD x6 (Rep, Other): Additional information received from a manufacturer representative reported that the engineer backed up the case files on the PC being replaced. Once backed up, the PC was upgraded to a new PC. Software 4.2.1 was put on the new PC. Once the software was installed per the installation documents, the case files were put back on the PC using the import feature of the software. The system was then checkout via the checklist ATP for the Mazor X surgical system type 2. Patient information was also updated. 2020-04-17 IUD x4 (Other): No new information. 2020-05-13 IUD x2 (Other): No new information. 2020-06-07 IUD x3 (Other): No new information. | x xeye motorsurgeon screen issues spinal procedure xeye motor functioning surgeon screen displaying picture responding touch software responding properly took carm images said bodies labeled correctly past registration attempted test demo carm images did save attempted perform soft restart blue screen saying pc needs restart got stuck black screen perform hard restart procedure delayed hour email iud additional engineer backed case files pc replaced backed pc upgraded new pc software new pc software installed installation documents case files pc using import feature software checkout checklist atp mazor x surgical type updated iud new iud new iud new unk mazor x tablet | NAV4110 | NAV3262 | NAV3246 | NAV3248 | NAV2050 | NAV4110NAV3262NAV3246NAV3248NAV2050 | False
6915 | 704064183 | XSE ¦ Deviation | 2020-11-11T18:47:23.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE ¦ Deviation 2020-Nov-11 00700679 (HCP,rep); Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance navigation device being used in a Sacroiliac and Thoracolumbar procedure. It was reported that they had to perform the arm calibration test four times and would pass but would not give them the green check to move forward. During the L4 - L5 case on L4 left and L5 right they showed they were deviated 5mm superior. The plan was on target and it was performed with cortical ATS screws. The surgeon then backed all four screws out and then used stealth navigation to set the screws. He decided to do all four screws due to the fact there was not enough bone left for purchase and also wanted to swap to regular pedicle screws. Chris stated that it was a skive even though the plan was good, that area with cortical ATS screws created the skive. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-Dec-02 IUD (other): No new information. MAZOR X SYSTEM TPL0059 | 2020-Nov-11 00700679 (HCP,rep); Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance navigation device being used in a Sacroiliac and Thoracolumbar procedure. It was reported that they had to perform the arm calibration test four times and would pass but would not give them the green check to move forward. During the L4 - L5 case on L4 left and L5 right they showed they were deviated 5mm superior. The plan was on target and it was performed with cortical ATS screws. The surgeon then backed all four screws out and then used stealth navigation to set the screws. He decided to do all four screws due to the fact there was not enough bone left for purchase and also wanted to swap to regular pedicle screws. Chris stated that it was a skive even though the plan was good, that area with cortical ATS screws created the skive. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-Dec-02 IUD (other): No new information. | xse deviation navigation sacroiliac thoracolumbar procedure perform arm calibration test times pass green check forward case left right showed deviated superior plan target performed cortical ats screws surgeon backed screws stealth navigation set screws decided screws fact bone left purchase wanted swap regular pedicle screws chris stated skive plan good area cortical ats screws created skive additional complications anticipated iud new mazor x | NAV5000 | NAV2058 | NAV2054 | NAV3249 | NAV2047 | NAV5000NAV2058NAV2054NAV3249NAV2047 | True
474 | 703111548 | XSE - deviation | 2019-04-17T20:31:14.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - deviation 2019-Apr-12 SCL1900945 intake form, msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that after completing 3define and snapshot, the arm was sent to marker and star marker attached. An O2 was brought in and scan completed and imported via USB. The surgeon and reps planned together. The surgeon was adamant about 14 degree axial angles at L4 and L5, in order to avoid proximal facet and preexisting screw tracks, respectively. The rep communicated options for an open approach (axial angles <12 degree, transmusccular approach, or extending length of incision). Trajectories were sent and at L4 & L5 bilaterally, the surgeon elected to proceed transmuscular while maintaining bilateral retraction of the skin. All screws portrayed by Nav as accurate during execution. Upon stimulating screws, L4 left read at less than 10 and L5 right read at 13. Upon visual inspection of the screws, surgeon felt they looked medial. The screws were removed and a ball-tip probe confirmed medial deviation at L4 left and suspected medial at L5 left. Surgeon elected to proceed by using nav feature to freehand place a new trajectory at L4 left. They attempted the same at L5 left but kept skiving into the previous trajec tory. They then decided to leave trajectory as initially executed with the robot. Post-screw spin confirmed L5 left to be slightly more medial than the plan. Post spin, the surgeon asked, "Why did it show us as being to plan when we were not?" It was noted there was no patient harm and the surgery was delayed over an hour due to an unplanned O-arm spin and redirect 2 trajectories. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Apr-12 SCL1900945 intake form, msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that after completing 3define and snapshot, the arm was sent to marker and star marker attached. An O2 was brought in and scan completed and imported via USB. The surgeon and reps planned together. The surgeon was adamant about 14 degree axial angles at L4 and L5, in order to avoid proximal facet and preexisting screw tracks, respectively. The rep communicated options for an open approach (axial angles <12 degree, transmusccular approach, or extending length of incision). Trajectories were sent and at L4 & L5 bilaterally, the surgeon elected to proceed transmuscular while maintaining bilateral retraction of the skin. All screws portrayed by Nav as accurate during execution. Upon stimulating screws, L4 left read at less than 10 and L5 right read at 13. Upon visual inspection of the screws, surgeon felt they looked medial. The screws were removed and a ball-tip probe confirmed medial deviation at L4 left and suspected medial at L5 left. Surgeon elected to proceed by using nav feature to freehand place a new trajectory at L4 left. They attempted the same at L5 left but kept skiving into the previous trajec tory. They then decided to leave trajectory as initially executed with the robot. Post-screw spin confirmed L5 left to be slightly more medial than the plan. Post spin, the surgeon asked, "Why did it show us as being to plan when we were not?" It was noted there was no patient harm and the surgery was delayed over an hour due to an unplanned O-arm spin and redirect 2 trajectories. There were no additional complications reported or anticipated as a result of the event. | xse deviation msg spinal procedure completing snapshot arm sent marker star marker attached brought scan completed imported usb surgeon reps planned surgeon adamant degree axial angles order avoid proximal facet preexisting screw tracks respectively communicated options open approach trajectories sent bilaterally surgeon elected proceed transmuscular maintaining bilateral retraction skin screws portrayed nav accurate execution stimulating screws left read right read visual inspection screws surgeon felt looked medial screws removed balltip probe confirmed medial deviation left suspected medial left surgeon elected proceed using nav feature freehand place new trajectory left attempted left kept skiving previous trajec tory decided leave trajectory initially executed robot postscrew spin confirmed left slightly medial plan post spin surgeon asked did plan noted surgery delayed hour unplanned oarm spin redirect trajectories additional complications anticipated mazor x software installation kit | NAV5000 | NAV3235 | NAV4098 | NAV3243 | NAV3241 | NAV5000NAV3235NAV4098NAV3243NAV3241 | True
4622 | 703554143 | X - Registration Issue/Cross View Error | 2020-01-09T19:18:40.000000+0000 | MazorX | NAV3242 | REGISTRATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration Issue/Cross View Error 2018-03-23 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed on 2018-March-20 for an open procedure for fixation at L4-S1. The system was mounted the patient with a schanz pin int he the patient's right iliac crest. Upon first attempting registration, software gave note that the AP and OBL image did not match indicating a cross view error and registration failed. New images were obtained and registration was eventually successful. All final trajectories were placed accurately. No patient harm was reported and surgery was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2018-03-23 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that planning was completed on 2018-March-20 for an open procedure for fixation at L4-S1. The system was mounted the patient with a schanz pin int he the patient's right iliac crest. Upon first attempting registration, software gave note that the AP and OBL image did not match indicating a cross view error and registration failed. New images were obtained and registration was eventually successful. All final trajectories were placed accurately. No patient harm was reported and surgery was delayed less than an hour. | x registration issuecross view error ccr spinal procedure planning completed open procedure fixation mounted schanz pin int patients right iliac crest attempting registration software gave note ap obl image did match indicating cross view error registration failed new images obtained registration eventually successful final trajectories placed accurately surgery delayed hour mazor x | NAV3242 | NAV4112 | NAV3243 | NAV2047 | NAV3247 | NAV3242NAV4112NAV3243NAV2047NAV3247 | True
5328 | 703649298 | XSE - Schanz Bond Damaged | 2020-03-04T16:17:15.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 30 | MAS2271-04S | KIT SVC MAS2271-04S SCHANZ BOND | XSE - Schanz Bond Damaged 2020-03-04 00685619 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz bond was damaged/stripped and no longer functioning properly. There was no patient involved. 2020-03-12 IUD x4 (Other): No new information. 2020-03-28 Email 2020-03-28 13:30 (Rep): Additional information received from a manufacturer representative reported that the schanz bond was stripped from wear and tear overtime. KIT SVC MAS2271-04S SCHANZ BOND | 2020-03-04 00685619 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the Schanz bond was damaged/stripped and no longer functioning properly. There was no patient involved. 2020-03-12 IUD x4 (Other): No new information. 2020-03-28 Email 2020-03-28 13:30 (Rep): Additional information received from a manufacturer representative reported that the schanz bond was stripped from wear and tear overtime. | xse schanz bond damaged outside procedure schanz bond damagedstripped longer functioning properly involved iud new email additional schanz bond stripped wear tear overtime kit svc schanz bond | NAV3295 | NAV3299 | NAV3298 | NAV3250 | NAV4090 | NAV3295NAV3299NAV3298NAV3250NAV4090 | True
3554 | 703376161 | XSE - SA Failed Stress Test during PM | 2019-09-23T18:48:06.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - SA Failed Stress Test during PM 2019-09-19 SCL1902844, PDF x2 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during preventative maintenance, the system failed multiple stress tests at joint 4. The field service engineer (FSE) attempted to calibrate joint 4 multiple times. The system passed stress tests after two calibrations, but the current consumption was still abnormally high along with arm noise. No patient was involved. SURGICAL ARM | 2019-09-19 SCL1902844, PDF x2 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that during preventative maintenance, the system failed multiple stress tests at joint 4. The field service engineer (FSE) attempted to calibrate joint 4 multiple times. The system passed stress tests after two calibrations, but the current consumption was still abnormally high along with arm noise. No patient was involved. | xse sa failed stress test pm pdf outside procedure preventative maintenance failed multiple stress tests joint field service engineer attempted calibrate joint multiple times passed stress tests calibrations current consumption abnormally high arm noise involved surgical arm | NAV3249 | NAV2054 | NAV3264 | NAV2058 | NAV4101 | NAV3249NAV2054NAV3264NAV2058NAV4101 | True
3606 | 703387880 | XSE - shoulder would not unlock | 2019-09-30T14:57:49.000000+0000 | MazorX Stealth Edition | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 10 | TPL0059-06 | MAZOR X SYSTEM | XSE - shoulder would not unlock 2019-Oct-09 SCL1903041 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they did an entire case without issue, Surgeon went to unlock by pushing button on shoulder, pushed harder than normal, and the shoulder failed to unlock. The rep tried to electronically unlock shoulder, tried to manually unlock positioner on both axes, but still failed to move. Soft restart did not fix the issue. Hard restart fixed the issue. There was no patient harm or additional complications reported or anticipated as a result of the event. MAZOR X SYSTEM | 2019-Oct-09 SCL1903041 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that they did an entire case without issue, Surgeon went to unlock by pushing button on shoulder, pushed harder than normal, and the shoulder failed to unlock. The rep tried to electronically unlock shoulder, tried to manually unlock positioner on both axes, but still failed to move. Soft restart did not fix the issue. Hard restart fixed the issue. There was no patient harm or additional complications reported or anticipated as a result of the event. | xse shoulder unlock spinal procedure did entire case issue surgeon went unlock pushing button shoulder pushed harder normal shoulder failed unlock tried electronically unlock shoulder tried manually unlock positioner axes failed soft restart did fix issue hard restart fixed issue additional complications anticipated mazor x | NAV4107 | NAV5000 | NAV3242 | NAV2047 | NAV1012 | NAV4107NAV5000NAV3242NAV2047NAV1012 | True
7597 | 704235281 | XSE - Shoulder Out of Calibration | 2021-02-17T21:03:23.000000+0000 | MazorX Stealth Edition | NAV4106 | CALIBRATION ISSUES | 20 | KIT1282-01 | SVC KIT1282-01 STEPPER DRIVE MODULE | XSE - Shoulder Out of Calibration 2021-Feb-17 Salesforce C00221247 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the shoulder storage position is always a red X. No matter what he does to calibrate the shoulder it still is a red X. There was no patient present. 2021-03-01 IUD (Other): No new information. 2021-03-16 IUD (Other): No new information. SVC KIT1282-01 STEPPER DRIVE MODULE | 2021-Feb-17 Salesforce C00221247 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the shoulder storage position is always a red X. No matter what he does to calibrate the shoulder it still is a red X. There was no patient present. 2021-03-01 IUD (Other): No new information. 2021-03-16 IUD (Other): No new information. | xse shoulder calibration salesforce outside procedure shoulder storage position red x matter does calibrate shoulder red x present iud new iud new svc stepper drive module | NAV3252 | NAV4106 | NAV3299 | NAV2058 | NAV5001 | NAV3252NAV4106NAV3299NAV2058NAV5001 | True
127 | 703004876 | X - X-Eye Not Working | 2019-02-13T08:19:31.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - X-Eye Not Working 2019-02-07 SCL1900338, Intake Form 2019-02-13 02:24 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the X-Eye camera was not working. Troubleshooting included checking all connections were secure, but the camera still was not powering on. There was no green indicator light on the camera. There was no patient harm, the procedure was not delayed more than an hour and the case was completed. UNK MAZOR X X-EYE CAMERA | 2019-02-07 SCL1900338, Intake Form 2019-02-13 02:24 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the X-Eye camera was not working. Troubleshooting included checking all connections were secure, but the camera still was not powering on. There was no green indicator light on the camera. There was no patient harm, the procedure was not delayed more than an hour and the case was completed. | x xeye working spinal procedure xeye camera working troubleshooting included checking connections secure camera powering green indicator light camera procedure delayed hour case completed unk mazor x xeye camera | NAV4110 | NAV3244 | NAV4109 | NAV3254 | NAV3239 | NAV4110NAV3244NAV4109NAV3254NAV3239 | True
4593 | 703551956 | X - Registration & Communication Issues | 2020-01-08T20:35:24.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Registration & Communication Issues 2018-01-29 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a deformity with fixation at L3-L4. It was reported that registration was attempted three times. It was likely a user error due to one dot not being placed precisely enough for the first two attempts. It is noted that this system was upgraded to software version 3.1 due to a new arm replacement. Some errors were seen including: Error 295, fatal error, and CANBUS connection was lost to nodes 9 and 10 (the system was hard rebooted), and warning 4427. Additionally, an invalid shoulder encoder value was encountered. The manufacturer representative was unsure how to check the encoders and proceeded with the case normally and had no issues. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SYSTEM TPL0059 | 2018-01-29 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat a deformity with fixation at L3-L4. It was reported that registration was attempted three times. It was likely a user error due to one dot not being placed precisely enough for the first two attempts. It is noted that this system was upgraded to software version 3.1 due to a new arm replacement. Some errors were seen including: Error 295, fatal error, and CANBUS connection was lost to nodes 9 and 10 (the system was hard rebooted), and warning 4427. Additionally, an invalid shoulder encoder value was encountered. The manufacturer representative was unsure how to check the encoders and proceeded with the case normally and had no issues. There was no patient harm and the procedure was not delayed over an hour. | x registration communication issues ccr spinal procedure treat deformity fixation registration attempted times likely user error dot placed precisely attempts noted upgraded software version new arm replacement errors seen including error fatal error canbus connection lost nodes warning additionally invalid shoulder encoder value encountered unsure check encoders proceeded case normally issues procedure delayed hour mazor x | NAV2047 | NAV3242 | NAV4112 | NAV4091 | NAV3246 | NAV2047NAV3242NAV4112NAV4091NAV3246 | True
6570 | 703979435 | XSE - Inaccurate Registration | 2020-09-26T00:03:00.000000+0000 | MazorX Stealth Edition | NAV3242 | REGISTRATION ISSUES | 20 | UNK_MXSE_SW | UNKNOWN MAZOR X SE SW | XSE - Inaccurate Registration 2020-09-25 00697522, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with screws and rods. It was reported that there was an inaccurate registration noticed in verify registration (~3mm) at S1 (L5-S1) case. The manufacturer representative noticed that the registration showed as passing, however there was a clear shift on S1.The site took another oblique shot which they verified was correct. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-09-28 IUD x2 (Other): No new information. UNKNOWN MAZOR X SE SW | 2020-09-25 00697522, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative disc with screws and rods. It was reported that there was an inaccurate registration noticed in verify registration (~3mm) at S1 (L5-S1) case. The manufacturer representative noticed that the registration showed as passing, however there was a clear shift on S1.The site took another oblique shot which they verified was correct. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2020-09-28 IUD x2 (Other): No new information. | xse inaccurate registration ccr spinal procedure treat degenerative disc screws rods inaccurate registration noticed verify registration case noticed registration showed passing clear shift site took oblique shot verified correct issue extended surgical time hour impact outcome iud new unknown mazor x se sw | NAV3242 | NAV4112 | NAV5000 | NAV2056 | NAV3243 | NAV3242NAV4112NAV5000NAV2056NAV3243 | True
2980 | 703261257 | X - Workstation Issues | 2019-07-16T15:32:26.000000+0000 | MazorX | NAV4105 | UNEXPECTED POSITIONER MOVEMENT | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X - Workstation Issues Questions about repairs to the system and other issues they have seen including, Positioner sliding down the Z-Axis, Pneumatic valve, Manipulator solenoid not working. UNK MAZOR X POSITIONER | Questions about repairs to the system and other issues they have seen including, Positioner sliding down the Z-Axis, Pneumatic valve, Manipulator solenoid not working. | x workstation issues questions repairs issues seen including positioner sliding zaxis pneumatic valve manipulator solenoid working unk mazor x positioner | NAV3264 | NAV3260 | NAV4105 | NAV3251 | NAV2047 | NAV3264NAV3260NAV4105NAV3251NAV2047 | True
4255 | 703518248 | X - Deviation at right L5/Arm issues | 2019-12-13T19:35:16.000000+0000 | MazorX | NAV4096 | ARM HIT WORK VOLUME | 20 | ASM0206-01 | SURGICAL ARM | X - Deviation at right L5/Arm issues 2018-05-22 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was planned prior to surgery for two vertebrae on levels L4 - L5. Scan and plan was used for the case. The surgical system was attached to the bed and draped. A single schanz pin was used and the guidance system was attached to the patient. The guidance system was sent to a preset accurate position and manually adjusted. O-arm spin and star marker were found. Screws were planned. 3Define was sent to planned trajectories. Arm could not reach left side due to hitting bone mount bridge. L5 right was deviated lateral. Starting point was screw inside body deviated lateral. The screw was repositioned by navigation. Placement was confirmed with O-arm spin. No patient arm was reported. SURGICAL ARM | 2018-05-22 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was planned prior to surgery for two vertebrae on levels L4 - L5. Scan and plan was used for the case. The surgical system was attached to the bed and draped. A single schanz pin was used and the guidance system was attached to the patient. The guidance system was sent to a preset accurate position and manually adjusted. O-arm spin and star marker were found. Screws were planned. 3Define was sent to planned trajectories. Arm could not reach left side due to hitting bone mount bridge. L5 right was deviated lateral. Starting point was screw inside body deviated lateral. The screw was repositioned by navigation. Placement was confirmed with O-arm spin. No patient arm was reported. | x deviation right issues ccr spinal procedure case planned prior surgery vertebrae levels scan plan case surgical attached bed draped single schanz pin attached sent preset accurate position manually adjusted oarm spin star marker screws planned sent planned trajectories arm reach left hitting bone mount bridge right deviated lateral starting point screw inside body deviated lateral screw repositioned navigation placement confirmed oarm spin arm surgical arm | NAV4091 | NAV4095 | NAV3244 | NAV3299 | NAV3235 | NAV4091NAV4095NAV3244NAV3299NAV3235 | False
4596 | 703552035 | X - Registration, Arm, & X-eye Issues | 2020-01-08T21:06:38.000000+0000 | MazorX | NAV4110 | NO TRACKING | 30 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | X - Registration, Arm, & X-eye Issues 2018-02-28 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L5. It was reported that an arm shift occurred, and registration had to be attempted a second time. Draw spine did not work. The surgeon had to disengage the robot to tap and then a third registration was done. The case was completed. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X X-EYE CAMERA | 2018-02-28 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L5. It was reported that an arm shift occurred, and registration had to be attempted a second time. Draw spine did not work. The surgeon had to disengage the robot to tap and then a third registration was done. The case was completed. There was no patient harm and the procedure was not delayed over an hour. | x registration arm xeye issues ccr spinal procedure treat degenerative disk disease fixation arm shift occurred registration attempted second time draw spine did work surgeon disengage robot tap registration case completed procedure delayed hour unk mazor x xeye camera | NAV4110 | NAV4098 | NAV4095 | NAV3254 | NAV3242 | NAV4110NAV4098NAV4095NAV3254NAV3242 | True
7446 | 704207891 | XSE - System Stopped Responding | 2021-02-02T13:39:50.000000+0000 | MazorX Stealth Edition | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 70 | ASM0214-02S | KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | XSE - System Stopped Responding 2021-02-02 00705633 (Rep, FOR): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system stopped responding. The system was mounted to the table. The teeth had been engaged on the bed frame adapter. When they went to disconnect the manipulator, they were not able to. Diodes on manipulator where off. Diodes on the positioner where off. The shoulder switch was not working. The monitor up/down button was not working. The surgical arm was working. They used the emergency stop button and it was working, but did not resolve the issues. They performed a hard reboot three times, 5 minutes, with no resolution. After a reboot the system displayed CANBUS error 7. They were trying to use emergency tool to release the positioner, but only one of the teeth opened. They removed the positioner with the bed frame adapter from the table to give surgeon access to the patient. The surgery continued with guidewire. There was no patient harm and the procedure was delayed less than an hour. 2021-02-03 IUD x5, email 2021-02-21 12:15 (Rep, For): No new information. 2021-Feb-23 email 2/2 (Rep, For): It was found that the surgical system was damaged and had to be replaced. The electrical panel was working properly and there was no need to replace it. Bed frame adaptor and emergency pump were also replaced as recommended after the issue. 2021-04-08 IUD x2 (Other, For): No new information. 2021-06-8 e1, e2 (Other, For): No new information. 2021-06-16 IUD (Other, For): No new information. 2021-06-28 IUD x3 (Other), For): No new information. 2021-Aug-09 IUD (other): No new information. 2021-10-03 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. KIT SVC ASM0214-02S SURGICAL SYST TYPE 2 | 2021-02-02 00705633 (Rep, FOR): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the system stopped responding. The system was mounted to the table. The teeth had been engaged on the bed frame adapter. When they went to disconnect the manipulator, they were not able to. Diodes on manipulator where off. Diodes on the positioner where off. The shoulder switch was not working. The monitor up/down button was not working. The surgical arm was working. They used the emergency stop button and it was working, but did not resolve the issues. They performed a hard reboot three times, 5 minutes, with no resolution. After a reboot the system displayed CANBUS error 7. They were trying to use emergency tool to release the positioner, but only one of the teeth opened. They removed the positioner with the bed frame adapter from the table to give surgeon access to the patient. The surgery continued with guidewire. There was no patient harm and the procedure was delayed less than an hour. 2021-02-03 IUD x5, email 2021-02-21 12:15 (Rep, For): No new information. 2021-Feb-23 email 2/2 (Rep, For): It was found that the surgical system was damaged and had to be replaced. The electrical panel was working properly and there was no need to replace it. Bed frame adaptor and emergency pump were also replaced as recommended after the issue. 2021-04-08 IUD x2 (Other, For): No new information. 2021-06-8 e1, e2 (Other, For): No new information. 2021-06-16 IUD (Other, For): No new information. 2021-06-28 IUD x3 (Other), For): No new information. 2021-Aug-09 IUD (other): No new information. 2021-10-03 IUD (Other): No new information. 2021-12-08 IUD (Other): No new information. | xse stopped responding spinal procedure stopped responding mounted table teeth engaged bed frame adapter went disconnect manipulator able diodes manipulator diodes positioner shoulder switch working monitor updown button working surgical arm working emergency stop button working did resolve issues performed hard reboot times minutes resolution reboot displayed canbus error trying use emergency tool release positioner teeth opened removed positioner bed frame adapter table surgeon access surgery continued guidewire procedure delayed hour iud email new email surgical damaged replaced electrical panel working properly need replace bed frame adaptor emergency pump replaced recommended issue iud new new iud new iud new iud new iud new iud new kit svc surgical syst type | NAV4103 | NAV3260 | NAV3250 | NAV2047 | NAV3251 | NAV4103NAV3260NAV3250NAV2047NAV3251 | True
6569 | 703979040 | XSE - Deviated Trajectories | 2020-09-25T20:03:11.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Deviated Trajectories 2020-09-25 00697517 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that there was a deviated screw at L5 left. It appeared to be inferior to the plan. They made a midline skin incision and doing a transmuscular approach. They had bilateral retraction. The manufacturer representative believes it was due to soft tissue pressure. They removed the guidance system and did a spin and that when they saw the screw was deviated. The surgeon used navigation to correct the screw. There was no impact on the patient outcome. 2020-09-28 email 2020-09-28 14:00 (Rep): Additional information received from a manufacturer representative reported that the screws were deviated between 3.5 and 10 mm. 2020-10-05 IUD, e1, e2, e3, e4,e 5, e6 (Rep): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2020-09-25 00697517 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that there was a deviated screw at L5 left. It appeared to be inferior to the plan. They made a midline skin incision and doing a transmuscular approach. They had bilateral retraction. The manufacturer representative believes it was due to soft tissue pressure. They removed the guidance system and did a spin and that when they saw the screw was deviated. The surgeon used navigation to correct the screw. There was no impact on the patient outcome. 2020-09-28 email 2020-09-28 14:00 (Rep): Additional information received from a manufacturer representative reported that the screws were deviated between 3.5 and 10 mm. 2020-10-05 IUD, e1, e2, e3, e4,e 5, e6 (Rep): No new information. | xse deviated trajectories intraoperatively spinal procedure deviated screw left appeared inferior plan midline skin incision doing transmuscular approach bilateral retraction believes soft tissue pressure removed did spin saw screw deviated surgeon navigation correct screw impact outcome email additional screws deviated mm iud new mazor x software installation kit | NAV5000 | NAV3246 | NAV2056 | NAV3248 | NAV3247 | NAV5000NAV3246NAV2056NAV3248NAV3247 | True
7539 | 704223973 | XSE - Security Patch Failed | 2021-02-10T21:46:19.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 20 | ASM0215-06S | ASM0215-06S MAZOR X 5820 PC STATION | XSE - Security Patch Failed 2021-Feb-10 Salesforce C00220845 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the rep was installing 5.0 software and the security patch failed. There was no patient involved. 2021-02-16 IUD x5 (Other): No new information. 2021-04-08 IUD x4 (Other): No new information. ASM0215-06S MAZOR X 5820 PC STATION | 2021-Feb-10 Salesforce C00220845 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the rep was installing 5.0 software and the security patch failed. There was no patient involved. 2021-02-16 IUD x5 (Other): No new information. 2021-04-08 IUD x4 (Other): No new information. | xse security patch failed salesforce outside procedure installing software security patch failed involved iud new iud new mazor x pc station | NAV3262 | NAV2054 | NAV3246 | NAV3264 | NAV2047 | NAV3262NAV2054NAV3246NAV3264NAV2047 | True
2151 | 703226527 | X- SW Issues | 2019-06-25T19:23:18.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X- SW Issues AP 3D Marker position of midline. MAZOR X SOFTWARE KIT VER-2.3 | AP 3D Marker position of midline. | x sw issues ap marker position midline mazor x software kit | NAV3246 | NAV3242 | NAV4091 | NAV3247 | NAV3248 | NAV3246NAV3242NAV4091NAV3247NAV3248 | True
2286 | 703231310 | X- Arm Issues | 2019-06-27T17:54:15.000000+0000 | MazorX | NAV4095 | ARM DIDN'T REACH LOCATION | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X- Arm Issues Dr. Kelley-CNS 7.20.17 Two main issues with the system: 1st Issue: After successful 10 point check is completed and during the 3 point check that follows, arm will stop all movement and an error message will appear that says arm didn’t reach (image attached). If the 3 point test is restarted it will stop in the exact spot every time. The only way to get the arm to complete the 3 point test is to send the arm to storage position, then initiate 3 pt. test. The arm will the reach all 3 stages and complete the 3 point test. If the arm is sent to clear up, or any other position prior to 3 point check, it will still not work. This was tested and repeated 3 times and the only position that will allow a successful 3pt check to follow is the storage position. UNK MAZOR X ARM | Dr. Kelley-CNS 7.20.17 Two main issues with the system: 1st Issue: After successful 10 point check is completed and during the 3 point check that follows, arm will stop all movement and an error message will appear that says arm didn’t reach (image attached). If the 3 point test is restarted it will stop in the exact spot every time. The only way to get the arm to complete the 3 point test is to send the arm to storage position, then initiate 3 pt. test. The arm will the reach all 3 stages and complete the 3 point test. If the arm is sent to clear up, or any other position prior to 3 point check, it will still not work. This was tested and repeated 3 times and the only position that will allow a successful 3pt check to follow is the storage position. | x arm issues dr kelleycns main issues issue successful point check completed point check follows arm stop movement error message appear says arm didn’t reach point test restarted stop exact spot time way arm complete point test send arm storage position initiate pt test arm reach stages complete point test arm sent clear position prior point check work tested repeated times position allow successful check follow storage position unk mazor x arm | NAV4095 | NAV4101 | NAV3244 | NAV2058 | NAV4099 | NAV4095NAV4101NAV3244NAV2058NAV4099 | True
5542 | 703749491 | XSE - L5 Deviation | 2020-05-12T18:46:45.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - L5 Deviation 2020-05-11 00688709, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the left and right L5 screws were low and below the pedicle compared to the plan during the L2-L5 procedure. The surgeon completed the left side of the construct first, followed by the right side. All over screws were accurate except for L5. The manufacturer representative and surgeon did not know the cause of the deviation. There were no skive potentials on the plan and the docking surface was flat. No patient shifts were seen before, during or after placement of the screws. The extenders were close at L4, but the surgeon extended the incision to help with the issue. The L2, L3 and L4 extenders were moved out of the way for the L5 trajectory and the surgeon verified that they did not place any pressure on the vertebral bodies as the incision was extended. The representative reviewed registration and it was accurate. Navigation was checked with the passive planar at the L4 divot and it was accurate. The body of L5 was also checked and it was accurate, but the screw was still low. The surgeon decided to re-instrument L5 free hand with guide wires. There was no patient harm and the procedure was delayed less than an hour. 2020-05-18 IUD (Other): No new information. 2020-05-26 email 2020-05-26 16:30, e1 (Rep): Additional information received from a manufacturer representative reported there was no patient harm and the trajectories were deviated less than 5 mm. 2020-06-04 e2 (Rep): Additional information received from a manufacturer representative reported that there was no change in surgery. During the procedure, the surgeon make one skin incision per side with separate facial incisions. The surgeon verified that they were not putting pressure on the extenders to move them out of the way. The surgeon removed the screws and placed them freehand. MAZOR X SYSTEM TPL0059 | 2020-05-11 00688709, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the left and right L5 screws were low and below the pedicle compared to the plan during the L2-L5 procedure. The surgeon completed the left side of the construct first, followed by the right side. All over screws were accurate except for L5. The manufacturer representative and surgeon did not know the cause of the deviation. There were no skive potentials on the plan and the docking surface was flat. No patient shifts were seen before, during or after placement of the screws. The extenders were close at L4, but the surgeon extended the incision to help with the issue. The L2, L3 and L4 extenders were moved out of the way for the L5 trajectory and the surgeon verified that they did not place any pressure on the vertebral bodies as the incision was extended. The representative reviewed registration and it was accurate. Navigation was checked with the passive planar at the L4 divot and it was accurate. The body of L5 was also checked and it was accurate, but the screw was still low. The surgeon decided to re-instrument L5 free hand with guide wires. There was no patient harm and the procedure was delayed less than an hour. 2020-05-18 IUD (Other): No new information. 2020-05-26 email 2020-05-26 16:30, e1 (Rep): Additional information received from a manufacturer representative reported there was no patient harm and the trajectories were deviated less than 5 mm. 2020-06-04 e2 (Rep): Additional information received from a manufacturer representative reported that there was no change in surgery. During the procedure, the surgeon make one skin incision per side with separate facial incisions. The surgeon verified that they were not putting pressure on the extenders to move them out of the way. The surgeon removed the screws and placed them freehand. | xse deviation ccr spinal procedure treat deformity fixation left right screws low pedicle compared plan procedure surgeon completed left construct followed right screws accurate surgeon did know cause deviation skive potentials plan docking surface flat shifts seen placement screws extenders close surgeon extended incision help issue extenders moved way trajectory surgeon verified did place pressure vertebral bodies incision extended reviewed registration accurate navigation checked passive planar divot accurate body checked accurate screw low surgeon decided reinstrument free hand guide wires procedure delayed hour iud new email additional trajectories deviated mm additional change surgery procedure surgeon make skin incision separate facial incisions surgeon verified putting pressure extenders way surgeon removed screws placed freehand mazor x | NAV5000 | NAV4091 | NAV3253 | NAV4112 | NAV2051 | NAV5000NAV4091NAV3253NAV4112NAV2051 | True
5430 | 703687459 | XSE - Inaccurate Screw Placement | 2020-03-28T02:25:44.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Inaccurate Screw Placement 2020-03-27 00686979 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the procedure started with the lower lumbar and executed L3-L5 (surgeon made intra-op decision not to do S1). Navigation showed off and they completed all checks needed and the surgeon wanted to proceed. The c arm shot showed all screws were accurate. Remounted and clamped onto T11-12- and all trajectories at this segment showed the left side medial and the right side lateral. All checks were completed again. The surgeon had his own theory that since the patient had a fracture, he thought that the patient's spine was different from the CT scan when positioned (not shift during case) and navigation was a couple mm off, so he had the manufacturer representative move all screws out lateral. The manufacturer representative explained to him that the s oftware looks at each body independently, but he insisted she move the screws to make navigation look better. Navigation was off, but the surgeon continued. The screws stemmed above 25 and all were accurate when they took the c arm shot. At the end of the case, the second robot had a warning message come up on top of the screen with a yellow exclamation point symbol for approximately 10 seconds and then went off that something along the lines of, “System performance is low, needs technical support immediately.” There was no known impact on the patient outcome. 2020-04-07 IUD (Other): No new information. 2020-04-01 email 2020-04-11 11:30 (Rep): Medtronic received additional information that the screws were deviated 3.5 millimeters to 5.0 millimeters. It was reported that the deviated screws were fixed. There was an estimated delay to the procedure of an hour and half due to this issue. It was reported that the procedure was completed successfully. 2020-04-17 IUD (Other): No new information. 2020-04-24 IUD x2 (Other): No new information. 2020-07-21 CCR, Attachment: No new information. MAZOR X SYSTEM TPL0059 | 2020-03-27 00686979 (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that the procedure started with the lower lumbar and executed L3-L5 (surgeon made intra-op decision not to do S1). Navigation showed off and they completed all checks needed and the surgeon wanted to proceed. The c arm shot showed all screws were accurate. Remounted and clamped onto T11-12- and all trajectories at this segment showed the left side medial and the right side lateral. All checks were completed again. The surgeon had his own theory that since the patient had a fracture, he thought that the patient's spine was different from the CT scan when positioned (not shift during case) and navigation was a couple mm off, so he had the manufacturer representative move all screws out lateral. The manufacturer representative explained to him that the s oftware looks at each body independently, but he insisted she move the screws to make navigation look better. Navigation was off, but the surgeon continued. The screws stemmed above 25 and all were accurate when they took the c arm shot. At the end of the case, the second robot had a warning message come up on top of the screen with a yellow exclamation point symbol for approximately 10 seconds and then went off that something along the lines of, “System performance is low, needs technical support immediately.” There was no known impact on the patient outcome. 2020-04-07 IUD (Other): No new information. 2020-04-01 email 2020-04-11 11:30 (Rep): Medtronic received additional information that the screws were deviated 3.5 millimeters to 5.0 millimeters. It was reported that the deviated screws were fixed. There was an estimated delay to the procedure of an hour and half due to this issue. It was reported that the procedure was completed successfully. 2020-04-17 IUD (Other): No new information. 2020-04-24 IUD x2 (Other): No new information. 2020-07-21 CCR, Attachment: No new information. | xse inaccurate screw placement intraoperatively spinal procedure procedure started lower lumbar executed navigation showed completed checks needed surgeon wanted proceed c arm shot showed screws accurate remounted clamped trajectories segment showed left medial right lateral checks completed surgeon theory fracture thought patients spine different ct scan positioned navigation couple mm screws lateral explained s oftware looks body independently insisted screws make navigation look better navigation surgeon continued screws stemmed accurate took c arm shot end case second robot warning message come screen yellow exclamation point symbol approximately seconds went lines “system performance low needs technical support immediately” known impact outcome iud new email additional screws deviated millimeters millimeters deviated screws fixed estimated delay procedure hour half issue procedure completed successfully iud new iud new ccr attachment new mazor x | NAV1012 | NAV2056 | NAV5000 | NAV5001 | NAV3249 | NAV1012NAV2056NAV5000NAV5001NAV3249 | True
2620 | 703244824 | X - Workstation Issue | 2019-07-06T10:56:40.000000+0000 | MazorX | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - Workstation Issue - Pneumatic lock is not fully stopping shoulder from rotating when SRA exterts a bending moment when SRA is positioned to the extreme right or left of center. - Bed Adaptor issue identified. Adding more clearance to the maximum aperture of the Bed adaptor should be explored. The newer versions of the Jackson bed rail adaptors make the distance to wide to fit our bed adaptor. ALso, the frame's maximum extension of 63cm makes it to narrow to fit Proaxis tables. Proaxis bed rail adaptors expand the total distance to 64cm. UNK MAZOR X SHOULDER | - Pneumatic lock is not fully stopping shoulder from rotating when SRA exterts a bending moment when SRA is positioned to the extreme right or left of center. - Bed Adaptor issue identified. Adding more clearance to the maximum aperture of the Bed adaptor should be explored. The newer versions of the Jackson bed rail adaptors make the distance to wide to fit our bed adaptor. ALso, the frame's maximum extension of 63cm makes it to narrow to fit Proaxis tables. Proaxis bed rail adaptors expand the total distance to 64cm. | x workstation issue pneumatic lock fully stopping shoulder rotating sra exterts bending moment sra positioned extreme right left center bed adaptor issue identified adding clearance maximum aperture bed adaptor explored newer versions jackson bed rail adaptors make distance wide fit bed adaptor frames maximum extension makes narrow fit proaxis tables proaxis bed rail adaptors expand total distance unk mazor x shoulder | NAV3255 | NAV3264 | NAV3250 | NAV4107 | NAV2058 | NAV3255NAV3264NAV3250NAV4107NAV2058 | True
4206 | 703514076 | XSE - Instrument Verification Issue | 2019-12-12T03:24:04.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 30 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Instrument Verification Issue 2019-12-11 00680141 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the instruments in the software were all verified and then the verification was lost in navigation. The instruments tool cards would go yellow and the site would have to re-verify their instruments. No patient was present. 2020-01-03 IUD (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-12-11 00680141 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the instruments in the software were all verified and then the verification was lost in navigation. The instruments tool cards would go yellow and the site would have to re-verify their instruments. No patient was present. 2020-01-03 IUD (Other): No new information. | xse instrument verification issue outside procedure instruments software verified verification lost navigation instruments tool cards yellow site reverify instruments present iud new mazor x software installation kit | NAV3246 | NAV2056 | NAV2053 | NAV3256 | NAV3248 | NAV3246NAV2056NAV2053NAV3256NAV3248 | True
7846 | 704291447 | XSE - Manipulator Froze/SW Unresponsive | 2021-03-18T20:02:11.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 30 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - Manipulator Froze/SW Unresponsive 2021-Mar-18 Salesforce C00222896 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while attempting to mount this Mazor XSE system to the bed it froze up and the manipulator was not able to be moved. The software showed a red circle in the top right and the keyboard, touchscreen, and mouse were unresponsive. There was no patient involved. 2021-03-18 IUD x5 (Other): No new information. 2021-04-05 IUD x2 (Other): No new information. 2021-05-23 IUD x4 (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2021-Mar-18 Salesforce C00222896 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while attempting to mount this Mazor XSE system to the bed it froze up and the manipulator was not able to be moved. The software showed a red circle in the top right and the keyboard, touchscreen, and mouse were unresponsive. There was no patient involved. 2021-03-18 IUD x5 (Other): No new information. 2021-04-05 IUD x2 (Other): No new information. 2021-05-23 IUD x4 (Other): No new information. | xse manipulator frozesw unresponsive salesforce outside procedure attempting mount mazor xse bed froze manipulator able moved software showed red circle right keyboard touchscreen mouse unresponsive involved iud new iud new iud new mazor x pc station refurbished | NAV3262 | NAV3248 | NAV3260 | NAV5000 | NAV1012 | NAV3262NAV3248NAV3260NAV5000NAV1012 | True
6546 | 703972711 | X - No Input on Monitor | 2020-09-23T12:01:00.000000+0000 | MazorX | NAV3262 | PC ISSUES | 30 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | X - No Input on Monitor 2020-09-23 00697290 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the manufacturer representative first attempted to start the system, a no input message was displayed and the screen was black. The PC power button was checked and it would flash orange twice and then go blank after flashing orange seven times. This repeated. There was no patient involvement. 2020-09-25 IUD x3 (Other): No new information. 2020-10-12 IUD x5 (Other): No new information. 2021-06-20 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2020-09-23 00697290 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that when the manufacturer representative first attempted to start the system, a no input message was displayed and the screen was black. The PC power button was checked and it would flash orange twice and then go blank after flashing orange seven times. This repeated. There was no patient involvement. 2020-09-25 IUD x3 (Other): No new information. 2020-10-12 IUD x5 (Other): No new information. 2021-06-20 IUD (Other): No new information. | x input monitor outside procedure attempted start input message displayed screen black pc power button checked flash orange twice blank flashing orange seven times repeated iud new iud new iud new mazor x pc station refurbished | NAV3262 | NAV2054 | NAV3261 | NAV1006 | NAV3246 | NAV3262NAV2054NAV3261NAV1006NAV3246 | True
8607 | 704448290 | XSE - Leak in Manipulator | 2021-06-14T19:48:16.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 40 | KIT0759-01 | PNEUMATIC REPAIR KIT | XSE - Leak in Manipulator 2021-Jun-14 Salesforce C00228336 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a leak in the manipulator right at the D5 joint. The manufacturer representative could hear it when moving the surgical arm around. No patient was present. 2021-06-16 IUD x3 (Other): No new information. PNEUMATIC REPAIR KIT | 2021-Jun-14 Salesforce C00228336 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was a leak in the manipulator right at the D5 joint. The manufacturer representative could hear it when moving the surgical arm around. No patient was present. 2021-06-16 IUD x3 (Other): No new information. | xse leak manipulator salesforce outside procedure leak manipulator right joint hear moving surgical arm present iud new pneumatic repair kit | NAV3264 | NAV3260 | NAV3246 | NAV3249 | NAV4101 | NAV3264NAV3260NAV3246NAV3249NAV4101 | True
9145 | 704558987 | XSE - SRA Failed BIST | 2021-08-16T17:41:47.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 20 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Failed BIST 2021-08-16 C00232996 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the BIST failed. The arm stopped during joint 2 and had a calibration failure. Overall, the system was making beeps as if it were moving and the CAD model looked like it was moving into each position, however, the arm was not physically moving. Troubleshooting involved the manufacturer representative redoing the BIST test, where the arm moved slowly. Then he received an Error 293 on joint 3. The manufacturer representative noted that arm did not seem like it was moving fully into position. Technical Services walked the manufacturer representative through resetting the motion in superuser and retried the BIST, however, he received the same Error 293. No patient was present. 2021-08-16 IUD x3, e1 (Other): No new information. 2021-08-27 IUD (Other): No new information. 2021-Sep-16 IUDx2 (other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-11-29 302648598, IUD (Other): No new information. 2021-12-01 IUD x5 (Other): No new information. 2022-01-03 IUD (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2021-08-16 C00232996 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the BIST failed. The arm stopped during joint 2 and had a calibration failure. Overall, the system was making beeps as if it were moving and the CAD model looked like it was moving into each position, however, the arm was not physically moving. Troubleshooting involved the manufacturer representative redoing the BIST test, where the arm moved slowly. Then he received an Error 293 on joint 3. The manufacturer representative noted that arm did not seem like it was moving fully into position. Technical Services walked the manufacturer representative through resetting the motion in superuser and retried the BIST, however, he received the same Error 293. No patient was present. 2021-08-16 IUD x3, e1 (Other): No new information. 2021-08-27 IUD (Other): No new information. 2021-Sep-16 IUDx2 (other): No new information. 2021-09-30 IUD (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-11-29 302648598, IUD (Other): No new information. 2021-12-01 IUD x5 (Other): No new information. 2022-01-03 IUD (Other): No new information. | xse sra failed bist outside procedure bist failed arm stopped joint calibration failure overall making beeps moving cad model looked like moving position arm physically moving troubleshooting involved redoing bist test arm moved slowly error joint noted arm did like moving fully position technical services walked resetting motion superuser retried bist error present iud new iud new new iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV4095 | NAV4101 | NAV2054 | NAV4098 | NAV3249NAV4095NAV4101NAV2054NAV4098 | True
2396 | 703234164 | X - ARM GUIDE ISSUE | 2019-06-29T13:29:35.000000+0000 | MazorX | NAV3294 | ARM GUIDE ISSUES | 20 | MAS2330-01 | ARM GUIDE | X - ARM GUIDE ISSUE The Arm Guide Attachment would not properly screw into the arm as the screw was cross threaded somehow. We had to flash and sterilize the Workstation one and have subsequently swapped the two. ARM GUIDE | The Arm Guide Attachment would not properly screw into the arm as the screw was cross threaded somehow. We had to flash and sterilize the Workstation one and have subsequently swapped the two. | x arm guide issue arm guide attachment properly screw arm screw cross threaded flash sterilize workstation subsequently swapped arm guide | NAV3294 | NAV2054 | NAV4098 | NAV3299 | NAV4101 | NAV3294NAV2054NAV4098NAV3299NAV4101 | True
5266 | 703632791 | XSE - Deviated Trajectories at L4/L5 | 2020-02-24T20:11:10.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Trajectories at L4/L5 2020-02-24 00684769 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during the L4-L5 perc case, a screw at L5 right was medial by 3 to 5mm. The arm was sent to L4 right and again the trajectory was medial by 3 to 4 mm. The robot was unmounted and screws free handed. There was no patient harm and the surgery was not delayed more than one hour. 2020-02-24 IUD x4 (Other) No new information. 2020-03-30 Email 2020-03-30 13:00, e1, e2, e3, e4, CCR (Rep, HCP): Additional information received from a manufacturer representative reported that the case was a L4-L5 TLIF. Elevate cage was placed first. There were no registration issues. Deviation happened at right L5. The surgeon placed 22 blade through arm guide and down to the bone. The cannula and dilator followed. The representative could see the dilator moving medial on navigation and pointed out to the surgeon. The surgeon continued. The drill guide was placed and malleted into bone. The surgeon drilled to 30mm and removed the drill guide. Hand tap did not advance easily. The surgeon re-drilled and then went back to tap. The surgeon tapped and then placed screw. Navigation showed trajectories were medial but within the pedicle. Multiple lateral x rays were taken of knife, dilator, drill, tap and screw and they looked accurate. The screw tested around 10-11 mA EMG testing. AP x ray was taken and showed the screw was at least 4-6mm medial. The surgeon was upset. There was no patient injury. 2020-04-13 e5, e6, e7 (Rep): Additional information received from a manufacturer representative reported the surgeon did not change sides during the procedure. The surgeon stood on the right side of the patient. MAZOR X SYSTEM TPL0059 | 2020-02-24 00684769 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during the L4-L5 perc case, a screw at L5 right was medial by 3 to 5mm. The arm was sent to L4 right and again the trajectory was medial by 3 to 4 mm. The robot was unmounted and screws free handed. There was no patient harm and the surgery was not delayed more than one hour. 2020-02-24 IUD x4 (Other) No new information. 2020-03-30 Email 2020-03-30 13:00, e1, e2, e3, e4, CCR (Rep, HCP): Additional information received from a manufacturer representative reported that the case was a L4-L5 TLIF. Elevate cage was placed first. There were no registration issues. Deviation happened at right L5. The surgeon placed 22 blade through arm guide and down to the bone. The cannula and dilator followed. The representative could see the dilator moving medial on navigation and pointed out to the surgeon. The surgeon continued. The drill guide was placed and malleted into bone. The surgeon drilled to 30mm and removed the drill guide. Hand tap did not advance easily. The surgeon re-drilled and then went back to tap. The surgeon tapped and then placed screw. Navigation showed trajectories were medial but within the pedicle. Multiple lateral x rays were taken of knife, dilator, drill, tap and screw and they looked accurate. The screw tested around 10-11 mA EMG testing. AP x ray was taken and showed the screw was at least 4-6mm medial. The surgeon was upset. There was no patient injury. 2020-04-13 e5, e6, e7 (Rep): Additional information received from a manufacturer representative reported the surgeon did not change sides during the procedure. The surgeon stood on the right side of the patient. | xse deviated trajectories spinal procedure perc case screw right medial arm sent right trajectory medial mm robot unmounted screws free handed surgery delayed hour iud new email ccr additional case tlif elevate cage placed registration issues deviation happened right surgeon placed blade arm guide bone cannula dilator followed dilator moving medial navigation pointed surgeon surgeon continued drill guide placed malleted bone surgeon drilled removed drill guide hand tap did advance easily surgeon redrilled went tap surgeon tapped placed screw navigation showed trajectories medial pedicle multiple lateral x rays taken knife dilator drill tap screw looked accurate screw tested ma emg testing ap x ray taken showed screw medial surgeon upset injury additional surgeon did change sides procedure surgeon stood right mazor x | NAV5000 | NAV4091 | NAV5001 | NAV3242 | NAV3241 | NAV5000NAV4091NAV5001NAV3242NAV3241 | True
9666 | 704692413 | XSE ¦ Disc Stuck & Unexpected Shutdown | 2021-11-01T22:27:38.000000+0000 | MazorX Stealth Edition | NAV3257 | CD DRIVE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE ¦ Disc Stuck & Unexpected Shutdown 2021-Nov-01 Salesforce C00237144 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a disc got stuck in the external disc reader. The manufacturer representative believes there may be two discs in the drive that caused the issue even though they pressed the eject button before inserting the disc. It was also reported that the system shutdown on its own. It rebooted and they were able to use the system, but upon shutdown after a case the manufacturer representative got a blue screen with a sad face that stated the PC ran into issues and was collecting error information. No patient was present. 2021-11-02 IUD x5 (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Nov-01 Salesforce C00237144 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a disc got stuck in the external disc reader. The manufacturer representative believes there may be two discs in the drive that caused the issue even though they pressed the eject button before inserting the disc. It was also reported that the system shutdown on its own. It rebooted and they were able to use the system, but upon shutdown after a case the manufacturer representative got a blue screen with a sad face that stated the PC ran into issues and was collecting error information. No patient was present. 2021-11-02 IUD x5 (Other): No new information. | xse disc stuck unexpected shutdown salesforce outside procedure disc got stuck external disc reader believes discs drive caused issue pressed eject button inserting disc shutdown rebooted able use shutdown case got blue screen sad face stated pc ran issues collecting error present iud new unk mazor x workstation | NAV2058 | NAV3262 | NAV3257 | NAV2047 | NAV3253 | NAV2058NAV3262NAV3257NAV2047NAV3253 | True
495 | 703123685 | XSE - missing new version of drape | 2019-04-25T19:05:52.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | KIT0574 | MAZOR X SPINE DISPOSABLE KIT KIT0574 | XSE - missing new version of drape 2019-Apr-23 SCL1901013 jpg (rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that KIT0574-06 was found to be missing the newer version of the drape. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. MAZOR X SPINE DISPOSABLE KIT KIT0574 | 2019-Apr-23 SCL1901013 jpg (rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that KIT0574-06 was found to be missing the newer version of the drape. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. | xse missing new version drape jpg spinal procedure missing newer version drape additional complications anticipated mazor x spine disposable kit | NAV2053 | NAV2054 | NAV3232 | NAV5000 | NAV3246 | NAV2053NAV2054NAV3232NAV5000NAV3246 | True
7289 | 704167104 | XSE - System Arrived Damaged | 2021-01-12T01:27:24.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 30 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | XSE - System Arrived Damaged 2021-01-11 00704078 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the system arrived for a demo with the x-eye cover missing and it was not in storage position. The bottom cover of the positioner was missing and not in place. The system arrived not in the correct position. No patient was present. 2021-01-22 IUD (Other): No new information. UNK MAZOR X POSITIONER | 2021-01-11 00704078 (Rep): Medtronic received information regarding a guidance system, outside of procedure. It was reported that the system arrived for a demo with the x-eye cover missing and it was not in storage position. The bottom cover of the positioner was missing and not in place. The system arrived not in the correct position. No patient was present. 2021-01-22 IUD (Other): No new information. | xse arrived damaged outside procedure arrived demo xeye cover missing storage position cover positioner missing place arrived correct position present iud new unk mazor x positioner | NAV2053 | NAV2055 | NAV2054 | NAV3251 | NAV4104 | NAV2053NAV2055NAV2054NAV3251NAV4104 | True
8143 | 704361152 | XSE - Inaccuracy | 2021-04-26T16:22:16.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Inaccuracy 2021-Apr-26 Salesforce C00225292, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a L4-L5 case, the dilator was placed at right L5 and it was superior and lateral. Dual Schanz arms and Schanz pins were used to mount the surgical system to the patient. A MIST bridge was also used. The manufacturer representative mentioned a modified hover-T platform was used as there was multiple points of fixation. The inaccuracy was confirmed through fluoro. An accuracy check was done using the divot and anatomical landmarks and the system passed. The surgeon decided to re-register the patient and the same issue occurred. Left L4 was checked and it was inferior and lateral. The surgeon decided to abort the use of the guidance system. An advanced accuracy check was done after the case and there were no issues. The representative noted that there was a low chance of skiving with the plan. There was no patient harm and the procedure was delayed less than an hour. 2021-05-19 IUD x2 (Other): No new information. 2021-06-21 email (Rep): Additional information received from a manufacturer representative reported the trajectory was deviated between 3.5 and 10 mm. The exact amount of deviation was not measured since the surgeon did not proceed with instrumentation once they saw how much navigation and fluoro was off. MAZOR X SYSTEM TPL0059 | 2021-Apr-26 Salesforce C00225292, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a L4-L5 case, the dilator was placed at right L5 and it was superior and lateral. Dual Schanz arms and Schanz pins were used to mount the surgical system to the patient. A MIST bridge was also used. The manufacturer representative mentioned a modified hover-T platform was used as there was multiple points of fixation. The inaccuracy was confirmed through fluoro. An accuracy check was done using the divot and anatomical landmarks and the system passed. The surgeon decided to re-register the patient and the same issue occurred. Left L4 was checked and it was inferior and lateral. The surgeon decided to abort the use of the guidance system. An advanced accuracy check was done after the case and there were no issues. The representative noted that there was a low chance of skiving with the plan. There was no patient harm and the procedure was delayed less than an hour. 2021-05-19 IUD x2 (Other): No new information. 2021-06-21 email (Rep): Additional information received from a manufacturer representative reported the trajectory was deviated between 3.5 and 10 mm. The exact amount of deviation was not measured since the surgeon did not proceed with instrumentation once they saw how much navigation and fluoro was off. | xse inaccuracy salesforce iud spinal procedure case dilator placed right superior lateral dual schanz arms schanz pins mount surgical mist bridge mentioned modified hovert platform multiple points fixation inaccuracy confirmed fluoro accuracy check using divot anatomical landmarks passed surgeon decided reregister issue occurred left checked inferior lateral surgeon decided abort use advanced accuracy check case issues noted low chance skiving plan procedure delayed hour iud new email additional trajectory deviated mm exact deviation measured surgeon did proceed instrumentation saw navigation fluoro mazor x | NAV5000 | NAV1012 | NAV4093 | NAV2047 | NAV3299 | NAV5000NAV1012NAV4093NAV2047NAV3299 | True
1416 | 703199904 | X - Pressure Issue/Error Message | 2019-06-11T14:25:20.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 30 | ASM0263-01 | RNX-WS-1040 | X - Pressure Issue/Error Message Pneumatic pressure was cycling in and out. Required 4 hard re starts just to get the system on and off the bed. Canbus 4 error. Pneumatic pressure faltered in and out during and before case. RNX-WS-1040 | Pneumatic pressure was cycling in and out. Required 4 hard re starts just to get the system on and off the bed. Canbus 4 error. Pneumatic pressure faltered in and out during and before case. | x pressure issueerror message pneumatic pressure cycling required hard starts just bed canbus error pneumatic pressure faltered case | NAV3264 | NAV2054 | NAV4095 | NAV4091 | NAV1006 | NAV3264NAV2054NAV4095NAV4091NAV1006 | False
6493 | 703950976 | XSE - Deviation/SRA Shaking | 2020-09-11T01:51:30.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - Deviation/SRA Shaking 2020-09-10 00696222, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a L3-L4 posterior spinal fusion, the screws at left L3 and L4 were deviated laterally by 2-3 mm. The surgical arm was vibrating while the surgeon was removing the driver after placing the screws. The manufacturer representative believed that the surgical arm was having issues causing the deviation. There was no patient harm and the procedure was delayed less than an hour. 2020-09-14 IUD x2, email 2020-10-01 07:00 (Rep): No new information. 2020-10-01 email 2020-10-01 07:00, email 2020-10-02 20:02, IUD x2 (Rep): Additional information received from a manufacturer representative reported the surgical arm was resent to the trajectory, but the arm continued to vibrate in an agitated manor. The surgeon aborted the use of the guidance system and continued the procedure freehand. There was no patient harm. 2020-10-27 email 2020-10-28 07:11 (Rep): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-09-10 00696222, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a L3-L4 posterior spinal fusion, the screws at left L3 and L4 were deviated laterally by 2-3 mm. The surgical arm was vibrating while the surgeon was removing the driver after placing the screws. The manufacturer representative believed that the surgical arm was having issues causing the deviation. There was no patient harm and the procedure was delayed less than an hour. 2020-09-14 IUD x2, email 2020-10-01 07:00 (Rep): No new information. 2020-10-01 email 2020-10-01 07:00, email 2020-10-02 20:02, IUD x2 (Rep): Additional information received from a manufacturer representative reported the surgical arm was resent to the trajectory, but the arm continued to vibrate in an agitated manor. The surgeon aborted the use of the guidance system and continued the procedure freehand. There was no patient harm. 2020-10-27 email 2020-10-28 07:11 (Rep): No new information. | xse deviationsra shaking iud spinal procedure posterior spinal fusion screws left deviated laterally mm surgical arm vibrating surgeon removing driver placing screws believed surgical arm having issues causing deviation procedure delayed hour iud email new email email iud additional surgical arm resent trajectory arm continued vibrate agitated manor surgeon aborted use continued procedure freehand email new kit svc surgical arm refurb | NAV4101 | NAV5000 | NAV3249 | NAV4098 | NAV4096 | NAV4101NAV5000NAV3249NAV4098NAV4096 | True
5058 | 703593195 | X- SW Froze | 2020-02-02T22:37:27.000000+0000 | MazorX | NAV3248 | UNRESPONSIVE | 20 | KIT0878-01 | MAZOR X SOFTWARE KIT VER-4.0.0 | X- SW Froze 2018-8-15 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the patient was positioned and the surgical system was mounted to the bed. The software froze, while on the accuracy check screen. The screen froze while processing, so a soft restart was completed and an accuracy check was redone. A schanz pin was inserted into the right PSIS and an O-arm spin was conducted. Planning was completed, and all trajectories were sent. All screws were verified with a confirmation O-arm spin and neuromonitoring. There was no patient harm and the procedure was not delayed over an hour. MAZOR X SOFTWARE KIT VER-4.0.0 | 2018-8-15 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the patient was positioned and the surgical system was mounted to the bed. The software froze, while on the accuracy check screen. The screen froze while processing, so a soft restart was completed and an accuracy check was redone. A schanz pin was inserted into the right PSIS and an O-arm spin was conducted. Planning was completed, and all trajectories were sent. All screws were verified with a confirmation O-arm spin and neuromonitoring. There was no patient harm and the procedure was not delayed over an hour. | x sw froze ccr spinal procedure positioned surgical mounted bed software froze accuracy check screen screen froze processing soft restart completed accuracy check redone schanz pin inserted right psis oarm spin conducted planning completed trajectories sent screws verified confirmation oarm spin neuromonitoring procedure delayed hour mazor x software kit | NAV3248 | NAV3247 | NAV3235 | NAV2047 | NAV4091 | NAV3248NAV3247NAV3235NAV2047NAV4091 | True
5973 | 703850832 | XSE - CANBUS Error | 2020-07-14T15:12:59.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | KIT1317 | SVC KIT1317 SYSTEM MAIN CABLE | XSE - CANBUS Error 2020-07-14 00692646, CCR x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a Scan & Plan case, the surgeon instrumented from the right side to the left side. An error was displayed, which indicated the surgical arm could not reach the trajectory. The manufacturer representative tried to send the surgical arm to multiple predefined locations, but the surgical arm would no respond. The emergency stop button also did not respond. A soft restart of the system was done and a few errors were displayed: Fatal CANBUS error (Error 295) and Gains lost on driver (Error 295). The light showed green on the controller connector and still would not respond to sending the surgical arm to predefined locations. A hard reboot of the system was done and they were able to continue the procedure with a new spin. There were no further issues. There was no patient harm and the procedure was delayed less than an hour. 2020-07-22 IUD (Other): No new information. 2020-07-30 IUD x5 (Other): No new information. 2020-08-18 IUD (Other): No new information. 2020-08-28 IUD (Other): No new information 2020-08-30 IUD x2 (Other): No new information. SVC KIT1317 SYSTEM MAIN CABLE | 2020-07-14 00692646, CCR x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a Scan & Plan case, the surgeon instrumented from the right side to the left side. An error was displayed, which indicated the surgical arm could not reach the trajectory. The manufacturer representative tried to send the surgical arm to multiple predefined locations, but the surgical arm would no respond. The emergency stop button also did not respond. A soft restart of the system was done and a few errors were displayed: Fatal CANBUS error (Error 295) and Gains lost on driver (Error 295). The light showed green on the controller connector and still would not respond to sending the surgical arm to predefined locations. A hard reboot of the system was done and they were able to continue the procedure with a new spin. There were no further issues. There was no patient harm and the procedure was delayed less than an hour. 2020-07-22 IUD (Other): No new information. 2020-07-30 IUD x5 (Other): No new information. 2020-08-18 IUD (Other): No new information. 2020-08-28 IUD (Other): No new information 2020-08-30 IUD x2 (Other): No new information. | xse canbus error ccr spinal procedure scan plan case surgeon instrumented right left error displayed indicated surgical arm reach trajectory tried send surgical arm multiple predefined locations surgical arm respond emergency stop button did respond soft restart errors displayed fatal canbus error gains lost driver light showed green controller connector respond sending surgical arm predefined locations hard reboot able continue procedure new spin issues procedure delayed hour iud new iud new iud new iud new iud new svc main cable | NAV4095 | NAV2047 | NAV1006 | NAV3249 | NAV4104 | NAV4095NAV2047NAV1006NAV3249NAV4104 | True
8692 | 704463060 | XSE ¦ Unable Move Manipulator | 2021-06-22T19:54:22.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 20 | UNK_MAZORX_MNPTR | UNK MAZOR X MANIPULATOR | XSE ¦ Unable Move Manipulator 2021-Jun-22 Salesforce C00228901 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manipulator would not move. The green LEDs on the manipulator handle were not illuminating. Pressure on the system was good. Hood was open. A hard reboot did not change this behavior. The LED on the back of the positioner also was mot illuminated, but the manufacturer representative was able to unlock using the software. Service tools tab in superuser showed that the controller was communicating with all the nodes it should. The positioner also would not release the teeth in the override panel, so the whole surgical system and bed frame adapter had to be removed from the bed. This is a training system with no patient present. 2021-06-23 IUD x6 (Other): No new information 2021-07-26 IUD (Other): No new information. 2021-Aug-25 IUD (other): No new information. UNK MAZOR X MANIPULATOR | 2021-Jun-22 Salesforce C00228901 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manipulator would not move. The green LEDs on the manipulator handle were not illuminating. Pressure on the system was good. Hood was open. A hard reboot did not change this behavior. The LED on the back of the positioner also was mot illuminated, but the manufacturer representative was able to unlock using the software. Service tools tab in superuser showed that the controller was communicating with all the nodes it should. The positioner also would not release the teeth in the override panel, so the whole surgical system and bed frame adapter had to be removed from the bed. This is a training system with no patient present. 2021-06-23 IUD x6 (Other): No new information 2021-07-26 IUD (Other): No new information. 2021-Aug-25 IUD (other): No new information. | xse unable manipulator salesforce outside procedure manipulator green leds manipulator handle illuminating pressure good hood open hard reboot did change behavior led positioner mot illuminated able unlock using software service tools tab superuser showed controller communicating nodes positioner release teeth override panel surgical bed frame adapter removed bed training present iud new iud new iud new unk mazor x manipulator | NAV3260 | NAV4103 | NAV3259 | NAV3250 | NAV2047 | NAV3260NAV4103NAV3259NAV3250NAV2047 | True
8283 | 704388006 | XSE - Deviated Trajectory | 2021-05-11T11:49:16.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Deviated Trajectory 2021-May-06 Salesforce C00226210 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the planned trajectory of the screws was not in line with the patient anatomy. The use of the guidance system was aborted. There was no patient harm and the procedure was delayed less than an hour. 2021-05-25 email 2021-05-25 02:15 (Rep, For): Additional information received from a manufacturer representative reported the trajectories were deviated between 3-4 mm during a thoracic procedure. No troubleshooting was done. The trajectory on the workstation was according to plan and no error messages were displayed. The surgical arm was mispositioned. The cause of the deviations was not determined. The use of the guidance system was aborted and the case was completed freehand. MAZOR X SYSTEM TPL0059 | 2021-May-06 Salesforce C00226210 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the planned trajectory of the screws was not in line with the patient anatomy. The use of the guidance system was aborted. There was no patient harm and the procedure was delayed less than an hour. 2021-05-25 email 2021-05-25 02:15 (Rep, For): Additional information received from a manufacturer representative reported the trajectories were deviated between 3-4 mm during a thoracic procedure. No troubleshooting was done. The trajectory on the workstation was according to plan and no error messages were displayed. The surgical arm was mispositioned. The cause of the deviations was not determined. The use of the guidance system was aborted and the case was completed freehand. | xse deviated trajectory salesforce spinal procedure planned trajectory screws line anatomy use aborted procedure delayed hour email additional trajectories deviated mm thoracic procedure troubleshooting trajectory workstation according plan error messages displayed surgical arm mispositioned cause deviations determined use aborted case completed freehand mazor x | NAV5000 | NAV4112 | NAV3243 | NAV3246 | NAV1012 | NAV5000NAV4112NAV3243NAV3246NAV1012 | True
4664 | 703559881 | XSE - Deviation, 3Define & Nav Issues | 2020-01-14T00:02:45.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Deviation, 3Define & Nav Issues 2020-01-13 00681983, IUD x4, CCR, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L4-S1. It was reported that the plan was reviewed and accepted by the surgeon. After powering the system for the first time, software showed "Error 3489: No connection established due to controller error". The workstation was properly shutdown and main power cable was unplugged from the power outlet. After letting the system recover after 15 minutes, the workstation was powered ON and system functioned normally. A 10 point accuracy check was performed with success. A Schanz arm was placed in the right PSIS. A bone mount bridge was attached with a Schanz bond. The 3Define failed to initiate 3 times before proceeding with generic work volume. Navigation was enabled by snapshot tracker and passive planner was used to verify safe working volume. Draw spine was completed and the arm was sent to the AP position. Arm position was adjusted and AP C-arm image was taken. Arm was sent to the OBL and the position of the 3D marker was adjusted. It was found that the virtual position of the arm on the workstation did not match the physical position of the arm. The virtual CAD model was seen to be in the "clear up" position along with the target/3D marker CAD models. OBL C-arm images were taken to continue with registration. L4 was labeled and complex was successful. Green registrations were achieved for all levels. Surgeon reviewed and approved. The arm was sent to each level to mark skin incision. The arm height required adjustment because arm the trajectory distance was too far and required LONG tools regardless of the safe working volume accepted. Arm height was lowered to enable SHORT tools for every trajectory. Navigation was checked for accuracy by placing passive planner prove into the arm guide divot and anatomical accuracy checked. Arm was sent to Left L4 and surgeon used long handle blade to make incision down to the bone. Cannula and navigated dilator were introduced and surgeon noted that the navigation was accurate. Drill guide was lightly tapped into place and short 3x30mm navigated drill was used. Implant was driven into place with the navigated driver. After placing the screw, navigation showed that the screw was slightly superior. As a result, surgeon requested to take an AP and lateral image to check final placement. (review left L4: AP & OBL images attached). Surgeon confirmed that the screw deviated and suspected patient movement. Surgeon removed left L4 screw and called for more c-arm imaging for a second registration. Arm was sent back to the AP position and the 3D marker was adjusted for C-arm image. Arm was sent to OBL position and the CAD/virtual model on the workstation continued to show a "clear up" position. C-arm images were taken and labeling was successful at L4. Registration was successful with green values at L4 and L5. Yellow value achieved at S1 and surgeon approved. Arm was sent back to left L4 and height was adjusted. Outer cannula and dilator was used. Navigation showed that the dilator was accurate and matched plan trajectory. Pedicle was drilled again and screw was introduced. Lateral image was taken and the screw was deviated superior again. Surgeon decided to leave the screw out and move onto left L5. Arm was sent to left L5 and skin incision was made followed by outer cannula/dilator. Navigation appeared to be accurate and surgeon introduced the drill guide. After placing the drill guide, placed the drill without drilling into the bone and used C-arm imaging to confirm accuracy. Surgeon did not feel the drill path would not be accurate and questioned accuracy of the system. Lateral image was taken and drill guide showed a superior and unplanned trajectory (see image attached "LEFT L5 DRILL). Surgeon decided to abort use of the system and place screws freehand with C-arm imaging. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X WORKSTATION | 2020-01-13 00681983, IUD x4, CCR, e1 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat Degenerative Disk Disease with fixation at L4-S1. It was reported that the plan was reviewed and accepted by the surgeon. After powering the system for the first time, software showed "Error 3489: No connection established due to controller error". The workstation was properly shutdown and main power cable was unplugged from the power outlet. After letting the system recover after 15 minutes, the workstation was powered ON and system functioned normally. A 10 point accuracy check was performed with success. A Schanz arm was placed in the right PSIS. A bone mount bridge was attached with a Schanz bond. The 3Define failed to initiate 3 times before proceeding with generic work volume. Navigation was enabled by snapshot tracker and passive planner was used to verify safe working volume. Draw spine was completed and the arm was sent to the AP position. Arm position was adjusted and AP C-arm image was taken. Arm was sent to the OBL and the position of the 3D marker was adjusted. It was found that the virtual position of the arm on the workstation did not match the physical position of the arm. The virtual CAD model was seen to be in the "clear up" position along with the target/3D marker CAD models. OBL C-arm images were taken to continue with registration. L4 was labeled and complex was successful. Green registrations were achieved for all levels. Surgeon reviewed and approved. The arm was sent to each level to mark skin incision. The arm height required adjustment because arm the trajectory distance was too far and required LONG tools regardless of the safe working volume accepted. Arm height was lowered to enable SHORT tools for every trajectory. Navigation was checked for accuracy by placing passive planner prove into the arm guide divot and anatomical accuracy checked. Arm was sent to Left L4 and surgeon used long handle blade to make incision down to the bone. Cannula and navigated dilator were introduced and surgeon noted that the navigation was accurate. Drill guide was lightly tapped into place and short 3x30mm navigated drill was used. Implant was driven into place with the navigated driver. After placing the screw, navigation showed that the screw was slightly superior. As a result, surgeon requested to take an AP and lateral image to check final placement. (review left L4: AP & OBL images attached). Surgeon confirmed that the screw deviated and suspected patient movement. Surgeon removed left L4 screw and called for more c-arm imaging for a second registration. Arm was sent back to the AP position and the 3D marker was adjusted for C-arm image. Arm was sent to OBL position and the CAD/virtual model on the workstation continued to show a "clear up" position. C-arm images were taken and labeling was successful at L4. Registration was successful with green values at L4 and L5. Yellow value achieved at S1 and surgeon approved. Arm was sent back to left L4 and height was adjusted. Outer cannula and dilator was used. Navigation showed that the dilator was accurate and matched plan trajectory. Pedicle was drilled again and screw was introduced. Lateral image was taken and the screw was deviated superior again. Surgeon decided to leave the screw out and move onto left L5. Arm was sent to left L5 and skin incision was made followed by outer cannula/dilator. Navigation appeared to be accurate and surgeon introduced the drill guide. After placing the drill guide, placed the drill without drilling into the bone and used C-arm imaging to confirm accuracy. Surgeon did not feel the drill path would not be accurate and questioned accuracy of the system. Lateral image was taken and drill guide showed a superior and unplanned trajectory (see image attached "LEFT L5 DRILL). Surgeon decided to abort use of the system and place screws freehand with C-arm imaging. There was no patient harm and the procedure was not delayed over an hour. | xse deviation nav issues iud ccr spinal procedure treat degenerative disk disease fixation plan reviewed accepted surgeon powering time software showed error connection established controller error workstation properly shutdown main power cable unplugged power outlet letting recover minutes workstation powered functioned normally point accuracy check performed success schanz arm placed right psis bone mount bridge attached schanz bond failed initiate times proceeding generic work volume navigation enabled snapshot tracker passive planner verify safe working volume draw spine completed arm sent ap position arm position adjusted ap carm image taken arm sent obl position marker adjusted virtual position arm workstation did match physical position arm virtual cad model seen clear position marker cad models obl carm images taken continue registration labeled complex successful green registrations achieved levels surgeon reviewed approved arm sent level mark skin incision arm height required adjustment arm trajectory distance far required long tools regardless safe working volume accepted arm height lowered enable short tools trajectory navigation checked accuracy placing passive planner prove arm guide divot anatomical accuracy checked arm sent left surgeon long handle blade make incision bone cannula navigated dilator introduced surgeon noted navigation accurate drill guide lightly tapped place short navigated drill implant driven place navigated driver placing screw navigation showed screw slightly superior surgeon requested ap lateral image check final placement surgeon confirmed screw deviated suspected movement surgeon removed left screw called carm imaging second registration arm sent ap position marker adjusted carm image arm sent obl position cadvirtual model workstation continued clear position carm images taken labeling successful registration successful green values yellow value achieved surgeon approved arm sent left height adjusted outer cannula dilator navigation showed dilator accurate matched plan trajectory pedicle drilled screw introduced lateral image taken screw deviated superior surgeon decided leave screw left arm sent left skin incision followed outer cannuladilator navigation appeared accurate surgeon introduced drill guide placing drill guide placed drill drilling bone carm imaging confirm accuracy surgeon did feel drill path accurate questioned accuracy lateral image taken drill guide showed superior unplanned trajectory surgeon decided abort use place screws freehand carm imaging procedure delayed hour unk mazor x workstation | NAV5001 | NAV3244 | NAV3237 | NAV3299 | NAV3263 | NAV5001NAV3244NAV3237NAV3299NAV3263 | False
9307 | 704606186 | XSE - OOBF/Installation Issues | 2021-09-13T05:55:40.000000+0000 | MazorX Stealth Edition | NAV2054 | OUT OF BOX FAILURES | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE - OOBF/Installation Issues 2021-Aug-07 Salesforce C00234297 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the system failed to install software, the manipulator handle cannot firmly old the weight of the Surgical system. X eye camera software not installed. Surgical system bin damaged from Bad packing during transit. Surgeon monitor Doesn't function correctly in application. There was no patient involved. 2021-09-23 e1 (Rep, For): Additional information received from a manufacturer representative reported that the the firmware file for the X-Eye camera was not on the PC. The surgical system had rub markings from bad packaging that caused damage to the plastic covers in the storage area. The surgeon monitor was unable to be linked to the application correctly. Multiple displays had to be engaged several times for the screen to register as a display and have touch screen input into the application. 2021-10-04 IUDx2, e2, e3 (Other, For): No new information. 2021-10-18 e4 (Other, For): No new information. 2021-10-28 e5, e6 (Other, For): No new information. 2021-11-07 e7, e8 (Other, For): No new information. 2021-11-08 e9 (Rep, For): Additional information received from a manufacturer representative reported the electrical panel looked like it worked, but failed to accept the installed software. The covers in the storage area were not replaced. The issue with the manipulator was not resolved. The shoulder was unable to be moved more than 10 degrees from center when on the manipulator. 2021-11-15 e10 (Other, For): No new information. 2021-11-22 e11 (Rep, For): Additional information received from a manufacturer representative reported that the surgical arm was passible, but the shoulder could not be moved a lot while the surgical system was on the manipulator, otherwise the system would tilt from 0 to 10 degrees. The X-Eye calibration file was not on the system so they had to get the correct file and load it to restore function. After several disconnections in the Windows multi-monitor display window, function was restored to the surgeon screen. The best result was to switch the monitor to duplicate, test the touch screen input in the software, and then set the software back to multiple displays. 2021-12-07 e12 (Rep, For): No new information. 2021-12-13 e13, e14 (Rep, For): No new information. 2022-01-11 e15, e16, e17, IUD (Rep, For): No new information. 2022-02-10 302721036 (Other, For): No new information. 2022-02-11 IUD x2 (Other, For): No new information. 2022-02-13 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Aug-07 Salesforce C00234297 (Rep, For): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the system failed to install software, the manipulator handle cannot firmly old the weight of the Surgical system. X eye camera software not installed. Surgical system bin damaged from Bad packing during transit. Surgeon monitor Doesn't function correctly in application. There was no patient involved. 2021-09-23 e1 (Rep, For): Additional information received from a manufacturer representative reported that the the firmware file for the X-Eye camera was not on the PC. The surgical system had rub markings from bad packaging that caused damage to the plastic covers in the storage area. The surgeon monitor was unable to be linked to the application correctly. Multiple displays had to be engaged several times for the screen to register as a display and have touch screen input into the application. 2021-10-04 IUDx2, e2, e3 (Other, For): No new information. 2021-10-18 e4 (Other, For): No new information. 2021-10-28 e5, e6 (Other, For): No new information. 2021-11-07 e7, e8 (Other, For): No new information. 2021-11-08 e9 (Rep, For): Additional information received from a manufacturer representative reported the electrical panel looked like it worked, but failed to accept the installed software. The covers in the storage area were not replaced. The issue with the manipulator was not resolved. The shoulder was unable to be moved more than 10 degrees from center when on the manipulator. 2021-11-15 e10 (Other, For): No new information. 2021-11-22 e11 (Rep, For): Additional information received from a manufacturer representative reported that the surgical arm was passible, but the shoulder could not be moved a lot while the surgical system was on the manipulator, otherwise the system would tilt from 0 to 10 degrees. The X-Eye calibration file was not on the system so they had to get the correct file and load it to restore function. After several disconnections in the Windows multi-monitor display window, function was restored to the surgeon screen. The best result was to switch the monitor to duplicate, test the touch screen input in the software, and then set the software back to multiple displays. 2021-12-07 e12 (Rep, For): No new information. 2021-12-13 e13, e14 (Rep, For): No new information. 2022-01-11 e15, e16, e17, IUD (Rep, For): No new information. 2022-02-10 302721036 (Other, For): No new information. 2022-02-11 IUD x2 (Other, For): No new information. 2022-02-13 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. | xse oobfinstallation issues salesforce outside proc edure failed install software manipulator handle firmly old weight surgical x eye camera software installed surgical bin damaged bad packing transit surgeon monitor doesnt function correctly application involved additional firmware file xeye camera pc surgical rub markings bad packaging caused damage plastic covers storage area surgeon monitor unable linked application correctly multiple displays engaged times screen register display touch screen input application new new new new additional electrical panel looked like worked failed accept installed software covers storage area replaced issue manipulator resolved shoulder unable moved degrees center manipulator new additional surgical arm passible shoulder moved lot surgical manipulator tilt degrees xeye calibration file correct file load restore function disconnections windows multimonitor display window function restored surgeon screen best switch monitor duplicate test touch screen input software set software multiple displays new new iud new new iud new iud new iud new mazorx sw installation kit | NAV3246 | NAV2054 | NAV2056 | NAV2057 | NAV3260 | NAV3246NAV2054NAV2056NAV2057NAV3260 | True
6752 | 704024054 | XSE - Screw Deviation | 2020-10-20T22:35:09.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Screw Deviation 2020-10-20 00699169 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a T9-T4 fusion, two screws were deviated. Left T7 and T8 were medial by 1-2 mm. The surgeon believed that the deviations could have been caused by the patient's high BMI or the resident could have been leaning against the patient. The screws were repositioned. There was no patient harm and the procedure was delayed less than an hour. 2020-10-27 IUD (Other): No new information. 2020-11-03 IUD (Other): No new information. 2020-12-14 email 2020-12-14 08:45 (Rep): The screws were redirected freehand and there was no patient harm that occurred as a result of deviated screws. MAZOR X SYSTEM TPL0059 | 2020-10-20 00699169 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that during a T9-T4 fusion, two screws were deviated. Left T7 and T8 were medial by 1-2 mm. The surgeon believed that the deviations could have been caused by the patient's high BMI or the resident could have been leaning against the patient. The screws were repositioned. There was no patient harm and the procedure was delayed less than an hour. 2020-10-27 IUD (Other): No new information. 2020-11-03 IUD (Other): No new information. 2020-12-14 email 2020-12-14 08:45 (Rep): The screws were redirected freehand and there was no patient harm that occurred as a result of deviated screws. | xse screw deviation spinal procedure fusion screws deviated left medial mm surgeon believed deviations caused patients high bmi resident leaning screws repositioned procedure delayed hour iud new iud new email screws redirected freehand occurred deviated screws mazor x | NAV5000 | NAV1012 | NAV4091 | NAV4112 | NAV2058 | NAV5000NAV1012NAV4091NAV4112NAV2058 | True
9277 | 704599512 | XSE - Damaged Bone Mount Bridge | 2021-09-08T18:13:26.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2188-04S | KIT SVC MAS2188-04S BM BRIDGE | XSE - Damaged Bone Mount Bridge 2021-Sep-08 Salesforce C00234094 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the screws the bone mount bridge became cross threaded. The proceeded with another bone mount bridge. However, when trying to disassemble the cross-threaded BMB, they were unable to remove the screw. There was no patient impact or delay. 2021-09-16 IUD x2 (Other): No new information. 2021-09-28 email 2021-09-28 09:03 (Rep): The manufacturer representative did not know the cause of the damage. 2021-10-01 IUD (Other): No new information. 2021-10-07 email 2021-10-07 14:01 (Rep): No new information. 2021-11-02 email 2021-11-02 09:47 (Rep): No new information. 2021-11-03 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 2021-12-12 IUD (Other): No new information. KIT SVC MAS2188-04S BM BRIDGE | 2021-Sep-08 Salesforce C00234094 (rep, hcp): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the screws the bone mount bridge became cross threaded. The proceeded with another bone mount bridge. However, when trying to disassemble the cross-threaded BMB, they were unable to remove the screw. There was no patient impact or delay. 2021-09-16 IUD x2 (Other): No new information. 2021-09-28 email 2021-09-28 09:03 (Rep): The manufacturer representative did not know the cause of the damage. 2021-10-01 IUD (Other): No new information. 2021-10-07 email 2021-10-07 14:01 (Rep): No new information. 2021-11-02 email 2021-11-02 09:47 (Rep): No new information. 2021-11-03 IUD x2 (Other): No new information. 2021-11-15 IUD (Other): No new information. 2021-12-12 IUD (Other): No new information. | xse damaged bone mount bridge salesforce spinal p rocedure screws bone mount bridge cross threaded proceeded bone mount bridge trying disassemble crossthreaded bmb unable remove screw impact delay iud new email did know cause damage iud new email new email new iud new iud new iud new kit svc bm bridge | NAV3299 | NAV3298 | NAV3250 | NAV2058 | NAV4090 | NAV3299NAV3298NAV3250NAV2058NAV4090 | True
1493 | 703202584 | X - SW ISSUE | 2019-06-12T15:56:26.000000+0000 | MazorX | NAV4099 | PLASTIC COVER ISSUES | 40 | ASM0206-01 | SURGICAL ARM | X - SW ISSUE No Connection error when starting WS – See Image on CCR. Table sync didn’t work as expected – display was duplicating the main screen instead of extend. SW Crash when using a Bovie (See SCL1800576)/1st Case MZX ver. 3.0.2.1. SZMC -T10-S2 Planning was done prior to case by the surgeon on the WS. Existing Previous L2-S1 instrumentation was planned to leave in except L2 screws was planned to be replaced by larger diameter screws. Surgeon Planned Ilium Screws. During planning the S2 VTK was cut, and didn't show the whole ROI selected for the Iliac Crest. Controller issue reported -See images. Controller persisted about an hour. Hard restart attempted and Issue was solved. (Message didn't appear again). Surgeon Screen connection was attempted, and the Default was define to duplicate the Main Screen, (The correct display on surgeon screen didn't appear). Using Windows environment the settings were adjusted to show the correct display on each screen.12:00 Skin Cut 12:25 -SW Crash due to bovie use -SW recovered to Summary details (Previous location) Following that -Firewire was pulled out to prevent additional crashes. Screws and rods were left in, for the whole Mazor X portion. 13:18 -1st Segement [T10-L1]. 13:20 -Mount Locked, Firewire was reconnected -X-eye didn't refresh. Few attempts to test camera on utilities didn't show display. SW restart was attempted, recovered and then X-eye display was refreshed. 13:22. -Mount Locked. -3Define Started. 13:24 -Draw spine -Labeled T12 3D Marker position had to be adjusted right to left to be positioned above midline, even though Draw Spine was marked correctly. Registration Red dot, Basic, Straight forward. 13:30 -Registration completed for T10-L1 and approved by Surgeon.Operation 13:32 -Right T10 Sent. 13:55 -Left L1 plan was adjusted, after surgeon commented, entry point is too medial. Axial angle convergence was increased. 2nd Segment -[S2] Ilium Screws. Surgeon Clamped on L4 Left Tulip+Rod(Previous Hardware). 14:00-3Define scan initiated, during the Scan the (3Define) Lens cover fell off. Scan didn't complete successfully because of that. After clipping it back -3Define restarted, with goo scan results. Draw Spine -again -3D-Marker position was off midline to the right side. Acquisition and registration 14:04 -AP sent and Acquired. Basic and straight forward registration for S2. Operation [S2]. Right L2 Sent -Bone felt very soft, surgeon suspected misplacement. Left L2 Sent and executed -bone dock felt much better. Surgeon took confirmatory images, and was pleased with results see attached. SURGICAL ARM | No Connection error when starting WS – See Image on CCR. Table sync didn’t work as expected – display was duplicating the main screen instead of extend. SW Crash when using a Bovie (See SCL1800576)/1st Case MZX ver. 3.0.2.1. SZMC -T10-S2 Planning was done prior to case by the surgeon on the WS. Existing Previous L2-S1 instrumentation was planned to leave in except L2 screws was planned to be replaced by larger diameter screws. Surgeon Planned Ilium Screws. During planning the S2 VTK was cut, and didn't show the whole ROI selected for the Iliac Crest. Controller issue reported -See images. Controller persisted about an hour. Hard restart attempted and Issue was solved. (Message didn't appear again). Surgeon Screen connection was attempted, and the Default was define to duplicate the Main Screen, (The correct display on surgeon screen didn't appear). Using Windows environment the settings were adjusted to show the correct display on each screen.12:00 Skin Cut 12:25 -SW Crash due to bovie use -SW recovered to Summary details (Previous location) Following that -Firewire was pulled out to prevent additional crashes. Screws and rods were left in, for the whole Mazor X portion. 13:18 -1st Segement [T10-L1]. 13:20 -Mount Locked, Firewire was reconnected -X-eye didn't refresh. Few attempts to test camera on utilities didn't show display. SW restart was attempted, recovered and then X-eye display was refreshed. 13:22. -Mount Locked. -3Define Started. 13:24 -Draw spine -Labeled T12 3D Marker position had to be adjusted right to left to be positioned above midline, even though Draw Spine was marked correctly. Registration Red dot, Basic, Straight forward. 13:30 -Registration completed for T10-L1 and approved by Surgeon.Operation 13:32 -Right T10 Sent. 13:55 -Left L1 plan was adjusted, after surgeon commented, entry point is too medial. Axial angle convergence was increased. 2nd Segment -[S2] Ilium Screws. Surgeon Clamped on L4 Left Tulip+Rod(Previous Hardware). 14:00-3Define scan initiated, during the Scan the (3Define) Lens cover fell off. Scan didn't complete successfully because of that. After clipping it back -3Define restarted, with goo scan results. Draw Spine -again -3D-Marker position was off midline to the right side. Acquisition and registration 14:04 -AP sent and Acquired. Basic and straight forward registration for S2. Operation [S2]. Right L2 Sent -Bone felt very soft, surgeon suspected misplacement. Left L2 Sent and executed -bone dock felt much better. Surgeon took confirmatory images, and was pleased with results see attached. | x sw issue connection error starting ws – image ccr table sync didn’t work expected – display duplicating main screen instead extend sw crash using bovie case mzx ver szmc planning prior case surgeon ws existing previous instrumentation planned leave screws planned replaced larger diameter screws surgeon planned ilium screws planning vtk cut didnt roi selected iliac crest controller issue images controller persisted hour hard restart attempted issue solved surgeon screen connection attempted default define duplicate main screen using windows environment settings adjusted correct display skin cut sw crash bovie use sw recovered summary details following firewire pulled prevent additional crashes screws rods left mazor x portion segement mount locked firewire reconnected xeye didnt refresh attempts test camera utilities didnt display sw restart attempted recovered xeye display refreshed mount locked started draw spine labeled marker position adjusted right left positioned midline draw spine marked correctly registration red dot basic straight forward registration completed approved surgeonoperation right sent left plan adjusted surgeon commented entry point medial axial angle convergence increased segment ilium screws surgeon clamped left tuliprod scan initiated scan lens cover fell scan didnt complete successfully clipping restarted goo scan results draw spine position midline right acquisition registration ap sent acquired basic straight forward registration operation right sent bone felt soft surgeon suspected misplacement left sent executed bone dock felt better surgeon took confirmatory images pleased results attached surgical arm | NAV4099 | NAV4091 | NAV2047 | NAV3247 | NAV3246 | NAV4099NAV4091NAV2047NAV3247NAV3246 | True
477 | 703113158 | XSE - failed 3Define/deviation/revision | 2019-04-18T13:45:39.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 30 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE - failed 3Define/deviation/revision 2019-Apr-15 SCL1900958 CCR intake form e1a (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that 3Define failed during the procedure. a second 3Define was successful. Registration was performed quickly. Upon instrumenting first trajectory (Left L4), the workstation beeped and gave an instruction to lock the arm in the software. The surgeon did not have tools in the arm guide when this notification occurred. The arm was relocked in the software and another 3Define was performed. The 3Define failed again on the first attempt. The second attempt for 3Define was successful. Registration was performed quickly, and surgeon instrumented all 4 trajectories. AP and Lateral X-ray images verified wires were accurately placed. Two days following the case, the rep received notification that the Right L4 screw had been misplaced lateral and superior. Revision procedure was scheduled for Saturday Apr-13-19 to remove screw and place new one. It was also noted that the workstation would not allow import or export of any data. The patient complained of right leg pain post operatively. Additional information received from the rep indicated that the screw inaccuracy was greater than 3.5mm but less than 10mm. The patient's symp toms had resolved after the successful revision surgery on 2019-Apr-13. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X 3DEFINE CAMERA | 2019-Apr-15 SCL1900958 CCR intake form e1a (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that 3Define failed during the procedure. a second 3Define was successful. Registration was performed quickly. Upon instrumenting first trajectory (Left L4), the workstation beeped and gave an instruction to lock the arm in the software. The surgeon did not have tools in the arm guide when this notification occurred. The arm was relocked in the software and another 3Define was performed. The 3Define failed again on the first attempt. The second attempt for 3Define was successful. Registration was performed quickly, and surgeon instrumented all 4 trajectories. AP and Lateral X-ray images verified wires were accurately placed. Two days following the case, the rep received notification that the Right L4 screw had been misplaced lateral and superior. Revision procedure was scheduled for Saturday Apr-13-19 to remove screw and place new one. It was also noted that the workstation would not allow import or export of any data. The patient complained of right leg pain post operatively. Additional information received from the rep indicated that the screw inaccuracy was greater than 3.5mm but less than 10mm. The patient's symp toms had resolved after the successful revision surgery on 2019-Apr-13. There were no additional complications reported or anticipated as a result of the event. | xse failed ccr spinal procedure failed procedure second successful registration performed quickly instrumenting trajectory workstation beeped gave instruction lock arm software surgeon did tools arm guide notification occurred arm relocked software performed failed attempt second attempt successful registration performed quickly surgeon instrumented trajectories ap lateral xray images verified wires accurately placed days following case notification right screw misplaced lateral superior revision procedure scheduled saturday remove screw place new noted workstation allow import export data complained right leg pain post operatively additional indicated screw inaccuracy greater patients symp toms resolved successful revision surgery additional complications anticipated unk mazor x camera | NAV3244 | NAV1012 | NAV5000 | NAV2058 | NAV2047 | NAV3244NAV1012NAV5000NAV2058NAV2047 | True
9096 | 704550393 | XSE - Navigation Mismatch | 2021-08-11T01:38:38.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | M728894B252 | NAVIGATION SOFTWARE MODULE 1.2 | XSE - Navigation Mismatch 2021-Aug-06 Salesforce C00232661, IUD x2, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with placement of screws and rods. It was reported that there was a mismatch between the navigation and guidance system during a L5-S1 ALIF procedure. The surgical system was mounted to the patient using a Schanz screw placed in the right PSIS and a bone mount bridge. After the patient was registered, the surgical arm was sent to the first trajectory at left L5. Navigation was found to not be matching the trajectory of the surgical arm and the planned implant on the workstation screen. The inaccuracy was only visible on the axial image and was 4-5 mm lateral to plan. Troubleshooting by the surgeon involved reintroducing the long handle blade and widen the skin incision to relieve soft tissue pressure, resent the surgical arm to the trajectory with a different elbow configuration, reverify the dilator, drill and tap, and capture a new snapshot. The surgeon decided to check the instruments at other trajectories of left S1, right L5 and right S1. Navigation was accurate at those trajectories and screws were successfully placed. The surgical arm was sent to left L5 again and navigation was still inaccurate. The surgeon decided to change the screw diameter and length to increase the safety margin. The screw was able to be successfully place. An O-arm spi n was done to check screw positions and the screw was placed in the same area where the navigation showed it was 4-5 mm lateral. The surgeon deemed the screw to be safe with acceptable placement. There was no patient harm and the procedure was delayed less than an hour. 2021-Sep-02 IUDx4 (rep): Additional information reported indicated that the system has been used since the reported complaint without issue. 2021-12-03 IUD (Other): No new information. 2021-12-20 IUD (Other): No new information. NAVIGATION SOFTWARE MODULE 1.2 | 2021-Aug-06 Salesforce C00232661, IUD x2, CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with placement of screws and rods. It was reported that there was a mismatch between the navigation and guidance system during a L5-S1 ALIF procedure. The surgical system was mounted to the patient using a Schanz screw placed in the right PSIS and a bone mount bridge. After the patient was registered, the surgical arm was sent to the first trajectory at left L5. Navigation was found to not be matching the trajectory of the surgical arm and the planned implant on the workstation screen. The inaccuracy was only visible on the axial image and was 4-5 mm lateral to plan. Troubleshooting by the surgeon involved reintroducing the long handle blade and widen the skin incision to relieve soft tissue pressure, resent the surgical arm to the trajectory with a different elbow configuration, reverify the dilator, drill and tap, and capture a new snapshot. The surgeon decided to check the instruments at other trajectories of left S1, right L5 and right S1. Navigation was accurate at those trajectories and screws were successfully placed. The surgical arm was sent to left L5 again and navigation was still inaccurate. The surgeon decided to change the screw diameter and length to increase the safety margin. The screw was able to be successfully place. An O-arm spi n was done to check screw positions and the screw was placed in the same area where the navigation showed it was 4-5 mm lateral. The surgeon deemed the screw to be safe with acceptable placement. There was no patient harm and the procedure was delayed less than an hour. 2021-Sep-02 IUDx4 (rep): Additional information reported indicated that the system has been used since the reported complaint without issue. 2021-12-03 IUD (Other): No new information. 2021-12-20 IUD (Other): No new information. | xse navigation mismatch salesforce iud ccr spinal procedure treat degenerative discs placement screws rods mismatch navigation alif procedure surgical mounted using schanz screw placed right psis bone mount bridge registered surgical arm sent trajectory left navigation matching trajectory surgical arm planned implant workstation screen inaccuracy visible axial image mm lateral plan troubleshooting surgeon involved reintroducing long handle blade widen skin incision relieve soft tissue pressure resent surgical arm trajectory different elbow configuration reverify dilator drill tap capture new snapshot surgeon decided check instruments trajectories left right right navigation accurate trajectories screws successfully placed surgical arm sent left navigation inaccurate surgeon decided change screw diameter length increase safety margin screw able successfully place oarm spi n check screw positions screw placed area navigation showed mm lateral surgeon deemed screw safe acceptable placement procedure delayed hour additional indicated complaint issue iud new iud new navigation software module | NAV5001 | NAV1012 | NAV5000 | NAV3256 | NAV5002 | NAV5001NAV1012NAV5000NAV3256NAV5002 | True
1918 | 703221082 | X - Surgical arm/software issue | 2019-06-21T19:28:59.000000+0000 | MazorX | NAV3246 | DEFECTS | 30 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X - Surgical arm/software issue It was reported that upon starting a case, the system was idle with hood open only. Joint 5 shift occurred w/o any movement of the arm or touching the screen. Reset movement didn’t fix the problem. Called service: was guided to reset gains. After this didn’t resolve the issue, motor field calibration was done for joint 5. Wasn’t effective. Trying to manipulate the motors one by one from the SW, we weren’t able to complete the movement of joints 5 and 6. Switched to RNS. Error message ‐ Mazor group error at the surgical arm MAZOR X SOFTWARE KIT VER-3.0 | It was reported that upon starting a case, the system was idle with hood open only. Joint 5 shift occurred w/o any movement of the arm or touching the screen. Reset movement didn’t fix the problem. Called service: was guided to reset gains. After this didn’t resolve the issue, motor field calibration was done for joint 5. Wasn’t effective. Trying to manipulate the motors one by one from the SW, we weren’t able to complete the movement of joints 5 and 6. Switched to RNS. Error message ‐ Mazor group error at the surgical arm | x surgical armsoftware issue starting case idle hood open joint shift occurred wo movement arm touching screen reset movement didn’t fix problem called service guided reset gains didn’t resolve issue motor field calibration joint wasn’t effective trying manipulate motors sw weren’t able complete movement joints switched rns error message ‐ mazor group error surgical arm mazor x software kit | NAV4098 | NAV3246 | NAV3247 | NAV3242 | NAV4091 | NAV4098NAV3246NAV3247NAV3242NAV4091 | True
8638 | 704455494 | XSE - Shoulder/Positioner Issues | 2021-06-17T19:51:24.000000+0000 | MazorX Stealth Edition | NAV4107 | SHOULDER NOT LOCKING/UNLOCKING | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | XSE - Shoulder/Positioner Issues 2021-Jun-17 Salesforce C00228685, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative had to use the override panel during the procedure to use the shoulder and positioner. The shoulder would not open or close when the button was pressed. The representative had to manually open and close the shoulder. The positioner would not move in the Y or Z axis and the teeth had to be opened with force. The workstation cover would not close. There was no patient harm and the procedure was not delayed. 2021-06-24 IUD (Other): No new information. UNK MAZOR X SHOULDER | 2021-Jun-17 Salesforce C00228685, IUD x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the manufacturer representative had to use the override panel during the procedure to use the shoulder and positioner. The shoulder would not open or close when the button was pressed. The representative had to manually open and close the shoulder. The positioner would not move in the Y or Z axis and the teeth had to be opened with force. The workstation cover would not close. There was no patient harm and the procedure was not delayed. 2021-06-24 IUD (Other): No new information. | xse shoulderpositioner issues salesforce iud spinal procedure use override panel procedure use shoulder positioner shoulder open close button pressed manually open close shoulder positioner y z axis teeth opened force workstation cover close procedure delayed iud new unk mazor x shoulder | NAV2058 | NAV4103 | NAV4107 | NAV4104 | NAV4106 | NAV2058NAV4103NAV4107NAV4104NAV4106 | True
1864 | 703220104 | X - Accessories Issue | 2019-06-21T09:52:17.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS2345-01 | TARGET EXTENDER SCREW | X - Accessories Issue 1.Screw that connect BMB to the schanz bond is stripped. 2. Target Extender Screw stripped as well. TARGET EXTENDER SCREW | 1.Screw that connect BMB to the schanz bond is stripped. 2. Target Extender Screw stripped as well. | x accessories issue connect bmb schanz bond stripped target extender screw stripped target extender screw | NAV3299 | NAV3295 | NAV4090 | NAV3233 | NAV4100 | NAV3299NAV3295NAV4090NAV3233NAV4100 | True
6733 | 704017186 | XSE - Unable To Detach Positioner | 2020-10-16T16:16:12.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | XSE - Unable To Detach Positioner 2020-10-16 00698960, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative was unable to detach the positioner from the bed. One tooth was locked on the bed and stuck so the positioner could not be detached or reattached. There was no patient involvement. UNK MAZOR X POSITIONER | 2020-10-16 00698960, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative was unable to detach the positioner from the bed. One tooth was locked on the bed and stuck so the positioner could not be detached or reattached. There was no patient involvement. | xse unable detach positioner iud outside procedure unable detach positioner bed tooth locked bed stuck positioner detached reattached unk mazor x positioner | NAV4103 | NAV3251 | NAV3250 | NAV4104 | NAV4105 | NAV4103NAV3251NAV3250NAV4104NAV4105 | True
7866 | 704299168 | XSE - X-eye Not Functioning | 2021-03-23T17:45:48.000000+0000 | MazorX Stealth Edition | NAV3262 | PC ISSUES | 50 | ASM0215-05R | MAZOR X PC STATION - REFURBISHED | XSE - X-eye Not Functioning 2021-Mar-19 Salesforce C00222991 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the X-eye was not working properly. No patient was present. 2021-03-31 IUD x6 (Other): No new information. 2021-06-02 IUD x2 (Other): No new information. 2021-06-13 IUD (Other): No new information. MAZOR X PC STATION - REFURBISHED | 2021-Mar-19 Salesforce C00222991 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the X-eye was not working properly. No patient was present. 2021-03-31 IUD x6 (Other): No new information. 2021-06-02 IUD x2 (Other): No new information. 2021-06-13 IUD (Other): No new information. | xse xeye functioning salesforce outside procedure xeye working properly present iud new iud new iud new mazor x pc station refurbished | NAV4110 | NAV3262 | NAV2054 | NAV2058 | NAV3242 | NAV4110NAV3262NAV2054NAV2058NAV3242 | True
3598 | 703383605 | XSE - 3Define Camera/SW Issue | 2019-09-27T12:22:50.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - 3Define Camera/SW Issue 2019-09-23 SCL1902921, SCL1902929 Intake Form x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera did not work and the plan for the surgical order went back to default when going back to planning from the operate mode. The surgical order also went back to default between planning on the laptop prior to the procedure and loading the plan onto the workstation. There was no patient harm and the procedure was not delayed more than an hour. 2019-09-27 email 2019-09-27 09:00 (Rep): Additional information received from a manufacturer representative reported that no issues were found with the 3Define scan when completing multiple surface scans. The representative noted that for every case that they planned out or attempted to change the screw placement order, the system would always default back to zig-zag starting from the top right. The cause was not determined. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-09-23 SCL1902921, SCL1902929 Intake Form x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the 3Define camera did not work and the plan for the surgical order went back to default when going back to planning from the operate mode. The surgical order also went back to default between planning on the laptop prior to the procedure and loading the plan onto the workstation. There was no patient harm and the procedure was not delayed more than an hour. 2019-09-27 email 2019-09-27 09:00 (Rep): Additional information received from a manufacturer representative reported that no issues were found with the 3Define scan when completing multiple surface scans. The representative noted that for every case that they planned out or attempted to change the screw placement order, the system would always default back to zig-zag starting from the top right. The cause was not determined. | xse camerasw issue spinal procedure camera did work plan surgical order went default going planning operate mode surgical order went default planning laptop prior procedure loading plan workstation procedure delayed hour email additional issues scan completing multiple surface scans noted case planned attempted change screw placement order default zigzag starting right cause determined mazor x software installation kit | NAV5000 | NAV3247 | NAV3244 | NAV3246 | NAV3236 | NAV5000NAV3247NAV3244NAV3246NAV3236 | True
9900 | 704744846 | XSE - SRA Joint Shift | 2021-12-02T19:10:47.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-03S | SURGICAL ARM SPARE PART ASM0206-03S | XSE - SRA Joint Shift 2021-Dec-02 Salesforce C00239066 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a joint shift error was displayed for joint 3. The manufacturer representative was able to reset the movement, but the error messages continued to be displayed. The surgical arm also stuttered sometimes when moving. There was no patient involvement. 2021-12-08 IUD (Other): No new information. 2021-12-15 IUD (Other): No new information. 2022-01-09 IUD (Other): No new information. 2022-01-26 302704718 (Other): No new information. 2022-02-09 IUD x3 (Other): No new information. SURGICAL ARM SPARE PART ASM0206-03S | 2021-Dec-02 Salesforce C00239066 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a joint shift error was displayed for joint 3. The manufacturer representative was able to reset the movement, but the error messages continued to be displayed. The surgical arm also stuttered sometimes when moving. There was no patient involvement. 2021-12-08 IUD (Other): No new information. 2021-12-15 IUD (Other): No new information. 2022-01-09 IUD (Other): No new information. 2022-01-26 302704718 (Other): No new information. 2022-02-09 IUD x3 (Other): No new information. | xse sra joint shift salesforce outside procedure joint shift error displayed joint able reset movement error messages continued displayed surgical arm stuttered moving iud new iud new iud new new iud new surgical arm spare | NAV4098 | NAV2054 | NAV4101 | NAV3249 | NAV3251 | NAV4098NAV2054NAV4101NAV3249NAV3251 | True
5097 | 703597626 | XSE - Schanz Arm | 2020-02-05T00:12:41.000000+0000 | MazorX Stealth Edition | NAV3298 | PLATFORM ISSUES | 20 | MAS2351-04 | SCHANZ ARM | XSE - Schanz Arm 2020-2-4 00683395(Rep): Medtronic received information regarding a guidance system being used. It was reported that the schanz arm would not lock with the system. The was no patient involved. 2020-02-06 Email 2020-02-06 1115 (Rep): Additional information received from the manufacturer representative reporting that troubleshooting was done with the schanz arm during the case - attempts were made to loosen and tighten up again. The cause of the issue was determined to be wear and tear. The representative stated the item could be returned when the replacement and shipping label arrived. The email with request for shipping information was forwarded to the case owner. 2020-02-17 Email 2020-02-24 13:00, IUD x2 (Rep, Other): No new information. 2020-03-16 IUD (Other): No new information. SCHANZ ARM | 2020-2-4 00683395(Rep): Medtronic received information regarding a guidance system being used. It was reported that the schanz arm would not lock with the system. The was no patient involved. 2020-02-06 Email 2020-02-06 1115 (Rep): Additional information received from the manufacturer representative reporting that troubleshooting was done with the schanz arm during the case - attempts were made to loosen and tighten up again. The cause of the issue was determined to be wear and tear. The representative stated the item could be returned when the replacement and shipping label arrived. The email with request for shipping information was forwarded to the case owner. 2020-02-17 Email 2020-02-24 13:00, IUD x2 (Rep, Other): No new information. 2020-03-16 IUD (Other): No new information. | xse schanz arm schanz arm lock involved email additional reporting troubleshooting schanz arm case attempts loosen tighten cause issue determined wear tear stated item returned replacement shipping label arrived email request shipping forwarded case owner email iud new iud new schanz arm | NAV3299 | NAV3298 | NAV3295 | NAV2054 | NAV2055 | NAV3299NAV3298NAV3295NAV2054NAV2055 | True
1609 | 703207910 | X- Workstation Issues | 2019-06-14T16:44:54.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X- Workstation Issues Plan was done by Dr. Francis the morning of the surgery. L5-S1 were planned for an open procedure. There were no concerns discussed or changes made to the plan. A single schanz pin was utilized and placed in the right PSIS. there were no complications mounting the platform. Although the pictures were poor quality, registration was achieved on the second attempt. the 3d marker had to be adjusted and a new image was necessary for clarity. There were no complications throughout he procedure. all wires and screws were placed accurately andconfirmed with x-ray. after the procedure when removing the system from the bed, there was some difficulty. after attaching the manipulator, unlocking the air pressure and hitting the unlock button, there was no air pressure assisting with removing from the bed. all of the weight was present when removing from the bed. after the patient had left the room, I re attached the robot to the bed, again with no air pressure assisting. after multiple trouble shooting it was noted that thehood was stuck in between the open and closed state due to friction against the side of the workstation. the door where the image adapter lives seemed to be loose and sat in a position that rubbed against the hood when opening and closing. because of this, the hood was not fully open, causing there to be no air pressure distributed to release the robot from the bed. the hood was adjusted and reset to the full open position and this was recognized by the green dot lit up under the sensors. the robot was then removed. the right door on the workstation still remains in an awkward position but doesn't seem to cause any friction anymore when opening and closing the hood. UNK MAZOR X WORKSTATION | Plan was done by Dr. Francis the morning of the surgery. L5-S1 were planned for an open procedure. There were no concerns discussed or changes made to the plan. A single schanz pin was utilized and placed in the right PSIS. there were no complications mounting the platform. Although the pictures were poor quality, registration was achieved on the second attempt. the 3d marker had to be adjusted and a new image was necessary for clarity. There were no complications throughout he procedure. all wires and screws were placed accurately andconfirmed with x-ray. after the procedure when removing the system from the bed, there was some difficulty. after attaching the manipulator, unlocking the air pressure and hitting the unlock button, there was no air pressure assisting with removing from the bed. all of the weight was present when removing from the bed. after the patient had left the room, I re attached the robot to the bed, again with no air pressure assisting. after multiple trouble shooting it was noted that thehood was stuck in between the open and closed state due to friction against the side of the workstation. the door where the image adapter lives seemed to be loose and sat in a position that rubbed against the hood when opening and closing. because of this, the hood was not fully open, causing there to be no air pressure distributed to release the robot from the bed. the hood was adjusted and reset to the full open position and this was recognized by the green dot lit up under the sensors. the robot was then removed. the right door on the workstation still remains in an awkward position but doesn't seem to cause any friction anymore when opening and closing the hood. | x workstation issues plan dr francis morning surgery planned open procedure concerns discussed changes plan single schanz pin utilized placed right psis complications mounting platform pictures poor quality registration achieved second attempt marker adjusted new image necessary clarity complications procedure wires screws placed accurately andconfirmed xray procedure removing bed difficulty attaching manipulator unlocking air pressure hitting unlock button air pressure assisting removing bed weight present removing bed left room attached robot bed air pressure assisting multiple trouble shooting noted thehood stuck open closed state friction workstation door image adapter lives loose sat position rubbed hood opening closing hood fully open causing air pressure distributed release robot bed hood adjusted reset open position recognized green dot lit sensors robot removed right door workstation remains awkward position doesnt cause friction anymore opening closing hood unk mazor x workstation | NAV3264 | NAV2058 | NAV3250 | NAV3243 | NAV3263 | NAV3264NAV2058NAV3250NAV3243NAV3263 | True
7269 | 704161005 | XSE - Unable To Boot/Comm Error | 2021-01-07T21:25:30.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Unable To Boot/Comm Error 2021-Jan-06 Salesforce C00218917 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was turned on and a message stating the sensors were not communicating was displayed. A hard restart of the system was done. When the system was powered back on, the screens remained black. Another hard reboot of the system was done with no change. The representative also tried hooking the system up to an external monitor, but there was no information showed. After a couple of additional hard restarts of the system, the system booted normally. The surgeon screen was disconnected as a temporary fix. There was no patient involvement. 2021-01-11 IUD x3 (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Jan-06 Salesforce C00218917 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the guidance system was turned on and a message stating the sensors were not communicating was displayed. A hard restart of the system was done. When the system was powered back on, the screens remained black. Another hard reboot of the system was done with no change. The representative also tried hooking the system up to an external monitor, but there was no information showed. After a couple of additional hard restarts of the system, the system booted normally. The surgeon screen was disconnected as a temporary fix. There was no patient involvement. 2021-01-11 IUD x3 (Other): No new information. | xse unable bootcomm error salesforce outside procedure turned message stating sensors communicating displayed hard restart powered screens remained black hard reboot change tried hooking external monitor showed couple additional hard restarts booted normally surgeon screen disconnected temporary fix iud new unk mazor x workstation | NAV3261 | NAV2058 | NAV3253 | NAV2047 | NAV4104 | NAV3261NAV2058NAV3253NAV2047NAV4104 | True
8207 | 704370831 | XSE ¦ Boot Up Error Messages | 2021-04-30T12:24:42.000000+0000 | MazorX Stealth Edition | NAV2056 | SOFTWARE ISSUES | 20 | KIT1382 | KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | XSE ¦ Boot Up Error Messages 2021-Apr-30 Salesforce C00225582 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the sites system booted up with the following error messages: 1. Failed to load navigation module nav mode will be disable 9254 2. MDT DB installation failed MDT implants will not be supported 10145 3. DB version mismatch The manufacturer representative stated the system was hard rebooted and booted up as expected after. No patient was present. 2021-05-12 IUD x2 (Other): No new information. 2021-05-25 IUD x4 (Other): No new information. KIT1382 MAZORX SW INSTALLATION KIT VR5.0 | 2021-Apr-30 Salesforce C00225582 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the sites system booted up with the following error messages: 1. Failed to load navigation module nav mode will be disable 9254 2. MDT DB installation failed MDT implants will not be supported 10145 3. DB version mismatch The manufacturer representative stated the system was hard rebooted and booted up as expected after. No patient was present. 2021-05-12 IUD x2 (Other): No new information. 2021-05-25 IUD x4 (Other): No new information. | xse boot error messages salesforce outside procedure sites booted following error messages failed load navigation module nav mode disable mdt db installation failed mdt implants supported db version mismatch stated hard rebooted booted expected present iud new iud new mazorx sw installation kit | NAV3246 | NAV2047 | NAV3248 | NAV2056 | NAV3247 | NAV3246NAV2047NAV3248NAV2056NAV3247 | True
5539 | 703749481 | X - Pneumatic Leak in the Shoulder | 2020-05-12T18:46:50.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 30 | ASM0207-02S | KIT SVC ASM0207-02S SHOULDER | X - Pneumatic Leak in the Shoulder 2020-05-12 00688706 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that when attempting to lock the shoulder the pressure drops quickly. The compressor does come on and can balance out the pressure between 5.8 and 6 but it never reaches the full 7.2 required to turn off. Once the shoulder is unlocked the pressure builds up and compressor turns off. The pressure will stay around 7 until the shoulder is locked again. Also the shoulder button never stops blinking making it appear as if the brake is not opening and closing completely. Also the pressure drops when moving the manipulator with the positioner attached but seem to be related to the shoulder. No patient was present. 2020-05-13 IUD x8 (Other): No new information. 2020-05-27 IUD x2 (Other): No new information. 2020-07-05 IUD x2 (Other): No new information. 2020-07-09 IUD (Other): No new information. 2020-07-20 IUD (Other): No new information. KIT SVC ASM0207-02S SHOULDER | 2020-05-12 00688706 (Rep): Medtronic received information regarding a guidance system being used outside of procedure. It was reported that when attempting to lock the shoulder the pressure drops quickly. The compressor does come on and can balance out the pressure between 5.8 and 6 but it never reaches the full 7.2 required to turn off. Once the shoulder is unlocked the pressure builds up and compressor turns off. The pressure will stay around 7 until the shoulder is locked again. Also the shoulder button never stops blinking making it appear as if the brake is not opening and closing completely. Also the pressure drops when moving the manipulator with the positioner attached but seem to be related to the shoulder. No patient was present. 2020-05-13 IUD x8 (Other): No new information. 2020-05-27 IUD x2 (Other): No new information. 2020-07-05 IUD x2 (Other): No new information. 2020-07-09 IUD (Other): No new information. 2020-07-20 IUD (Other): No new information. | x pneumatic leak shoulder outside procedure attempting lock shoulder pressure drops quickly compressor does come balance pressure reaches required turn shoulder unlocked pressure builds compressor turns pressure stay shoulder locked shoulder button stops blinking making appear brake opening closing completely pressure drops moving manipulator positioner attached related shoulder present iud new iud new iud new iud new iud new kit svc shoulder | NAV3264 | NAV3252 | NAV4107 | NAV3259 | NAV1006 | NAV3264NAV3252NAV4107NAV3259NAV1006 | True
9663 | 704692009 | XSE - Difficulties Mounting SSY | 2021-11-01T18:27:39.000000+0000 | MazorX Stealth Edition | NAV4104 | UNABLE TO MOVE | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | XSE - Difficulties Mounting SSY 2021-Oct-05 Salesforce C00237115, Intake Form (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon had trouble mounting the surgical system to the bed frame adapter. The positioner was very difficult to maneuver up and down and from side to side. The manufacturer representative noted that attaching the surgical system took a long time. There was no patient harm and the procedure was delayed less than an hour. 2021-11-14 email 2021-11-14 06:00 (Rep, For): Additional information received from a manufacturer representative reported that the pressure was high, but they were still able to mount the surgical system to the bed. UNK MAZOR X POSITIONER | 2021-Oct-05 Salesforce C00237115, Intake Form (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the surgeon had trouble mounting the surgical system to the bed frame adapter. The positioner was very difficult to maneuver up and down and from side to side. The manufacturer representative noted that attaching the surgical system took a long time. There was no patient harm and the procedure was delayed less than an hour. 2021-11-14 email 2021-11-14 06:00 (Rep, For): Additional information received from a manufacturer representative reported that the pressure was high, but they were still able to mount the surgical system to the bed. | xse difficulties mounting ssy salesforce spinal procedure surgeon trouble mounting surgical bed frame adapter positioner difficult maneuver noted attaching surgical took long time procedure delayed hour email additional pressure high able mount surgical bed unk mazor x positioner | NAV3250 | NAV3264 | NAV4105 | NAV3251 | NAV4103 | NAV3250NAV3264NAV4105NAV3251NAV4103 | False
6519 | 703963517 | XSE - SRS Mounting Issues | 2020-09-17T19:58:18.000000+0000 | MazorX Stealth Edition | NAV4103 | POSITIONER NOT LOCKING/UNLOCKING | 20 | ASM0214-02 | SSY | XSE - SRS Mounting Issues 2020-09-17 00696929 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the surgical system was being put on the bed by the surgeon, one of the teeth locked on and the other did not. The surgeon was unable to unlock the tooth in the override panel and the pressure would not increase. The bed frame adapter was manually detached from the bed and then the surgical system was able to be removed. The pressure in the system then started to build and the representative was able to mount the system to the bed. The surgeon did not feel comfortable using the system so they decided to abort the case and reschedule it for a later date. The procedure was delayed over an hour. The patient was no patient harm, but the patient was under anesthesia. 2020-09-23 IUD x4 (Other): The manufacturer representative was able to run through system checks and was able to place the robot on and off the bed 10 times with little effort, the system had no issues moving after attached and had no issues going back into storage mode. Ran through all accuracy checks with no issues, all cameras were working, everything seemed to be in working order. Leak was due to teeth not being mounted properly on BFA. SSY | 2020-09-17 00696929 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when the surgical system was being put on the bed by the surgeon, one of the teeth locked on and the other did not. The surgeon was unable to unlock the tooth in the override panel and the pressure would not increase. The bed frame adapter was manually detached from the bed and then the surgical system was able to be removed. The pressure in the system then started to build and the representative was able to mount the system to the bed. The surgeon did not feel comfortable using the system so they decided to abort the case and reschedule it for a later date. The procedure was delayed over an hour. The patient was no patient harm, but the patient was under anesthesia. 2020-09-23 IUD x4 (Other): The manufacturer representative was able to run through system checks and was able to place the robot on and off the bed 10 times with little effort, the system had no issues moving after attached and had no issues going back into storage mode. Ran through all accuracy checks with no issues, all cameras were working, everything seemed to be in working order. Leak was due to teeth not being mounted properly on BFA. | xse srs mounting issues spinal procedure surgical bed surgeon teeth locked did surgeon unable unlock tooth override panel pressure increase bed frame adapter manually detached bed surgical able removed pressure started build able mount bed surgeon did feel comfortable using decided abort case reschedule later date procedure delayed hour anesthesia iud able run checks able place robot bed times little effort issues moving attached issues going storage mode ran accuracy checks issues cameras working working order leak teeth mounted properly bfa ssy | NAV3250 | NAV3264 | NAV4103 | NAV3246 | NAV3253 | NAV3250NAV3264NAV4103NAV3246NAV3253 | True
4186 | 703511020 | X - Error 4059 | 2019-12-10T19:23:31.000000+0000 | MazorX | NAV3262 | PC ISSUES | 30 | ASM0215-04 | MAZOR-X PC STATION | X - Error 4059 2019-12-10 00679974, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a 4059 error message was displayed every few seconds when the system was powered on. A hard restart was done and the issue was resolved. There was no patient involvement. 2019-12-12 IUD x4 (Other): No new information 2020-1-14 IUD (Other): No new information. 2020-02-10 IUD x2 (Rep): No new information. 2020-03-12 IUD (Other): No new information. MAZOR-X PC STATION | 2019-12-10 00679974, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a 4059 error message was displayed every few seconds when the system was powered on. A hard restart was done and the issue was resolved. There was no patient involvement. 2019-12-12 IUD x4 (Other): No new information 2020-1-14 IUD (Other): No new information. 2020-02-10 IUD x2 (Rep): No new information. 2020-03-12 IUD (Other): No new information. | x error iud outside procedure error message displayed seconds powered hard restart issue resolved iud new iud new iud new iud new mazorx pc station | NAV2047 | NAV3246 | NAV3262 | NAV4095 | NAV3261 | NAV2047NAV3246NAV3262NAV4095NAV3261 | True
9286 | 704602096 | XSE ¦ 3Define Issues | 2021-09-09T18:14:50.000000+0000 | MazorX Stealth Edition | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | XSE ¦ 3Define Issues 2021-Sep-09 Salesforce C00234200 (rep, hcp): Medtronic received information regarding a guidance system. It was reported that during a case 3Define was unsuccessful. They attempted 3Define twice. They moved forward with a generic work volume. There was no patient harm. 2021-09-26 IUD (Other): No new information. 2021-10-11 IUD (Other): No new information. 2021-10-18 IUD (Other): The issue occurred during a spinal procedure. There was about a 2 minute delay to the procedure. 2021-10-18 email 2021-10-18 10:16 (Rep): The cause of the issue was stated to be not determined and has not occurred again since. 2021-11-15 IUD (Other): No new information. UNK MAZOR X 3DEFINE CAMERA | 2021-Sep-09 Salesforce C00234200 (rep, hcp): Medtronic received information regarding a guidance system. It was reported that during a case 3Define was unsuccessful. They attempted 3Define twice. They moved forward with a generic work volume. There was no patient harm. 2021-09-26 IUD (Other): No new information. 2021-10-11 IUD (Other): No new information. 2021-10-18 IUD (Other): The issue occurred during a spinal procedure. There was about a 2 minute delay to the procedure. 2021-10-18 email 2021-10-18 10:16 (Rep): The cause of the issue was stated to be not determined and has not occurred again since. 2021-11-15 IUD (Other): No new information. | xse issues salesforce case unsuccessful attempted twice moved forward generic work volume iud new iud new iud issue occurred spinal procedure minute delay procedure email cause issue stated determined occurred iud new unk mazor x camera | NAV3244 | NAV2058 | NAV3261 | NAV4110 | NAV2047 | NAV3244NAV2058NAV3261NAV4110NAV2047 | True
2433 | 703235767 | X - Workstation Issue | 2019-07-01T12:09:23.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - Workstation Issue 3Define not working, not recognized by system. UNK MAZOR X 3DEFINE CAMERA | 3Define not working, not recognized by system. | x workstation issue working recognized unk mazor x camera | NAV3244 | NAV2058 | NAV2047 | NAV2054 | NAV3264 | NAV3244NAV2058NAV2047NAV2054NAV3264 | True
10322 | 704853819 | XSE - Joint Shift Error | 2022-02-09T01:08:30.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 30 | ASM0206-01 | SURGICAL ARM | XSE - Joint Shift Error 2022-Feb-08 Salesforce C00242907 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgeon was using Bovie during the procedure and the Bovie tool made contact with an instrument placed down the arm guide while it was attached to the surgical arm. The system displayed a joint shift error. The manufacturer representative tried to send the surgical arm to the up position and back to the trajectory, but the error message remained. The surgeon decided to abort the use of the guidance system and complete the procedure freehand. After the procedure, a hard reboot of the system was done and the error message was cleared. The surgical arm successfully passed an accuracy and stress test. There was no patient harm and the procedure was delayed less than an hour. 2022-02-15 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. 2022-03-13 302755530, IUD (Other): No new information. 2022-03-18 IUD x2 (Other): No new information. 2022-03-30 IUD x4 (Other): No new information. SURGICAL ARM | 2022-Feb-08 Salesforce C00242907 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the surgeon was using Bovie during the procedure and the Bovie tool made contact with an instrument placed down the arm guide while it was attached to the surgical arm. The system displayed a joint shift error. The manufacturer representative tried to send the surgical arm to the up position and back to the trajectory, but the error message remained. The surgeon decided to abort the use of the guidance system and complete the procedure freehand. After the procedure, a hard reboot of the system was done and the error message was cleared. The surgical arm successfully passed an accuracy and stress test. There was no patient harm and the procedure was delayed less than an hour. 2022-02-15 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. 2022-03-13 302755530, IUD (Other): No new information. 2022-03-18 IUD x2 (Other): No new information. 2022-03-30 IUD x4 (Other): No new information. | xse joint shift error salesforce spinal p rocedure surgeon using bovie procedure bovie tool contact instrument placed arm guide attached surgical arm displayed joint shift error tried send surgical arm position trajectory error message remained surgeon decided abort use complete procedure freehand procedure hard reboot error message cleared surgical arm successfully passed accuracy stress test procedure delayed hour iud new iud new iud new iud new iud new surgical arm | NAV4098 | NAV3249 | NAV3246 | NAV2047 | NAV4095 | NAV4098NAV3249NAV3246NAV2047NAV4095 | True
5372 | 703662862 | XSE - L4 Right Trajectory 3-4mm Lateral | 2020-03-11T23:06:02.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - L4 Right Trajectory 3-4mm Lateral 2020-03-11 00686227 (rep, hcp): Medtronic received information regarding a guidance system being used for a spinal procedure. It was reported that there was an alleged inaccuracy during an L4-S1 case. On the L4 right trajectory, the arm was 3-4 millimeters lateral. The surgeon reported feeling no skiving. The surgeon checked the trajectory by holding the dilator on plan and it showed as lateral to the intended plan. Troubleshooting was performed by adjusting the plan to be medial from the intended trajectory, which resulted in an acceptable trajectory. There was a less than 1 hour delay to the procedure and no impact to patient outcome. 2020-03-13 email, image 2020-03-16 17:48, CCR x2 (Rep, HCP): AAdditional information received from a manufacturer representative reported a single position lateral fusion case for L4-S1 was being performed. The RBT set up and registration was performed with out any problems. Registration was achieved using flouro with basic. Green values were achieved for all levels. Since a lateral case was being performed, the surgeon was placing all screws percutaneously. The surgical arm was sent to all right sided trajectories first in order to mark the incision points. Once all incisions were marked, the arm was cleared up and the surgeon performed a Wiltse incision. The surgeon requested to drill holes at all levels on each side first, then place screws. Right S1 was sent first and drilled with no issues. Right L5 was then sent and drilled with no issues. Right L4 was then sent and while positioning the cannula and dilator, the trajectory was shown on navigation to be lateral. The surgeon removed the tools and used the long scalpel to make a clear incision all the way down to bone. The surgeon then reintroduced the dilator and cannula and the dilator still showed a lateral trajec tory. The surgeon removed the tools from the end effector and navigation was checked with the pointer probe in the end effector. Navigation was shown to be accurate. The surgeon placed the dilator and outer cannula back through the arm guide and navigation showed a lateral trajectory again. The surgeon reported feeling no skiving and decided to proceed with drilling and screw placement. Right L4 was drilled and the screw was placed. The arm was then sent to right L5 again and the screw was placed. Right S1 was then sent again and the screw was placed. The surgeon then sent the arm to all Left sided trajectories and repeated the same workflow. Incision sites were marked first then the arm was sent again to each trajectory, starting with left S1, and drilled. After drilling the left L4 trajectory, the screw was placed. Then the left L5 trajectory was sent and screw placed, followed by left S1 trajectory and screw placed. The surgical arm was detached from the patient and the surgeon proceeded to perform his interbody cage procedure at this point. Once completed, everything was checked using flouro. It was found that all screws were placed accurately except for the right L4. Right L4 was resting lateral to the pedicle. The surgeon decided to bring the arm back in and perform a scan and plan to correct the right L4 screw. Platform used was the same as before. Scan and plan registration was performed successfully and right L4 was planned for the same position as before. The arm was sent to the right L4 trajectory. The surgeon used the scalpel first to ensure a clear path to bone, then introduced the dilator and outer cannula. Again the navigation showed a lateral trajectory. The surgeon removed the tools from the incision and held them through the end effector above the skin. This was done to see if the robotic trajectory was lateral or if the instruments were being pushed lateral inside of the incision. When held above the skin, through the end effector, navigation showed the green dashed line lateral to the planned trajectory. The right L4 screw plan was then shifted medial and the arm was resent. Surgeon then reintroduced the dilator and outer cannula into the incision. Navigation showed the tools lateral to the new plan but acceptable in terms of alignment with the pedicle. The surgeon proceeded to drill the trajectory and place the screw. Once finished, flouro was brought in to check and it was found that the screw was placed lateral to plan, but acceptable in the pedicle. Robotic arm was then detached and surgeon continued to place rods and close the incisions. 2020-03-18 IUD x3 (Other): Additional information received from a manufacturer representative reported that the system has been used in multiple prone cases since inaccuracy and cases have had no issues. Also rep ran and passed 10 point accuracy test, and 3 point test. Inaccuracy could be possible due to compression during lateral case. This is not a system issue. MAZOR X SYSTEM TPL0059 | 2020-03-11 00686227 (rep, hcp): Medtronic received information regarding a guidance system being used for a spinal procedure. It was reported that there was an alleged inaccuracy during an L4-S1 case. On the L4 right trajectory, the arm was 3-4 millimeters lateral. The surgeon reported feeling no skiving. The surgeon checked the trajectory by holding the dilator on plan and it showed as lateral to the intended plan. Troubleshooting was performed by adjusting the plan to be medial from the intended trajectory, which resulted in an acceptable trajectory. There was a less than 1 hour delay to the procedure and no impact to patient outcome. 2020-03-13 email, image 2020-03-16 17:48, CCR x2 (Rep, HCP): AAdditional information received from a manufacturer representative reported a single position lateral fusion case for L4-S1 was being performed. The RBT set up and registration was performed with out any problems. Registration was achieved using flouro with basic. Green values were achieved for all levels. Since a lateral case was being performed, the surgeon was placing all screws percutaneously. The surgical arm was sent to all right sided trajectories first in order to mark the incision points. Once all incisions were marked, the arm was cleared up and the surgeon performed a Wiltse incision. The surgeon requested to drill holes at all levels on each side first, then place screws. Right S1 was sent first and drilled with no issues. Right L5 was then sent and drilled with no issues. Right L4 was then sent and while positioning the cannula and dilator, the trajectory was shown on navigation to be lateral. The surgeon removed the tools and used the long scalpel to make a clear incision all the way down to bone. The surgeon then reintroduced the dilator and cannula and the dilator still showed a lateral trajec tory. The surgeon removed the tools from the end effector and navigation was checked with the pointer probe in the end effector. Navigation was shown to be accurate. The surgeon placed the dilator and outer cannula back through the arm guide and navigation showed a lateral trajectory again. The surgeon reported feeling no skiving and decided to proceed with drilling and screw placement. Right L4 was drilled and the screw was placed. The arm was then sent to right L5 again and the screw was placed. Right S1 was then sent again and the screw was placed. The surgeon then sent the arm to all Left sided trajectories and repeated the same workflow. Incision sites were marked first then the arm was sent again to each trajectory, starting with left S1, and drilled. After drilling the left L4 trajectory, the screw was placed. Then the left L5 trajectory was sent and screw placed, followed by left S1 trajectory and screw placed. The surgical arm was detached from the patient and the surgeon proceeded to perform his interbody cage procedure at this point. Once completed, everything was checked using flouro. It was found that all screws were placed accurately except for the right L4. Right L4 was resting lateral to the pedicle. The surgeon decided to bring the arm back in and perform a scan and plan to correct the right L4 screw. Platform used was the same as before. Scan and plan registration was performed successfully and right L4 was planned for the same position as before. The arm was sent to the right L4 trajectory. The surgeon used the scalpel first to ensure a clear path to bone, then introduced the dilator and outer cannula. Again the navigation showed a lateral trajectory. The surgeon removed the tools from the incision and held them through the end effector above the skin. This was done to see if the robotic trajectory was lateral or if the instruments were being pushed lateral inside of the incision. When held above the skin, through the end effector, navigation showed the green dashed line lateral to the planned trajectory. The right L4 screw plan was then shifted medial and the arm was resent. Surgeon then reintroduced the dilator and outer cannula into the incision. Navigation showed the tools lateral to the new plan but acceptable in terms of alignment with the pedicle. The surgeon proceeded to drill the trajectory and place the screw. Once finished, flouro was brought in to check and it was found that the screw was placed lateral to plan, but acceptable in the pedicle. Robotic arm was then detached and surgeon continued to place rods and close the incisions. 2020-03-18 IUD x3 (Other): Additional information received from a manufacturer representative reported that the system has been used in multiple prone cases since inaccuracy and cases have had no issues. Also rep ran and passed 10 point accuracy test, and 3 point test. Inaccuracy could be possible due to compression during lateral case. This is not a system issue. | xse right trajectory lateral spinal procedure alleged inaccuracy case right trajectory arm millimeters lateral surgeon feeling skiving surgeon checked trajectory holding dilator plan showed lateral intended plan troubleshooting performed adjusting plan medial intended trajectory resulted acceptable trajectory hour delay procedure impact outcome email image ccr aadditional single position lateral fusion case performed rbt set registration performed problems registration achieved using flouro basic green values achieved levels lateral case performed surgeon placing screws percutaneously surgical arm sent right sided trajectories order mark incision points incisions marked arm cleared surgeon performed wiltse incision surgeon requested drill holes levels place screws right sent drilled issues right sent drilled issues right sent positioning cannula dilator trajectory shown navigation lateral surgeon removed tools long scalpel make clear incision way bone surgeon reintroduced dilator cannula dilator showed lateral trajec tory surgeon removed tools end effector navigation checked pointer probe end effector navigation shown accurate surgeon placed dilator outer cannula arm guide navigation showed lateral trajectory surgeon feeling skiving decided proceed drilling screw placement right drilled screw placed arm sent right screw placed right sent screw placed surgeon sent arm left sided trajectories repeated workflow incision sites marked arm sent trajectory starting left drilled drilling left trajectory screw placed left trajectory sent screw placed followed left trajectory screw placed surgical arm detached surgeon proceeded perform interbody cage procedure point completed checked using flouro screws placed accurately right right resting lateral pedicle surgeon decided bring arm perform scan plan correct right screw platform scan plan registration performed successfully right planned position arm sent right trajectory surgeon scalpel ensure clear path bone introduced dilator outer cannula navigation showed lateral trajectory surgeon removed tools incision held end effector skin robotic trajectory lateral instruments pushed lateral inside incision held skin end effector navigation showed green dashed line lateral planned trajectory right screw plan shifted medial arm resent surgeon reintroduced dilator outer cannula incision navigation showed tools lateral new plan acceptable terms alignment pedicle surgeon proceeded drill trajectory place screw finished flouro brought check screw placed lateral plan acceptable pedicle robotic arm detached surgeon continued place rods close incisions iud additional multiple prone cases inaccuracy cases issues ran passed point accuracy test point test inaccuracy possible compression lateral case issue mazor x | NAV5000 | NAV3243 | NAV4095 | NAV4091 | NAV4094 | NAV5000NAV3243NAV4095NAV4091NAV4094 | True
3730 | 703431541 | XSE - Shoulder Shifts | 2019-10-24T19:59:59.000000+0000 | MazorX Stealth Edition | NAV4108 | SHOULDER SHIFTS | 20 | ASM0214-01 | SURGICAL SYSTEM | XSE - Shoulder Shifts 2019-Oct-24 00676786 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were three shoulder shifts in one case. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-10-30 IUD x4 (Other): Additional information received from a field serviced in the field on 2020-10-28 and the complaint was confirmed. The shoulder brake was not within manufacturer specifications. The field service engineer was able to tighten the shoulder brake. The system then passed the 10pt accuracy test and a stress test. The system functioned as designed. The engineer reported the most likely cause of the issue was loss of pressure over time due to use/wear and tear. SURGICAL SYSTEM | 2019-Oct-24 00676786 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that there were three shoulder shifts in one case. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-10-30 IUD x4 (Other): Additional information received from a field serviced in the field on 2020-10-28 and the complaint was confirmed. The shoulder brake was not within manufacturer specifications. The field service engineer was able to tighten the shoulder brake. The system then passed the 10pt accuracy test and a stress test. The system functioned as designed. The engineer reported the most likely cause of the issue was loss of pressure over time due to use/wear and tear. | xse shoulder shifts spinal procedure shoulder shifts case additional complications anticipated iud additional field serviced field complaint confirmed shoulder brake specifications field service engineer able tighten shoulder brake passed accuracy test stress test functioned designed engineer likely cause issue loss pressure time usewear tear surgical | NAV3252 | NAV4108 | NAV5000 | NAV2054 | NAV4106 | NAV3252NAV4108NAV5000NAV2054NAV4106 | True
9986 | 704764611 | XSE ¦ CD Drive will not Eject | 2021-12-14T20:28:23.000000+0000 | MazorX Stealth Edition | NAV3257 | CD DRIVE ISSUES | 30 | KIT1241-01 | SVC KIT1241-01 DVD RW GP70NS50 | XSE ¦ CD Drive will not Eject 2021-Dec-14 Salesforce C00239692 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative was unable to eject a CD from the drive. The rep put a disc into the system when there already was a disc in the drive. There was no patient involved. 2021-12-17 IUD x2 (Other): No new information. 2021-12-17 IUD x2 (Other): No new information. SVC KIT1241-01 DVD RW GP70NS50 | 2021-Dec-14 Salesforce C00239692 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative was unable to eject a CD from the drive. The rep put a disc into the system when there already was a disc in the drive. There was no patient involved. 2021-12-17 IUD x2 (Other): No new information. 2021-12-17 IUD x2 (Other): No new information. | xse cd drive eject salesforce outside procedure unable eject cd drive disc disc drive involved iud new iud new svc dvd rw | NAV3257 | NAV3299 | NAV3249 | NAV3253 | NAV4095 | NAV3257NAV3299NAV3249NAV3253NAV4095 | True
9726 | 704706451 | XSE - Damaged Surgeon Screen | 2021-11-10T00:05:06.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01S | KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | XSE - Damaged Surgeon Screen 2021-Nov-08 Salesforce C00237629, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgeon screen was damaged. The joint holding the arm for the screen had to be taped together. The issue was found during inspection after the system arrived from a lab. There was no patient involvement. 2021-11-16 IUD (Other): No new information. 2021-11-30 IUD (Other): No new information. 2021-12-06 IUD (Other): No new information. 2021-12-23 IUD (Other): No new information. KIT SVC ASM0357-01S SRGN SCREEN WITH ARM | 2021-Nov-08 Salesforce C00237629, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgeon screen was damaged. The joint holding the arm for the screen had to be taped together. The issue was found during inspection after the system arrived from a lab. There was no patient involvement. 2021-11-16 IUD (Other): No new information. 2021-11-30 IUD (Other): No new information. 2021-12-06 IUD (Other): No new information. 2021-12-23 IUD (Other): No new information. | xse damaged surgeon screen salesforce iud outside proc edure surgeon screen damaged joint holding arm screen taped issue inspection arrived lab iud new iud new iud new iud new kit svc srgn screen arm | NAV3253 | NAV2055 | NAV2054 | NAV2058 | NAV3249 | NAV3253NAV2055NAV2054NAV2058NAV3249 | True
2374 | 703234005 | X - SW, x-eye, WS, & 3Define issue | 2019-06-28T19:55:54.000000+0000 | MazorX | NAV2058 | WORKSTATION ISSUES | 40 | ASM0219-01 | WORKSTATION | X - SW, x-eye, WS, & 3Define issue 2017-6-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on T5-T8. It was reported that the bed frame was mounted too low and the angle of the shoulder was too open. A 3Define scan was done, but it had poor quality so a second scan had to be done. The x-eye could not see the reference marker, so predefined AP and OBL trajectories were sent for the acquisition. The predefined trajectories were manually adjusted to get the optimal images. Registration was completed on the first attempt off of T8. Fluoro verification was accurate. All trajectories were sent and drilled, and the surgeon was pleased with the outcome. The manufacturer representative attempted to test the 3Define from utilities. A timeout pop-up appeared after 20 seconds, and the 3Define image froze. They refreshed the 3Define image and the timeout pop-up appeared again. While dismounting the surgical system from the bed the pneumatic system stopped working. The arm almost fell on the patient but the clinical engineer held it up. The representative performed a hard restart to try and fix the pressure issue, but the software crashed en dlessly. The representative had to return the surgical system to the workstation manually. During this process, the hood unexpectedly fell, so the clinical engineer inserted a drape near the hydraulic cylinders to keep the hood open. There was no patient harm a nd the procedure was not delayed over an hour. WORKSTATION | 2017-6-11 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure on T5-T8. It was reported that the bed frame was mounted too low and the angle of the shoulder was too open. A 3Define scan was done, but it had poor quality so a second scan had to be done. The x-eye could not see the reference marker, so predefined AP and OBL trajectories were sent for the acquisition. The predefined trajectories were manually adjusted to get the optimal images. Registration was completed on the first attempt off of T8. Fluoro verification was accurate. All trajectories were sent and drilled, and the surgeon was pleased with the outcome. The manufacturer representative attempted to test the 3Define from utilities. A timeout pop-up appeared after 20 seconds, and the 3Define image froze. They refreshed the 3Define image and the timeout pop-up appeared again. While dismounting the surgical system from the bed the pneumatic system stopped working. The arm almost fell on the patient but the clinical engineer held it up. The representative performed a hard restart to try and fix the pressure issue, but the software crashed en dlessly. The representative had to return the surgical system to the workstation manually. During this process, the hood unexpectedly fell, so the clinical engineer inserted a drape near the hydraulic cylinders to keep the hood open. There was no patient harm a nd the procedure was not delayed over an hour. | x sw xeye ws issue ccr spinal procedure bed frame mounted low angle shoulder open scan poor quality second scan xeye reference marker predefined ap obl trajectories sent acquisition predefined trajectories manually adjusted optimal images registration completed attempt fluoro verification accurate trajectories sent drilled surgeon pleased outcome attempted test utilities timeout popup appeared seconds image froze refreshed image timeout popup appeared dismounting surgical bed pneumatic stopped working arm fell clinical engineer held performed hard restart try fix pressure issue software crashed en dlessly return surgical workstation manually process hood unexpectedly fell clinical engineer inserted drape near hydraulic cylinders hood open nd procedure delayed hour workstation | NAV3264 | NAV1006 | NAV3248 | NAV4101 | NAV4110 | NAV3264NAV1006NAV3248NAV4101NAV4110 | False
3662 | 703407415 | XSE - Reference frame stripped thread | 2019-10-10T23:13:48.000000+0000 | MazorX Stealth Edition | NAV3295 | BROKEN ACCESSORY | 20 | KIT0899 | ROBOTIC FRAME CONNECTION | XSE - Reference frame stripped thread 2019-10-10 00675860 (HCP,Rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the screw threading on the robotic frame connection where it attaches to the shoulder was stripped. Used the spine reference frame attached to the patient instead to proceed with the case. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-10-21 IUD x3, Email 2019-10-29 13:41 (Other): No new information. ROBOTIC FRAME CONNECTION | 2019-10-10 00675860 (HCP,Rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the screw threading on the robotic frame connection where it attaches to the shoulder was stripped. Used the spine reference frame attached to the patient instead to proceed with the case. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-10-21 IUD x3, Email 2019-10-29 13:41 (Other): No new information. | xse reference frame stripped thread spinal procedure screw threading robotic frame connection attaches shoulder stripped spine reference frame attached instead proceed case additional complications anticipated iud email new robotic frame connection | NAV3299 | NAV4100 | NAV2054 | NAV3295 | NAV2047 | NAV3299NAV4100NAV2054NAV3295NAV2047 | True
9590 | 704677947 | XSE ¦ Missing Accessories | 2021-10-22T22:21:55.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MAS2271-04S | KIT SVC MAS2271-04S SCHANZ BOND | XSE ¦ Missing Accessories 2021-Oct-22 Salesforce C00236708 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the schanz bond and target extender screw for this system were missing. No patient was present. 2021-11-02 IUD x3 (Other): No new information. KIT SVC MAS2271-04S SCHANZ BOND | 2021-Oct-22 Salesforce C00236708 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the schanz bond and target extender screw for this system were missing. No patient was present. 2021-11-02 IUD x3 (Other): No new information. | xse missing accessories salesforce outside procedure schanz bond target extender screw missing present iud new kit svc schanz bond | NAV2053 | NAV3299 | NAV4090 | NAV3298 | NAV3295 | NAV2053NAV3299NAV4090NAV3298NAV3295 | True
9371 | 704619712 | XSE - Damaged Part\Stripped Screws | 2021-09-20T18:00:57.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 50 | MEC5196-01S | KIT SVC MEC5196-01S RNX-ACC-063 | XSE - Damaged Part\Stripped Screws 2021-Sep-20 Salesforce C00234697 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter clamp was cross threaded. It was also reported that while inspecting this site's instruments, they found 3 bone mount bride screws were stripped. No patient was present. 2021-09-27 IUD x3 (Other): No new information. 2021-10-13 IUD (Other): No new information. 2021-11-09 email 2021-11-09 12:45 (Rep): Additional information received from a manufacturer representative reported that the cause of the image adapter being cross threaded was overtightening of the clamp since it needs to be tight to make sure it is securely attached. The suspected cause of the stripped screws was overtightening and loosening over time. KIT SVC MEC5196-01S RNX-ACC-063 | 2021-Sep-20 Salesforce C00234697 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the image adapter clamp was cross threaded. It was also reported that while inspecting this site's instruments, they found 3 bone mount bride screws were stripped. No patient was present. 2021-09-27 IUD x3 (Other): No new information. 2021-10-13 IUD (Other): No new information. 2021-11-09 email 2021-11-09 12:45 (Rep): Additional information received from a manufacturer representative reported that the cause of the image adapter being cross threaded was overtightening of the clamp since it needs to be tight to make sure it is securely attached. The suspected cause of the stripped screws was overtightening and loosening over time. | xse damaged partstripped screws salesforce outside procedure image adapter clamp cross threaded inspecting sites instruments bone mount bride screws stripped present iud new iud new email additional cause image adapter cross threaded overtightening clamp needs tight make sure securely attached suspected cause stripped screws overtightening loosening time kit svc | NAV3299 | NAV3258 | NAV2054 | NAV3256 | NAV3252 | NAV3299NAV3258NAV2054NAV3256NAV3252 | True
1536 | 703205104 | X - Deviation Issue | 2019-06-13T12:31:18.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 20 | KIT0759-01 | PNEUMATIC REPAIR KIT | X - Deviation Issue 2018-01-12 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a global deformity with fixation at T3-L1. It was reported that planning was completed the day before the procedure. The system went through 10 point and 3 point accuracy tests and shoulder calibration checked prior to being brought the the OR. The manufacturer representative then noticed when moving the RBT device by the positioner in the OR, that the positioner would not extend, and the manipulator only tilted side to side. The representatives present discussed that the positioner issue could complicate attachment of the system to the bed. When attaching the system to the bed, the representative requested the anesthetist to rotated the bed slightly to line up the positioner to attach and the system was attached to the bed on the first try. However, the manipulator was still stuck in the same position, allowing only side-to-side tilting movement and no completely free movement. In some positions a hissing sound from the pneumatic lines was heard. For surgery, the surgeon first completed dissection. Then the dual clamp was attached to T11 and T9. The 3Define scan and Draw Spine were completed. AP and OBL images were taken and registration was achieved off of T8 using the basic algorithm with good values achieved at L1-T7. The surgeon started at left L1 and worked down to left L7. When instrumenting L9, the tools were running into the clamp; the trajectory was, therefore, not drawn the first time. There was a similar issue at T8 and the surgeon asked one of the representatives to rotate the trajectory slightly lateral. While this was being completed, a repre sentative notified the surgeon that the patient's spine was definitely moving while unilateral retraction was being applied. The surgeon was aware but stated retraction would be released once the tools docked on bones. T7-L1 were completed by the fellow without immediately noted issue. The platform mount was unlocked to register for the second segment of surgery. The dual clamp was attached to T4 because the spinous processes of T5 and T6 were too small to allow for proper bone mounting. The 3Define scan and Draw Spine were completed without issue. AP and OBL images were obtained and registration was achieved on automatic off of T5 using the basic algorithm. T5 and T6 trajectories were sent. T4 was running directly into the clamp to that trajectory was not drilled. There were similar issues at T3. T5 and T6 were sent and placed first on the right, then on the left. T3 was the last trajectory sent and this trajectory was well outside of the open incision. A vertical trajectory was created to interface with the starting point of the screw trajectory and this was sent to screw the starting point of the screw. Once this was completed, the screw trajectory was sent and line up with the starting point taking care to avoid manipulating the spine to much. Upon probing the pilot holes and placing the screws, it was noticed that the initial screws on the patient's lift side were slightly medial and were redirected freehand. The surgeon stated this was because of the unilateral retractioin. Motos were lost during the coronal manipulation to correct the spine. Screws were all tested and found to be in good standing order and the spine was released. The surgeon felt that the issue was related to the patient's BP, not the system. PNEUMATIC REPAIR KIT | 2018-01-12 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a global deformity with fixation at T3-L1. It was reported that planning was completed the day before the procedure. The system went through 10 point and 3 point accuracy tests and shoulder calibration checked prior to being brought the the OR. The manufacturer representative then noticed when moving the RBT device by the positioner in the OR, that the positioner would not extend, and the manipulator only tilted side to side. The representatives present discussed that the positioner issue could complicate attachment of the system to the bed. When attaching the system to the bed, the representative requested the anesthetist to rotated the bed slightly to line up the positioner to attach and the system was attached to the bed on the first try. However, the manipulator was still stuck in the same position, allowing only side-to-side tilting movement and no completely free movement. In some positions a hissing sound from the pneumatic lines was heard. For surgery, the surgeon first completed dissection. Then the dual clamp was attached to T11 and T9. The 3Define scan and Draw Spine were completed. AP and OBL images were taken and registration was achieved off of T8 using the basic algorithm with good values achieved at L1-T7. The surgeon started at left L1 and worked down to left L7. When instrumenting L9, the tools were running into the clamp; the trajectory was, therefore, not drawn the first time. There was a similar issue at T8 and the surgeon asked one of the representatives to rotate the trajectory slightly lateral. While this was being completed, a repre sentative notified the surgeon that the patient's spine was definitely moving while unilateral retraction was being applied. The surgeon was aware but stated retraction would be released once the tools docked on bones. T7-L1 were completed by the fellow without immediately noted issue. The platform mount was unlocked to register for the second segment of surgery. The dual clamp was attached to T4 because the spinous processes of T5 and T6 were too small to allow for proper bone mounting. The 3Define scan and Draw Spine were completed without issue. AP and OBL images were obtained and registration was achieved on automatic off of T5 using the basic algorithm. T5 and T6 trajectories were sent. T4 was running directly into the clamp to that trajectory was not drilled. There were similar issues at T3. T5 and T6 were sent and placed first on the right, then on the left. T3 was the last trajectory sent and this trajectory was well outside of the open incision. A vertical trajectory was created to interface with the starting point of the screw trajectory and this was sent to screw the starting point of the screw. Once this was completed, the screw trajectory was sent and line up with the starting point taking care to avoid manipulating the spine to much. Upon probing the pilot holes and placing the screws, it was noticed that the initial screws on the patient's lift side were slightly medial and were redirected freehand. The surgeon stated this was because of the unilateral retractioin. Motos were lost during the coronal manipulation to correct the spine. Screws were all tested and found to be in good standing order and the spine was released. The surgeon felt that the issue was related to the patient's BP, not the system. | x deviation issue ccr spinal procedure treat global deformity fixation planning completed day procedure went point point accuracy tests shoulder calibration checked prior brought noticed moving rbt positioner positioner extend manipulator tilted representatives present discussed positioner issue complicate attachment bed attaching bed requested anesthetist rotated bed slightly line positioner attach attached bed try manipulator stuck position allowing sidetoside tilting movement completely free movement positions hissing sound pneumatic lines heard surgery surgeon completed dissection dual clamp attached scan draw spine completed ap obl images taken registration achieved using basic algorithm good values achieved surgeon started left worked left instrumenting tools running clamp trajectory drawn time similar issue surgeon asked representatives rotate trajectory slightly lateral completed repre sentative notified surgeon patients spine definitely moving unilateral retraction applied surgeon aware stated retraction released tools docked bones completed fellow immediately noted issue platform mount unlocked register second segment surgery dual clamp attached spinous processes small allow proper bone mounting scan draw spine completed issue ap obl images obtained registration achieved automatic using basic algorithm trajectories sent running directly clamp trajectory drilled similar issues sent placed right left trajectory sent trajectory outside open incision vertical trajectory created interface starting point screw trajectory sent screw starting point screw completed screw trajectory sent line starting point taking care avoid manipulating spine probing pilot holes placing screws noticed initial screws patients lift slightly medial redirected freehand surgeon stated unilateral retractioin motos lost coronal manipulation correct spine screws tested good standing order spine released surgeon felt issue related patients bp pneumatic repair kit | NAV3260 | NAV3243 | NAV4091 | NAV3250 | NAV4103 | NAV3260NAV3243NAV4091NAV3250NAV4103 | False
3113 | 703266091 | X - Bumped in Crate/OOBF | 2019-07-18T15:03:27.000000+0000 | MazorX | NAV4092 | INACCURACY DURING SETUP | 20 | ASM0206-01 | SURGICAL ARM | X - Bumped in Crate/OOBF SRA removed from crate and shows some minor signs of being bumped around in the crate. When installed the SRA failed accuracy tests and is consider an OTBF SURGICAL ARM | SRA removed from crate and shows some minor signs of being bumped around in the crate. When installed the SRA failed accuracy tests and is consider an OTBF | x bumped crateoobf sra removed crate shows minor signs bumped crate installed sra failed accuracy tests consider otbf surgical arm | NAV3249 | NAV2055 | NAV4093 | NAV4092 | NAV1006 | NAV3249NAV2055NAV4093NAV4092NAV1006 | True
9324 | 704607655 | XSE - CANBUS Error | 2021-09-13T21:56:26.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 30 | CBL0115-01 | CABLE, UNSH, MFM, 14_POSITION, L=1.5M,FR | XSE - CANBUS Error 2021-Sep-13 Salesforce C00234357 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system threw an error, 10251 fatal error CANBUS connection lost to node number 7. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-09-15 IUD x6 (Other): No new information. 2021-09-24 email 2021-09-24 20:30 (Rep): No new information. 2021-09-27 IUD x2 (Other): No new information. 2021-10-18 IUD x4 (Other): No new information. 2021-11-05 email 2021-11-05 13:46 (Rep): No new information. 2022-04-05 IUD (Other): No new information. CABLE, UNSH, MFM, 14_POSITION, L=1.5M,FR | 2021-Sep-13 Salesforce C00234357 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that the system threw an error, 10251 fatal error CANBUS connection lost to node number 7. The issue extended the surgical time by less than 1 hour. There was no impact on the patient outcome. 2021-09-15 IUD x6 (Other): No new information. 2021-09-24 email 2021-09-24 20:30 (Rep): No new information. 2021-09-27 IUD x2 (Other): No new information. 2021-10-18 IUD x4 (Other): No new information. 2021-11-05 email 2021-11-05 13:46 (Rep): No new information. 2022-04-05 IUD (Other): No new information. | xse canbus error salesforce spinal p rocedure threw error fatal error canbus connection lost node number issue extended surgical time hour impact outcome iud new email new iud new iud new email new iud new cable unsh mfm | NAV2047 | NAV3249 | NAV2054 | NAV4107 | NAV3253 | NAV2047NAV3249NAV2054NAV4107NAV3253 | True
5083 | 703595102 | XSE - CT Scans not Loading to Mazor | 2020-02-04T00:12:05.000000+0000 | MazorX Stealth Edition | NAV3236 | CT IMAGES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CT Scans not Loading to Mazor 2020-01-30 00683298 (REP, HCP): Medtronic received information regarding a guidance system in use outside of procedure. It was reported that while loading the CT they were receiving import errors from the CDs. They were attempting to use different CDs and sometimes it worked and other times it did not. There was no patient involvement. 2020-02-17 Email 2020-02-17 1752 (Rep): Additional information received form the rep reporting that the CT exams were burned to multiple CDs without success in loading the exam to the system. At the site, CT scan slice inconsistencies were found in each scan and it was determined that the most likely cause was from the CT Scanner or the CD burning process, not the guidance system. UNK MAZOR X WORKSTATION | 2020-01-30 00683298 (REP, HCP): Medtronic received information regarding a guidance system in use outside of procedure. It was reported that while loading the CT they were receiving import errors from the CDs. They were attempting to use different CDs and sometimes it worked and other times it did not. There was no patient involvement. 2020-02-17 Email 2020-02-17 1752 (Rep): Additional information received form the rep reporting that the CT exams were burned to multiple CDs without success in loading the exam to the system. At the site, CT scan slice inconsistencies were found in each scan and it was determined that the most likely cause was from the CT Scanner or the CD burning process, not the guidance system. | xse ct scans loading mazor use outside procedure loading ct receiving import errors cds attempting use different cds worked times did email additional reporting ct exams burned multiple cds success loading exam site ct scan slice inconsistencies scan determined likely cause ct scanner cd burning process unk mazor x workstation | NAV3236 | NAV3244 | NAV2058 | NAV3257 | NAV2047 | NAV3236NAV3244NAV2058NAV3257NAV2047 | True
6015 | 703859548 | XSE - SRA Freezing | 2020-07-18T22:58:34.000000+0000 | MazorX Stealth Edition | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01 | SURGICAL ARM | XSE - SRA Freezing 2020-06-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the generic AP and oblique trajectories were used to acquire images for registration. The first segment of the construct was successfully registered and all trajectories were sent to be executed. The surgical arm was sent to the clear up position, but the arm began to freeze after minimal movement and a 8798 error was displayed. A soft restart of the system resolved the issue. Registration for the second segment had to be done twice due to mislabeling T6. After the second segment was completed, the surgeon tried to unlock the shoulder and move it out of the way, but the button did not work. The manufacturer representative unlocked the button in the software to move the shoulder. The surgeon opted to freehand the bottom of the construct from L3-L1 to reduce registration segments. The surgeon had decided to hook the top of the construct and skip right T4. All scr ews were accurately placed. There was no patient harm and the procedure was delayed less than an hour. SURGICAL ARM | 2020-06-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat deformity with fixation. It was reported that the generic AP and oblique trajectories were used to acquire images for registration. The first segment of the construct was successfully registered and all trajectories were sent to be executed. The surgical arm was sent to the clear up position, but the arm began to freeze after minimal movement and a 8798 error was displayed. A soft restart of the system resolved the issue. Registration for the second segment had to be done twice due to mislabeling T6. After the second segment was completed, the surgeon tried to unlock the shoulder and move it out of the way, but the button did not work. The manufacturer representative unlocked the button in the software to move the shoulder. The surgeon opted to freehand the bottom of the construct from L3-L1 to reduce registration segments. The surgeon had decided to hook the top of the construct and skip right T4. All scr ews were accurately placed. There was no patient harm and the procedure was delayed less than an hour. | xse sra freezing ccr spinal procedure treat deformity fixation generic ap oblique trajectories acquire images registration segment construct successfully registered trajectories sent executed surgical arm sent clear position arm began freeze minimal movement error displayed soft restart resolved issue registration second segment twice mislabeling second segment completed surgeon tried unlock shoulder way button did work unlocked button software shoulder surgeon opted freehand construct reduce registration segments surgeon decided hook construct skip right scr ews accurately placed procedure delayed hour surgical arm | NAV4107 | NAV3242 | NAV3248 | NAV3249 | NAV4095 | NAV4107NAV3242NAV3248NAV3249NAV4095 | False
8216 | 704372749 | X - Deviated Trajectories | 2021-05-02T19:44:42.000000+0000 | MazorX | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviated Trajectories 2021-May-02 Salesforce C00225642, e1, Intake Form, email 2021-05-04 01:15, CCR, email 2021-05-06 04:16 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported a T4-T12 open scoliosis case was done using a bridge and dual clamps placed on the spinous processes of T8 and T10/T11. Prior to the operation, the manufacturer representative ran a 10 point accuracy check, checked all of the cameras, and checked all of the connections to ensure there were no loose connections or foreign objects that could cause issues. During the procedure, the surgeon secured the mount. The representative noted that the plan was to reach T4 and T5 while the clamp was over T8, but they never got to those trajectories. The 3Define scan was done and registration was completed without any noticeable movement. One surgeon worked on the patient's left side and another surgeon worked on the right side of the patient. The surgeon started at T12 and moved up in a zig zag pattern. The surgeon working on the left side decided to skip a few trajectories as the mount was obstructing the trajectories. When at right T6, the surgeon felt the trajectory was medial to plan by less than 3.5 mm. The surgical arm was sent to the clear up position and back to the traject ory, but it went to the same position. Planning was adjusted to make the screw more lateral. The surgical arm was sent to the trajectory, but it was again medial to plan. The surgeon moved on to left T8. When the surgical arm reached the location, it looked like it was at the trajectory of a previous screw. The surgical arm was then set to right T5 and showed medial to the plan so the arm was sent to left T5 and was lateral to plan. The representative suggested the patient be re-registered due to the deviations, but the surgeon decided to abort the use of the guidance system and complete the case freehand. Once the surgical system was unmounted from the bed, fluoro imaging showed deviations at left T7 and left T8. The screws were placed inferior and hit the disk space. The cause of the deviations was not determined. The surgical team suspected it was the guidance system. There was no patient harm and the procedure was delayed over an hour. MAZOR X SYSTEM TPL0059 | 2021-May-02 Salesforce C00225642, e1, Intake Form, email 2021-05-04 01:15, CCR, email 2021-05-06 04:16 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported a T4-T12 open scoliosis case was done using a bridge and dual clamps placed on the spinous processes of T8 and T10/T11. Prior to the operation, the manufacturer representative ran a 10 point accuracy check, checked all of the cameras, and checked all of the connections to ensure there were no loose connections or foreign objects that could cause issues. During the procedure, the surgeon secured the mount. The representative noted that the plan was to reach T4 and T5 while the clamp was over T8, but they never got to those trajectories. The 3Define scan was done and registration was completed without any noticeable movement. One surgeon worked on the patient's left side and another surgeon worked on the right side of the patient. The surgeon started at T12 and moved up in a zig zag pattern. The surgeon working on the left side decided to skip a few trajectories as the mount was obstructing the trajectories. When at right T6, the surgeon felt the trajectory was medial to plan by less than 3.5 mm. The surgical arm was sent to the clear up position and back to the traject ory, but it went to the same position. Planning was adjusted to make the screw more lateral. The surgical arm was sent to the trajectory, but it was again medial to plan. The surgeon moved on to left T8. When the surgical arm reached the location, it looked like it was at the trajectory of a previous screw. The surgical arm was then set to right T5 and showed medial to the plan so the arm was sent to left T5 and was lateral to plan. The representative suggested the patient be re-registered due to the deviations, but the surgeon decided to abort the use of the guidance system and complete the case freehand. Once the surgical system was unmounted from the bed, fluoro imaging showed deviations at left T7 and left T8. The screws were placed inferior and hit the disk space. The cause of the deviations was not determined. The surgical team suspected it was the guidance system. There was no patient harm and the procedure was delayed over an hour. | x deviated trajectories salesforce email ccr email spinal procedure open scoliosis case using bridge dual clamps placed spinous processes prior operation ran point accuracy check checked cameras checked connections ensure loose connections foreign objects cause issues procedure surgeon secured mount noted plan reach clamp got trajectories scan registration completed noticeable movement surgeon worked patients left surgeon worked right surgeon started moved zig zag pattern surgeon working left decided skip trajectories mount obstructing trajectories right surgeon felt trajectory medial plan mm surgical arm sent clear position traject ory went position planning adjusted make screw lateral surgical arm sent trajectory medial plan surgeon moved left surgical arm reached location looked like trajectory previous screw surgical arm set right showed medial plan arm sent left lateral plan suggested reregistered deviations surgeon decided abort use complete case freehand surgical unmounted bed fluoro imaging showed deviations left left screws placed inferior hit disk space cause deviations determined surgical team suspected procedure delayed hour mazor x | NAV4091 | NAV5000 | NAV3243 | NAV4095 | NAV4096 | NAV4091NAV5000NAV3243NAV4095NAV4096 | True
2439 | 703235802 | X - SOFTWARE ISSUE | 2019-07-01T13:15:06.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0639-08 | MAZOR X SOFTWARE KIT VER-2.3 | X - SOFTWARE ISSUE Reference Marker on the wrong side. The surgical arm was attached to the bed on the right side of the patient. The single Schanz pin was inserted into the right PSIS with plenty of room to extend the shoulder. When locking the mount, the 3D reference marker was identified, however it was displayed on the screen on the left side of the surgical arm, even though it was attached to the right side. The stylus was correctly recognized and the green line was accurately drawn which provided an exceptional AP and OBL image. The case ended up going perfectly. Slow SW performance Dr. Poelstra confirmed that the 2.3 software seemed to be running slower than the previous version. In particular, when selecting one vertebral body to another it took 3-4 seconds to switch screens and did not provide any feedback that it was working. MAZOR X SOFTWARE KIT VER-2.3 | Reference Marker on the wrong side. The surgical arm was attached to the bed on the right side of the patient. The single Schanz pin was inserted into the right PSIS with plenty of room to extend the shoulder. When locking the mount, the 3D reference marker was identified, however it was displayed on the screen on the left side of the surgical arm, even though it was attached to the right side. The stylus was correctly recognized and the green line was accurately drawn which provided an exceptional AP and OBL image. The case ended up going perfectly. Slow SW performance Dr. Poelstra confirmed that the 2.3 software seemed to be running slower than the previous version. In particular, when selecting one vertebral body to another it took 3-4 seconds to switch screens and did not provide any feedback that it was working. | x software issue reference marker wrong surgical arm attached bed right single schanz pin inserted right psis plenty room extend shoulder locking mount reference marker identified displayed screen left surgical arm attached right stylus correctly recognized green line accurately drawn provided exceptional ap obl image case ended going perfectly slow sw performance dr poelstra confirmed software running slower previous version particular selecting vertebral body took seconds switch screens did provide feedback working mazor x software kit | NAV3248 | NAV3246 | NAV4091 | NAV3247 | NAV3250 | NAV3248NAV3246NAV4091NAV3247NAV3250 | True
5728 | 703790706 | XSE - Damaged Green Screw Driver | 2020-06-08T17:13:14.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS1650-03S | KIT SVC MAS1650-03S UNIV SCREWDRIVER | XSE - Damaged Green Screw Driver 2020-06-08 00690303 (Rep): Medtronic received information regarding a guidance system accessory being used outside of procedure. It was reported that the green screw driver is stripped making it difficult to tighten the screws down. No patient was present. 2020-06-15 IUD x2 (Other): No new information. 2020-06-24 IUD (Other): No new information. 2020-07-06 email 2020-07-06 09:45 (Rep): Additional information received from a manufacturer representative reported that the part was replaced. 2020-08-05 IUD (Other): No new information. 2020-08-28 e2, e3 (Rep): No new information. 2020-09-09 email 2020-09-09 10:31, e1 (Rep): No new information. KIT SVC MAS1650-03S UNIV SCREWDRIVER | 2020-06-08 00690303 (Rep): Medtronic received information regarding a guidance system accessory being used outside of procedure. It was reported that the green screw driver is stripped making it difficult to tighten the screws down. No patient was present. 2020-06-15 IUD x2 (Other): No new information. 2020-06-24 IUD (Other): No new information. 2020-07-06 email 2020-07-06 09:45 (Rep): Additional information received from a manufacturer representative reported that the part was replaced. 2020-08-05 IUD (Other): No new information. 2020-08-28 e2, e3 (Rep): No new information. 2020-09-09 email 2020-09-09 10:31, e1 (Rep): No new information. | xse damaged green screw driver accessory outside procedure green screw driver stripped making difficult tighten screws present iud new iud new email additional replaced iud new new email new kit svc univ screwdriver | NAV3299 | NAV3295 | NAV2054 | NAV4090 | NAV3249 | NAV3299NAV3295NAV2054NAV4090NAV3249 | True
6805 | 704036037 | X - Missing Screw/Targ Ext Not Aligned | 2020-10-27T14:38:48.000000+0000 | MazorX | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2310-03 | TARGET EXTENDER | X - Missing Screw/Targ Ext Not Aligned 2020-09-29 00699587, e1, e2, e3 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the screw for the target extender was lost and the alignment pin was not attached to the target extender. There was no patient harm and the procedure was not delayed. TARGET EXTENDER | 2020-09-29 00699587, e1, e2, e3 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the screw for the target extender was lost and the alignment pin was not attached to the target extender. There was no patient harm and the procedure was not delayed. | x missing screwtarg ext aligned spinal procedure screw target extender lost alignment pin attached target extender procedure delayed target extender | NAV2053 | NAV2054 | NAV3295 | NAV3299 | NAV3242 | NAV2053NAV2054NAV3295NAV3299NAV3242 | True
9192 | 704571253 | XSE ¦ Screws for DVI Port Fell Loose | 2021-08-23T17:46:12.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 20 | KIT1266-01 | SVC KIT1266-01 D-SUB CONNECTOR SCW | XSE ¦ Screws for DVI Port Fell Loose 2021-Aug-23 Salesforce C00233331 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws on the DVI port fell out. This was found outside of a case, no patient present. 2021-Aug-30 IUDx2 (other): No new information. SVC KIT1266-01 D-SUB CONNECTOR SCW | 2021-Aug-23 Salesforce C00233331 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the screws on the DVI port fell out. This was found outside of a case, no patient present. 2021-Aug-30 IUDx2 (other): No new information. | xse screws dvi port fell loose salesforce outside procedure screws dvi port fell outside case present new svc dsub connector scw | NAV2058 | NAV3299 | NAV3249 | NAV3253 | NAV4099 | NAV2058NAV3299NAV3249NAV3253NAV4099 | True
10055 | 704792527 | XSE - Controller Error/No Pressure | 2022-01-05T20:31:13.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Controller Error/No Pressure 2021-Dec-29 Salesforce C00240261, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that an error message was displayed when the guidance system was booted up. The error message stated the controller was not connected and pressure would not build. The system was rebooted, but the issue was not resolved. The manufacturer rep resentative checked the pressure and it was at 0 and the compressor would not turn on. There was no patient involvement. 2022-01-12 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-02 IUD x3 (Other): No new information. UNK MAZOR X WORKSTATION | 2021-Dec-29 Salesforce C00240261, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that an error message was displayed when the guidance system was booted up. The error message stated the controller was not connected and pressure would not build. The system was rebooted, but the issue was not resolved. The manufacturer rep resentative checked the pressure and it was at 0 and the compressor would not turn on. There was no patient involvement. 2022-01-12 IUD (Other): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-02 IUD x3 (Other): No new information. | xse controller errorno pressure salesforce iud outside proc edure error message displayed booted error message stated controller connected pressure build rebooted issue resolved resentative checked pressure compressor turn iud new iud new iud new unk mazor x workstation | NAV3264 | NAV2047 | NAV3262 | NAV4104 | NAV4095 | NAV3264NAV2047NAV3262NAV4104NAV4095 | True
4427 | 703543131 | X-Arm/SW/Positioner/Platform issues | 2020-01-03T15:21:48.000000+0000 | MazorX | NAV4096 | ARM HIT WORK VOLUME | 40 | ASM0206-01 | SURGICAL ARM | X-Arm/SW/Positioner/Platform issues 2018-01-01 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for fourteen vertebrae on level T2 - L3. The case was planned prior to the case on workstation. During planning more than one vertebral body where not aligned by auto PBRS. Advanced setup was completed successfully. Following that first attempt to send arm to surgical arm setup ended with sound but no movement. Action was stopped using stop button. Second attempt to send arm moved to complete surgical arm set (3pts) successfully. Surgeon requested to import the study for the next case on the workstation. Import initiated and then failed after attempting to attach 'SRS' to OR table. Software failed to multitask. Second attempt to import succeeded successfully. Surgeon decided to freehand the bottom of the construct L1 - L3 prior to MX portion. Surgeon stated he would skip T2 bilateral trajectories execution. Advanced (10pts) setup was completed on OR table. First segment was T2 - T9. 3Define scan was completed successfully. Draw spine was performed with good positioning. Slight adjustment to lower 3D marker along t rajectory to allow capture of long segment (T2 - T9). Right T9 was sent. The arm was lowered, and it touched BMB lever slightly. Arm returned up along trajectory and surgeon didn't feel additional registration is required because no movement was observed. Between segments mount unlocked. Shoulder button pressed but only shoulder brakes were released and Y, Z axis on positioner was still locked. Second segment T10 - T12. 3Define of second segment was successful on first attempt. Draw spine was completed. Left T12 the last trajectory was sent. The guidance system was removed from the OR table. No patient harm was reported. SURGICAL ARM | 2018-01-01 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for fourteen vertebrae on level T2 - L3. The case was planned prior to the case on workstation. During planning more than one vertebral body where not aligned by auto PBRS. Advanced setup was completed successfully. Following that first attempt to send arm to surgical arm setup ended with sound but no movement. Action was stopped using stop button. Second attempt to send arm moved to complete surgical arm set (3pts) successfully. Surgeon requested to import the study for the next case on the workstation. Import initiated and then failed after attempting to attach 'SRS' to OR table. Software failed to multitask. Second attempt to import succeeded successfully. Surgeon decided to freehand the bottom of the construct L1 - L3 prior to MX portion. Surgeon stated he would skip T2 bilateral trajectories execution. Advanced (10pts) setup was completed on OR table. First segment was T2 - T9. 3Define scan was completed successfully. Draw spine was performed with good positioning. Slight adjustment to lower 3D marker along t rajectory to allow capture of long segment (T2 - T9). Right T9 was sent. The arm was lowered, and it touched BMB lever slightly. Arm returned up along trajectory and surgeon didn't feel additional registration is required because no movement was observed. Between segments mount unlocked. Shoulder button pressed but only shoulder brakes were released and Y, Z axis on positioner was still locked. Second segment T10 - T12. 3Define of second segment was successful on first attempt. Draw spine was completed. Left T12 the last trajectory was sent. The guidance system was removed from the OR table. No patient harm was reported. | xarmswpositionerplatform issues ccr spinal procedure case fourteen vertebrae level case planned prior case workstation planning vertebral body aligned auto pbrs advanced setup completed successfully following attempt send arm surgical arm setup ended sound movement action stopped using stop button second attempt send arm moved complete surgical arm set successfully surgeon requested import study case workstation import initiated failed attempting attach srs table software failed multitask second attempt import succeeded successfully surgeon decided freehand construct prior mx portion surgeon stated skip bilateral trajectories execution advanced setup completed table segment scan completed successfully draw spine performed good positioning slight adjustment lower marker t rajectory allow capture long segment right sent arm lowered touched bmb lever slightly arm returned trajectory surgeon didnt feel additional registration required movement observed segments mount unlocked shoulder button pressed shoulder brakes released y z axis positioner locked second segment second segment successful attempt draw spine completed left trajectory sent removed table surgical arm | NAV4095 | NAV3249 | NAV4101 | NAV3246 | NAV4104 | NAV4095NAV3249NAV4101NAV3246NAV4104 | False
6437 | 703941393 | XSE - Stripped Screws on Flat Panel | 2020-09-04T17:49:09.000000+0000 | MazorX Stealth Edition | NAV3258 | IMAGE ADAPTOR ISSUES | 20 | KIT0685 | FLAT PANEL ADAPTOR KIT B KIT0685 | XSE - Stripped Screws on Flat Panel 2020-09-04 00695905 (Rep): Medtronic received information regarding guidance system, outside of procedure. It was reported that the square adapter for the image adapter has stripped screws and it has to be tightened very hard in order for it to stay on the c-arm. No patient was present. 2020-09-15 IUD (Other): No new information. 2020-10-06 email 2020-10-06 13:15 (Rep): No new information. 2020-11-01 email 2020-11-01 15:15 (Rep): No new information. FLAT PANEL ADAPTOR KIT B KIT0685 | 2020-09-04 00695905 (Rep): Medtronic received information regarding guidance system, outside of procedure. It was reported that the square adapter for the image adapter has stripped screws and it has to be tightened very hard in order for it to stay on the c-arm. No patient was present. 2020-09-15 IUD (Other): No new information. 2020-10-06 email 2020-10-06 13:15 (Rep): No new information. 2020-11-01 email 2020-11-01 15:15 (Rep): No new information. | xse stripped screws flat panel outside procedure square adapter image adapter stripped screws tightened hard order stay carm present iud new email new email new flat panel adaptor kit b | NAV3258 | NAV2054 | NAV3246 | NAV4099 | NAV3299 | NAV3258NAV2054NAV3246NAV4099NAV3299 | True
5804 | 703809100 | XSE - Unable to Build Pressure | 2020-06-18T15:55:49.000000+0000 | MazorX Stealth Edition | NAV3264 | PRESSURE ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Unable to Build Pressure 2020-06-18 00691093 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was unable to build pressure to the desired 7.2 PSI. This caused the compressor to time out and a display an error message. There was no patient involvement. 2020-06-26 IUD (Other): No new information. 2020-06-29 00691742 (Rep): No new information. 2020-07-02 IUD x3 (Other): No new information. 2020-07-15 IUD x2 (Rep): Additional information received from a manufacturer representative reported the issue was resolved. UNK MAZOR X WORKSTATION | 2020-06-18 00691093 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the system was unable to build pressure to the desired 7.2 PSI. This caused the compressor to time out and a display an error message. There was no patient involvement. 2020-06-26 IUD (Other): No new information. 2020-06-29 00691742 (Rep): No new information. 2020-07-02 IUD x3 (Other): No new information. 2020-07-15 IUD x2 (Rep): Additional information received from a manufacturer representative reported the issue was resolved. | xse unable build pressure outside procedure unable build pressure desired psi caused compressor time display error message iud new new iud new iud additional issue resolved unk mazor x workstation | NAV3264 | NAV3261 | NAV3244 | NAV2047 | NAV2054 | NAV3264NAV3261NAV3244NAV2047NAV2054 | True
206 | 703038325 | XSE - Navlock would not verify | 2019-03-05T16:28:05.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 30 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Navlock would not verify 2019-Mar-01 SCL1900498 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the black NavLock would not verify once the mount was locked. They attempted to use a dilator and drill, and did not use the black NavLock for the case. It was able to verify once the mount was unlocked. It was noted the issue occurred during a second case where the system was not turned off between cases. When the HCP clicked on the next patient on the home screen, the surgical arm setup, kitcode, and keycode were required. However, all instruments were already verified when entering instruments setup. For due diligence, the HCP verified all NavLocks again. Following entrance to operate procedure and lock of mount on current patient, all instrument verifications previously performed were erased. Passive planar needed to be re-verified before draw spine could be performed. Once it was determined after registration that they had lost instrument verification, re-verification of the Violet, Blue, and Black Navlocks were attempted. When the Violet/ Blue/Black Navlocks were brought close to the robotic reference frame, the banners indicated that the Navlocks had already been verified, but when an attempt was made to place the Navlock into the surgical field, the banner would read as unverified. The Violet and Blue Navlocks were verified using the dilator. Many attempts were made to verify the BlackNavlock without success by using the drill, using the dilator, trying all divots, and verifying spheres were securely attached. For the rest of the procedure only the Blue and Violet Navlocks were used. The Violet Navlock stayed on the dilator throughout the procedure. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Information unrelated to this event was included in a separate PE; see PE 703038322-trajectories deep. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Mar-01 SCL1900498 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the black NavLock would not verify once the mount was locked. They attempted to use a dilator and drill, and did not use the black NavLock for the case. It was able to verify once the mount was unlocked. It was noted the issue occurred during a second case where the system was not turned off between cases. When the HCP clicked on the next patient on the home screen, the surgical arm setup, kitcode, and keycode were required. However, all instruments were already verified when entering instruments setup. For due diligence, the HCP verified all NavLocks again. Following entrance to operate procedure and lock of mount on current patient, all instrument verifications previously performed were erased. Passive planar needed to be re-verified before draw spine could be performed. Once it was determined after registration that they had lost instrument verification, re-verification of the Violet, Blue, and Black Navlocks were attempted. When the Violet/ Blue/Black Navlocks were brought close to the robotic reference frame, the banners indicated that the Navlocks had already been verified, but when an attempt was made to place the Navlock into the surgical field, the banner would read as unverified. The Violet and Blue Navlocks were verified using the dilator. Many attempts were made to verify the BlackNavlock without success by using the drill, using the dilator, trying all divots, and verifying spheres were securely attached. For the rest of the procedure only the Blue and Violet Navlocks were used. The Violet Navlock stayed on the dilator throughout the procedure. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. Information unrelated to this event was included in a separate PE; see PE 703038322-trajectories deep. | xse navlock verify spinal procedure black navlock verify mount locked attempted use dilator drill did use black navlock case able verify mount unlocked noted issue occurred second case turned cases hcp clicked home screen surgical arm setup kitcode keycode required instruments verified entering instruments setup diligence hcp verified navlocks following entrance operate procedure lock mount current instrument verifications previously performed erased passive planar needed reverified draw spine performed determined registration lost instrument verification reverification violet blue black navlocks attempted violet blueblack navlocks brought close robotic reference frame banners indicated navlocks verified attempt place navlock surgical field banner read unverified violet blue navlocks verified using dilator attempts verify blacknavlock success using drill using dilator trying divots verifying spheres securely attached rest procedure blue violet navlocks violet navlock stayed dilator procedure noted additional complications anticipated unrelated included separate pe pe deep mazor x software installation kit | NAV3247 | NAV3246 | NAV1006 | NAV5002 | NAV2051 | NAV3247NAV3246NAV1006NAV5002NAV2051 | True
5707 | 703785579 | X - Pneumatic Coupler Damaged | 2020-06-04T14:21:24.000000+0000 | MazorX | NAV3264 | PRESSURE ISSUES | 30 | KIT0759-01 | PNEUMATIC REPAIR KIT | X - Pneumatic Coupler Damaged 2020-06-04 00690104 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the pneumatic coupler was warped leading into the pneumatic brick of the manipulator. There was no patient involvement. 2020-06-11 IUD x2 (Other): No new information. 2020-06-26 IUD (Other): No new information. 2020-07-02 IUD x2 (Other): No new information. 2002-07-17 IUD x2 (Other): No new information. 2020-07-31 IUD (Other): No new information. 2020-08-05 IUD x3 (Other): No new information. 2020-08-28 IUD (Other): The system is currently functioning as designed. PNEUMATIC REPAIR KIT | 2020-06-04 00690104 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the pneumatic coupler was warped leading into the pneumatic brick of the manipulator. There was no patient involvement. 2020-06-11 IUD x2 (Other): No new information. 2020-06-26 IUD (Other): No new information. 2020-07-02 IUD x2 (Other): No new information. 2002-07-17 IUD x2 (Other): No new information. 2020-07-31 IUD (Other): No new information. 2020-08-05 IUD x3 (Other): No new information. 2020-08-28 IUD (Other): The system is currently functioning as designed. | x pneumatic coupler damaged outside procedure pneumatic coupler warped leading pneumatic brick manipulator iud new iud new iud new iud new iud new iud new iud currently functioning designed pneumatic repair kit | NAV3264 | NAV3260 | NAV2054 | NAV1006 | NAV2057 | NAV3264NAV3260NAV2054NAV1006NAV2057 | True
1724 | 703210641 | X - encoder not working | 2019-06-17T17:18:15.000000+0000 | MazorX | NAV4106 | CALIBRATION ISSUES | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | X - encoder not working 2019-Jun-12 SCL190716 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the front encoder was not working properly, <20 degree of error. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. TPL0059 License Information: We got no license information for this product. Sold Date: Refer to SAP UNK_MAZORX_SHLDER License Information: We got no license information for this product. Sold Date: Refer to SAP zhangl83 2019-6-18 UNK MAZOR X SHOULDER | 2019-Jun-12 SCL190716 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the front encoder was not working properly, <20 degree of error. No patient involvement was reported. There were no additional complications reported or anticipated as a result of the event. TPL0059 License Information: We got no license information for this product. Sold Date: Refer to SAP UNK_MAZORX_SHLDER License Information: We got no license information for this product. Sold Date: Refer to SAP zhangl83 2019-6-18 | x encoder working spinal procedure encoder working properly degree error additional complications anticipated license got license product sold date refer sap unkmazorxshlder license got license product sold date refer sap unk mazor x shoulder | NAV3244 | NAV2054 | NAV3295 | NAV3249 | NAV3264 | NAV3244NAV2054NAV3295NAV3249NAV3264 | False
919 | 703180799 | X - positioner issue | 2019-05-30T13:41:33.000000+0000 | MazorX | NAV3251 | POSITIONER ISSUES | 20 | ASM0293-01 | RNX-PO-119 | X - positioner issue During FSE visit onsite, the left positioner handle was found broken. The handle was replaced, system passed all tests and found working as designed. It was also found during FSE visit that there were side swipes on the left side of cart, same side of the broken handle. Such damage, according to FSE, indicate improper handling of the system. RNX-PO-119 | During FSE visit onsite, the left positioner handle was found broken. The handle was replaced, system passed all tests and found working as designed. It was also found during FSE visit that there were side swipes on the left side of cart, same side of the broken handle. Such damage, according to FSE, indicate improper handling of the system. | x positioner issue fse visit onsite left positioner handle broken handle replaced passed tests working designed fse visit swipes left cart broken handle damage according fse indicate improper handling | NAV3251 | NAV2058 | NAV2055 | NAV3295 | NAV4091 | NAV3251NAV2058NAV2055NAV3295NAV4091 | True
9410 | 704632172 | XSE - Pneumatic System Not Pressurizing | 2021-09-27T22:05:08.000000+0000 | MazorX Stealth Edition | NAV2057 | SURGICAL SYSTEM ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - Pneumatic System Not Pressurizing 2021-Sep-27 Salesforce C00235236, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the system was unable to pressurize. After several reboots, the system was able to perform as intended until dismounting. While dismounting, the pneumatic system went down again making it difficult to put away. This was for a demo, there was no patient present. 2021-10-01 IUD x2 (Other): No new information. 2021-10-11 e1 (Other): No new information. 2021-10-13 e2, e3 (Rep): Additional information received from a manufacturer representative reported that the pneumatic manifold was replaced and the issue was resolved. 2021-10-27 e4, e5 (Rep): No new information. UNK MAZOR X WORKSTATION | 2021-Sep-27 Salesforce C00235236, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the system was unable to pressurize. After several reboots, the system was able to perform as intended until dismounting. While dismounting, the pneumatic system went down again making it difficult to put away. This was for a demo, there was no patient present. 2021-10-01 IUD x2 (Other): No new information. 2021-10-11 e1 (Other): No new information. 2021-10-13 e2, e3 (Rep): Additional information received from a manufacturer representative reported that the pneumatic manifold was replaced and the issue was resolved. 2021-10-27 e4, e5 (Rep): No new information. | xse pneumatic pressurizing salesforce iud outside proc edure unable pressurize reboots able perform intended dismounting dismounting pneumatic went making difficult away demo present iud new new additional pneumatic manifold replaced issue resolved new unk mazor x workstation | NAV3264 | NAV3244 | NAV2058 | NAV2054 | NAV3262 | NAV3264NAV3244NAV2058NAV2054NAV3262 | False
1982 | 703223457 | X- SW Issues | 2019-06-24T15:56:00.000000+0000 | MazorX | NAV3246 | DEFECTS | 20 | KIT0766-01 | MAZOR X SOFTWARE KIT VER-3.0 | X- SW Issues This has happened on 3 different trajectories. Before sending the Arm to T12 R I checked the solution in the VTK. The time that was written was 19 seconds. I practice , The movement took much longer (~30 sec) the movement was different then the VTK. and a 360 deg rotation of joint 6 was add to the movement. I think this is a SW bug in version 3. MAZOR X SOFTWARE KIT VER-3.0 | This has happened on 3 different trajectories. Before sending the Arm to T12 R I checked the solution in the VTK. The time that was written was 19 seconds. I practice , The movement took much longer (~30 sec) the movement was different then the VTK. and a 360 deg rotation of joint 6 was add to the movement. I think this is a SW bug in version 3. | x sw issues happened different trajectories sending arm r checked solution vtk time written seconds practice movement took longer movement different vtk deg rotation joint add movement think sw bug version mazor x software kit | NAV3246 | NAV3249 | NAV4091 | NAV3247 | NAV3242 | NAV3246NAV3249NAV4091NAV3247NAV3242 | True
6127 | 703871026 | X - Deviation | 2020-07-26T23:09:00.000000+0000 | MazorX | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | X - Deviation 2020-03-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs and deformity with fixation. It was reported that during a T10-L5 procedure, right L3 was deviated in the lateral direction due to skiving. There were no other issues during the procedure. There was no patient harm and the procedure was delayed less than an hour. MAZOR X SYSTEM TPL0059 | 2020-03-04 CCR (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs and deformity with fixation. It was reported that during a T10-L5 procedure, right L3 was deviated in the lateral direction due to skiving. There were no other issues during the procedure. There was no patient harm and the procedure was delayed less than an hour. | x deviation ccr spinal procedure treat degenerative discs deformity fixation procedure right deviated lateral direction skiving issues procedure procedure delayed hour mazor x | NAV4091 | NAV5000 | NAV3242 | NAV2047 | NAV4112 | NAV4091NAV5000NAV3242NAV2047NAV4112 | True
9318 | 704607088 | XSE - Difficult to Maneuver Positioner | 2021-09-13T17:55:52.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 40 | KIT0925 | MANIPULATOR SOLENOID | XSE - Difficult to Maneuver Positioner 2021-Sep-13 Salesforce C00234319 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was difficulty moving manipulator from bed. Attempted to mount then unmount, pressure is at 6.6. There was no patient involved. 2021-09-27 IUD x2 (Other): No new information. 2021-10-08 IUD x2 (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-10-27 IUD x2 (Other): No new information. 2021-12-15 IUD (Other): No new information. 2021-12-19 IUD (Other): No new information. MANIPULATOR SOLENOID | 2021-Sep-13 Salesforce C00234319 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that there was difficulty moving manipulator from bed. Attempted to mount then unmount, pressure is at 6.6. There was no patient involved. 2021-09-27 IUD x2 (Other): No new information. 2021-10-08 IUD x2 (Other): No new information. 2021-10-17 IUD (Other): No new information. 2021-10-27 IUD x2 (Other): No new information. 2021-12-15 IUD (Other): No new information. 2021-12-19 IUD (Other): No new information. | xse difficult maneuver positioner salesforce outside procedure difficulty moving manipulator bed attempted mount unmount pressure involved iud new iud new iud new iud new iud new iud new manipulator solenoid | NAV3260 | NAV3250 | NAV3264 | NAV3251 | NAV1006 | NAV3260NAV3250NAV3264NAV3251NAV1006 | True
8765 | 704478395 | XSE - Surgeon Screen Loose | 2021-06-30T23:59:43.000000+0000 | MazorX Stealth Edition | NAV3253 | TABLET ISSUES | 20 | ASM0357-01 | SURGEON SCREEN | XSE - Surgeon Screen Loose 2021-Jun-30 Salesforce C00229366, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgeon screen was loose on the system. There was no patient involvement. 2021-07-07 IUD (Other): No new information. SURGEON SCREEN | 2021-Jun-30 Salesforce C00229366, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgeon screen was loose on the system. There was no patient involvement. 2021-07-07 IUD (Other): No new information. | xse surgeon screen loose salesforce iud outside proc edure surgeon screen loose iud new surgeon screen | NAV3253 | NAV4109 | NAV2057 | NAV2058 | NAV2054 | NAV3253NAV4109NAV2057NAV2058NAV2054 | True
260 | 703056814 | XSE - stylus marker worn/dvd drive | 2019-03-15T14:24:15.000000+0000 | MazorX Stealth Edition | NAV3257 | CD DRIVE ISSUES | 40 | DSD0038-01 | DVD, PORTABLE DVD REWRITER, GP70NS50 | XSE - stylus marker worn/dvd drive 2019-Mar-13 SCL1900579, intake form, SCL1900580 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the old stylus markers were worn. It was also noted that the DVD drive was found to be non-functional. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. DVD, PORTABLE DVD REWRITER, GP70NS50 | 2019-Mar-13 SCL1900579, intake form, SCL1900580 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the old stylus markers were worn. It was also noted that the DVD drive was found to be non-functional. It was noted there was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse stylus marker worndvd drive spinal procedure old stylus markers worn noted dvd drive nonfunctional noted additional complications anticipated dvd portable dvd rewriter | NAV3257 | NAV3295 | NAV3299 | NAV5000 | NAV2054 | NAV3257NAV3295NAV3299NAV5000NAV2054 | True
2718 | 703247330 | X - Arm joint shifts/CANBUS error | 2019-07-08T20:54:56.000000+0000 | MazorX | NAV2047 | COMMUNICATION ISSUES | 30 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | X - Arm joint shifts/CANBUS error Malfunction occuring on Arm #12 Arm is shaking during several movement, and the issue should be inspected. Arm Shift Detected on two different occasions 1 st time Arm Shift detected in Joint #2 (Happened twice) – 79th Procedure @ SHHEC. 2 nd time Arm Shift detected in joint #4 (once) followed by Fatal Error: CANBUS connection error message – 80th procedure @ SHHEC UNK MAZOR X WORKSTATION | Malfunction occuring on Arm #12 Arm is shaking during several movement, and the issue should be inspected. Arm Shift Detected on two different occasions 1 st time Arm Shift detected in Joint #2 (Happened twice) – 79th Procedure @ SHHEC. 2 nd time Arm Shift detected in joint #4 (once) followed by Fatal Error: CANBUS connection error message – 80th procedure @ SHHEC | x arm joint shiftscanbus error malfunction occuring arm arm shaking movement issue inspected arm shift detected different occasions st time arm shift detected joint – procedure shhec nd time arm shift detected joint followed fatal error canbus connection error message – procedure shhec unk mazor x workstation | NAV4098 | NAV4101 | NAV4095 | NAV2054 | NAV2047 | NAV4098NAV4101NAV4095NAV2054NAV2047 | True
1477 | 703202232 | X - accessory issue | 2019-06-12T12:38:07.000000+0000 | MazorX | NAV1013 | MAZOR X/STEALTH EDITION | 20 | KIT0737-02 | SHORT DRILL GUIDE AND SHORT CANNULA | X - accessory issue Brian called and requested short drill guide and cannula SHORT DRILL GUIDE AND SHORT CANNULA | Brian called and requested short drill guide and cannula | x accessory issue brian called requested short drill guide cannula short drill guide short cannula | NAV3299 | NAV3295 | NAV4091 | NAV3242 | NAV4090 | NAV3299NAV3295NAV4091NAV3242NAV4090 | False
6448 | 703941951 | XSE - Damaged/Lost Instruments | 2020-09-04T21:48:56.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 40 | MAS2351-04S | KIT SVC MAS2351-04S SCHANZ ARM | XSE - Damaged/Lost Instruments 2020-09-04 00695938 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the accuracy pointer, screw driver, and Schanz arm was damaged or missing. There was no patient involvement. 2020-09-08 e1 (Rep): All were missing, none were damaged. 2020-09-15 IUD x2 (Other): No new information. 2020-09-29 IUD x2 (Other): No new information. KIT SVC MAS2351-04S SCHANZ ARM | 2020-09-04 00695938 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the accuracy pointer, screw driver, and Schanz arm was damaged or missing. There was no patient involvement. 2020-09-08 e1 (Rep): All were missing, none were damaged. 2020-09-15 IUD x2 (Other): No new information. 2020-09-29 IUD x2 (Other): No new information. | xse damagedlost instruments outside procedure accuracy pointer screw driver schanz arm damaged missing missing damaged iud new iud new kit svc schanz arm | NAV2053 | NAV3299 | NAV2054 | NAV2055 | NAV3295 | NAV2053NAV3299NAV2054NAV2055NAV3295 | True
5326 | 703648908 | XSE - Surgical Arm Error | 2020-03-04T12:16:28.000000+0000 | MazorX Stealth Edition | NAV4098 | ARM JOINT SHIFTS | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Surgical Arm Error 2020-MAR-04 00685607 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a quote for a replacement arm was requested due to some difficult arm error received. There was no patient present when this issue was identified. 2020-03-04 IUD x3, Email 2020-03-09 15:01 (Rep): Additional information received from a manufacturer representative reported that there was an error during initial Startup where they initially had a Joint 2 error which was fixed by a hard restart. The next issue was the system was working fine and a nurse unplugged the system while working. When rebooted the system was taking in excess of 5-8 minutes to start up and the software for Mazor was not working. Multiple hard restarts were tried but did not fix the issue. The representative tried another outlet and a restart, but this too did not fix the issue. The representative finally let the system stay on for an extended period of time and they were finally able to get the software up and running. 2020-03-30 Email 2020-03-30 14:15 (Rep): Additional information received from a manufacturer representative reported that the issues occurred before the case and the error code was 4059. 2020-04-16 email 2020-04-16 11:30 x2, e1, e2 (Rep): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-MAR-04 00685607 (REP): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that a quote for a replacement arm was requested due to some difficult arm error received. There was no patient present when this issue was identified. 2020-03-04 IUD x3, Email 2020-03-09 15:01 (Rep): Additional information received from a manufacturer representative reported that there was an error during initial Startup where they initially had a Joint 2 error which was fixed by a hard restart. The next issue was the system was working fine and a nurse unplugged the system while working. When rebooted the system was taking in excess of 5-8 minutes to start up and the software for Mazor was not working. Multiple hard restarts were tried but did not fix the issue. The representative tried another outlet and a restart, but this too did not fix the issue. The representative finally let the system stay on for an extended period of time and they were finally able to get the software up and running. 2020-03-30 Email 2020-03-30 14:15 (Rep): Additional information received from a manufacturer representative reported that the issues occurred before the case and the error code was 4059. 2020-04-16 email 2020-04-16 11:30 x2, e1, e2 (Rep): No new information. | xse surgical arm error outside procedure quote replacement arm requested difficult arm error present issue identified iud email additional error initial startup initially joint error fixed hard restart issue working fine nurse unplugged working rebooted taking excess minutes start software mazor working multiple hard restarts tried did fix issue tried outlet restart did fix issue finally let stay extended period time finally able software running email additional issues occurred case error code email new kit svc surgical arm | NAV2047 | NAV3249 | NAV3246 | NAV2054 | NAV3242 | NAV2047NAV3249NAV3246NAV2054NAV3242 | False
8962 | 704516084 | XSE - CANBUS Error Messages | 2021-07-22T08:11:45.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | UNK_MAZORX_WKSTN | UNK MAZOR X WORKSTATION | XSE - CANBUS Error Messages 2021-Jul-21 Salesforce C00230591 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that two error messages were displayed during the procedure after registration was completed. An error 295 (CANBUS connection lost on knot 1,2,3,4,5,6,7 Positioner 8 Cart I0) and error 295 (enforcement on driver lost. Maintenance needed) was displayed. A hard restart of the system was done and the manufacturer representative waited five minutes before booting up the system. There was no patient harm and the procedure was delayed less than an hour.\ 2021-07-23 email 2021-07-23 07:45 (Rep): Additional information was received stating that a hard reboot solved the issue. The arm could move after the hard reboot. The cause or suspected cause of the error messages could have been due to the fact that there was still the old connection cable from workstation to the arm. This should be changed in time to avoid this error. The surgeon finished the operation freehand. UNK MAZOR X WORKSTATION | 2021-Jul-21 Salesforce C00230591 (Rep, HCP, For): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that two error messages were displayed during the procedure after registration was completed. An error 295 (CANBUS connection lost on knot 1,2,3,4,5,6,7 Positioner 8 Cart I0) and error 295 (enforcement on driver lost. Maintenance needed) was displayed. A hard restart of the system was done and the manufacturer representative waited five minutes before booting up the system. There was no patient harm and the procedure was delayed less than an hour.\ 2021-07-23 email 2021-07-23 07:45 (Rep): Additional information was received stating that a hard reboot solved the issue. The arm could move after the hard reboot. The cause or suspected cause of the error messages could have been due to the fact that there was still the old connection cable from workstation to the arm. This should be changed in time to avoid this error. The surgeon finished the operation freehand. | xse canbus error messages salesforce spinal procedure error messages displayed procedure registration completed error error displayed hard restart waited minutes booting procedure delayed hour email additional stating hard reboot solved issue arm hard reboot cause suspected cause error messages fact old connection cable workstation arm changed time avoid error surgeon finished operation freehand unk mazor x workstation | NAV2047 | NAV2058 | NAV4095 | NAV3264 | NAV3263 | NAV2047NAV2058NAV4095NAV3264NAV3263 | True
1931 | 703221165 | X - ARM ISSUE | 2019-06-21T20:12:12.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - ARM ISSUE Arm bent back on its self when sent to AP. This repeated when sending to a trajectory during a case. When moving to the last trajectory the Arm began shaking and stop shaking when sent again. When sending to a trajectory the arm collided with the emergency stop button./ Checks Preformed. 1. Advanced Accuracy Check to include a check to that joint 6 is stable 2. Robot X Eyes 3. SA Cameras 4. Shoulder calibration 5. Pressure Bar: Issue with pressure, preformed soft and hard restart which corrected the issue. Pre-op Setup 1. Imaged adaptor attached 2. Bed rail adaptors mounted 3. Bed Mount placed on Jackson Table 4. Accuracy check preformed 6. Tentative plan created 7. Surgeon checked plan 8. Adjusted bed mount 9. Attached the SA to the bed frame 10. Prepped the scrub tech, place items in order of attachment 11. Calibrated C Arm 12. Walked the OR team through draping the SA 12:24. Time out: 12:26. Intra-op 1. Skin incision for mount 2. Mounted SA to the patient without issue 3. X Eye immediately recognized the reference marker 4. CSR preformed the 3 Define scan without any issue 5. Draw spine was conducted without issue, CSR labeled L5, checked OB and sent to AP. At this point the SA began rotating and folding on itself in a very awkward position. See pic 6. CSR sent the SA to the OB without issue, one AP was taken, and two OB were taken 7. CSR sent to the SA to right side L5, then S1. The surgeon lengthen the incision 8. The CSR sent to clear up while the surgeon used his bovie. 9. The CSR sent back to right side trajectories and the surgeon instrumented without issue. 10. The CSR sent the RBT to Left side trajectories for skin incisions. The surgeon lengthen the incisions. The CSR sent the SA to clear up during the bovie use. 11. The CSR sent back to the Left side, L5 was instrumented without issues. During the instrumentation on S1 the SA began to shake violently prior to the use of the mallet. The CSR instructed the surgeon to remove the tools, CSR sent the SA to clear up and back to the trajectory. 12. All trajectories were accurate./ RBT arm started violently shaking on last trajectory of the procedure. In conjunction with the shaking, rbt also attempted to hit its on e-stop when sending trajectories. Following the case, RTC performed a stress test and tested individual joints. Joints 2 and 3 had severe malfunction. UNK MAZOR X ARM | Arm bent back on its self when sent to AP. This repeated when sending to a trajectory during a case. When moving to the last trajectory the Arm began shaking and stop shaking when sent again. When sending to a trajectory the arm collided with the emergency stop button./ Checks Preformed. 1. Advanced Accuracy Check to include a check to that joint 6 is stable 2. Robot X Eyes 3. SA Cameras 4. Shoulder calibration 5. Pressure Bar: Issue with pressure, preformed soft and hard restart which corrected the issue. Pre-op Setup 1. Imaged adaptor attached 2. Bed rail adaptors mounted 3. Bed Mount placed on Jackson Table 4. Accuracy check preformed 6. Tentative plan created 7. Surgeon checked plan 8. Adjusted bed mount 9. Attached the SA to the bed frame 10. Prepped the scrub tech, place items in order of attachment 11. Calibrated C Arm 12. Walked the OR team through draping the SA 12:24. Time out: 12:26. Intra-op 1. Skin incision for mount 2. Mounted SA to the patient without issue 3. X Eye immediately recognized the reference marker 4. CSR preformed the 3 Define scan without any issue 5. Draw spine was conducted without issue, CSR labeled L5, checked OB and sent to AP. At this point the SA began rotating and folding on itself in a very awkward position. See pic 6. CSR sent the SA to the OB without issue, one AP was taken, and two OB were taken 7. CSR sent to the SA to right side L5, then S1. The surgeon lengthen the incision 8. The CSR sent to clear up while the surgeon used his bovie. 9. The CSR sent back to right side trajectories and the surgeon instrumented without issue. 10. The CSR sent the RBT to Left side trajectories for skin incisions. The surgeon lengthen the incisions. The CSR sent the SA to clear up during the bovie use. 11. The CSR sent back to the Left side, L5 was instrumented without issues. During the instrumentation on S1 the SA began to shake violently prior to the use of the mallet. The CSR instructed the surgeon to remove the tools, CSR sent the SA to clear up and back to the trajectory. 12. All trajectories were accurate./ RBT arm started violently shaking on last trajectory of the procedure. In conjunction with the shaking, rbt also attempted to hit its on e-stop when sending trajectories. Following the case, RTC performed a stress test and tested individual joints. Joints 2 and 3 had severe malfunction. | x arm issue arm bent self sent ap repeated sending trajectory case moving trajectory arm began shaking stop shaking sent sending trajectory arm collided emergency stop button checks preformed advanced accuracy check include check joint stable robot x eyes sa cameras shoulder calibration pressure bar issue pressure preformed soft hard restart corrected issue preop setup imaged adaptor attached bed rail adaptors mounted bed mount placed jackson table accuracy check preformed tentative plan created surgeon checked plan adjusted bed mount attached sa bed frame prepped scrub tech place items order attachment calibrated c arm walked team draping sa time intraop skin incision mount mounted sa issue x eye immediately recognized reference marker csr preformed define scan issue draw spine conducted issue csr labeled checked ob sent ap point sa began rotating folding awkward position pic csr sent sa ob issue ap taken ob taken csr sent sa right surgeon lengthen incision csr sent clear surgeon bovie csr sent right trajectories surgeon instrumented issue csr sent rbt left trajectories skin incisions surgeon lengthen incisions csr sent sa clear bovie use csr sent left instrumented issues instrumentation sa began shake violently prior use mallet csr instructed surgeon remove tools csr sent sa clear trajectory trajectories accurate rbt arm started violently shaking trajectory procedure conjunction shaking rbt attempted hit estop sending trajectories following case rtc performed stress test tested individual joints joints severe malfunction unk mazor x arm | NAV4101 | NAV3299 | NAV3250 | NAV3263 | NAV4097 | NAV4101NAV3299NAV3250NAV3263NAV4097 | True
7294 | 704169362 | XSE - SRA Accuracy Test Not Saving | 2021-01-12T21:27:52.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT0830 | USA REGIONAL KIT KIT0830 | XSE - SRA Accuracy Test Not Saving 2021-Jan-11 Salesforce C00219146 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that prior to a procedure, the 3 point surgical arm accuracy test was not displaying as complete. The manufacturer representative repeated the test three times with no change. A soft reboot of the system was done. The test was completed again and the software displayed it as completed. There was no patient involvement. 2021-01-14 e1, e2 (Rep): Additional information received from a manufacturer representative reported that no error messages were dis played. The test would finish like it normally does, the system would not say the test was complete. 2021-01-20 IUD (Other): No new information. USA REGIONAL KIT KIT0830 | 2021-Jan-11 Salesforce C00219146 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that prior to a procedure, the 3 point surgical arm accuracy test was not displaying as complete. The manufacturer representative repeated the test three times with no change. A soft reboot of the system was done. The test was completed again and the software displayed it as completed. There was no patient involvement. 2021-01-14 e1, e2 (Rep): Additional information received from a manufacturer representative reported that no error messages were dis played. The test would finish like it normally does, the system would not say the test was complete. 2021-01-20 IUD (Other): No new information. | xse sra accuracy test saving salesforce outside procedure prior procedure point surgical arm accuracy test displaying complete repeated test times change soft reboot test completed software displayed completed additional error messages dis played test finish like normally does say test complete iud new usa regional kit | NAV3246 | NAV4092 | NAV3249 | NAV4093 | NAV3248 | NAV3246NAV4092NAV3249NAV4093NAV3248 | True
9635 | 704687304 | XSE - Unable Recognize Star Marker | 2021-10-28T18:25:40.000000+0000 | MazorX Stealth Edition | NAV3235 | 3D SCANS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Unable Recognize Star Marker 2021-Oct-28 Salesforce C00236990, C00237082 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after transferring the O-arm spin to the guidance system, an error message 295 was displayed stating there was an issue with detecting the star marker. The procedure was changed to use CT-Fluoro workflow. After the case, the manufacturer representative performed rad/fluoro gain calibrations on the O-arm. When using the same star marker, the representative was able to proceed through a demo case. There was no patient harm and the procedure was delayed less than an hour. 2021-11-05 IUD (Other): No new information. 2022-01-07 IUD (Other): No new information. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2021-Oct-28 Salesforce C00236990, C00237082 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that after transferring the O-arm spin to the guidance system, an error message 295 was displayed stating there was an issue with detecting the star marker. The procedure was changed to use CT-Fluoro workflow. After the case, the manufacturer representative performed rad/fluoro gain calibrations on the O-arm. When using the same star marker, the representative was able to proceed through a demo case. There was no patient harm and the procedure was delayed less than an hour. 2021-11-05 IUD (Other): No new information. 2022-01-07 IUD (Other): No new information. | xse unable recognize star marker salesforce spinal procedure transferring oarm spin error message displayed stating issue detecting star marker procedure changed use ctfluoro workflow case performed radfluoro gain calibrations oarm using star marker able proceed demo case procedure delayed hour iud new iud new mazor x software installation kit | NAV3235 | NAV3246 | NAV2047 | NAV4112 | NAV5000 | NAV3235NAV3246NAV2047NAV4112NAV5000 | True
4506 | 703548192 | X - Arm and 3Define camera issues | 2020-01-07T15:10:23.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X - Arm and 3Define camera issues 2018-07-20 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on level L3 - S1. The surgeon planned the case day of surgery. The system was attached to the bed and draped. The schanz pin was Inserted into a stable spinous process. 3Define failed. 3Define went through the right side of the surface scan perfectly but when it looked like it was almost completing finishing the left side of the surface scan, the guidance system would beep with a couldn't reach trajectory beep but would have a blank image on the workstation. Surgeon wanted to continue with a generic workstation. Four Ap and three oblique images were taken. Registration was achieved and operation started. Skin was cut and fluoro images taken. The operation started from left L3 - S1 and then Right L3 - S1. Confirmation images were taken, and suture done. During the procedure; the Arm shook in a finished sent trajectory position. Arm would not finish sending to any predefine trajectories. 3Define would not finish and would have no images of the surface scan. L3 and S1 trajectories were accurate. No patient harm was reported. UNK MAZOR X 3DEFINE CAMERA | 2018-07-20 CCR (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was for four vertebrae on level L3 - S1. The surgeon planned the case day of surgery. The system was attached to the bed and draped. The schanz pin was Inserted into a stable spinous process. 3Define failed. 3Define went through the right side of the surface scan perfectly but when it looked like it was almost completing finishing the left side of the surface scan, the guidance system would beep with a couldn't reach trajectory beep but would have a blank image on the workstation. Surgeon wanted to continue with a generic workstation. Four Ap and three oblique images were taken. Registration was achieved and operation started. Skin was cut and fluoro images taken. The operation started from left L3 - S1 and then Right L3 - S1. Confirmation images were taken, and suture done. During the procedure; the Arm shook in a finished sent trajectory position. Arm would not finish sending to any predefine trajectories. 3Define would not finish and would have no images of the surface scan. L3 and S1 trajectories were accurate. No patient harm was reported. | x arm camera issues ccr spinal procedure case vertebrae level surgeon planned case day surgery attached bed draped schanz pin inserted stable spinous process failed went right surface scan perfectly looked like completing finishing left surface scan beep reach trajectory beep blank image workstation surgeon wanted continue generic workstation ap oblique images taken registration achieved operation started skin cut fluoro images taken operation started left right confirmation images taken suture procedure arm shook finished sent trajectory position arm finish sending predefine trajectories finish images surface scan trajectories accurate unk mazor x camera | NAV3244 | NAV4095 | NAV2058 | NAV4110 | NAV3261 | NAV3244NAV4095NAV2058NAV4110NAV3261 | True
4104 | 703500187 | XSE - Firmware Update Needed/OOBF | 2019-12-04T15:20:21.000000+0000 | MazorX Stealth Edition | NAV2056 | SOFTWARE ISSUES | 30 | KIT0749-01 | KIT SPARE PART FOR THE 3D CAMERA | XSE - Firmware Update Needed/OOBF 2019-12-04 00679546, IUD x2, PDF, 2019-12-08 14:51 x2 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the 3Define camera firmware was incorrect during install of the system. There was no patient involvement. 2019-12-13 IUD x2 (Other): No new information. KIT SPARE PART FOR THE 3D CAMERA | 2019-12-04 00679546, IUD x2, PDF, 2019-12-08 14:51 x2 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that the 3Define camera firmware was incorrect during install of the system. There was no patient involvement. 2019-12-13 IUD x2 (Other): No new information. | xse firmware update neededoobf iud pdf prior procedure camera firmware incorrect install iud new kit spare camera | NAV2054 | NAV3244 | NAV3299 | NAV3249 | NAV3246 | NAV2054NAV3244NAV3299NAV3249NAV3246 | False
3611 | 703388356 | X - target extender screw missing | 2019-09-30T19:59:05.000000+0000 | MazorX | NAV2053 | MISSING PARTS | 20 | MAS2345-01 | TARGET EXTENDER SCREW | X - target extender screw missing 2019-Sep-27 SCL1902932 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the target extender screw was not in the instrument tray. There was no patient harm or additional complications reported or anticipated as a result of the event. TARGET EXTENDER SCREW | 2019-Sep-27 SCL1902932 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the target extender screw was not in the instrument tray. There was no patient harm or additional complications reported or anticipated as a result of the event. | x target extender screw missing spinal procedure target extender screw instrument tray additional complications anticipated target extender screw | NAV2053 | NAV3299 | NAV3295 | NAV2054 | NAV4090 | NAV2053NAV3299NAV3295NAV2054NAV4090 | True
3164 | 703268322 | XSE - Deviation | 2019-07-19T15:04:50.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - Deviation 2019-Jul-18 SCL1902131 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that L1-S1 trajectories were planned with an open approach. Upon placement of the guide wire, they moved directly to the screw. Once they took x-rays, all screws looked good but the L1 & L2 screws looked lateral. Surgeon and PA looked at the screws and deemed them to be acceptable. They still seemed lateral to the rep but with the surgeons approval they moved forward with the operation. The docking sites may have not been prepared enough before drilling for both trajectories but, all other screws were accurate. Plan was done in conjunction with the rep and approved of by both the surgeon and PA. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-07-19 email 2019-07-19 14:45 (Rep): Additional information received from a manufacturer representative reported the amount of deviation was unknown. The surgeon indicated the screws looked accurate, but the representative indicated they were slightly off. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-Jul-18 SCL1902131 intake form (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that L1-S1 trajectories were planned with an open approach. Upon placement of the guide wire, they moved directly to the screw. Once they took x-rays, all screws looked good but the L1 & L2 screws looked lateral. Surgeon and PA looked at the screws and deemed them to be acceptable. They still seemed lateral to the rep but with the surgeons approval they moved forward with the operation. The docking sites may have not been prepared enough before drilling for both trajectories but, all other screws were accurate. Plan was done in conjunction with the rep and approved of by both the surgeon and PA. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. 2019-07-19 email 2019-07-19 14:45 (Rep): Additional information received from a manufacturer representative reported the amount of deviation was unknown. The surgeon indicated the screws looked accurate, but the representative indicated they were slightly off. | xse deviation spinal procedure trajectories planned open approach placement guide wire moved directly screw took xrays screws looked good screws looked lateral surgeon pa looked screws deemed acceptable lateral surgeons approval moved forward operation docking sites prepared drilling trajectories screws accurate plan conjunction approved surgeon pa additional complications anticipated email additional deviation unknown surgeon indicated screws looked accurate indicated slightly mazor x software installation kit | NAV5000 | NAV4091 | NAV3246 | NAV2056 | NAV4112 | NAV5000NAV4091NAV3246NAV2056NAV4112 | True
1906 | 703220978 | X - Arm joint issue | 2019-06-21T17:16:16.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 20 | UNK_MAZORX_ARM | UNK MAZOR X ARM | X - Arm joint issue Joint #1 failure after SW 3.0.1 upgrade UNK MAZOR X ARM | Joint #1 failure after SW 3.0.1 upgrade | x arm joint issue joint failure sw upgrade unk mazor x arm | NAV4098 | NAV4101 | NAV3249 | NAV2054 | NAV2058 | NAV4098NAV4101NAV3249NAV2054NAV2058 | True
3615 | 703390455 | XSE - Controller Error/Arm Joint Error | 2019-10-01T16:46:13.000000+0000 | MazorX Stealth Edition | NAV2047 | COMMUNICATION ISSUES | 20 | ASM0206-01 | SURGICAL ARM | XSE - Controller Error/Arm Joint Error 2019-09-27 SCL1902939, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the controller lost connection to nodes 1-6 during the procedure. The use of the guidance system was aborted and the case was completed with navigation. A hard restart of the system was done later and connection was reestablished. There was no patient harm and the procedure was not delayed more than an hour. SURGICAL ARM | 2019-09-27 SCL1902939, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the controller lost connection to nodes 1-6 during the procedure. The use of the guidance system was aborted and the case was completed with navigation. A hard restart of the system was done later and connection was reestablished. There was no patient harm and the procedure was not delayed more than an hour. | xse controller errorarm joint error spinal procedure controller lost connection nodes procedure use aborted case completed navigation hard restart later connection reestablished procedure delayed hour surgical arm | NAV2047 | NAV3249 | NAV4095 | NAV4112 | NAV5000 | NAV2047NAV3249NAV4095NAV4112NAV5000 | True
1151 | 703191392 | X- Software Issues | 2019-06-05T19:56:41.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | UNK_MAZORX_3DCAM | UNK MAZOR X 3DEFINE CAMERA | X- Software Issues Surgeon reviewed plan prior to procedure.Surgeon decided to use the single schanz in the right psis with the bmb. 3Define scan @1323 would not scan. The Firmware Version # disappeared. RTC stopped and restarted scan before unlocking and checking cameras and noticed version # was gone. RTC did a soft restart and the firmware version came back. RTC then did a 3Define with good success. Drawspine was done from L4-S1 and L5 was labeled. Registration first attempted failed. Imaging was rough due to high BMI patient and cross view error seemed to be the issue. RTC then took a new AP and was able to register. This was achieved on basic algorithm with good values except L3 values which were Red but wasnt going to be operated on regardless. Surgeon drilled all trajectories to planned position and was satisfied with the outcome. UNK MAZOR X 3DEFINE CAMERA | Surgeon reviewed plan prior to procedure.Surgeon decided to use the single schanz in the right psis with the bmb. 3Define scan @1323 would not scan. The Firmware Version # disappeared. RTC stopped and restarted scan before unlocking and checking cameras and noticed version # was gone. RTC did a soft restart and the firmware version came back. RTC then did a 3Define with good success. Drawspine was done from L4-S1 and L5 was labeled. Registration first attempted failed. Imaging was rough due to high BMI patient and cross view error seemed to be the issue. RTC then took a new AP and was able to register. This was achieved on basic algorithm with good values except L3 values which were Red but wasnt going to be operated on regardless. Surgeon drilled all trajectories to planned position and was satisfied with the outcome. | x software issues surgeon reviewed plan prior proceduresurgeon decided use single schanz right psis bmb scan scan firmware version disappeared rtc stopped restarted scan unlocking checking cameras noticed version gone rtc did soft restart firmware version came rtc did good success drawspine labeled registration attempted failed imaging rough high bmi cross view error issue rtc took new ap able register achieved basic algorithm good values values red wasnt going operated regardless surgeon drilled trajectories planned position satisfied outcome unk mazor x camera | NAV3244 | NAV3250 | NAV3242 | NAV3249 | NAV3235 | NAV3244NAV3250NAV3242NAV3249NAV3235 | True
6692 | 704007657 | XSE ¦ Missing Clamp Link Screw | 2020-10-12T20:13:14.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MEC5206-01S | KIT SVC MEC5206-01S RNX-ACC-021 | XSE ¦ Missing Clamp Link Screw 2020-Oct-12 00698607 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in outside of a procedure. It was reported that they were missing the screw to the clamp link. They grabbed a clamp link from another tray to use for the procedure. There was no patient involved. KIT SVC MEC5206-01S RNX-ACC-021 | 2020-Oct-12 00698607 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in outside of a procedure. It was reported that they were missing the screw to the clamp link. They grabbed a clamp link from another tray to use for the procedure. There was no patient involved. | xse missing clamp link screw outside procedure missing screw clamp link grabbed clamp link tray use procedure involved kit svc | NAV2053 | NAV2054 | NAV3299 | NAV4090 | NAV3243 | NAV2053NAV2054NAV3299NAV4090NAV3243 | True
10030 | 704775614 | XSE - Lost X-Eye Function | 2021-12-21T20:34:51.000000+0000 | MazorX Stealth Edition | NAV3256 | CABLE ISSUES | 60 | UNK_MAZOR_COMP | UNKNOWN MAZOR COMPONENT | XSE - Lost X-Eye Function 2021-Dec-21 Salesforce C00240056 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that they lost XEye function, the cover will not open without override, and they experienced black screen on startup. There was no patient involved. 2021-12-24 IUD x2 (Other): No new information. 2022-01-19 email 2022-01-19 18:02 (Rep): Additional information received from a manufacturer representative reported that the PC was going to be replaced. 2022-01-25 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. 2022-02-01 email 2022-02-01 15:33 (Rep): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. UNKNOWN MAZOR COMPONENT | 2021-Dec-21 Salesforce C00240056 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that they lost XEye function, the cover will not open without override, and they experienced black screen on startup. There was no patient involved. 2021-12-24 IUD x2 (Other): No new information. 2022-01-19 email 2022-01-19 18:02 (Rep): Additional information received from a manufacturer representative reported that the PC was going to be replaced. 2022-01-25 IUD (Other): No new information. 2022-01-31 IUD (Other): No new information. 2022-02-01 email 2022-02-01 15:33 (Rep): No new information. 2022-02-23 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. 2022-03-15 IUD (Other): No new information. | xse lost xeye function salesforce outside procedure lost xeye function cover open override experienced black screen startup involved iud new email additional pc going replaced iud new iud new email new iud new iud new iud new unknown mazor component | NAV4110 | NAV2053 | NAV2054 | NAV3246 | NAV2058 | NAV4110NAV2053NAV2054NAV3246NAV2058 | False
1559 | 703205332 | X- Arm Issues | 2019-06-13T15:00:39.000000+0000 | MazorX | NAV4101 | UNEXPECTED ARM MOVEMENT | 20 | ASM0206-01 | SURGICAL ARM | X- Arm Issues 2018-01-22 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a degenerative global deformity with fixation at T10-S2. It was reported that the manufacturer representative completed planning the day of the procedure and asked the surgeon to review the surgical plans prior to the case but the surgeon did not review the plans. Trajectories were planned from S2 to T10. L1-L3 were noted by the representative to have small pedicles bilaterally suggesting potential in-out-in trajectories as well as potential for skiving. L5 right was noted by the representative to have a large anatomical gap between facet joints. The called out these notes during the procedure. The system pass both a 10 point and 3 point test during set up. The system was then mounted to the bed and draped. At 1057 the system was mounted to the patient using schanz arm platform attached with two schanz pins, one in each patient's iliac crests. The first 3Define scan was not acceptable the first time due to the PA's hand being in the surgical field during the scan. The second 3Define scan was successful and draw spine was acquired without i ssue. During registration automatic registration failed and the scan did not match up; the OBL images appeared too lateral. The C-arm was adjusted and new images obtained; automatic registration was successful. The flow of the surgery was: right S2, left S2, left S1, left L5, left L4, left L3, rigth S1, right L5, right L4, right L3. At that point a second 3Define scan and draw spine were completed for a new segment. Registration was notably difficult. Automatic registration was unsuccessful twice. Manual registration was attempted and still unsuccessful. The C-arm was adjusted and registration continued to fail. Eventually registration was accomplished with the c-arm angle differing from the 3D marker auto placement. Eventually, automatic registration was achieved but the efforts took 23 minutes and did not include T10 vertebrae. At 1201 left L2 trajectory was sent and arm drop was observed but stopped. The arm was re-sent to the trajectory and the problem did not recur at that point and did not recur at any other time during the case. Case flow for the second segment of surgery was: left L2, left L1. left T12, left T11, right L2, right L1, right T12, right T11. The system was dismounted from the patient at 1224 and dismounted from the bed at 1228. Fluoroscopic images were acquired for confirmation visualization. Left T11, left L1 and right L3 were all deemed to be inferior to plan in the pedicles. No patient harm was reported and surgery was delayed less than an hour. During case review, the surgeon was unhappy with the speed, workflow, and mechanical failures of the surgical arm and the schanz arm not fitting over the pin. The surgeon specifically noted the extended registration time and T10 not being included in the second registration during case review. SURGICAL ARM | 2018-01-22 CCR (Rep, HCP) Medtronic received information regarding a guidance system being used during a spinal procedure to treat a degenerative global deformity with fixation at T10-S2. It was reported that the manufacturer representative completed planning the day of the procedure and asked the surgeon to review the surgical plans prior to the case but the surgeon did not review the plans. Trajectories were planned from S2 to T10. L1-L3 were noted by the representative to have small pedicles bilaterally suggesting potential in-out-in trajectories as well as potential for skiving. L5 right was noted by the representative to have a large anatomical gap between facet joints. The called out these notes during the procedure. The system pass both a 10 point and 3 point test during set up. The system was then mounted to the bed and draped. At 1057 the system was mounted to the patient using schanz arm platform attached with two schanz pins, one in each patient's iliac crests. The first 3Define scan was not acceptable the first time due to the PA's hand being in the surgical field during the scan. The second 3Define scan was successful and draw spine was acquired without i ssue. During registration automatic registration failed and the scan did not match up; the OBL images appeared too lateral. The C-arm was adjusted and new images obtained; automatic registration was successful. The flow of the surgery was: right S2, left S2, left S1, left L5, left L4, left L3, rigth S1, right L5, right L4, right L3. At that point a second 3Define scan and draw spine were completed for a new segment. Registration was notably difficult. Automatic registration was unsuccessful twice. Manual registration was attempted and still unsuccessful. The C-arm was adjusted and registration continued to fail. Eventually registration was accomplished with the c-arm angle differing from the 3D marker auto placement. Eventually, automatic registration was achieved but the efforts took 23 minutes and did not include T10 vertebrae. At 1201 left L2 trajectory was sent and arm drop was observed but stopped. The arm was re-sent to the trajectory and the problem did not recur at that point and did not recur at any other time during the case. Case flow for the second segment of surgery was: left L2, left L1. left T12, left T11, right L2, right L1, right T12, right T11. The system was dismounted from the patient at 1224 and dismounted from the bed at 1228. Fluoroscopic images were acquired for confirmation visualization. Left T11, left L1 and right L3 were all deemed to be inferior to plan in the pedicles. No patient harm was reported and surgery was delayed less than an hour. During case review, the surgeon was unhappy with the speed, workflow, and mechanical failures of the surgical arm and the schanz arm not fitting over the pin. The surgeon specifically noted the extended registration time and T10 not being included in the second registration during case review. | x arm issues ccr spinal procedure treat degenerative global deformity fixation completed planning day procedure asked surgeon review surgical plans prior case surgeon did review plans trajectories planned noted small pedicles bilaterally suggesting potential inoutin trajectories potential skiving right noted large anatomical gap facet joints called notes procedure pass point point test set mounted bed draped mounted using schanz arm platform attached schanz pins patients iliac crests scan acceptable time pas hand surgical field scan second scan successful draw spine acquired ssue registration automatic registration failed scan did match obl images appeared lateral carm adjusted new images obtained automatic registration successful flow surgery right left left left left left rigth right right right point second scan draw spine completed new segment registration notably difficult automatic registration unsuccessful twice manual registration attempted unsuccessful carm adjusted registration continued fail eventually registration accomplished carm angle differing marker auto placement eventually automatic registration achieved efforts took minutes did include vertebrae left trajectory sent arm drop observed stopped arm resent trajectory problem did recur point did recur time case case flow second segment surgery left left left left right right right right dismounted dismounted bed fluoroscopic images acquired confirmation visualization left left right deemed inferior plan pedicles surgery delayed hour case review surgeon unhappy speed workflow mechanical failures surgical arm schanz arm fitting pin surgeon specifically noted extended registration time included second registration case review surgical arm | NAV3298 | NAV3242 | NAV4091 | NAV3249 | NAV3244 | NAV3298NAV3242NAV4091NAV3249NAV3244 | False
9516 | 704661411 | XSE - Instrument Verification Cleared | 2021-10-13T22:15:48.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | M728894B252 | NAVIGATION SOFTWARE MODULE 1.2 | XSE - Instrument Verification Cleared 2021-Oct-13 Salesforce C00236215, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the instrument verification cleared and had to be redone during the procedure. During the procedure, instrument setup was completed, instruments were paired with Navlock trackers, and verification was done. The surgeon proceeded with the procedure and did a 3Define Scan, but Draw Spine was missing. Generic AP and oblique trajectories were used to acquire images and complete registration since the system was not tracking anything. Registration was completed without issue. The surgical arm was sent to the first trajectory, and no navigation was possible since the instruments were not verified. All of the instruments had to be reverified. There was no patient harm and the procedure was delayed less than an hour. 2021-22 IUD (Other): No new information. 2021-11-01 IUD x2 (Other): No new information. 2021-11-12 IUD (Other): No new information. 2021-11-30 IUD (Other): No new information. NAVIGATION SOFTWARE MODULE 1.2 | 2021-Oct-13 Salesforce C00236215, IUD (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the instrument verification cleared and had to be redone during the procedure. During the procedure, instrument setup was completed, instruments were paired with Navlock trackers, and verification was done. The surgeon proceeded with the procedure and did a 3Define Scan, but Draw Spine was missing. Generic AP and oblique trajectories were used to acquire images and complete registration since the system was not tracking anything. Registration was completed without issue. The surgical arm was sent to the first trajectory, and no navigation was possible since the instruments were not verified. All of the instruments had to be reverified. There was no patient harm and the procedure was delayed less than an hour. 2021-22 IUD (Other): No new information. 2021-11-01 IUD x2 (Other): No new information. 2021-11-12 IUD (Other): No new information. 2021-11-30 IUD (Other): No new information. | xse instrument verification cleared salesforce iud spinal procedure instrument verification cleared redone procedure procedure instrument setup completed instruments paired navlock trackers verification surgeon proceeded procedure did scan draw spine missing generic ap oblique trajectories acquire images complete registration tracking registration completed issue surgical arm sent trajectory navigation possible instruments verified instruments reverified procedure delayed hour iud new iud new iud new iud new navigation software module | NAV5001 | NAV3246 | NAV3299 | NAV4112 | NAV4110 | NAV5001NAV3246NAV3299NAV4112NAV4110 | True
8878 | 704501766 | XSE ¦ Lost Instruments | 2021-07-14T20:07:44.000000+0000 | MazorX Stealth Edition | NAV2053 | MISSING PARTS | 20 | MAS2310-04S | KIT SVC MAS2310-04S TARGET EXTENDER | XSE ¦ Lost Instruments 2021-Jul-14 Salesforce C00230205 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site lost the target extender and target extender screws for a set. No patient was present. 2021-07-20 IUD x5 (Other): No new information. 2021-08-10 IUD x4 (Other): No new information. 2021-Sep-08 IUDx2 (other): No new information. 2021-10-01 IUD (Other): No new information. KIT SVC MAS2310-04S TARGET EXTENDER | 2021-Jul-14 Salesforce C00230205 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the site lost the target extender and target extender screws for a set. No patient was present. 2021-07-20 IUD x5 (Other): No new information. 2021-08-10 IUD x4 (Other): No new information. 2021-Sep-08 IUDx2 (other): No new information. 2021-10-01 IUD (Other): No new information. | xse lost instruments salesforce outside procedure site lost target extender target extender screws set present iud new iud new new iud new kit svc target extender | NAV2053 | NAV3299 | NAV2054 | NAV2047 | NAV2058 | NAV2053NAV3299NAV2054NAV2047NAV2058 | True
3518 | 703366383 | X - Joint 6 Failure | 2019-09-17T17:46:15.000000+0000 | MazorX | NAV4098 | ARM JOINT SHIFTS | 20 | ASM0206-01 | SURGICAL ARM | X - Joint 6 Failure 2019-09-13 SCL1902780, Intake Form (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that after updating the 3Define camera and installing software, joint 6 of the surgical arm failed. The joint was grinding and resulted in joint shifts. There was no patient involved. SURGICAL ARM | 2019-09-13 SCL1902780, Intake Form (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that after updating the 3Define camera and installing software, joint 6 of the surgical arm failed. The joint was grinding and resulted in joint shifts. There was no patient involved. | x joint failure prior procedure updating camera installing software joint surgical arm failed joint grinding resulted joint shifts involved surgical arm | NAV3249 | NAV2054 | NAV4097 | NAV4098 | NAV4101 | NAV3249NAV2054NAV4097NAV4098NAV4101 | True
9917 | 704745309 | XSE - Shoulder Lock & X-Eye Issues | 2021-12-02T23:11:55.000000+0000 | MazorX Stealth Edition | NAV4109 | LOOSE | 50 | SCW0376-01 | X-EYE SCREW | XSE - Shoulder Lock & X-Eye Issues 2021-Dec-02 Salesforce C00239087 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative had difficulty getting the unit to recognize that the shoulder was locked on the surgical arm setup procedure. He moved the shoulder several times before he was able to get the system to recognize the shoulder was locked. The shoulder did physically lock every time they pressed the shoulder lock. It was also noticed that the X-eye camera was not rigidly attached to the post. No patient was present. 2021-12-14 IUD (Other): No new information. 2021-12-21 IUD (Other): No new information. 2022-01-06 IUD x2 (Other): No new information. 2022-02-09 IUD (Other): The X-eye bracket was completely off and the screw securing it was inside the positioner (the camera was dangling on arrival). The Firewire cable at the base of the positioner was unplugged. Also the cover on the bottom rear of the positioner is damaged and was off. The shoulder was overtightened resulting in the piston scraping on the side wall of the shoulder and throwing off the home position sensor. The front shoulder was assembled incorrectly. The cover was not flush and therefore the encoder that gives real time feedback of the shoulder position could not communicate. 2022-02-21 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. X-EYE SCREW | 2021-Dec-02 Salesforce C00239087 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the manufacturer representative had difficulty getting the unit to recognize that the shoulder was locked on the surgical arm setup procedure. He moved the shoulder several times before he was able to get the system to recognize the shoulder was locked. The shoulder did physically lock every time they pressed the shoulder lock. It was also noticed that the X-eye camera was not rigidly attached to the post. No patient was present. 2021-12-14 IUD (Other): No new information. 2021-12-21 IUD (Other): No new information. 2022-01-06 IUD x2 (Other): No new information. 2022-02-09 IUD (Other): The X-eye bracket was completely off and the screw securing it was inside the positioner (the camera was dangling on arrival). The Firewire cable at the base of the positioner was unplugged. Also the cover on the bottom rear of the positioner is damaged and was off. The shoulder was overtightened resulting in the piston scraping on the side wall of the shoulder and throwing off the home position sensor. The front shoulder was assembled incorrectly. The cover was not flush and therefore the encoder that gives real time feedback of the shoulder position could not communicate. 2022-02-21 IUD (Other): No new information. 2022-03-03 IUD (Other): No new information. | xse shoulder lock xeye issues salesforce outside procedure difficulty getting unit recognize shoulder locked surgical arm setup procedure moved shoulder times able recognize shoulder locked shoulder did physically lock time pressed shoulder lock noticed xeye camera rigidly attached post present iud new iud new iud new iud xeye bracket completely screw securing inside positioner firewire cable base positioner unplugged cover rear positioner damaged shoulder overtightened resulting piston scraping wall shoulder throwing home position sensor shoulder assembled incorrectly cover flush encoder gives real time feedback shoulder position communicate iud new iud new xeye screw | NAV3254 | NAV4106 | NAV4107 | NAV4110 | NAV4103 | NAV3254NAV4106NAV4107NAV4110NAV4103 | False
3035 | 703263642 | X - Arm Issues | 2019-07-17T16:02:29.000000+0000 | MazorX | NAV3244 | UNABLE TO DEFINE WORK VOLUME | 20 | ASM0206-01 | SURGICAL ARM | X - Arm Issues Monitor separated during shipping causing touchscreen to be non-functional. The slides were also broken. SURGICAL ARM | Monitor separated during shipping causing touchscreen to be non-functional. The slides were also broken. | x arm issues monitor separated shipping causing touchscreen nonfunctional slides broken surgical arm | NAV3261 | NAV3249 | NAV2055 | NAV4101 | NAV3295 | NAV3261NAV3249NAV2055NAV4101NAV3295 | False
3876 | 703460851 | X - X-Eye Not Working | 2019-11-11T20:07:17.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | ELB0057-01 | MAZOR X FIRE WIRE CARD | X - X-Eye Not Working 2019-Nov-11 00678061 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Mazor X eye was unable to work during his procedure. The rep had tried to do a soft restart of the system but was unable to get the eye to work. The site moved on with out it. Less than 5 minute delay in the case. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-19 IUD x3 (Other): No new information. MAZOR X FIRE WIRE CARD | 2019-Nov-11 00678061 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Mazor X eye was unable to work during his procedure. The rep had tried to do a soft restart of the system but was unable to get the eye to work. The site moved on with out it. Less than 5 minute delay in the case. There was no patient harm or additional complications reported or anticipated as a result of the event. 2019-11-19 IUD x3 (Other): No new information. | x xeye working spinal procedure mazor x eye unable work procedure tried soft restart unable eye work site moved minute delay case additional complications anticipated iud new mazor x wire card | NAV4110 | NAV4112 | NAV1012 | NAV5000 | NAV3242 | NAV4110NAV4112NAV1012NAV5000NAV3242 | True
8506 | 704430631 | XSE - Shoulder Loose | 2021-06-03T23:02:09.000000+0000 | MazorX Stealth Edition | NAV3252 | SHOULDER ISSUES | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | XSE - Shoulder Loose 2021-Jun-03 Salesforce C00227793 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the shoulder was locking and unlocking, but there was no tightness. A reboot of the system was done and each joint was manipulated in the software, but the issue was not resolved. There was no patient involvement. 2021-Jun-29 IUDx3 (other): No new information. 2021-07-27 IUD x2 (Other): No new information. 2021-Aug-18 email 8/18x2 IUD (rep): No new information. UNK MAZOR X SHOULDER | 2021-Jun-03 Salesforce C00227793 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the shoulder was locking and unlocking, but there was no tightness. A reboot of the system was done and each joint was manipulated in the software, but the issue was not resolved. There was no patient involvement. 2021-Jun-29 IUDx3 (other): No new information. 2021-07-27 IUD x2 (Other): No new information. 2021-Aug-18 email 8/18x2 IUD (rep): No new information. | xse shoulder loose salesforce outside procedure shoulder locking unlocking tightness reboot joint manipulated software issue resolved new iud new email iud new unk mazor x shoulder | NAV4107 | NAV3252 | NAV2058 | NAV4108 | NAV3299 | NAV4107NAV3252NAV2058NAV4108NAV3299 | True
3517 | 703366376 | XSE - Inferior Deviation | 2019-09-17T17:32:57.000000+0000 | MazorX Stealth Edition | NAV5000 | MXSE - INACCURACY | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Inferior Deviation 2019-09-12 SCL1902779, Intake Form, CCR, image 2019-09-24 12:55 x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using fixation. It was reported that during a Scan & Plan case for an L1-L2 open revision and extension of construct with placement of L1 and L2, the L1 screws were placed inferior to plan by 5 mm. During the procedure, the patient was scanned with the O-arm and the case was planned and approved by the surgeon. A disposable single spinous clamp was used for a mount and placed at L1. The surgical system was then mounted and positioned. A 3Define scan was completed successfully and a snap shot was captured. Accuracy was verified with the passive planar probe on the surgical arm guide divot. The surgical arm was sent to AP scanning position with passive planar position marked. During registration, the star marker was readjusted by moving north towards the patient's head and closer to the patient's skin. The region of interest was localized and centered with the O-arm. The spin was successful and the star marker was recognized. L1 and L2 were planned and the trajectories were approved by the surgeon. The surgical arm was sent to L1 and L2. Trajectories were drilled and screws were placed using long tools. Confirmation shots with the C-arm were taken and both the right and left L1 screws were too inferior. The L2 screws were accurate. The surgeon aborted the use of the guidance system and redirected the screws using navigation. The surgical system was removed and a spine air frame was attached . A second spin was done to register navigation and the screw was redirected. There was no patient harm and the procedure was not delayed more than an hour. The procedure was completed. 2019-09-24 email 2019-09-24 16:15 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved bringing in navigation and redirecting the bilateral screws at L1. The cause of the deviation was high skive potentials. MAZOR X SYSTEM TPL0059 | 2019-09-12 SCL1902779, Intake Form, CCR, image 2019-09-24 12:55 x2 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs using fixation. It was reported that during a Scan & Plan case for an L1-L2 open revision and extension of construct with placement of L1 and L2, the L1 screws were placed inferior to plan by 5 mm. During the procedure, the patient was scanned with the O-arm and the case was planned and approved by the surgeon. A disposable single spinous clamp was used for a mount and placed at L1. The surgical system was then mounted and positioned. A 3Define scan was completed successfully and a snap shot was captured. Accuracy was verified with the passive planar probe on the surgical arm guide divot. The surgical arm was sent to AP scanning position with passive planar position marked. During registration, the star marker was readjusted by moving north towards the patient's head and closer to the patient's skin. The region of interest was localized and centered with the O-arm. The spin was successful and the star marker was recognized. L1 and L2 were planned and the trajectories were approved by the surgeon. The surgical arm was sent to L1 and L2. Trajectories were drilled and screws were placed using long tools. Confirmation shots with the C-arm were taken and both the right and left L1 screws were too inferior. The L2 screws were accurate. The surgeon aborted the use of the guidance system and redirected the screws using navigation. The surgical system was removed and a spine air frame was attached . A second spin was done to register navigation and the screw was redirected. There was no patient harm and the procedure was not delayed more than an hour. The procedure was completed. 2019-09-24 email 2019-09-24 16:15 (Rep): Additional information received from a manufacturer representative reported that troubleshooting involved bringing in navigation and redirecting the bilateral screws at L1. The cause of the deviation was high skive potentials. | xse inferior deviation ccr image spinal procedure treat degenerative discs using fixation scan plan case open revision extension construct placement screws placed inferior plan mm procedure scanned oarm case planned approved surgeon disposable single spinous clamp mount placed surgical mounted positioned scan completed successfully snap shot captured accuracy verified passive planar probe surgical arm guide divot surgical arm sent ap scanning position passive planar position marked registration star marker readjusted moving north patients head closer patients skin region localized centered oarm spin successful star marker recognized planned trajectories approved surgeon surgical arm sent trajectories drilled screws placed using long tools confirmation shots carm taken right left screws inferior screws accurate surgeon aborted use redirected screws using navigation surgical removed spine air frame attached second spin register navigation screw redirected procedure delayed hour procedure completed email additional troubleshooting involved bringing navigation redirecting bilateral screws cause deviation high skive potentials mazor x | NAV5000 | NAV3235 | NAV2056 | NAV4098 | NAV3243 | NAV5000NAV3235NAV2056NAV4098NAV3243 | True
3580 | 703380481 | X - Unable to Mount Surgical System | 2019-09-25T14:31:45.000000+0000 | MazorX | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | UNK_MAZORX_PSTNR | UNK MAZOR X POSITIONER | X - Unable to Mount Surgical System 2019-09-24 SCL1902891, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was canceled due to not being able to mount the surgical system to the bed. Only one tooth of the positioner could be mounted to the bed frame adaptor. The surgical system had to be removed by taking off the bed frame adaptor with the system still attached. The patient had been placed under anesthesia. There was no patient harm. 2019-09-25 email 2019-09-25 11:15 (Rep): Additional information received from a manufacturer representative reported they tried to do a soft restart at first, but the locking mechanism was frozen and would not move. The representative called a field service engineer before proceeding with troubleshooting. They advised the representative to override the panel to force the teeth open, but that was unsuccessful after several attempts. The representative then tried to force the teeth closed, but that was unsuccessful after several attempts. A hard restart of the system was done, but the system was still frozen. The surgical system had to be removed with the bed frame adapter still attached. At this point the surgeon decided to abort the use of the guidance system and proceed. The system was removed from the operating table and the restarted. The bed frame adapter was able to be removed. After the restart, the system was able to be mounted multiple times with no issues. UNK MAZOR X POSITIONER | 2019-09-24 SCL1902891, Intake Form (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the case was canceled due to not being able to mount the surgical system to the bed. Only one tooth of the positioner could be mounted to the bed frame adaptor. The surgical system had to be removed by taking off the bed frame adaptor with the system still attached. The patient had been placed under anesthesia. There was no patient harm. 2019-09-25 email 2019-09-25 11:15 (Rep): Additional information received from a manufacturer representative reported they tried to do a soft restart at first, but the locking mechanism was frozen and would not move. The representative called a field service engineer before proceeding with troubleshooting. They advised the representative to override the panel to force the teeth open, but that was unsuccessful after several attempts. The representative then tried to force the teeth closed, but that was unsuccessful after several attempts. A hard restart of the system was done, but the system was still frozen. The surgical system had to be removed with the bed frame adapter still attached. At this point the surgeon decided to abort the use of the guidance system and proceed. The system was removed from the operating table and the restarted. The bed frame adapter was able to be removed. After the restart, the system was able to be mounted multiple times with no issues. | x unable mount surgical spinal procedure case canceled able mount surgical bed tooth positioner mounted bed frame adaptor surgical removed taking bed frame adaptor attached placed anesthesia email additional tried soft restart locking mechanism frozen called field service engineer proceeding troubleshooting advised override panel force teeth open unsuccessful attempts tried force teeth closed unsuccessful attempts hard restart frozen surgical removed bed frame adapter attached point surgeon decided abort use proceed removed operating table restarted bed frame adapter able removed restart able mounted multiple times issues unk mazor x positioner | NAV3250 | NAV3255 | NAV4103 | NAV2058 | NAV4107 | NAV3250NAV3255NAV4103NAV2058NAV4107 | True
9707 | 704702285 | XSE - Multiple System Issues | 2021-11-06T06:31:27.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE - Multiple System Issues 2021-Nov-05 Salesforce C00237478, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgical system would not attach to the bed frame adapter, the surgical arm moved freely when locked, the system monitor screws were loose, the surgical system teeth were engaged when not mounted to the bed and the hood was opening and closing slowly. There was no patient involvement. 2021-11-15 IUD (Other): No new information. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2021-Nov-05 Salesforce C00237478, IUD (Rep): Medtronic received information regarding a guidance system being used outside of a proc edure. It was reported that the surgical system would not attach to the bed frame adapter, the surgical arm moved freely when locked, the system monitor screws were loose, the surgical system teeth were engaged when not mounted to the bed and the hood was opening and closing slowly. There was no patient involvement. 2021-11-15 IUD (Other): No new information. | xse multiple issues salesforce iud outside proc edure surgical attach bed frame adapter surgical arm moved freely locked monitor screws loose surgical teeth engaged mounted bed hood opening closing slowly iud new kit svc surgical new | NAV4103 | NAV3250 | NAV4101 | NAV3264 | NAV3261 | NAV4103NAV3250NAV4101NAV3264NAV3261 | True
4470 | 703545559 | x- draw spine did not work | 2020-01-06T16:00:46.000000+0000 | MazorX | NAV4110 | NO TRACKING | 20 | UNK_MAZORX_XCAM | UNK MAZOR X X-EYE CAMERA | x- draw spine did not work 2018-4-13 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that draw spine did not work. There was no patient harm and the procedure was not delayed over an hour. UNK MAZOR X X-EYE CAMERA | 2018-4-13 CCR(Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that draw spine did not work. There was no patient harm and the procedure was not delayed over an hour. | x draw spine did work ccr spinal procedure draw spine did work procedure delayed hour unk mazor x xeye camera | NAV4110 | NAV3244 | NAV4107 | NAV3254 | NAV4095 | NAV4110NAV3244NAV4107NAV3254NAV4095 | True
4874 | 703583172 | XSE - BMB Stripped Screws | 2020-01-27T20:08:34.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 40 | MEC5196-01S | KIT SVC MEC5196-01S RNX-ACC-063 | XSE - BMB Stripped Screws 2020-01-27 00682825 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that the screws on the surgical systems bone mount bridge were stripped and they were looking to get replacements. The site was able to get the screws off the system. No patient was present and there was no delay of surgery. 2020-01-28 email 2020-01-28 1034 (Rep): Additional information received from the rep reporting that it is the opinion of those in the field that the heads of the screws strip out because they are used so frequently and related to the shape of the rounded hex screw driver used with them. This was a repeated issue. 2020-02-11 IUD (Other): No new information. 2020-02-14 Email (Rep): No new information. 2020-02-27 IUD (Other): No new information. 2020-02-27 IUD x4 (Other): No new information. 2020-04-08 IUD (Other): No new information. KIT SVC MEC5196-01S RNX-ACC-063 | 2020-01-27 00682825 (REP, HCP): Medtronic received information regarding a guidance system during use. It was reported that the screws on the surgical systems bone mount bridge were stripped and they were looking to get replacements. The site was able to get the screws off the system. No patient was present and there was no delay of surgery. 2020-01-28 email 2020-01-28 1034 (Rep): Additional information received from the rep reporting that it is the opinion of those in the field that the heads of the screws strip out because they are used so frequently and related to the shape of the rounded hex screw driver used with them. This was a repeated issue. 2020-02-11 IUD (Other): No new information. 2020-02-14 Email (Rep): No new information. 2020-02-27 IUD (Other): No new information. 2020-02-27 IUD x4 (Other): No new information. 2020-04-08 IUD (Other): No new information. | xse bmb stripped screws use screws surgical systems bone mount bridge stripped looking replacements site able screws present delay surgery email additional reporting opinion field heads screws strip frequently related shape rounded hex screw driver repeated issue iud new email new iud new iud new iud new kit svc | NAV3299 | NAV2054 | NAV2047 | NAV4090 | NAV3295 | NAV3299NAV2054NAV2047NAV4090NAV3295 | True
6081 | 703866657 | XSE ¦ Deviation | 2020-07-23T03:53:05.000000+0000 | MazorX Stealth Edition | NAV4091 | DEVIATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE ¦ Deviation 2020-04-08 CCR, Attachment (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that at L4 left free hand was tried after tapping and screw stimmed low at 5. The surgeon removed, decompressed and then L4 right-stimmed at 7. The screw was advanced too far by accident. The surgeon went and re-planned and resent screw. The screw then stimmed at 12. All planned screws were then executed accurately. There was no known impact on the patient outcome. MAZOR X SYSTEM TPL0059 | 2020-04-08 CCR, Attachment (Rep, HCP): Medtronic received information regarding a guidance system being used intra-operatively of a spinal procedure. It was reported that at L4 left free hand was tried after tapping and screw stimmed low at 5. The surgeon removed, decompressed and then L4 right-stimmed at 7. The screw was advanced too far by accident. The surgeon went and re-planned and resent screw. The screw then stimmed at 12. All planned screws were then executed accurately. There was no known impact on the patient outcome. | xse deviation ccr attachment intraoperatively spinal procedure left free hand tried tapping screw stimmed low surgeon removed decompressed rightstimmed screw advanced far accident surgeon went replanned resent screw screw stimmed planned screws executed accurately known impact outcome mazor x | NAV5000 | NAV3242 | NAV3243 | NAV4112 | NAV3244 | NAV5000NAV3242NAV3243NAV4112NAV3244 | False
3950 | 703472893 | X - E-Stop Button Pulled From SRA | 2019-11-18T16:11:08.000000+0000 | MazorX | NAV3249 | ARM ISSUES | 20 | ASM0206-01 | SURGICAL ARM | X - E-Stop Button Pulled From SRA 2019-11-18 00678552 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the emergency stop button was pulled out of the arm when the drape was being removed, disconnecting the emergency stop button. There was no patient harm and the procedure was not delayed over an hour. The manufacturer representative noted that the system is not currently functioning until a FSE can review and service it. 2019-11-18 IUD x7, e1, e2, e3 (Rep): No new information. 2020-02-26 IUD (Other): No new information. 2020-02-27 IUD (Other): No new information. SURGICAL ARM | 2019-11-18 00678552 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that the emergency stop button was pulled out of the arm when the drape was being removed, disconnecting the emergency stop button. There was no patient harm and the procedure was not delayed over an hour. The manufacturer representative noted that the system is not currently functioning until a FSE can review and service it. 2019-11-18 IUD x7, e1, e2, e3 (Rep): No new information. 2020-02-26 IUD (Other): No new information. 2020-02-27 IUD (Other): No new information. | x estop button pulled sra spinal procedure emergency stop button pulled arm drape removed disconnecting emergency stop button procedure delayed hour noted currently functioning fse review service iud new iud new iud new surgical arm | NAV3249 | NAV1006 | NAV4101 | NAV3246 | NAV4112 | NAV3249NAV1006NAV4101NAV3246NAV4112 | True
552 | 703150330 | XSE - shoulder hissing sound/blue screen | 2019-05-13T13:50:23.000000+0000 | MazorX Stealth Edition | NAV3246 | DEFECTS | 20 | KIT1048-02 | MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | XSE - shoulder hissing sound/blue screen 2019-May-08 SCL1901136 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that on the day of the report the screen went blue with the robot was plugged into the normal outlet and the screen was backlit but black and the dell logo did not come up or the software. They waited 5 to 7 minutes, turned it off, changed the outlet, and everything worked fine. They unlocked the shoulder with the shoulder button and heard an audible hissing noise coming for the shoulder. The system did indicate it was losing pressure. They had to perform a hard restart to get the system working properly. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. MAZOR X SOFTWARE INSTALLATION KIT VER4.2 | 2019-May-08 SCL1901136 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that on the day of the report the screen went blue with the robot was plugged into the normal outlet and the screen was backlit but black and the dell logo did not come up or the software. They waited 5 to 7 minutes, turned it off, changed the outlet, and everything worked fine. They unlocked the shoulder with the shoulder button and heard an audible hissing noise coming for the shoulder. The system did indicate it was losing pressure. They had to perform a hard restart to get the system working properly. No patient harm was reported. There were no additional complications reported or anticipated as a result of the event. | xse shoulder hissing soundblue screen spinal procedure day report screen went blue robot plugged normal outlet screen backlit black dell logo did come software waited minutes turned changed outlet worked fine unlocked shoulder shoulder button heard audible hissing noise coming shoulder did indicate losing pressure perform hard restart working properly additional complications anticipated mazor x software installation kit | NAV3252 | NAV3246 | NAV3247 | NAV4107 | NAV5000 | NAV3252NAV3246NAV3247NAV4107NAV5000 | True
1754 | 703213391 | XSE - drapes ripped | 2019-06-18T18:00:54.000000+0000 | MazorX Stealth Edition | NAV3232 | DRAPE ISSUES | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | XSE - drapes ripped 2019-Jun-13 SCL1901764 intake form, msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgical drape was tearing around the joints of the arm when sent to snapshot position. It was noted that it could be a result of the drape being rotated prior to the tech clipping in the 3define lense. The issue was occurring with both positioner types. The issue was being seen when the robotic arm moved to a heavy bend position such as clear right, snapshot, and 3define angles. In the technique guide it did not specify to ensure a specific position of how to align the 3align camera lens on step 5.8. The rep believe the issue was arising by the drape being rotated prior to the lens being clipped causing not enough slack for these elbow angles. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. UNK MAZOR X ACCESSORIES | 2019-Jun-13 SCL1901764 intake form, msg (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used in a spinal procedure. It was reported that the Surgical drape was tearing around the joints of the arm when sent to snapshot position. It was noted that it could be a result of the drape being rotated prior to the tech clipping in the 3define lense. The issue was occurring with both positioner types. The issue was being seen when the robotic arm moved to a heavy bend position such as clear right, snapshot, and 3define angles. In the technique guide it did not specify to ensure a specific position of how to align the 3align camera lens on step 5.8. The rep believe the issue was arising by the drape being rotated prior to the lens being clipped causing not enough slack for these elbow angles. There was no patient harm. There were no additional complications reported or anticipated as a result of the event. | xse drapes ripped msg spinal procedure surgical drape tearing joints arm sent snapshot position noted drape rotated prior tech clipping lense issue occurring positioner types issue seen robotic arm moved heavy bend position clear right snapshot angles technique guide did specify ensure specific position align camera lens step believe issue arising drape rotated prior lens clipped causing slack elbow angles additional complications anticipated unk mazor x accessories | NAV3232 | NAV2054 | NAV2044 | NAV3299 | NAV3244 | NAV3232NAV2054NAV2044NAV3299NAV3244 | True
7002 | 704098147 | XSE - SRA Failures/OOBF | 2020-12-01T18:58:50.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 30 | ASM0206-01R | KIT SVC ASM0206-01R SURGICAL ARM REFURB | XSE - SRA Failures/OOBF 2020-Dec-01 00701777 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the system is failing the BIST test and needs replaced. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-03 IUD (Rep): Additional information received from a manufacturer representative reported that the replacement surgical arm failed at joint 2. 2020-12-07 IUD x3, e1 (Rep): Additional information received from a manufacturer representative reported the replacement surgical arm failed again. An error 4059 arm shift detected in joints 2 and 3 was displayed. 2020-12-21 IUD x5 (Other): No new information. 2021-01-17 IUD x5, e2 (Rep): No new information. 2021-02-23 IUD x14 (Other): No new information. 2021-03-01 e3, e4 (Other): No new information. 2021-03-04 e5 (Rep): Additional information received from a manufacturer representative reported that all of the surgical arm failed the BIST during upgrade of the software. 2021-03-08 IUD x2, e6 (Other): No new information. 2021-04-06 IUD (Other): No new information. 2021-04-29 IUD (Other): No new information. 2021-06-15 IUD x3 (Other): No new information. 2021-11-29 IUD (Other): No new information. KIT SVC ASM0206-01R SURGICAL ARM REFURB | 2020-Dec-01 00701777 (HCP,rep): Medtronic received information from a healthcare professional (HCP) via a manufacturer representative (rep) regarding a guidance device being used outside of a procedure. It was reported that the system is failing the BIST test and needs replaced. There was no patient harm or additional complications reported or anticipated as a result of the event. 2020-12-03 IUD (Rep): Additional information received from a manufacturer representative reported that the replacement surgical arm failed at joint 2. 2020-12-07 IUD x3, e1 (Rep): Additional information received from a manufacturer representative reported the replacement surgical arm failed again. An error 4059 arm shift detected in joints 2 and 3 was displayed. 2020-12-21 IUD x5 (Other): No new information. 2021-01-17 IUD x5, e2 (Rep): No new information. 2021-02-23 IUD x14 (Other): No new information. 2021-03-01 e3, e4 (Other): No new information. 2021-03-04 e5 (Rep): Additional information received from a manufacturer representative reported that all of the surgical arm failed the BIST during upgrade of the software. 2021-03-08 IUD x2, e6 (Other): No new information. 2021-04-06 IUD (Other): No new information. 2021-04-29 IUD (Other): No new information. 2021-06-15 IUD x3 (Other): No new information. 2021-11-29 IUD (Other): No new information. | xse sra failuresoobf outside procedure failing bist test needs replaced additional complications anticipated iud additional replacement surgical arm failed joint iud additional replacement surgical arm failed error arm shift detected joints displayed iud new iud new iud new new additional surgical arm failed bist upgrade software iud new iud new iud new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV2054 | NAV1006 | NAV4098 | NAV4092 | NAV3249NAV2054NAV1006NAV4098NAV4092 | True
109 | 702999770 | XSE - Shoulder Encoder Issue | 2019-02-10T07:17:47.000000+0000 | MazorX Stealth Edition | NAV4106 | CALIBRATION ISSUES | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | XSE - Shoulder Encoder Issue 2019-02-05 SCL1900315, SCL1900317, Image 2019-02-11 05:39 (Rep): Medtronic received information regarding a guidance system. It was reported that the shoulder encoder was not seated properly during installation and the screws were found to be stripped when trying to reseat the encoder. There was no patient involvement. 2019-02-10 e1 (Rep): Additional information received from a manufacturer representative reported the shoulder encoder issue was identified during the shoulder calibration portion of the system setup. The encoder was outputting degrees at the extremes of the range of motion for the front should joint that kept shifting with every sweep. The issue was resolved by disassembling the encoder plate from the shoulder, disassembling the encoder from the plate and reseating it flush. The encoder was re-calibrated and the values were verified to be consistent. UNK MAZOR X SHOULDER | 2019-02-05 SCL1900315, SCL1900317, Image 2019-02-11 05:39 (Rep): Medtronic received information regarding a guidance system. It was reported that the shoulder encoder was not seated properly during installation and the screws were found to be stripped when trying to reseat the encoder. There was no patient involvement. 2019-02-10 e1 (Rep): Additional information received from a manufacturer representative reported the shoulder encoder issue was identified during the shoulder calibration portion of the system setup. The encoder was outputting degrees at the extremes of the range of motion for the front should joint that kept shifting with every sweep. The issue was resolved by disassembling the encoder plate from the shoulder, disassembling the encoder from the plate and reseating it flush. The encoder was re-calibrated and the values were verified to be consistent. | xse shoulder encoder issue image shoulder encoder seated properly installation screws stripped trying reseat encoder additional shoulder encoder issue identified shoulder calibration portion setup encoder outputting degrees extremes range motion joint kept shifting sweep issue resolved disassembling encoder plate shoulder disassembling encoder plate reseating flush encoder recalibrated values verified consistent unk mazor x shoulder | NAV2054 | NAV4106 | NAV4100 | NAV3252 | NAV2058 | NAV2054NAV4106NAV4100NAV3252NAV2058 | True
3495 | 703361490 | XSE - RRF Loose/Drape Residue in Hole | 2019-09-15T20:41:37.000000+0000 | MazorX Stealth Edition | NAV3252 | SHOULDER ISSUES | 20 | UNK_MAZORX_SHLDER | UNK MAZOR X SHOULDER | XSE - RRF Loose/Drape Residue in Hole 2019-09-11 email 2019-09-15 15:53 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that when connecting the robotic reference frame, the connecting screw felt like it was tightened fully, but did not seem to go fully into the hole. This allowed the reference frame to have some freedom to move. The representative first thought the metal threaded insert was loose, but when checking further they saw drape residues in the holes from previous procedures, which limited the ability of the screw to fully tighten. The representative thought that this happening in the filed during cases could lead to mismatch and variances. There was no patient involvement. 2019-09-16 email 2019-10-08 11:53 (Rep): No new information. UNK MAZOR X SHOULDER | 2019-09-11 email 2019-09-15 15:53 (Rep): Medtronic received information regarding a guidance system being used prior to a procedure. It was reported that when connecting the robotic reference frame, the connecting screw felt like it was tightened fully, but did not seem to go fully into the hole. This allowed the reference frame to have some freedom to move. The representative first thought the metal threaded insert was loose, but when checking further they saw drape residues in the holes from previous procedures, which limited the ability of the screw to fully tighten. The representative thought that this happening in the filed during cases could lead to mismatch and variances. There was no patient involvement. 2019-09-16 email 2019-10-08 11:53 (Rep): No new information. | xse rrf loosedrape residue hole email prior procedure connecting robotic reference frame connecting screw felt like tightened fully did fully hole allowed reference frame freedom thought metal threaded insert loose checking saw drape residues holes previous procedures limited ability screw fully tighten thought happening filed cases lead mismatch variances email new unk mazor x shoulder | NAV2054 | NAV4100 | NAV3299 | NAV3251 | NAV4108 | NAV2054NAV4100NAV3299NAV3251NAV4108 | False
2334 | 703233263 | X - Accessory issue | 2019-06-28T14:33:47.000000+0000 | MazorX | NAV3295 | BROKEN ACCESSORY | 20 | UNK_MAZORX_ACCY | UNK MAZOR X ACCESSORIES | X - Accessory issue It was reported that the dual clamp broke during DeVito case. UNK MAZOR X ACCESSORIES | It was reported that the dual clamp broke during DeVito case. | x accessory issue dual clamp broke devito case unk mazor x accessories | NAV3295 | NAV3299 | NAV4090 | NAV2044 | NAV3233 | NAV3295NAV3299NAV4090NAV2044NAV3233 | True
8660 | 704458250 | XSE - Damaged Manipulator Switch | 2021-06-18T23:52:22.000000+0000 | MazorX Stealth Edition | NAV3260 | MANIPULATOR ISSUES | 30 | KIT0928-01 | SVC KIT0928-01 MANIPULATOR SENSOR | XSE - Damaged Manipulator Switch 2021-Jun-18 Salesforce C00228777 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the switch around the manipulator was damaged. There was no patient involvement. 2021-06-21 IUD x5 (Other): No new information. 2021-07-14 IUD x2 (Other): No new information. 2021-08-04 IUD x2 (Other): No new information. SVC KIT0928-01 MANIPULATOR SENSOR | 2021-Jun-18 Salesforce C00228777 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the switch around the manipulator was damaged. There was no patient involvement. 2021-06-21 IUD x5 (Other): No new information. 2021-07-14 IUD x2 (Other): No new information. 2021-08-04 IUD x2 (Other): No new information. | xse damaged manipulator switch salesforce outside procedure switch manipulator damaged iud new iud new iud new svc manipulator sensor | NAV3260 | NAV2057 | NAV3249 | NAV3264 | NAV3299 | NAV3260NAV2057NAV3249NAV3264NAV3299 | True
5186 | 703613916 | XSE - Arm failed to Reach Destination | 2020-02-13T20:05:04.000000+0000 | MazorX Stealth Edition | NAV4095 | ARM DIDN'T REACH LOCATION | 30 | ASM0206-01S | KIT SVC ASM0206-01S SURGICAL ARM | XSE - Arm failed to Reach Destination 2020-02-13 00684070 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when sending the arm to trajectory, errors failed to reach destination and failed under max closing tolerance were received. The representative attempted a soft restart and a hard restart and also tried to reset movements with no success. The representative was able to manually move the arm at joint 5 and send the arm to clear up. Outside of the room, the representative attempted to send the arm to predefined locations and again it failed. There was no patient harm and the procedure was delayed less than an hour. 2020-02-14 IUD x5, Email 2020-02-20 21:45 (Rep): Additional information was received from the manufacturer's representative. The field service engineer (FSE) reported that he could not move joints 4,5, and 6. The FSE ran field motor calibrations and power cycled the system. He was able to move joints 4 and 5 but joint 6 is still not moving. 2020-03-19 IUD (Other): No new information. 2020-05-19 IUD (Other): No new information. KIT SVC ASM0206-01S SURGICAL ARM | 2020-02-13 00684070 (REP, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure. It was reported that when sending the arm to trajectory, errors failed to reach destination and failed under max closing tolerance were received. The representative attempted a soft restart and a hard restart and also tried to reset movements with no success. The representative was able to manually move the arm at joint 5 and send the arm to clear up. Outside of the room, the representative attempted to send the arm to predefined locations and again it failed. There was no patient harm and the procedure was delayed less than an hour. 2020-02-14 IUD x5, Email 2020-02-20 21:45 (Rep): Additional information was received from the manufacturer's representative. The field service engineer (FSE) reported that he could not move joints 4,5, and 6. The FSE ran field motor calibrations and power cycled the system. He was able to move joints 4 and 5 but joint 6 is still not moving. 2020-03-19 IUD (Other): No new information. 2020-05-19 IUD (Other): No new information. | xse arm failed reach destination spinal procedure sending arm trajectory errors failed reach destination failed max closing tolerance attempted soft restart hard restart tried reset movements success able manually arm joint send arm clear outside room attempted send arm predefined locations failed procedure delayed hour iud email additional manufacturers field service engineer joints fse ran field motor calibrations power cycled able joints joint moving iud new iud new kit svc surgical arm | NAV4095 | NAV3264 | NAV3299 | NAV4092 | NAV1006 | NAV4095NAV3264NAV3299NAV4092NAV1006 | True
9532 | 704665143 | XSE - Damaged Screws | 2021-10-15T14:16:27.000000+0000 | MazorX Stealth Edition | NAV2058 | WORKSTATION ISSUES | 30 | MEC5142-01S | KIT SVC MEC5142-01S CENTRAL CURTAIN COV | XSE - Damaged Screws 2021-Oct-15 Salesforce C00236319 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the cover and positioner have two screws that are damaged. No patient was present. 2021-10-15 IUD x4 (Other): No new information. KIT SVC MEC5142-01S CENTRAL CURTAIN COV | 2021-Oct-15 Salesforce C00236319 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that the cover and positioner have two screws that are damaged. No patient was present. 2021-10-15 IUD x4 (Other): No new information. | xse damaged screws salesforce outside procedure cover positioner screws damaged present iud new kit svc central curtain cov | NAV3251 | NAV3299 | NAV3249 | NAV2058 | NAV3264 | NAV3251NAV3299NAV3249NAV2058NAV3264 | True
4488 | 703546261 | XSE - Setup Screw Driver Broken | 2020-01-06T20:00:09.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 30 | MAS1650-01 | UNIVERSAL SCREWDRIVER | XSE - Setup Screw Driver Broken 2020-01-06 00681425 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the green setup screw driver broke in half. No patient was present. 2020-01-02 email 2020-01-07 14:18, IUD x2 (Rep): No new information. UNIVERSAL SCREWDRIVER | 2020-01-06 00681425 (REP): Medtronic received information regarding a guidance system outside of procedure. It was reported that the green setup screw driver broke in half. No patient was present. 2020-01-02 email 2020-01-07 14:18, IUD x2 (Rep): No new information. | xse setup screw driver broken outside procedure green setup screw driver broke half present email iud new universal screwdriver | NAV3295 | NAV3299 | NAV4090 | NAV3249 | NAV3246 | NAV3295NAV3299NAV4090NAV3249NAV3246 | True
7072 | 704114278 | XSE - SRA Failed BIST Testing | 2020-12-09T15:03:55.000000+0000 | MazorX Stealth Edition | NAV3249 | ARM ISSUES | 40 | ASM0206 | KIT SVC ASM0206 SURGICAL ARM REFURB | XSE - SRA Failed BIST Testing 2020-12-09 00702348 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while upgrading the software for the surgical system, the surgical arm failed the built in self test. There was no patient involvement. 2020-12-15 IUD x7 (Other): No new information. 2021-01-17 IUD x3 (Other): No new information. 2021-02-14 IUD x2, 302351062 (Other): No new information. 2021-03-02 IUD, 302351888 (Other): No new information. 2021-03-03 302353608 (Other): No new information. 2021-03-03 IUD x5 (Other): No new information. 2021-Mar-15 IUD (other): No new information. 2021-03-21 Attachment 2021-03-21 07:31 (Other): No new information. 2021-04-12 IUD x3 (Other): No new information. 2021-06-15 IUD x5 (Other): No new information. KIT SVC ASM0206 SURGICAL ARM REFURB | 2020-12-09 00702348 (Rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that while upgrading the software for the surgical system, the surgical arm failed the built in self test. There was no patient involvement. 2020-12-15 IUD x7 (Other): No new information. 2021-01-17 IUD x3 (Other): No new information. 2021-02-14 IUD x2, 302351062 (Other): No new information. 2021-03-02 IUD, 302351888 (Other): No new information. 2021-03-03 302353608 (Other): No new information. 2021-03-03 IUD x5 (Other): No new information. 2021-Mar-15 IUD (other): No new information. 2021-03-21 Attachment 2021-03-21 07:31 (Other): No new information. 2021-04-12 IUD x3 (Other): No new information. 2021-06-15 IUD x5 (Other): No new information. | xse sra failed bist testing outside procedure upgrading software surgical surgical arm failed built self test iud new iud new iud new iud new new iud new iud new attachment new iud new iud new kit svc surgical arm refurb | NAV3249 | NAV2054 | NAV4095 | NAV1006 | NAV4093 | NAV3249NAV2054NAV4095NAV1006NAV4093 | True
7317 | 704174471 | XSE - Damaged Schanz Arm | 2021-01-14T21:29:17.000000+0000 | MazorX Stealth Edition | NAV3299 | UNUSABLE ACCESSORIES | 20 | MAS2351-01 | SCHANZ ARM | XSE - Damaged Schanz Arm 2021-01-14 00704304 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the site's schanz arm was noted to be damaged. The screw that holds the clamp pin in place is stuck in the the engaged position and will not allow the pin to be inserted. No patient was present. 2021-01-26 IUD x2 (Other): No new information. 2021-02-18 email 2021-02-18 12:45 (Rep): Additional information received from a manufacturer representative reported that the cause was frequent use. The pin got stuck and would not fit due to the that. 2021-03-17 IUD x2 (Other): No new information. 2021-03-31 email 2021-03-31 15:16 x2, image 2021-03-31 15:16 (Rep): No new information. SCHANZ ARM | 2021-01-14 00704304 (Rep): Medtronic received information regarding a guidance system accessory, outside of procedure. It was reported that the site's schanz arm was noted to be damaged. The screw that holds the clamp pin in place is stuck in the the engaged position and will not allow the pin to be inserted. No patient was present. 2021-01-26 IUD x2 (Other): No new information. 2021-02-18 email 2021-02-18 12:45 (Rep): Additional information received from a manufacturer representative reported that the cause was frequent use. The pin got stuck and would not fit due to the that. 2021-03-17 IUD x2 (Other): No new information. 2021-03-31 email 2021-03-31 15:16 x2, image 2021-03-31 15:16 (Rep): No new information. | xse damaged schanz arm accessory outside procedure sites schanz arm noted damaged screw holds clamp pin place stuck engaged position allow pin inserted present iud new email additional cause frequent use pin got stuck fit iud new email image new schanz arm | NAV3299 | NAV2054 | NAV3295 | NAV4090 | NAV3244 | NAV3299NAV2054NAV3295NAV4090NAV3244 | True
9994 | 704766993 | XSE ¦ System Beeping when attaching arm | 2021-12-15T20:29:06.000000+0000 | MazorX Stealth Edition | NAV3250 | MOUNTING/DISMOUNTING ISSUES | 20 | ASM0214 | KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | XSE ¦ System Beeping when attaching arm 2021-Dec-15 Salesforce C00239768 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that prior to a case, after removing the surgical arm to reposition the bed frame adapter, the system was beeping when the manufacturer representative tried to mount the arm again. She stated that when trying to mount the arm the system would give an error beep but no error would appear in the software. She was able to successfully mount the arm. She stated that she could hear the pneumatics kick on when moving the arm. There was no patient involved. 2021-12-20 IUD x3 (Other): No new information. KIT SVC ASM0214 SURGICAL SYSTEM 03S NEW | 2021-Dec-15 Salesforce C00239768 (rep): Medtronic received information regarding a guidance system being used outside of a procedure. It was reported that prior to a case, after removing the surgical arm to reposition the bed frame adapter, the system was beeping when the manufacturer representative tried to mount the arm again. She stated that when trying to mount the arm the system would give an error beep but no error would appear in the software. She was able to successfully mount the arm. She stated that she could hear the pneumatics kick on when moving the arm. There was no patient involved. 2021-12-20 IUD x3 (Other): No new information. | xse beeping attaching arm salesforce outside procedure prior case removing surgical arm reposition bed frame adapter beeping tried mount arm stated trying mount arm error beep error appear software able successfully mount arm stated hear pneumatics kick moving arm involved iud new kit svc surgical new | NAV3249 | NAV4095 | NAV2054 | NAV3246 | NAV4099 | NAV3249NAV4095NAV2054NAV3246NAV4099 | False
7631 | 704239710 | XSE - Unable To Register | 2021-02-19T21:04:42.000000+0000 | MazorX Stealth Edition | NAV4112 | NO REGISTRATION | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Unable To Register 2021-Feb-19 Salesforce C00221407 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that registration could not be completed during the case. The manufacturer representative tried taking new images, recalibrating the C-arm, and trying both automatic and semi-automatic registration methods. The procedure had to be rescheduled at a different facility. There was no patient harm and the procedure was delayed over an hour. 2021-02-26 IUD x3, email 2021-02-28 17:45 (Rep): Additional information received from a manufacturer representative reported that the cause of the registration issue was poor image quality and the patient having a BMI of 48. The representative noted that the operating room only had a Ziem solo 9 inch C-arm. The case was reschedule for the main operating room since they have an O-arm that could be used for Scan & Plan if needed. The schedule was completed with a different C-arm with no registration issues at all. MAZOR X SYSTEM TPL0059 | 2021-Feb-19 Salesforce C00221407 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal p rocedure. It was reported that registration could not be completed during the case. The manufacturer representative tried taking new images, recalibrating the C-arm, and trying both automatic and semi-automatic registration methods. The procedure had to be rescheduled at a different facility. There was no patient harm and the procedure was delayed over an hour. 2021-02-26 IUD x3, email 2021-02-28 17:45 (Rep): Additional information received from a manufacturer representative reported that the cause of the registration issue was poor image quality and the patient having a BMI of 48. The representative noted that the operating room only had a Ziem solo 9 inch C-arm. The case was reschedule for the main operating room since they have an O-arm that could be used for Scan & Plan if needed. The schedule was completed with a different C-arm with no registration issues at all. | xse unable register salesforce spinal p rocedure registration completed case tried taking new images recalibrating carm trying automatic semiautomatic registration methods procedure rescheduled different facility procedure delayed hour iud email additional cause registration issue poor image quality having bmi noted operating room ziem solo inch carm case reschedule main operating room oarm scan plan needed schedule completed different carm registration issues mazor x | NAV4112 | NAV3242 | NAV3237 | NAV2058 | NAV3264 | NAV4112NAV3242NAV3237NAV2058NAV3264 | True
7747 | 704263742 | XSE - Unable To Place Screw | 2021-03-04T21:12:49.000000+0000 | MazorX Stealth Edition | NAV3243 | TRAJECTORY GUIDANCE ISSUES | 10 | TPL0059 | MAZOR X SYSTEM TPL0059 | XSE - Unable To Place Screw 2021-Mar-04 Salesforce C00222114, CCR, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with screws and rod placement.. It was reported that the site was able to complete 3 out of the 4 screw placements during the procedure. The rep noted that the right side of the patient was completed as intended how ever the left side they were able to complete one of the screws before noting that the second screw was showing skived potential. The site had tried to re plan the screw placement and send the system to the correct position but were still unable to place the screw due to the patient's anatomy. No noted issue was with the system. Patient was present. 20 min delay noted in the case. There was no patient harm. MAZOR X SYSTEM TPL0059 | 2021-Mar-04 Salesforce C00222114, CCR, IUD x3 (Rep, HCP): Medtronic received information regarding a guidance system being used during a spinal procedure to treat degenerative discs with screws and rod placement.. It was reported that the site was able to complete 3 out of the 4 screw placements during the procedure. The rep noted that the right side of the patient was completed as intended how ever the left side they were able to complete one of the screws before noting that the second screw was showing skived potential. The site had tried to re plan the screw placement and send the system to the correct position but were still unable to place the screw due to the patient's anatomy. No noted issue was with the system. Patient was present. 20 min delay noted in the case. There was no patient harm. | xse unable place screw salesforce ccr iud spinal procedure treat degenerative discs screws rod placement site able complete screw placements procedure noted right completed intended left able complete screws noting second screw showing skived potential site tried plan screw placement send correct position unable place screw patients anatomy noted issue present min delay noted case mazor x | NAV4112 | NAV3244 | NAV5000 | NAV3243 | NAV3299 | NAV4112NAV3244NAV5000NAV3243NAV3299 | True